## General
Pre-qualification Data Summary of Sector Assembly: General.
Sector ID | 20212000019005
Loading Site | LBNL
Loading date start date | 2004-05-04
Loading date end date | 2004-06-09
Thermal transfer fluid type | DI water, C6F14

## Survey
Pre-qualification Data Summary of Sector Assembly: Inspection - Survey.
Note | Category |  | Mod. #1 | Mod. #2 | Mod. #3 | Mod. #4 | Mod. #5 | Mod. #6
| Module ID |  | 510209 | 510679 | 510674 | 510216 | 510916 | 510174
Module Loading data
| Loading date | (dd/mm/yy) | 2004-05-04 | 2004-05-05 | 2004-05-05 | 2004-06-07 | 2004-06-08 | 2004-06-08
| Module weight | (g) | 5.75 | 5.73 | 5.75 | 5.74
| Module potted? | (yes/no) | yes | yes | yes | yes | yes
| Flex trim from sensor edge | FE-0/F (µm) | N/A | N/A | N/A | N/A | N/A
|  | FE-7/8 (µm) | N/A | N/A | N/A | N/A | N/A
| Glue shear test | (g) | N/A | N/A | N/A | 250g
A | Any damage visible? | (yes/no) | no | no | no | no
B | Loading proc. Logsheet | (yes/no) | no | no | no | yes | yes
Survey data
| Survey date | (dd/mm/yy) | 2004-05-11 | 2004-05-11 | 2004-05-11 | 2004-06-09 | 2004-06-09 | 2004-06-09
| Sensor mark#1 (FE-0) | Visible? (Y/N) | yes | yes | yes | yes | yes | yes
|  | Delta X (µm) | -13 | -18 | -12 | 5 | -3 | -7
|  | Delta Y (µm) | -11 | -10 | -14 | -8 | -2 | -3
|  | Delta Z (µm) | 41 | 11 | 22
| Sensor mark#2 (FE-7) | Visible? (Y/N) | yes | yes | yes | yes | yes | no
|  | Delta X (µm) | 82 | -3 | -16 | 15 | -8
|  | Delta Y (µm) | 87 | -3 | -5 | -3 | -1
D |  | Delta Z (µm) | 0 | 0 | 0
| Sensor mark#3 (FE-8) | Visible? (Y/N) | yes | yes | yes | yes | yes | yes
|  | Delta X (µm) | -87 | 4 | -18 | 14 | -12 | -2
|  | Delta Y (µm) | -30 | 3 | -10 | 1 | -2 | -1
D |  | Delta Z (µm) | 0 | 0 | 0
| Sensor mark#4 (FE-F) | Visible? (Y/N) | yes | yes | yes | yes | yes | yes
|  | Delta X (µm) | -10 | -11 | -8 | 7 | 0 | -6
|  | Delta Y (µm) | -4 | 1 | -14 | -7 | -4 | -1
|  | Delta Z (µm) | 41 | 11 | 22
| Horiz. tilt angle | (deg) | 0.14 | -0.06 | 0.06
C | Module bow Before/after loading | Before (µm) | -1 | 78 | 34 | -5 | 28 | 8
|  | After (µm) | -73 | 37 | 49
Note:
A | Attach explanation if the damage was noticed during initial visual inspection or it was done during mounting
B | Has the "ATLAS Pixel Module Loading on Stave Procedure Logsheet" been filled?
C | Module bow compared with the one from FlexModule
D | Delta Z for one end of module defined to be zero by measurement.

## Load
Pre-qualification Data Summary of Sector Assembly: Initial Loading Tests
Note | Category |  | Mod. #1 | Mod. #2 | Mod. #3 | Mod. #4 | Mod. #5 | Mod. #6
A | Module ID |  | 510209 | 510679 | 510674 | 510216 | 510916 | 510174
Thermal test with single module powered at time
| Test Date | (dd/mm/yy) | 2004-07-07 | 2004-07-08 | 2004-07-09 | 2004-07-09 | 2004-07-09 | 2004-07-09
| T power off | NTC (ºC) | 22.653 | 22.701 | 22.654 | 22.873 | 22.695 | 22.736
IR image (ºC)
B | Module V/I supply | IDDA (I) | 1.196 | 1.176 | 1.188 | 1.219 | 1.232 | 1.191
C |  | IDDD (I) | 0.94 | 0.971 | 0.9 | 0.904 | 0.897 | 0.899
|  | Power (W) | 3.7936 | 3.8236 | 3.7008 | 3.7584 | 3.7652 | 3.7036
| Temperature of fully configured module | NTC (ºC) | 25.12 | 24.84 | 26.02 | 25.96 | 24.97 | 25.05
IR image (ºC)
D | T input fluid | (ºC) | 22 | 22 | 22 | 22 | 22 | 22
| DT | (ºC) | 2.467 | 2.139 | 3.366 | 3.087 | 2.275 | 2.314
E | DT / W | (ºC/W) | 0.65031 | 0.55942 | 0.90953 | 0.82136 | 0.60422 | 0.6248
Threshold scan with single module powered at time
| Mean Threshold Before/After Loading | Before | 3128 | 3162 | 3169 | 3186 | 3134 | 3141
|  | After | 3133 | 3142 | 3194 | 3147 | 3143 | 3129
| Sigma Threshold Before/After Loading | Before | 61.1 | 55.7 | 60.77 | 49.73 | 57.49 | 60.45
|  | After | 54.68 | 60.22 | 57.23 | 59.51 | 50.65 | 62.95
| Mean Noise Before/After Loading | Before | 170.7 | 176.1 | 177.7 | 172.4 | 172.6 | 174.2
|  | After | 176 | 181 | 182.7 | 175.1 | 175 | 178.1
| Sigma Noise Before/After Loading | Before | 14.79 | 13.79 | 15.6 | 15.75 | 16.52 | 15.31
|  | After | 10.63 | 10.56 | 11.74 | 12.82 | 11.8 | 10.87
Cross-talk scan with single module powered at time
F | Mean Cross-talk Before/After Loading | Before | 1.111 | n/a | 1.929 | 0.9756 | 1.145 | 1.172
F |  | After | 1.113 | n/a | n/a | 0.9746 | 1.141 | 1.142
F | Sigma Cross-talk Before/After Loading | Before | 0.0663 | n/a | 1.473 | 0.1066 | 0.07186 | 0.06481
F |  | After | 0.05989 | n/a | n/a | 0.06952 | 0.06671 | 0.07009
| Mean Long Cross-talk Before/After Loading | Before | 3.04 | 1.908 | 1.81 | 2.26 | 3.034 | 3.083
|  | After | 3.115 | 1.893 | 1.817 | 2.279 | 3.13 | 3.085
| Sigma Long Cross-talk Before/After Loading | Before | 0.1891 | 0.2122 | 0.2807 | 0.325 | 0.1979 | 0.1792
|  | After | 0.2229 | 0.2626 | 0.3093 | 0.3238 | 0.2205 | 0.215
Disconnected Pixels
G | Exact Number | Before | 4 | 21 | 10 | 40 | 37 | 2
H | Estimate Before/After Loading | Before | 0 | 0 | 0 | 0 | 9 | 0
H |  | After | 0 | 0 | 0 | 0 | 1 | 0
Note:
A | Module ID from PDB.
B | AVDD = 1.6 V
C | DVDD = 2.0 V
D | Chiller set point.
E | Delta T (per watt) between the temperature of the module on and off divided by the input power.
F | n/a: Fit gives value less than 0.
G | From Source Scan
H | From VCAL Scan without sensor bias.

## Thermal
Pre-qualification Data Summary of Sector Assembly: Thermal - Burnin.
Note | Category |  | Mod. #1 | Mod. #2 | Mod. #3 | Mod. #4 | Mod. #5 | Mod. #6 | Environment
A | Module ID |  | 510209 | 510679 | 510674 | 510216 | 510916 | 510174
Thermal test/burn in with no module powered (min-max temperature)
| Test Date | (dd/mm/yy) | 2004-08-03 | 2004-08-03 | 2004-08-03 | 2004-08-03 | 2004-08-03 | 2004-08-03
B | Low temperature | NTC (ºC) | -16.72 | -16.59 | -16.6 | -16.38 | -16.87 | -16.9 | Cold box N temperature (ºC)
C | High temperature | NTC (ºC) | 20.36 | 20.38 | 20.32 | 20.46 | 20.37 | 20.33 | Cold box relative humidity (%) | 8.3
D | No. of thermal cycles: 3
Thermal Image test with single module powered at time (T = 10÷15 ºC)
| Test Date | (dd/mm/yy) | 2004-07-17 | 2004-07-17 | 2004-07-17 | 2004-07-17 | 2004-07-17 | 2004-07-17
| T power off | NTC (ºC) | 17.95 | 17.83 | 17.99 | 18.33 | 18 | 18.08 | Cold box N temperature (ºC)
|  | IR image (ºC) | 20 | 19.6 | 19.8 | 20.5 | 20.5 | 19.7 | Cold box relative humidity (%)
| Module V/I supply | IDDA (I) | 1.18 | 1.161 | 1.174 | 1.204 | 1.218 | 1.174 | AVDD (V) | 1.6
|  | IDDD (I) | 0.895 | 0.896 | 0.9 | 0.906 | 0.895 | 0.898 | DVDD (V) | 2
|  | Power (W) | 3.678 | 3.6496 | 3.6784 | 3.7384 | 3.7388 | 3.6744
| Temperature of fully configured module | NTC (ºC) | 20.67 | 19.93 | 21.31 | 21.09 | 20.24 | 20.31
|  | IR image (ºC) | 21.05 | 21 | 22.9 | 22.15 | 21.7 | 21.7
| T input fluid | (ºC) | 16 | 16 | 16 | 16 | 16 | 16
| ΔT | (ºC) | 2.72 | 2.1 | 3.32 | 2.76 | 2.24 | 2.23
E | DT / W | (ºC/W) | 0.73953 | 0.57541 | 0.90257 | 0.73828 | 0.59912 | 0.6069
Thermal test with three modules powered at time(T = -7÷-10 ºC)
| Test Date | (dd/mm/yy) | 2004-08-05 | 2004-08-05 | 2004-08-05 | 2004-08-06 | 2004-08-06 | 2004-08-06 | Cold box N temperature (ºC)
| T power off | NTC (ºC) | -13.95 | -12.45 | -12.55 | -13.55 | -12.65 | -12.85 | Cold box relative humidity (%) | 7.8
| Module V/I supply | IDDA (I) | 1.08 | 1.01 | 1.03 | 1.13 | 1.06 | 1.01 | AVDD (V) | 1.6
|  | IDDD (I) | 0.88 | 0.81 | 0.81 | 0.93 | 0.71 | 0.73 | DVDD (V) | 2
|  | Power (W) | 3.488 | 3.236 | 3.268 | 3.668 | 3.116 | 3.076
| T power on | NTC (ºC) | -9.45 | -8.25 | -7.15 | -8.15 | -8.35 | -9.25
| T input fluid | (ºC) | -17 | -17 | -17 | -17 | -17 | -17
| ΔT | (ºC) | 4.5 | 4.2 | 5.4 | 5.4 | 4.3 | 3.6
E | DT / W | (ºC/W) | 1.2901 | 1.2979 | 1.6524 | 1.4722 | 1.38 | 1.1704
Note:
A | Module ID from PDB
B | Thermal cycle with no power on module. Lowest temperature measured on module NTC (should be between -25 and -30 ºC)
C | Thermal cycle with no power on module. Highest temperature measured on module NTC (should be between 25 and 30 ºC)
D | Number of thermal cycles between minimum and maximum temperature
E | Delta T (per watt) between the temperature of the module on and off divided by the input power.

## Electrical
Pre-qualification Data Summary of Sector Assembly: Inspection - Survey.
Note | Category | Test Group | Units | Mod. #1 | Mod. #2 | Mod. #3 | Mod. #4 | Mod. #5 | Mod. #6
A | Module ID |  |  | 510209 | 510679 | 510674 | 510216 | 510916 | 510174
B | Root files on server for all test? | STAVE | Y/N
C | T (ºC) | STAVE
D | Sensor I_leakage @ 150 V | FLEX | I (mA)
T (ºC)
D | Sensor I_leakage @ 150 V | STAVE | I (mA)
T (ºC)
E | Source scan open bumps | FLEX | No.
|  | STAVE | No.
E | Test No. bad bump candidates | FLEX | No.
|  | STAVE | No.
| Analog scan average noise | FLEX | (e)
|  | STAVE | (e)
| Threshold dispersion | FLEX | (e)
|  | STAVE | (e)
| MCC VDD threshold | FLEX | (e)
|  | STAVE | (e)
Note:
A | Module ID from PDB
B | Root file must be put under the tree <ModuleID>/STAVE/01/
C | Temperature measured by the module NTC (supposed constant during the test). Temperature must be between -7 and -10ºC.
D | Sensor leakage current to be compared to the one of the FLEX test group. Specify temperature if made at different value with respect to the other tests.
E | From module analysis bad bump candidate cuts, to be released after 14 april 04.