FOOD AND DRUG ADMINISTRATION

+ + + + +

GENERAL AND PLASTIC SURGERY DEVICES PANEL

OF THE

MEDICAL DEVICES ADVISORY COMMITTEE

+ + + + +

EMPHYSEMA AND ABLATION DEVICES CLINICAL ISSUES

DISCUSSION SESSION

+ + + + +

FRIDAY,

FEBRUARY 28, 2003

The Panel was convened in the Grand Ballroom of the Gaithersburg Hilton Hotel, 620 Perry Parkway, Gaithersburg, Maryland at 1:30 p.m., Dr. Robert L. McCauley, Acting Chair, presiding.

PRESENT:

ROBERT L. McCAULEY, M.D.		Acting Chair

DAVID KRAUSE, Ph.D.				Executive Secretary

MICHAEL J. MILLER, M.D.			Voting Member

DEBORAH M. BROWN				Industry Representative

CHRISTINE MOORE				Consumer Representative

BRENT A. BLUMENSTEIN, Ph.D.		Temporary Voting Member

DEAN E. BRENNER, M.D.			Temporary Voting Member

THOMAS B. FERGUSON, M.D.			Temporary Voting Member

JOSEPH LoCICERO, III, M.D.		Temporary Voting Member

DEBORAH SHURE, M.D.				Temporary Voting Member

STEPHEN B. SOLOMON, M.D.			Temporary Voting Member

JAMES K. STOLLER, M.D.			Temporary Voting Member

CELIA WITTEN, Ph.D., M.D.		FDA

C-O-N-T-E-N-T-S

Call to Order	4

Panel Introductions	5

Robert L. McCauley, MD, Acting Chairman

FDA Presentation	7

Introduction	7

Dr. Binita Ashar

Industry Presentations	8

Broncus Technologies, Incorporated	7

Dr. Joel Cooper

Closure Medical Corporation	18

Dr. Thomas Stevens

Emphasys Medical, Incorporated	23

Dr. Geddes

Spiration, Incorporated	35

Dr. Rick Shea	35

Dr. Matthew Brenner	36

Dr. Douglas Wood	39

Open Public Comments	44

Dr. Joel Cooper, ATS	44

Dr. Joe Putnam	45

Panel Deliberations and Address	50

of FDA Questions

Discussion of Question 1	47

Discussion of Question 2	62

Discussion of Question 3	68

Discussion of Question 4	79

C-O-N-T-E-N-T-S (Continued)

FDA Presentation	85

Introduction	85

Dr. Binita Ashar

Industry Presentations	86

RITA Medical Systems, Incorporated	86

Dr. Barry Cheskin

Dr. Steven Rosenberg	88

Open Public Comments	105

Dr. Joe Putnam	106

Dr. Patrick Sewell	112

Dr. Joel Cooper	118

Panel Deliberations and Address	123

of FDA Questions

Discussion of Question No. 1	123

Discussion of Question No. 2	138

Discussion of Question No. 3	145

Discussion of Question No. 4	152

Discussion of Question No. 5	157

P-R-O-C-E-E-D-I-N-G-S

1:25 p.m.

DR. KRAUSE:  Okay, if you don't sit down in the next five minutes, we're going to make you stay until 10:00.  A little FDA humor there.

[Laughter.]

All right, I was going to read the conflict of interest and temporary voting status again, but I don't really need to because this afternoon there's no vote.  So these people don't need to deputized, which is good.  I read the conflict of interest this morning, so I'm not going to do it again.

So, in that case, I would like to turn the afternoon session over to Dr. McCauley, our Chairman.

CHAIRMAN McCAULEY:  Good afternoon.  My name is Dr. Robert McCauley.  I'm Professor of Surgery and Pediatrics at the University of Texas Medical Branch in Galveston and Chief of the Plastic Surgery Services for Shriners Burns Hospital in Galveston.

This afternoon the Panel will be making recommendations to the Food and Drug Administration regarding clinical concerns involving two classes of medical devices:  those intended to treat emphysema and those intended for ablation of lung tumors.

The next item of business is to introduce the Panel members, who are giving of their time to help the FDA in these matters, and the FDA staff at this table.  I'm going to ask each person to introduce him or herself stating his or her specialty, position, title, institution, and his or her status on the Panel, starting with my near left.

DR. SHURE:  My name is Deborah Shure.  I'm Professor of Medicine at the University of Texas Health Science Center at Houston, and I'm Director of a multidisciplinary program in pulmonary vascular disease there.  I'm a pulmonologist and I'm a consultant to the Panel.

DR. MILLER:  I'm Michael Miller.  I'm a Professor of Plastic Surgery at the University of Texas M.D. Anderson Cancer Center in Houston, and I'm a member of the Panel.

DR. BLUMENSTEIN:  My name is Brent Blumenstein, and I'm not from M.D. Anderson.

[Laughter.]

I'm a private consultant on the Panel, a biostatistician.

DR. STOLLER:  My name is Jamie Stoller.  I'm a lung doctor.  I'm at the Cleveland Clinic, where I am a member of Pulmonary and Critical Care Group.

DR. BROWN:  I'm Debera Brown.  I'm with Broncus Technologies in California, the Vice President of Regulatory Affairs, and I'm a regular Panel member as the industry representative.

MS. MOORE:  My name is Christine Moore.  I'm the consumer member of the Panel, and I'm the former Dean of Student Services at the Community College of Baltimore.

DR. WITTEN:  I'm Dr. Celia Witten, the Division Director of GRND in FDA.

DR. FERGUSON:  I'm Tom Ferguson from Washington University in St. Louis, a Meritus Professor of Surgery there, and I'm a consultant to the Panel.

DR. BRENNER:  I'm Dean Brenner.  I'm a Professor of Internal Medicine and Professor of Pharmacology at the University of Michigan.  I'm a medical oncologist, and I presume I'm supposed to advise on oncologic things here, and I'm a consultant.

DR. LoCICERO:  I'm Joseph LoCicero.  I am a thoracic surgeon.  I'm currently the Chair of the Department of Surgery at the University of South Alabama in Mobile, and I'm a Panel member.

CHAIRMAN McCAULEY:  I would like to remind public observers that while this portion of the meeting is open to public observation, public attendees may not participate except at the specific request of the Panel.

We are now ready to hear the FDA presentation.  The FDA is going to give us some background information for this topic.

DR. ASHAR:  Thank you, Dr. McCauley.

Good afternoon.  My name is Binita Ashar.  I'm a Medical Officer at FDA.  We have asked you, the Panel, the General and Plastic Surgery Devices Panel, to consider clinical trial issues concerning the application of devices placed in local areas of the lung for the treatment of emphysema.

As you know, the present treatment options for patients are evolving.  Common treatment includes medical management, occasionally lung transplantation for severe cases, and potentially lung volume reduction, which is still under study.

There has been some interest in developing devices placed in areas of the lung to achieve local physiological changes in the lung as a treatment for emphysema.  Device designs, as discussed in the literature, include the use of fibrin-based glues, occluded stents, medical adhesives, and inter-bronchial valves.

The purpose of this session is to obtain the Panel's thoughts on clinical trial issues to be considered when studying devices that produce local pulmonary physiological effects for the treatment of patients with emphysema.

Following the presentations made by the sponsors this afternoon, we will be asking you, the Panel, to comment on a number of questions, including an appropriate control group, safety parameters, efficacy endpoints, and duration of followup.

Thank you very much and, Dr. McCauley, I will hand over the questions to you for review by the Panel after the presentations made by the sponsors.

CHAIRMAN McCAULEY:  We will now hear from some members of industry in regards to this topics.  Please come up in this order:  The first would be Broncus Technology; followed by Closure; followed by Emphasys Medical, Incorporated; followed by Spiration Incorporated.

Bronchus Technologies, Incorporated.

DR. COOPER:  Good afternoon, Mr. Chairman.  I'm Joel Cooper.  I'm a thoracic surgeon.  I am Chief of Cardiotheracic Surgery at Washington University School of Medicine.  I'm here in several capacities.  Primarily, I'm as a representative for Broncus, a company to which I took a proposed intervention and a project which I have been working with them to develop.

I also have other sort of roles.  I consult for Emphasys, one of the companies which will be presenting.  I have a relationship with a company called Synovis, which has a product we use in volume reduction surgery, which is a staple line reinforcement which I developed some years ago for volume reduction surgery, so I could be perceived as having an interest in volume reduction surgery.  I'm also here as Vice President and President-Elect of the American Association for Thoracic Surgery, as their representative.

The items that I will be discussing are shown here.  These are the headings of the subsequent slides, some of the issues that you wished discussed.

I need to introduce this with a few words about emphysema.  Chronic obstructive pulmonary disease, the fourth leading cause of death in the United States, is quite prevalent.  However, the term "COPD," or chronic obstructive pulmonary disease really carries two different disease populations under that rubric.  One is emphysema and one is airways disease, or what is called chronic bronchitis or asthmatic bronchitis.

For many of the therapies for COPD, the distinction is not too important, but for device therapies, which are proposed to treat emphysema, it is a critical distinction to be able to make.  Even among the pulmonary population I must say that this distinction is not always perceived, and many patients are referred for treatment of emphysema who, in fact, do not have significant emphysema but who have the other form of COPD, the airways disease, the chronic bronchitis.

Emphysema specifically refers to destruction of the fabric of the lung, a loss of its physical elastic properties, and a loss of the gas exchange surface on the alveoli.  As a person with emphysema has progressive disease, they are limited by the gas exchange problems; oxygen and carbon dioxide exchange is limited, and that may limit their activities.  But, in the end, when they are crippled by emphysema, it is the mechanics of breathing, the persistent overinflation of the hyperexpanded chest, the distorted diaphragm and chest wall, the effort of breathing, and the associated dyspnea or sense of pervasive breathlessness, which cripples these individuals.

And the goal of treating emphysema by any means has always been an attempt to reduce what is called dynamic hyperinflation, the oversize of the chest, which these patients have.  They are trapped in the state of permanent hyperinflation, and no amount of forced expiratory maneuver can empty the air out of their chest and allow them to take a deep breath.

The treatment options for emphysema are, and will continue to be, mainly medical management, antibiotics, bronchodilators, steroids, all designed to try to improve expiratory flow from the patient with emphysema and with asthmatic bronchitis, and it will continue to be the mainstay.

We were fortunate in 1983 to be able to be successful with the initial lung transplants in humans which were used for a variety of diseases, including emphysema, and emphysema continues to be one of the main indications for lung transplantation.  For those patients, any pattern of emphysema is suitable because you are removing the lungs.

But when you look at specific therapies, the pattern in which the destruction occurs is extremely germane to the treatment being considered.  Volume reduction surgery, something we developed that began 10 years ago, was based upon observations in transplant patients and requires that the patient have a particular pattern of emphysema.  We call that heterogeneity.  One part of the lung very badly destroyed; that's the part that's removed, to allow the remaining lung, less destroyed, to function better and to improve the respiratory mechanics.

The volume reduction surgery, five presented randomized trials already, two more in press, in addition to the NETT trial, 200 publications, all showing its value, has generated a fair amount of interest.  The NETT results, soon to be published, will confirm the previous findings, I understand.

Because of the success of volume reduction surgery, which taught us that you can change the mechanics of a patient who is crippled from severe emphysema, and by changing those mechanics, you can alter the quality of life dramatically, that has led some to look at alternative proposals to try to achieve the same benefit -- you will hear from some of those this afternoon -- an attempt to achieve the benefit of volume reduction surgery by in some ways reducing the volume of the targeted, overextended, overdistended lung.

The program that I am working with at Broncus is a totally different concept, and that is it's an attempt to provide, through yet another group of patients, an alternative to medical therapy when there is no existing one.  Those patients are those who have a fairly homogeneous pattern, top to bottom, of severe emphysematous destruction, who are not candidates for lung transplant because they are too old or have co-existing morbidities, who are not candidates for volume reduction surgery or one of its substitutes because they do not have that heterogeneous pattern of emphysema.

So for an intervention which is designed for people who have no alternative, we would argue that the only possible control you have is continued medical management when evaluating these new therapies.  If you are proposing a therapy as an alternative to an existing therapy, then you may, in fact, come up with a difference as to what should be the control group.

I want to emphasize again that for homogeneous emphysema, currently, medical management is the only option.  These are severely-disabled patients.  They cannot take a shower easily or at all.  It takes an hour to get dressed.  They cannot visit the grandkids.  They are trapped with a hyperinflated state.  They just want to put a gun to their head, and they will accept almost any risk if they could be provided a period of relief from the disabling dyspnea which they have.

One of the questions that you are addressing this afternoon is, how do you evaluate therapies?  I would argue that it depends on who is your target population.  What alternatives do they have?  What's the nature of your intervention?

The history of the last century is filled with devices and procedures promoted for the relief of breathlessness in patients with severe emphysema which ended upon the waste heap of history, and for one reason:  All of these procedures, glomectomy, phrenicectomy, chostochrondrectomy, radical denervation, all promoted to relieve symptoms, had only subjective benefit, no objective, verifiable benefit to correlate with the subjective benefit.

And the Santayana has taught us those who fail to learn the lessons of history are condemned to repeat them.  When you are dealing with dyspnea and breathlessness, and you propose to improve it, you must have something objective, at least a function measure which is physiologically logical in terms of the intervention and, ideally, one quality-of-life or subjective improvement.

The patient doesn't come to you and say, "Dr. Cooper, my FEV1 is killing me and my residual volume is out of sight."  They say, "I can't get dressed.  I can't take a shower.  I can't breathe."  Nothing is more important than subjective benefit, but it cannot alone be trusted because of the sham effect that you can have, unless there is a physiologically-relevant, objective measurement which can also be monitored.

In looking at what you might choose, one spirometric measurement, whether it's the vital capacity, the force vital capacity, or the FEV1, I think is very relevant, and a 15 percent improvement in any of those, I would argue, is a relevant improvement.

I personally believe that the residual volume [Dr. Cooper blows air out of his lungs], the amount of air that's left after the maximum forced improvement, is a very difficult measurement to make in anybody.  You have to have a body box.  In these patients, it is very difficult to make.

But I would argue that, if you could do it and do it accurately, it is probably the most relevant because the size of your chest when you're breathing is what really determines your work of breathing and your sense of breathlessness.  So I believe that is important to measure, but I recognize it takes specially-trained technicians, and not all patients can do it.  So I would argue for a spirometric measurement as perhaps a surrogate.

And then at least one measurement of function or quality of life or subjective benefit, be it a well-verified, quality-of-life questionnaire, a dismal using the MRC scale, or a functional test.  Some of these are more eloquent, but, again, very complicated.

Again, for homogeneous emphysema, looking at safety, when you discuss a safety issue, and what is the risk that a patient should be subjected to, again, I would argue that you must look at whether or not there's an acceptable alternative to the proposed new treatment.

If there is no alternative, and if the patient is severely disabled, and particularly if the patient has a limited life expectancy, then your standard may be one thing.  If the patient isn't as disabled, if there is an alternative therapy, but maybe this is being proposed because it's less risky, then I think you may have to set a different standard.

I would only argue that you must look at the target population for these different devices and interventions before you decide where the risk/benefit ratio lies.  One thing is in your favor, in discussing these things:  The natural history of emphysema is very well-described.

Nobody stays the same.  Everybody gets worse at a predictable rate.  Life expectancy for the patients who will be considered for these interventions, medium survival is less than three years, and for the homogeneous, severe patients, probably about a year and a half.

In terms of safety -- others will talk about this more than I have time to -- there are obvious things you need to look for in terms of these interventions, whether it's hemoptysis, which can be serious; pneumonia; sacerbation of their respiratory status, intubation, air leaks, and death, but death is not the worst endpoint in our patients who are subject to these operations, be it transplant or volume reduction; know that.  The worst thing is to be made worse and to end up ventilator-dependent and more miserable than you start out with.

So it is important, again, to look at the disease you are treating, the intervention, the target population for these in making your decision.  I don't think, in terms of followup, you are addressing that.  I would argue the same thing.  if you have no alternative, if life expectancy is short, as in severe, homogeneous emphysema, three to six months of benefit may be very worthwhile.  It's like a patient with cancer of the esophagus when we put in a stent, so they can swallow and have a reasonable dignity and quality of life, even though it may be only for a matter of months.

On the other hand, if you are looking at an alternative to an existing therapy, then your period of followup clearly has to be longer to justify an alternative in comparison with the current standard of care.

At any rate, I thank you very much for the opportunity of presenting, and if there is an opportunity to answer any questions later in the day, I will be pleased to do so.  Thank you.

CHAIRMAN McCAULEY:  Are there any questions from the Panel members for Dr. Cooper?

[No response.]

We will now proceed with the presentation by Closure Medical, Incorporated.

DR. STEVENS:  Hello.  My name is Tom Stevens.  I'm Director of Regulatory Affairs for Closure Medical Corporation.  We are here today to talk about our endobronchial occlusion device.

With me today I have Joe Barefoot, my boss, Vice President of Regulatory Affairs and Quality Assure at Closure Medical; Jeff Clark, Vice President of R&D for Closure; Kevin Kovitz, interventional pulmonologist from Tulane University Medical Center could not be with us today, and David West, Vice President of Qintas Consulting.

We have heard just now about the benefits of lung volume reduction surgery.  The current modalities of achieving this are median sternotomy, as well as video-assisted thoracic surgery, or VATS.  Currently, surgical mortalities of 16 to 25 percent have been reported for the lung volume reduction procedure; zero to 9 percent for the VATS procedure up to a period of a year after the procedure has been performed.

We are looking at potential alternatives to the surgical procedure.  Bronchoscopic procedures to achieve ongoing reduction are believed to possess less risk associated with morbidity and mortality compared to the surgical intervention.

The Closure medical device depends on placement of occlusive plug using bronchoscopy and is expected to result in segment-specific atelectasis, reducing lung volume.  Closure's endobronchial occlusion device is expected to offer, through atelectasis, a less-invasive, lower-risk modality of achieving ongoing reduction than the surgical therapy.

The Closure medical device is a liquid that is delivered through the catheter and polymerizes in place within the segment in the lung.  It is composed of a proprietary cyano acrylate formulation that is combined with an activator or a polymerization initiator, viscosity modifier, and a radio-opaque agent.

Delivery of procedure:  The bronchoscope is introduced into the desired lobe with conscious sedation.  The balloon catheter is positioned within the target bronch segment, and the delivery syringe is attached and the liquid is delivered through the catheter.  Subsequently, the catheter and bronchoscope are removed, once polymerization has been achieved.

We have evaluated this device in a pre-clinical model.  Our objectives:  to provide proof of concept for an occlusive medical device, to achieve safe and effective atelectasis at the segmental level.  Our objectives also included development of an associated procedure for placement of the device as well as pre- and post-operative therapy to manage the risk of infection.

Pre-clinical study design was a goat model used to provide an accessible tracheal bronchus and an isolated upper lobe, rigid bronchoscope and six French Berenstein occlusion balloon catheters were used, and the occlusive plug was placed into the proximal tracheal bronchus.

The design included a seven-day antibiotic treatment pre- and post-placement, endoscopic and x-ray analysis at approximately seven days and ninety days post-treatment, after which the animals were euthanized for lung explant, culture, and histopathology.

This is a bronchoscopic view of the tracheal bronchus to the lower right, and you can see the occlusion device in place at day ninety.  In this picture it is in place and intact.

This is a photograph of the explanted lung.  The segment here is where the occlusion device is placed, and we were happy with this result.

So our animal study, what we concluded from that was that we were able to demonstrate a concept of an occlusive device to achieve atelectasis.  We were successful in developing a procedure and device formulation to achieve the desired placement and retention of the plug, and also were successful in managing the infection risk, using the sterile product and the antibiotic therapy.

The issues for FDA for the study design, from our point of view, need to focus on the necessary and sufficient information to demonstrate safety and effectiveness; define the expectations for clinical utility, and balance to achieve the appropriate level of scientific rigor with consideration of the practical restraints of undertaking these studies.

Specifically, questions for FDA, and with the help and guidance of the Panel, are:  Are there appropriate objective performance criteria or other scientific means for designing studies of safety and effectiveness that obviate the requirement for active control studies; that is, the surgical intervention?

If control studies are required, what would be the appropriate control, surgical lung volume reduction or medical management?  Answer the question of:  How are the subjects to be recruited and retained?  Are the inclusion and exclusion criteria to be based strictly on the NETT protocol or are there other criteria that can be based on the NETT experience?  What are the appropriate clinical endpoints:  pulmonary function, quality of life, or others?  And what is the appropriate study duration?

That concludes the presentation of Closure Medical.

CHAIRMAN McCAULEY:  Does anyone on the Panel have any questions for Dr. Stevens?

DR. LoCICERO:  Yes, I do.  While we're not here today to discuss the pre-clinical trial as such, you did present something that might affect the length of followup or the type of followup.  Can you comment on how the cyano acrylate is molded?  Is this done outside of the body or is this done inside of the body?

DR. STEVENS:  It's done inside of the body.  The cyano acrylate is mixed with the other components after the bronchoscope has been placed, then introduced through the catheter as a liquid into the segment, where it then polymerizes within a period of a few minutes.  So it's formed in place.

CHAIRMAN McCAULEY:  Any other questions?

[No response.]

We will now proceed with Emphasys.

DR. GEDDES:  Good afternoon.  My name is Dr. Geddes.  I'm a Professor of Respiratory Medicine and Pharmacology at the Royal Brompton Hospital in London.  I am speaking on behalf of Emphasys, though I am not part of Emphasys.  I'm not a company person.  I'm not a regulated person.

I run an emphysema clinic, and we have experience in both volume reduction surgery and early experience with the endobronchial valve system that Emphasys has developed.  It is for that reason that I have been asked to talk about my views as to the future ideal clinical trial.

I am just starting a combination of all the randomized control trials published for volume reduction surgery to date.  This does not include the NETT study and does not include the other two that Joel Cooper mentioned in press.

Just to give an indication of what can be achieved in these patients, an improvement of FEV1 measured between 15 and 30 percent, while the control, the medical control group, are declining.  One has to remember that these patients are getting worse during periods of study.  Six-minute walking distance improved or declining in the control group, but a significant mortality in the surgical group, a significant mortality in both groups, but more in the surgical group.  So any treatment that could provide similar or possibly less benefit with less morbidity and mortality is certainly worth exploring.

The endobronchial devices that we have used, have been using, and I just want to show this slide as an example to make one simple point.  Look at the x-ray top left and compare it with the x-ray bottom right.  There you see emphysema in the upper lobes, a very black lung at the top right.  After placement of the valves, look at this one.  You see the diaphragm has come up, and this lung shows loss of volume and expansion of some of the lung that was not expanded before.

The reason I am showing that is just to make the simple point that endobronchial devices can work at volume reduction.  So the question then is:  Is that going to provide patient benefit and are they safe?

The safety experience with these valves is now based on 91 patients with 365 valve insertions.  These are in pilot, uncontrolled studies, allowing a period of observation of 14,000-odd patient-days.  So there is building already quite a lot of information about these devices.

I have summarized the adverse effects that have so far been reported:  one death that was probably procedure-related, infection in the occluded segment; three important episodes of infection, and eight episodes of pneumothorax, three of which lasted more than a week.

There have been some late events, which I have classified separately.  For example, a death at ninety days from interstitial perforation is unlikely to be related to the device.

So this begins to define some of the safety measures that should certainly be assessed as any trial goes on.

And I just want to share with you the experience we have in a very limited group of nine patients in whom we have tried this treatment as a real treatment.  These are people who we selected for volume reduction surgery, but did not go forward for surgery, either because they were too sick or because they refused surgery.

Looking in green, you can see the changes that followed bronchoscopic placement of the valve as compared with our own controlled clinical trial of volume reduction surgery.  Now this is not a randomized group.  I just want to make the simple point that the types of changes that we are getting with the bronchoscopic approach are comparable with those that can be achieved with the surgical approach.

The patients are not the same and our entry criteria are different.  So I don't want to say that the treatments are in any way equivalent; just that similar sorts of measures of improvements can be achieved.  They are not always achieved.  We have four start out of nine; three didn't get any benefit.  So we are still learning.

So if we then go on to what studies might be done, and what outcome measures might be appropriate to use, this has largely been covered.  They are already quite conventional.  You can take these from the volume reduction studies and others:  A physiological measure, an FEV1 always is the conventional one, although it has its faults.  Some measure of exercise performance.  The six-minute distance has been most commonly used, but other maximum exercise or endurance exercise tests may be better.

Some measure of quality of life, we use the St. Georges Respiratory Questionnaire.  It's probably the best available and is more appropriate for these patients than a global quality-of-life measure.

Then I would propose that a tertiary measure of hospital days would give information that I would find useful for what the patients are having to endure following the randomization to whatever the control and active groups are, and also it is a surrogate for costs as well as patient unpleasantness.

I think the outcome measures are relatively straightforward.  What I find more difficult is what the exact trial design should be.  As a clinician, these are the questions that I want to answer:

Does the system, whether it is a valve or any of the others, work?  Is it safe?  And then we also need to know -- and these come in smaller print because they are secondary questions -- we need to know what are the ideal patients.  I think we only learn that by experience, not by theory.  It may well be that we will want to do a comparison with surgery at some stage.  But until we know who the ideal patients for each group, it would be quite inappropriate to use surgery as a comparison group.  The results will depend on selection of patients rather than the treatment itself.

Already we have some information.  I know that these can work.  I don't know if they work well enough because we don't have an adequate control group.

Is it safe?  I know the Emphasys valve is probably safe with a number of 92; a previous number of 91 you saw.  It's come up to 92 because we did another patient yesterday morning.  The long-term safety and the rare event safety events, obviously, will need larger groups.

Selection criteria will depend on more experience and probably will be gained from a longer period than the initial controlled trial, and a comparison with surgery, I'm sure, will be interesting at some stage in the future, but not yet.

My own view is that these devices, if they do work and are safe, are likely to find their place as a trial of treatment before surgery but not be used instead of surgery.  The patients, when you talk to them about it, they say, "If there is a way that might be simpler and safer, then I would like to try it first, but if it doesn't work, can I go on to the other?"  And that is probably what any of us would say.

So there are three clinical trial designs that I think are possibly worth considering.  One would be bronchoscopic volume reduction against surgical.  I think it's premature, but that would be a possibility.  Another would be just to use the patients as their own controls and compare pre and against post, and the third would be to use bronchoscopic volume reduction against best medical therapy, following the tradition set by volume reduction surgery.

My own personal view is that the first one is premature until we know better what the selection criteria should be and who are the ideal patients for each.  Then possibly randomize those in whom it is a real option, one or the other.

The pre and post without an adequate other non-treated control group I think is probably inadequate because we won't be able to make valid conclusions about all the outcome measures.  My own preference would be to go for the bronchoscopic insertion of the device against best medical treatment, but there are problems.

And the problems are that you have a treatment with surgery, which already has a lot of controlled trial data to suggest that it works in a patient group who do not have a lot of time.  So I think any trial design would have to take that into account, and we would need to define entry criteria well, so that those who were clearly suitable for surgery didn't go into this.

We would include surgical candidates and those in whom it was equivocal as to which was the best treatment, which would be the majority.  We would do a primary comparison at six months, by which time you would know whether the thing has worked, and also you would have given the patients enough time for the benefit to be worthwhile.  Even three months is worthwhile to these patients.

But in those who did not improve, then surgery would need to be allowed.  I think it would be unethical to ask these patients to remain in a control trial for years, during which they are likely to die without ever having access to a treatment which has a possible benefit.  There would need, then, to be a long surveillance period after the trial to make sure that late and rare side effects were properly picked up.

I would hope that we would be able to do these things relatively soon because there are millions of patients like that for whom we do not have a treatment at the moment, and these devices do, in our own experience, offer real potential benefits.

We have one patient who, having had a bronchoscopic insertion, refused to go on the television to report the benefits because he would now lose his disability benefit because he was able to do so much more.

[Laughter.]

So, at a purely anecdotal level, this can work.  I would like to be able to get to the position where we know that it does work and for enough people.  Thank you.  That's all I have to say.

CHAIRMAN McCAULEY:  Thank you.

Are there any questions from the Panel for Dr. Geddes?  Dr. Ferguson?

DR. FERGUSON:  Yes, that was very nice.  All of these presentations do not mention in the protocol one of the things that I have been most impressed with in the lung volume reduction group that I've seen at Washington University, and that is the positive benefits that you get with several months of intense medical therapy over and above what we call medical therapy, in terms of getting the patients ready for whatever they're going to have.

I think it's important to emphasize that that is a standard for lung volume reduction surgery, and I would not like to see that lost as we talk about other devices.

DR. GEDDES:  Could I just respond?

CHAIRMAN McCAULEY:  Do you want to comment on that?

DR. GEDDES:  I completely agree with that, and it is essential because these patients are sometimes neglected.  In our own trials, both volume reduction and of the endoscopic approach, they have all been through a period of months of optimization treatment, including formal physical rehabilitation.  So I think that's essential.  Some of the patients no longer are interested at the end of that time.

CHAIRMAN McCAULEY:  Other questions?  Dr. Miller?

DR. MILLER:  Yes, I'm not familiar with the device you're talking about.  Is it a physical device that totally occludes the airway?

DR. GEDDES:  It's equivalent to a metal stent with a plastic surround but a duck-billed valve in the center to allow air to come out of the occluded segment, but not to go in.  So it is not complete occlusion.  It is occlusion on inspiration, but airflow possible on expiration.  So mucous or air can get out of the occluded segment.

DR. MILLER:  Does it cause atelectasis in the distal lung the way the occlusive device does?

DR. GEDDES:  What it has done in half of the patients, formulate atelectasis of the lobe we've been targeting in some partial atelectasis; in others none at all.  Patient selection to achieve atelectasis is an important consideration.

DR. MILLER:  Can it migrate or can you take it out once you put it in?

DR. GEDDES:  You can take it out.  They have been taken out as late as, I think the latest is three months after insertion.  So far, migration has not been a problem.  They expand on placement, but those are a good example of the long-term surveillance requirements because I think, in fact, there will be those sort of problems occasionally.  It hasn't occurred yet.

DR. MILLER:  Thank you.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  Most foreign bodies would have significant granulation tissue and fibrosis by three months.  Is this Teflon or something else that would prevent that?

DR. GEDDES:  Yes.

DR. SHURE:  You say it can be removed at three months?

DR. GEDDES:  Well, that's right.  Again, that is the experience so far with the longest period of insertion that I personally have looked at, is now nine months afterwards, and there is no granulation tissue at that time.

DR. SHURE:  No granulation tissue?

DR. GEDDES:  No.

CHAIRMAN McCAULEY:  Thank you.

We'll now hear from Spiration, Incorporated.

DR. SHEA:  Good afternoon.  My name is Rick Shea.  I am President and CEO of Spiration, Inc.  Spiration was founded in 1999 to develop medical technologies for the treatment of lung diseases.

For this project, Spiration has assembled a scientific advisory board of over 20 experts.  These slides that Lori is showing here depict some of them.  We have arranged for two of these advisors to address the questions posed to the panel.

Dr. Matthew Brenner is Chief of Pulmonary Medicine and a Professor of Medicine at the University of California in Irvine.  Dr. Brenner has published extensively regarding his research on emphysema and the effects of lung volume reduction surgery in various models, including experimental models of emphysema.

Dr. Douglas Wood is Professor and Chief of General Thoracic Surgery at the University of Washington in Seattle.  Dr. Wood has substantial experience in both lung volume reduction surgery and interventional bronchoscopy, and has published extensively on the clinical aspects of lung volume surgery, lung cancer, and the use of tracheobronchial devices to treat airway obstruction.

Dr. Wood participated as a co-principal investigator for the NIH/CMS-sponsored National Emphysema Treatment Trial, as well as the companion study of cost-effectiveness of lung volume reduction surgery, sponsored by the Agency for Health Care Policy and Research.

Spiration has also two physicians on staff who are here today to address the Panel's questions.  Dr. Steven Springmeyer, our Medical Director, is a pulmonologist and interventional bronchoscopist and a Clinical Professor of Medicine at the University of Washington.

Dr. Xavier Gonzalez is our Vice President of Research.  Dr Gonzalez has supervised all the pre-clinical in vivo and ex vivo testing for Spiration's intra-bronchial valve system.

I would now like to introduce Dr. Matthew Brenner.

DR. BRENNER:  I thank you, and good afternoon.  I will begin with a short animation showing the use of the Spiration intra-bronchial valve system, which I'll now refer to as the IVV.  This animation shows delivery, deployment, and removability of the IVV device.

The IVV device is a one-way valve delivered under topical anesthesia with conscious sedation to selected regions of the bronchial tree using a flexible bronchoscope.  It is designed for rapid, simple deployment.

Once the device is in place, it's designed to limit flow of air into the distal diseased portions of the lung.  The valve is designed to allow mucous to exit the targeted area, to minimize the potential obstruction that might lead to pneumonia.  Pneumonia is a concern as a possible complication of airway obstruction with any intra-bronchial device.

The valve can easily be removed via flexible bronchoscopy.  The ease of removability was designed specifically to increase patient safety in the event of complications, incorrect placement, or lack of therapeutic benefit.

My comments will specifically address Question No. 2 regarding the safety parameters to be evaluated in the clinical trials of these devices.  The process of pre-clinical investigations of new devices is important to analyze and predict the safety factors needed to be addressed in clinical studies.

I and the other members of the scientific advisory board of Spiration have recommended a rigorous and complete characterization of the safety profile of the IVV device in pre-clinical studies, including delivery, deployment, and removability, before the company proceeds to clinical trials in these very fragile patients.

Spiration designs studies in healthy, in vivo animal models to provide specific information that will help predict responses that may be observed in human airways with the IVV device.  The company is conducting progressive, acute, and chronic studies.  In this pre-clinical work, Spiration has implanted over 500 devices using three different species for periods of one, three, and six months.  The company has safety data regarding the removal of over 100 devices for periods of up to six months.

In summary, pre-clinical studies remain an important step to establish the safety of devices before moving to clinical trials.  As a result of conducting these pre-clinical studies, the companies identified and analyzed key safety parameters listed in this slide that should be required prior to clinical studies of new intra-bronchial devices in emphysematous patients.

Thank you, and I would like to now introduce Dr. Douglas Wood, who will discuss clinical study design of intra-bronchial devices in emphysema patients.

DR. WOOD:  Thank you, and good afternoon.  In this presentation I am representing the members of the Spiration Scientific Advisory Board.

Under current eligibility criteria, only about 10 to 20 percent of emphysema patients are eligible for surgical treatment.  There are no surgical alternatives for the remaining 80 to 90 percent of emphysema patients who are suffering from a progressive disease with no known cure.

In considering the target population for initial pilot studies, the population may differ based upon differences in device design, risk profile, and physiologic mechanism of action.  For the pilot clinical study of the Spiration intra-bronchial valve, it is the opinion of the Scientific Advisory Board that the initial pilot studies should be confined to patients with hyperinflation, severe heterogeneous, upper lobe emphysema who do not qualify for LVRS or transplantation and who are able to tolerate bronchoscopy.

Patients will be required to complete the program in pulmonary rehabilitation prior to enrollment to match the majority of LVRS studies and to establish a post-rehabilitation baseline to serve as the most legitimate pre-treatment control.

To address Question No. 1, we believe that the appropriate control group would be patients receiving best medical management, including pulmonary rehabilitation.  LVRS would be appropriate as a control group if the device risk profile approaches that of surgery or if device implantation precludes subsequent LVRS.

The results of pilot studies and the published results of the NETT randomized clinical trial will help guide the design of pivotal clinical studies.

Regarding Question 2, the most important aspect of the initial clinical studies for these devices is the need to demonstrate safety.  There is now an extensive body of literature regarding safety of LVRS.  The published data should serve as a guide for evaluating new, minimally-invasive procedures for the palliation of end-stage emphysema.

As scientific advisors, we agree with the adverse outcomes defined in Question 2 and recommend adding the need for device removable and/or surgical intervention as additional adverse outcomes.

With regard to the measures of effectiveness asked in Question 3, we agree with the FDA's proposed measures, including improvement in FEV1, six-minute walk, maximum exercise capacity as measured by cycle ergometry, and quality-of-life and dyspnea scales.

We recommend including measurements of lung volumes, a standard test that provides objective evidence of reduced lung volume.   However, the proposed variable of length of hospital stay is unrelated to device effectiveness; it may be a surrogate for device safety, but is largely influenced by the study protocol.

The degree of clinical improvement achievable with these intra-bronchial devices has not been established and should first be evaluated in a well-controlled pilot clinical study.  It is anticipated that the NETT will provide a consensus of the investigators regarding the degree of improvement or decline that will be considered clinically-significant.  The pilot study data can then be combined with these results to establish reasonable endpoints for a pivotal clinical study.

Relief of dyspnea is the primary clinical motivation for treatment in most patients, and these studies should include subjective measurements of dyspnea and quality of life.

Regarding Question 4, followup is needed to assure acute and chronic safety as well as to evaluate persistence of efficacy.  Duration of followup is very much dependent upon the device design, the risk profile, and the reversibility of the procedure.

For the IVV device that has established a substantial body of data from pre-clinical studies and has the ability to be removed, we think acute and chronic safety can be determined in three months.  Devices or procedures that have not been demonstrated to be reversible may require longer followup before determining safety.

Regarding the followup for effectiveness, we think the benefits of the IVV device may be evaluable within one month of treatment.  This should be confirmed at three months after treatment.  The chronic safety and effectiveness can then be evaluated during post-approval followup that conform to the guidelines for implantable devices.

In conclusion, I hope this has adequately answered the questions submitted to the Panel.  The IVV device and other minimally-invasive technologies may offer improvement of symptoms in patients not eligible for surgery and in the future may provide effective palliation with lower risk than surgery.

We thank you for the opportunity to provide input into your deliberations, and we offer to address any additional questions that you may have.

I also just wanted to introduce two additional members of our Scientific Advisory Board that are with us in the audience.  Dr. Roger Yusen is a pulmonologist and an Assistant Professor of Medicine at Washington University School of Medicine in St. Louis, and Dr. Yusen in the audience has authored several publications on patient selection for LVRS and long-term outcomes after surgery.

Dr. Dan Sterman is the Director of Interventional Pulmonology and an Assistant Professor of Medicine at the University of Pennsylvania.  He has authored a number of publications in the area of lung cancer and interventional pulmonology.

Thank you.

CHAIRMAN McCAULEY:  Are there any questions from the Panel for the members of Spiration, Incorporated?

[No response.]

We will now proceed with the open public comment session.  All persons addressing the Panel speak clearly into the microphone, as the transcriptionist is dependent upon this means of providing an accurate record of this meeting.

We are requesting that all persons making statements during the open public hearing session of the meeting disclose whether they have financial interest in any medical device company.  Before making your presentation to the Panel, in addition to stating your name and affiliation, please state the nature of your financial interest, if any, and disclose if anyone else besides yourselves paid for your transportation or accommodations.

We will begin with those individuals who have notified the FDA of their request to present in this forum.  Dr. Cooper?

DR. COOPER:  Thank you.  The ATS, which I am representing, appreciates the opportunity of being invited here to participate and to be able to answer questions.

I don't really have any specific comments to make other than to say that the organization is very interested in the process, very willing to help in any way, shape, or form with the design, the establishment of standards, and the evaluation of this, and is particularly concerned with the process of introduction of new technologies in a responsible fashion into the community.

We would like to assure that the process assures that, once devices are, in fact, made available, that that introduction of the device is something which is really beyond the FDA I believe, is done in the fashion which patients are fully informed and very aware of the risks and benefits which must be established, I believe, before the FDA finally approves these devices.

So I don't have other specific comments to make other than that we appreciate participating in the process and are happy to avail you of any expert advice that you may wish to seek from the cardiothoracic surgical community.  Thank you.

CHAIRMAN McCAULEY:  Dr. Putnam?

DR. PUTNAM:  Good afternoon.  Thank you, Mr. Chairman.  My name is Dr. Joe B. Putnam, Jr., and my practice is dedicated solely to general thoracic surgery in Houston at the University of Texas M.D. Anderson Cancer Center.

As a surgeon, Professor of Surgery, Deputy Chairman, Program Director, and Clinical Investigator, I have focused on the specific problems of patients with lung cancer and other pulmonary diseases for the past 15 years.  I serve as Chair of the Thoracic Oversight Committee for the American College of Surgeons Oncology Group, which is a National Cancer Institute-funded cooperative group improving the care of cancer patients through prospective, multi-institutional clinical trials, specifically examining the surgical and multidisciplinary treatment of cancer.

I also serve as Chair of the Work Force on Clinical Trials for the Society of Thoracic Surgery.  I have been involved in over 40 clinical trials related to thoracic surgery over the course of my career as primary investigator, co-primary investigator, and collaborator.

Thank you for the opportunity to be here today to address the Panel's posed questions for devices for the treatment of emphysema.  For disclosures, I have received non-restricted grant funding from Radiotherapeutics, recently purchased by Boston Scientific, and from RITA Medical Systems, for the conduct of pre-clinical and clinical trials, respectively.  My transportation has been subsidized by the Society of Thoracic Surgery.

Question 1:  the appropriate control group for clinical evaluation.  Because the NETT trial is the largest trial to date for the assessment and treatment of end-stage lung disease, the design and structure of this trial would be appropriate to follow.  A subset of patients within this trial should have some benefit and should pose the gold standard for which comparisons are made in any new clinical trial.  The results of this NETT trial should be available within the next 30 to 60 days, and the recommendations would be to assess the results of this trial carefully prior to establishing treatment and control groups.

It is imperative that the results of any and all trials, whether positive or negative, Phase 1 or greater, be published in a timely manner.  Trials must be optimally performed on a national level at multiple centers.

Finally, educational programs for devices must be made available to the participants in the trial to ensure consistency of care and application.

Question 2:  clinical devices for safety.  The endpoints presented, rehospitalization, COPD exacerbation, and so forth, are excellent endpoints to monitor.  Again, because the NETT trial is the largest to date, endpoints monitored in this trial must also be considered for future trials if a meaningful comparison is to be done.

Future endpoints must be objectively defined prior to the study origination through peer-reviewed publications or consensus conferences addressing such endpoints.  Common procedure-related toxicities may be developed for use by all clinical investigations, such as has been proposed by the common surgical toxicities through the National Institutes of Health.

Question 3:  the merits of the parameters in evaluating effectiveness of such devices.  These included pulmonary function tests, six-minute walk tests, and so forth.  The clinical significance of percent changes relate to the endpoint in question, so it is a bit premature to discuss what would be appropriate changes in the absence of the context of the trial.

The maximal oxygen consumption, a management of the cardiopulmonary axis, how well the body absorbs, transports, and uses oxygen, as well as the diffusion capacity, the DLCO, are currently easily available and objective measures of pulmonary reserve.

Question No. 4:  the duration of followup.  Short-term followup, less than six months, is usually presented.  However, 12 months, 18 months, and longer really define the issues of long-term safety and efficacy.

As the NETT trial has proposed, a minimum of six months is required.  Many patients in the NETT trial have been followed for five years.  Long-term followup is obviously required for late events.

From the Society of Thoracic Surgeons, I would like to mention to the Panel that our National Cardiac Database, the largest in the world with over 2 million patients entered to date, including over 525 hospitals and 1,500 individual surgeons, has allowed thoracic surgeons to improve their patients' care.

Recently, a post-market evaluation of transmyocardial revascularization techniques, funded by the FDA, in association with the Society of Thoracic Surgeons, was highly successful.  The TMR procedure was tracked in over 2,500 patients over a two-and-a-half year period alone or in conjunction with cardiac bypass.  This study is now closed, and the outcomes are pending.

The Society of Thoracic Surgeons and its Work Force on Clinical Trials welcomes the opportunity to assist our patients and the FDA in its quest for improvements in health care and to balance the considerations of clinical, ethical, and economic outcomes.

Thank you for the opportunity to discuss these issues.

CHAIRMAN McCAULEY:  Is there anyone else wishing to address the Panel?

[No response.]

Thank you.  Since there are no other requests to speak in the open public hearing, we will provide with the general Panel discussion, including discussion of the FDA's questions.

Are there any general comments from the Panel at this time?

[No response.]

We will proceed with the FDA questions.

Question No. 1:  "What is the appropriate control group for clinical evaluation of these devices?" -- has two separate parts to it.  "For which patients would LDRS be an appropriate control group, and for which patients would medical management be an appropriate control group?"

Dr. Ferguson, you have been asked to lead the discussion of this question.

DR. FERGUSON:  First, let me say that the presentations from the companies and the thought that they have put into it, in my view, is right on target in terms of appropriateness and discussions of safety and efficacy.

To answer, however, the question, I think it is much too soon to decide what are the appropriate control groups.  This was brought up by Dr. Wood and some others.

To me, the control groups would, for a long time in the future, have to be those individuals who are not candidates for any other type of surgical procedure, LVRS, transplant, or others that might come on the scene, but those are the ones here.

I think that my answer to this would be that the appropriate control group for a long-term time in the future would be those patients who have received maximum medical benefit under medical therapy.

So I have skirted the first part of that and gone directly to the second.

CHAIRMAN McCAULEY:  Dr. Brenner, any comments?

DR. BRENNER:  From an oncologist's point of view?  No.

CHAIRMAN McCAULEY:  Dr. LoCicero?

DR. LoCICERO:  Pretty soon we're going to have information concerning lung volume reduction from the National Emphysema Treatment Trial which may actually eventually affect what we decide as far as using this as a control group.

However, it's probably not appropriate to use these devices for -- have a control group of lung volume reduction against these devices because these devices are not intended to be a complete treatment as lung volume reduction, as they're not affecting -- most of them, as far as I gather, do not affect all lobes, and only portions of the lung, and may not be equivalent to lung volume reduction surgery.  So, in that case, it would be inadvisable to use lung volume reduction as the control.

Additionally, the death rate or the mortality rate for the operation is measurable and these devices show a much lower mortality rate.  This would set up an unduly-difficult -- or an unduly-easy hurdle for these particular products in comparison to that approach and, again, may be an inappropriate match.

Concerning the second point, best medical therapy versus this type of device might be an appropriate control.  All of these patients will deteriorate at the same rate.  It allows us an opportunity to look at divergence based on their deterioration.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  Well, I think there are several aspects of this, and some of it may overlap with effectiveness.  I'll try not to do that too much.

At least some of these devices are said to be an alternative.  That's how they are proposed in the presentation, as an alternative to LVRS.  I think one has to ask the question of, how are these areas being targeted?  Are they all heterogeneous cases.  Except for Dr. Cooper's, they seem to be.  Well, how do we know that in a segmental obstruction you're getting the areas that you would want to get?  I don't think that was clarified and would be a source of concern.

I think, ideally, NETT would be a control.  NETT is not going to repeated.  One could consider cohort matching of these patients in a large multicenter trial of all the devices.  I think it's important that whatever be done here be multicenter because there will be differences in the patient population and levels of skill.  But I think that would be ideal.

I'm not sure that that's possible, though, because there are going to be more factors about here, and retrospective cohort matching may be difficult.  So I think we are left with a medical control.  We do have good data for what the deterioration is in this patient population.  So that's good.

I think I agree with whoever said that we need optimal medical management for at least six months prior to this.  So I think that would serve as the best control.

I think a number of things need to be put in here that would make it analogous to NETT, if we are looking at them, which is CT grading, excluding pulmonary hypertension, measurement of FEV1, and so forth.

I think in terms of starting any of these studies you need to define crossover to LVRS, when that would be considered, when it would be considered appropriate.  So I think those are the issues, and I think just because of the feasibility, you would need a multicenter, randomized, controlled trial using medical group as a control.

CHAIRMAN McCAULEY:  Dr. Miller?

DR. MILLER:  This area is not one that I am particularly expert in, but from what I understand of this, it would seem that you have not a homogeneous group of patients here.  The patients who are at various stages of disease who require different treatments at whatever stage they are, and I think that the proper control group would be that group that would ordinarily get a treatment for whatever stage they are when you propose to use one of these devices.

So the competing alternative to the device is what they would get at that level of disease.  So I guess the answer to the first question is, it depends on the patient and what their alternatives are and what you are comparing to in terms of the treatment.

CHAIRMAN McCAULEY:  Dr. Blumenstein?

DR. BLUMENSTEIN:  If the results from NETT are to be known in a couple of months from now -- is that what I heard? -- it seems like we are having this meeting two months early, because I would imagine that a great deal will be learned from NETT, probably not the first two-month presentation, but some of the other analyses that would follow the initial analysis and the announcement of the trial, and so forth.

In particular, some of the analyses of subgroups -- I don't mean that in the sense of isolating them, but the relationship between treatment effect, the interactions between treatment effect and group identification.

It seems to me there are analogies here also with advanced stage cancer, for which there's no alternative, and that there may be some of the regulations that govern these cancer trials today, the accelerated approval type of things or some of the single-arm, acceptance of single-arm studies in situations where there are no alternative treatments, some of the models there might apply in this situation.

Again, all of those depend on how well you know your historical data, and the historical data, at least to some extent, may come out of NETT, and it may be worthwhile to wait a bit.

But one thing that occurred to be here also is that, even if you are in a group, if you're studying a group for which lung reduction volume surgery -- did I say that right?  I'm not so facile with this.  But for which that is inappropriate, that a randomization between early use of these devices versus a later use of these devices would gain you something very important, and that is to give you an initial, at least an initial randomized data comparing the immediate effect of the installation of these devices.

I assume that the procedure time is fairly short on these things, I'm gathering from the little animation that I saw.  I've never done one, so I can't tell.

So I would assume that randomization between early versus late might be something that could be done fairly quickly, and you wouldn't be delaying the group randomized to a delay treatment very long, but at least you would, then, develop that randomized data that would allow you to assess in an unbiased way the immediate effect of the device.

CHAIRMAN McCAULEY:  Dr. Stoller?

DR. STOLLER:  Well, I would echo Dr. Ferguson thanks to the presenters.  I think that the issues have been framed very carefully.

My own view is at the end of the day both effectiveness and efficacy of these devices in the minds of patients are going to be framed by the clinical choices that patients have.  And irrespective of the pending results of NETT, I think the appealing patient choice will be that of a non-surgical procedure unassociated, as Dr. LoCicero has pointed, presumably unassociated with the same short-term mortality rate, one would at least hope.  So I think that consideration frames a strong preference for a medically-managed control group.

That said, I think one of the potential appeals of a bronchoscopic approach is, as we have heard with regard to some of the devices, its reversibility.  So the reversibility at least allows the possibility of crossover at a later time, undoing and potentially crossover to a surgical intervention that would not preclude surgical candidates from the medically-managed groups.

So that my own sense of the inclusion criteria would be to be, in fact, broad and to not exclude potentially surgical lung volume reduction candidates who are being medically managed at the time from candidacy in the trials, recognizing that they  may be allocated to a control group.

CHAIRMAN McCAULEY:  Ms. Brown?

DR. BROWN:  I'm sympathetic to the FDA and the companies that are trying to design the studies today, given that the NETT trial results aren't out.  If the NETT trial results are coming out in two months, unfortunately, the FDA in general has 30 days to do almost anything, and they have to make decisions based on the information they have today, rather than two months from now.  So just to let the Panel know that I'm not sure the FDA gets to wait for future scientific data that may be helpful.

In addition, I think that it is important with these studies to carefully think about the options that the patients do have.  The homogeneous patient has a different option than the heterogeneous patient.  With the procedures that are bronchoscopic, they are going to carry more risk than the medically-managed patients, but, on the other hand, the surgical procedures carry a whole lot more risk, but the efficacy is going to fall somewhere probably in between.  So that is going to impact how the FDA thinks about the safety of these devices with respect to both medical management and surgery.

CHAIRMAN McCAULEY:  Ms. Moore?

MS. MOORE:  I have nothing to add to what's been said already, so I pass.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  One other comment that I forgot to make was the procedure that I think involves actually collapsing the lung, washing with surfactant, and so forth, and then the glue.  That's a fairly dramatic procedure for the patient, and I think one might want to consider sham bronchoscopy with a lavage as a control in that study as a possibility, just because they have gone through something similar.  It's analogous to having surgery and feeling you've gone through a lot.

CHAIRMAN McCAULEY:  Dr. Ferguson?

DR. FERGUSON:  I would like to reiterate what Dr. Stoller said.  I think pegging any of these devices to lung volume reduction is the wrong thing to do.  They're all different.  They're a different approach, and I don't, frankly, see how the results of the NETT study, whatever they show -- I think we know what they're going to show -- but whatever they show have much impact on what we're talking about here today.

CHAIRMAN McCAULEY:  Dr. LoCicero?

DR. LoCICERO:  I'll just expand on that a little bit.  To limit the indications to whether or not this works against lung volume reduction ignores the vast majority of patients who don't qualify for lung volume reduction.  This also may affect the company's ability to recruit patients and gather data in any kind of a timely manner.

Dr. Stoller?

DR. STOLLER:  Thank you.  I guess I would like to make one other comment with regard to candidacy in the context of homogeneous versus heterogeneous lung function.  There's a population -- and many who know me will anticipate an interest of mine in my comment -- but there's a population of patients with alpha 1 antitrypsin deficiency who, in fact, have hetero disease, but it's in the wrong direction and not easily amenable to surgical lung volume reduction.  In fact, the relatively embryonic experience with these patients, and not in a controlled way, suggests rather disappointing surgical results.

So it seems to me the opportunity exists, and I would urge careful thought about the potential inclusion of patients with alpha 1 antitrypsin deficiency, in which there is, in fact, heterogeneous lung disease, but it's basally distributed and, therefore, not easily amenable to the conventional surgical approaches that have been undertaken to address them.

So these devices, in a funny way, offer an opportunity to look at a heterogeneous emphysema population that's in a funny way been excluded from consideration and for whom benefit may be conferred.  So that would be something that I would urge attention to as these pivotal trials are being considered.

CHAIRMAN McCAULEY:  Dr. Witten, it appears that the Panel does not feel that lung volume reduction surgery would be an appropriate control group and that the appropriate control group for these devices would be patients that have had maximized medical therapy.  Does that satisfy your --

DR. WITTEN:  Yes.  Thanks for the discussion.

CHAIRMAN McCAULEY:  Thank you.

We will now move to the second question which has been posed to the Panel by the FDA.  "What are the most important safety parameters to be evaluated in clinical trials of these devices?"

Dr. Shure has agreed to lead off this discussion.

DR. SHURE:  I think many of the parameters were very reasonably discussed by the presenters.  I would like to make one comment at the beginning, before going on to detailing those.  That is that I think that what we are dealing with is we are dealing with a chronic illness.  So we have to look at both acute consequences of the intervention, plus chronic consequences, because these are chronically-ill patients.  So we need to look at both short-term and long-term endpoints.

I think these are probably analogous to what we'll talk about under efficacy, but we perhaps could look at composite endpoints or parameters of death; air leak greater than seven days; bleeding requiring transfusion or reoperation or a need for mechanical ventilation; pneumonia, looking at both the acute and chronic in terms of frequency of pneumonic episodes; need for hospitalization; number of hospital days; need for reoperation to remove the device if migrates; respiratory failure as a consequence of the device -- that would be long-term -- decrease in FEV1.  I mentioned reintervention because of migration.

I think people have said that these devices don't migrate, but I think we really need to assess that.  I don't know if they radio-opaque.  I'm assuming they are, but I think they need to be looked at bronchoscopically, probably at intervals of at least three, six, nine, and twelve months.  I would propose something on that order.

Secondarily, perhaps decreasing quality-of-life index or decrease in dyspnea, but I would view the other ones as I mentioned as primary safety issues.

CHAIRMAN McCAULEY:  Any comments, Dr. LoCicero?

DR. LoCICERO:  I agree pretty much with all of what Dr. Shure has said.  I would also add that the device that is formed within the body may pose slightly different safety endpoints, which might include local heating effects, which may be measured by bleeding or some other effects on the bronchus.  This may also pose some long-term effects, which may require us to look at this particular device for a much longer period of time.

CHAIRMAN McCAULEY:  Dr. Brenner?

DR. BRENNER:  I have no comments.

CHAIRMAN McCAULEY:  Dr. Ferguson?

DR. FERGUSON:  I agree with Dr. Shure.  It seems to me that one of the things that is going to become an important consideration here, and I'm not speaking for or against any device, believe me, but it is the fact that, is the device removable?  And all of the things that have been mentioned in terms of the easing, and so on, are going to be important parameters down the road affecting what patients are selected, just any number of things.

CHAIRMAN McCAULEY:  Ms. Moore?

MS. MOORE:  I think Dr. Ferguson just mentioned the point that I was going to mention.  One of the presenters did suggest that we add to this list the ease of administering and removing the device.  I would concur with him on that.

CHAIRMAN McCAULEY:  Ms. Brown?

DR. BROWN:  Just one comment about the serious adverse events:  With respect to patients who rehospitalized after the procedure is done and COPD exacerbation kinds of information, it is my understanding that, since these patients have an ongoing disease state and they're continuously deteriorating, that that's something that we expect to see in this population.  So that when we do a comparison between treatment and control, hopefully, we will be able to sort out what's device-related and what's not.  But we do expect to see those things in these patients.

CHAIRMAN McCAULEY:  Dr. Stoller?

DR. STOLLER:  Again, I agree with Dr. Shure's comments.  My notion is that an evaluation of the complications ought to carefully attend to complications of the presence in the airway of a foreign body, of which these all are.  That would obviously subsume outcomes such as, you know, hemoptysis infection, both acutely and chronically, as well as complications relating to the procedure itself.

Patients with emphysema are characteristically at higher risk, it seems to me, with regard to bronchoscopy.  These procedures are longer than the traditional diagnostic bronchoscopies, I would imagine.  So one would have to capture complications of the bronchoscopic intervention independent of the placement of a device in assessing the complications of the procedure.

Then, of course, I think as has been very nicely stated by all of the presenters and Dr. Shure and other members of the Panel, you know, the intermediate-term complications would surely to capture rehospitalization, intubation, mechanical ventilation both invasively and non-invasively, and death, frankly.  I agree with the notion about the need to remove the device by virtue of complications as a consequence of the placement of a foreign body in the airway.

Clearly, one would want to capture the crossover rate; that is to say, which is ultimately a testimony to clinical efficacy and patient preference.  Patients who go on to have a device and then ultimately find themselves electing based on results of the available clinical trials and the NETT trial, whatever that shows, could be forthcoming, who elect to go on to surgical intervention as a lung volume reduction opportunity.

So I think it has been nicely stated by all:  The assessment of complications needs to be very broad-reaching because this represents an intervention not conventionally available to patients with relatively severe emphysema.

CHAIRMAN McCAULEY:  Dr. Blumenstein?

DR. BLUMENSTEIN:  I don't have anything to add.

CHAIRMAN McCAULEY:  Dr. Miller?

DR. MILLER:  I think in terms of looking at the complication, I think one thing that should be considered is how difficult it is to treat the complication once it occurs, rather than just scoring it occurred.  Because I can certainly imagine pneumonia occurring behind an irremovable plug of glue in the airway is going to be much more difficult than just having a pneumonia behind a removable valve that suddenly you can drain that segment.  So I think some examination of that would be important.

DR. SHURE:  Days of hospitalization or need for surgery could be part of that.

CHAIRMAN McCAULEY:  Dr. Witten, it appears that two points have really come up relative to Question 2:  complications related to the foreign body in the airway in and of itself versus the complications associated with procedure.  These have been secondarily broken down to acute consequences and chronic consequences.  The parameters that have been identified include death, bleeding, mechanical ventilation, pneumonia, hospitalization or rehospitalization, reoperation, respiratory failure, decrease in FEV1, migration of the device, time periods for evaluation being three, six, nine, and twelve months, with secondary endpoints of decrease in dyspnea and improvement in quality of life.

Does that satisfy you?

DR. WITTEN:  Yes.  Thank you.

CHAIRMAN McCAULEY:  We'll now move to the third question.  "What endpoints, among those listed and any others not listed, would be important to demonstrate device effectiveness?  Pulmonary function, six-minute walk test, maximum exercise capacity, quality of life, dyspnea questionnaires, length of hospitalization?"

For this question, Dr. Stoller has agreed to be the lead person.

DR. STOLLER:  Well, thank you.  I appreciate the opportunity to comment.  Again, I think these issues have been very nicely laid out by the discussants.

My own view of this is that the assessment of primary and secondary outcomes really should take its lead from the design of available, in fact, surgical trials, including NETT, and which I think has been nicely discussed.  The spirit in which the outcomes were designed took into consideration that this is really a palliative procedure.  So one needs to capture outcomes that relate to functional capacity and patient's experience of this disease, while, of course, not neglecting the complications, as we've discussed, in long-term survival.

My own sense I think has been discussed really by many others:  that the primary outcome measures must include some measurement of exercise capacity.  We have been asked to comment on the minimally-important clinical difference, and I'll come back to that as I go through the individual outcomes, at least my best assessment of those.  I will try to gauge those comments by what information I'm aware of in the literature that guides some comment in that regard.

But I think that with regard to functional status, the major contenders, if you will, would be maximum exercise capacity is done on a bicycle exercise test, as has been done in NETT, as well as an assessment of the six-minute walk, for which there are, in fact, standards now published and some assessment of minimally-important clinical differences that comes from work by Redelmeier and Guyatt that are included with the recent ATS standards document on six-minute walk.

I think that a disease-specific quality-of-life measure needs to be included, as has been done in several of the studies, the randomized trials that have been discussed; certainly, in the St. Georges enjoys reasonably popularity, is self-administered, and would certainly satisfy such a measurement in my view.

In addition, it enjoys some assessment of the minimally-important clinical difference, that difference being in the literature as four units, although more conservatively other studies have taken a somewhat more aggressive delta, if you will, and have insisted on an eight-unit outcome measure as a clinically-important difference.  So I think there's some room for discussion about that, but I think that that's a reasonable disease-specific, quality-of-life measure.

Others should be collected, I think, but, of course, in the design of pivotal trials and the assessment of power calculations, one has to anchor these on the targeted outcome measures.

Then, finally, I think as has been amply discussed, no matter what the outcome, one needs to have some sense of physiologic plausibility. Physiologic outcome measures are not sufficient, I think as has been pointed out nicely by Dr. Cooper.  Patients don't come to see us with regard to, you know, FEV1 is better and I completely embrace that notion.

By the same token, if there is a benefit, one has to understand what are the physiologic underpinnings by which one could understand that.  So that requires some assessment, in my view, of airflow and lung volumes; in particular, spirometry because it's widely done, highly quality controlled.

My view is that it would need to be done in a maximally-bronchial-dilated state, recognizing the high prevalence of reversible airflow obstruction in patients with COBD, well demonstrated in the IPPB trials, and even in alpha 1 antitrypsin deficiency, one would have to look at post-bronchial dilator FEV1 in a standardized way, as well as a measurement of lung volumes, including residual volumes, as I think is pointed out, would be important.

Now, having said that the primary outcome measures should really be threefold -- exercise capacity in some dimension, quality of life, and physiology -- clearly, the other outcome measures that, again, have been included in many of the available randomized trials, including survival, should absolutely be included.  It would weigh towards the long-as-possible followup, as one may see different kind of temporal trends in mortality-related results, as well as to the natural history of the disease affected by the placement of the device.

I think that going down the actual list and trying to address the minimally-important clinical differences, with regard to pulmonary function, in particular FEV1, there are several choices here.  This has been approached in several ways.

My own view is that one would want to assure a magnitude of difference at least large enough to exclude the reversible component of airflow obstruction that one sees.  By conventional definition, in terms of defining a bronchodilated response in these patients, that number has characteristically been for the FEV1 200 ml and 12 percent, as a delta large enough to say that this difference exceeds what one might expect from the aggressive bronchodilated management of these patients, again recognizing that the prevalence of reversible airflow obstruction is very high.

So I would favor something along those lines in terms of a magnitude of change.  Having said that, the only paper of which I am aware that has addressed this head-on in terms of ascertaining what the minimally-important clinical difference that is recognizable to patients is the work by Redelmeier and Guyatt.  That speaks to about a 3 to 4 percent predicted rise in the FEV1 as the minimally-important difference that could be clinically recognized, subjectively recognized, by patients as they talk to their colleagues in a rehab program.  "Am I better than my colleague in terms of the impact of rehab on me?"

The delta that was required in their study was about 3 to 4 percent, depending on whether you look at the decrement and the increment or the increment alone.

With regard to the six-minute walk, again, the data that really come from the same group defined a delta, and that's been stated in the ATS summary, of about 54 meters, as I recall.  Again, if it's on the clinical improvement side, it's a little higher, about 70 meters of a six-minute walk increment to be the minimally-important clinical difference.  This is an embryonic literature, I would hasten to add, but what little exists I would bring to bear in this discussion.

With regard to the maximum exercise capacity, again, I would take guidance from the kind of dichotomous assessment of outcome that have been used in conventionally-available surgical trials, and in NETT, as I recall, that difference is a 10-watt increase in bicycle testing exercise capacity.  We're not commonly using the shuttle walk test in this country, but that, of course, has been looked at as well.

I don't have strong feelings about SF-36 in terms of the overall health dimension, although I am aware that that has been looked at.  For example, in Dr. Geddes' randomized trial, as I recall, the surgical lung volume reduction group experienced on average a 22-unit rise, which was statistically significantly different.

I'm not aware of literature that addresses the minimally-important clinical difference.  So I really can't comment there.

With regard to length of hospital stay, I would regard that as something important to capture, but it is certainly not something that one would power a trial of such an intervention.

CHAIRMAN McCAULEY:  Ms. Brown?

DR. BROWN:  With respect to the endpoints, I agree with Dr. Stoller with respect to that there are documented and validated quality-of-life measures, the St. Georges questionnaire, and SF-36.  There is literature to describe what's considered clinically-significant.

So I'm not going to talk about exactly what those numbers are, but I agree with the numbers that you have presented.  I believe that is pretty well supported in the literature.

With respect to FEV1, the ATS guidelines discuss two potential standards.  One is for bronchodilator changes.  That's the 12 percent, 200 millimeters.  An alternative is the 15 percent that's discussed with respect to what is considered clinically-significant over a year for normal patients.  So 15 percent is an alternate without the 200-millimeter criteria.

And the six-minute walk, the Redelmeier paper discusses 54 meters, but those are for patients who, on average, at baseline can walk 360 meters.  So I think that the baseline for the patients should be taken into consideration.  That's a 15 percent improvement, and so maybe some consideration should be made for other baseline walking distance for those patients and their severity.

CHAIRMAN McCAULEY:  Ms. Moore?

MS. MOORE:  Nothing to add.

CHAIRMAN McCAULEY:  Dr. Ferguson?

DR. FERGUSON:  Nothing to add.

CHAIRMAN McCAULEY:  Dr. Brenner?

DR. BRENNER:  Nothing to add.

CHAIRMAN McCAULEY:  Dr. LoCicero?

DR. LoCICERO:  One thing that we didn't mention so far is length of hospital stay, and that really is not an indicator at all of this particular device, and probably will be too inaccurate and too little a variation to include.

We really, Dr. Cooper, representing Broncus, really articulated it best, and that is, there really are two endpoints.  One is functional and one is physiologic.  Everyone measures FEV1, and so that makes it a good endpoint.  We do have, as has already been brought out, some reasonable data concerning that.

The other is functional, where there remain some differences of opinion, and it might be something that could be considered a secondary endpoint with two or three different potential variables as a measure.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  I agree with everything that Dr. Stoller said.  I think in terms of the six-minute versus the MVo 2, the MVo 2 is certainly better standardized, which makes it a little better perhaps, but the six-minute walk is something that most centers can do without any trouble.

The one thing that we might consider suggesting as well, though, is something I think was looked at in that, which is a decrease in oxygen need.  That would be a secondary endpoint, reflecting perhaps both an objective endpoint and quality of life.  That could be assessed by a room error Pa02 of greater than 60 with exercise, because that would allow patients not to walk around with oxygen, and it would perhaps make them happier than almost anything.

CHAIRMAN McCAULEY:  Dr. Miller?

DR. MILLER:  I wonder if it would be appropriate to include something, some measure of the pace of the disease, if you have affected the natural history of the disease.  I mean, you can insert one of these devices, and maybe measure one of these things at a point in time, but if the patient deteriorates over two years instead of six months, that may be worth noting.

CHAIRMAN McCAULEY:  Dr. Witten, it appears that --

DR. WITTEN:  You have one more.

CHAIRMAN McCAULEY:  Oh, did I miss one?  I'm sorry.  Dr. Blumenstein, forgive me.

DR. BLUMENSTEIN:  I'm not being nearly as contentious this time as I was in this morning's session.

[Laughter.]

I was going to suggest that the NETT trial, one of the things that struck me, as I was reading the description of it, is that they're employing quality-adjusted life-years.  I think that we'll learn a lot of good reference data and methodological techniques from how they handle the data there in this particular disease setting.

There are also other methodologies, like

QTWIST, and so forth, that may help us understand the tradeoff between quality of life versus shorter life.

CHAIRMAN McCAULEY:  Dr. Witten, it appears that there are parameters or classes of parameters that have been discussed:  those which are clinical parameters of improvement versus physiological parameters of improvement.

All of the parameters listed under Question 3 appear to be appropriate, except for length of hospital stay.  Additional parameters that may be included:  decreased need for supplemental oxygen.

Does that satisfy?

DR. WITTEN:  Thank you.

CHAIRMAN McCAULEY:  We'll now move to Question 4.  "The duration of followup should allow adequate safety and effectiveness assessment of these permanently-implanted devices in a chronically-ill population.  What is the appropriate duration of followup for a pivotal clinical study of these devices?"

Dr. LoCicero has agreed to lead off this question.

DR. LoCICERO:  There are really three different intervals of evaluation required.  One is the immediate evaluation, looking at post- or interprocedure complications.  That would be the day of the procedure, first to evaluate bleeding pneumothorax, et cetera, some of the things that have already been mentioned.

The second is potential post-procedure reactions which may occur within the first few days following the procedure, the implantation of the device.  This typically is the first two to three days, when we might develop atelectasis or other problems, and also potentially-delayed hemorrhage or pneumonia, which might develop at approximately seven to ten days.  So for procedure-related events, we need to look at the day of the procedure two to three days later, seven to ten days later.

In terms of evaluation of effectiveness, the first real time that we can evaluate these devices is, as has been presented today, would be at approximately one month.  The NETT trial has evaluation points at three months, six months, et cetera.

What we see from some of this data is that there is a plateau that occurs at around six months.  Since our control arm is going to deteriorate at the same rate and our patients are going to deteriorate at the same rate, we would probably be able to get a reasonable evaluation within six months.  So that this might be the time when we would have maximum benefit from the device prior to deterioration of the patient.

We also want to be able to give that patient the opportunity to go on to some other form of treatment.  So I would recommend that our endpoint for effectiveness be six months.

CHAIRMAN McCAULEY:  Dr. Brenner?

DR. BRENNER:  No comment.

CHAIRMAN McCAULEY:  Dr. Ferguson?

DR. FERGUSON:  No.

CHAIRMAN McCAULEY:  Ms. Moore?

MS. MOORE:  No.

CHAIRMAN McCAULEY:  Ms. Brown?

DR. BROWN:  I think that the length of followup should be considered separately, depending on the severity of the disease, so patients who have fewer options perhaps would have a shorter length of followup than patients who have more alternatives.

CHAIRMAN McCAULEY:  Dr. Stoller?

DR. STOLLER:  I agree with Dr. LoCicero.  I think that the consideration with regard to the length of followup has to take into account the need to assess both short-term and intermediate-term, if you will, outcomes with regard to assessing effectiveness, but also needs to take into consideration the fact that patients who fail this approach may be candidates for other approaches such as surgical volume reduction.

The length of followup cannot be so long that it actually puts these patients at risk with regard to the natural history of their disease, when they're already starting at a highly-obstructed level for which the five-year -- we heard numbers before.  The five-year survival of patients whose FEV1 is less than a liter, the five-year mortality rate is about 50 percent.

So that trading off the desire to have longer-term outcomes versus the need to capture efficacy, I would agree with what was said.  I think six months is a very reasonable balance between maximizing benefit, but not disallowing patients from crossing over to other interventions if they remain in such a pivotal trial for the full duration of expected outcome.

CHAIRMAN McCAULEY:  Dr. Blumenstein?

DR. BLUMENSTEIN:  I think it would be important to also explore the possibility of use of outcomes such as failure-free outcome -- I mean failure-free survival; that is, the duration of time until an alternative therapy has to be implemented, in which case it may require maybe an active follow-up period through the suggested six months, perhaps with a little less intense followup following that, if that were appropriate.

CHAIRMAN McCAULEY:  Dr. Miller?

DR. MILLER:  I think the efficacy and the safety should probably have different lengths of followup because the efficacy, you want to know that immediately, if it's working.  But if a patient gets stabilized and does well on one, I think you want to continue the safety study for a longer term, because they have this permanent device in their airway and there may be some long-term problems.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  I agree with Dr. Miller's comment.  I would feel more comfortable with a minimum of one-year followup in terms of safety and allowing crossover to LEVRS at six months.  I think there are different issues.

I'm personally concerned with the long-term issues of occult infections in a space that's been deliberately collapsed and having a foreign body in for that period of time, with issues of migration and proximal as well as distal infection and hemorrhage.

CHAIRMAN McCAULEY:  It appears that the two issues of safety and efficacy basically mean a few time intervals.  In terms of safety of these products, that a time period of seven to ten days, which would cover the operative and immediately post-operative and sort of peri-operative period within a hospital or outside of a hospital within seven to ten days would be needed.

In addition, most of the six-month followup which has been recommended I think is based on the NETT trial, where a patient's performance has plateaued out in six months.  However, some of the Panel members are concerned about longer-term followups of I think six to twelve months followup in this patient population is what I'm hearing.

Does that satisfy your answers to that question?

DR. WITTEN:  Yes, thank you.

CHAIRMAN McCAULEY:  Right now we'll take a 15-minute break.

(Whereupon, the foregoing matter went off the record at 3:15 p.m. and went back on the record at 3:34 p.m.)

CHAIRMAN McCAULEY:  This afternoon we will address clinical issues regarding medical devices used to ablate lung tumors.  We are now ready for the FDA presentation.

DR. ASHAR:  In this next session we would like to focus on the topic of lung tumor ablation devices.  Presently, there are a number of ablation technologies on the market that FDA has cleared with a general indication for soft tissue ablation.

For a device to obtain a more specific indication, however, we expect that a clinical study for this new indication be performed to demonstrate device safety and effectiveness.  At this time no ablation device has been cleared for the indication of treatment of lung tumors.

The purpose of this session is to obtain your recommendations regarding the role of thermal ablation devices in the treatment of lung cancers.  We would also like to obtain guidance from the Panel regarding the types of clinical studies that would need to be performed to demonstrate safe and effective use of local ablation technologies for the treatment of lung tumors.

Following the sponsor's presentation, FDA will be asking the Panel to comment on a number of questions pertaining to the current role of these types of local treatments, the appropriate efficacy endpoints, appropriate imaging modalities, the length of followup, and treatment of patients who would otherwise operative candidates using these types of devices.

Thank you.

CHAIRMAN McCAULEY:  Does anyone on the Panel have any questions?

[No response.]

We will now hear from some members of industry in regards to this topic.  The first presenter is RITA Medical Systems, Incorporated.

DR. CHESKIN:  Good afternoon.  My name is Barry Cheskin, and I'm the President and Chief Executive Officer of RITA Medical Systems.  I want to thank you for the opportunity to speak before the Panel today.

Our company, RITA Medical, is a company that is built around minimally-invasive surgical devices for treating patients with various kinds of tumors.  We have the device, as has been indicated, that has been marketed since 1997 for the ablation of soft tissue.  It was cleared by the FDA for the specific indications of unresetable liver lesions, as well as painful bone metastases, and is currently used by physicians around the world for a number of applications today, including lung tumors.  It is estimated that about 40,000 of these cases have been done to date, primarily in the area of treating unresectable liver lesions.

Just to give you a sense, the system consists of a source of energy, a radio frequency generator, as well as a single-use, minimally-invasive, surgical tool which you see pictured there, which is used to actually deliver the radio frequency energy to the tumor or to the site that is to be destroyed.

The device, the needle electrode is inserted directly into the tumor.  A secondary array of wires is deployed from the tip of that needle.  Radio frequency energy is used to energize those electrodes.  Temperature sensors are used in our device to sense the temperature of the tissue as it's being heated up to ensure that it reaches a sufficient temperature to destroy the tissue.  The result, as we have demonstrated, is a very repeatable, spherical area of tissue ablation.

Our purpose in being here today, as we have discussed, is to review the clinical issues related ultimately to marketing clearance of this device for this indication.  RITA has been discussing this indication with FDA for some time now, and we believe there's already strong clinical evidence to support safety and effectiveness of this device in this application.  We hope that today's meeting will lead us to a consensus on a clear path for FDA marketing clearance for these devices in the lung application.

I would like to now turn things over to Dr Steven Rosenberg, who will give a summary of our clinical experience and discuss some of the clinical issues, and, hopefully, answer some of the questions that FDA has raised.

DR. ROSENBERG:  Hi.  I'm Steve Rosenberg.  I'm the Chief of Surgery at the National Cancer Institute, a position I've had for the last 27 years.  I need to emphasize that I'm not here in any official capacity of the government.  I do this as an approved outside activity.  I have no equity in RITA, but I am a consultant to RITA.  This is the only company that I consult for.

I became involved with RITA several years ago because, as a surgeon who cares for cancer patients, I was really interested in finding innovative ways to treat patients that could not be dealt with adequately surgically, and that is patients that had metastatic cancers to the liver that were not surgical candidates.

I became very interested in this technology as a relatively low morbidity method for the ablation of metastatic cancers in the liver, helped RITA develop that application, and it is now virtually standard of care.  It has been used in well over 40,000 patients and represents a very effective modality for the control of liver metastatic and primary liver tumors in patients that are not surgical candidates.

It's with the interest in seeing that this modality can be brought to bear similarly for patients with tumors in the lung, who currently have no other treatment modalities available to them, that I would like to discuss some of the data for the use of this treatment in the lung.  I firmly believe that its potential applications in the lung are the equivalent potentially to those in the liver, and that it's a modality that patients who have no other alternatives deserve to have available to them.

So the targeted patient population here currently have very few options.  We're talking about patients with pulmonary metastases that are less 3.5 sonometers who are non-surgical candidates, generally because of co-morbidities, that have lung-predominant disease and refractory to other treatments.

So I do not think that this kind of treatment should be advocated now for patients who have surgical resections available to them.  This is for those patients who are non-surgical candidates and who do not have effective treatments available.

Similarly, patients with non-small cell lung cancer that are small, again, that are non-surgical candidates on a refractory or treatments or are not candidates for radiation, either because they have already had it or the heart is in the radiation field in patients with cardiac compromise.  So with the modality for the use in patients with cancer that currently do not have other effective options.

Now, as with any local treatment, the primary, and in fact the only, endpoint is the tumor control, the local control of that lesion.  This, of course, is exactly what we believe has been achieved in the liver as well.

We measure tumor control in oncology by either stabilization or decrease in tumor size.  This is a standard criteria that is used to evaluate any chemotherapy regimen and other local therapies such as radiation therapy.  We use the same exact criteria for this application as well.

Now the imaging modality to determine effectiveness is CT imaging, computerized axial tomography, which is recommended as the standard for evaluation of tumors in the lung, but basically tumors in virtually any location for cancer patients.  The standard methods that are used to measure tumor responses, there are two standard methods that are used:  the one now recommended by the World Health Organization for all studies in oncology as a response evaluation criteria in solid tumors or the RECIST criteria, which are guidelines which measure the longest diameter of a lesion and measure either its stability or decrease, based on that longest diameter.

Other criteria use the product, some of the product of perpendicular diameters of lesions, and that's another way of measuring tumor response and local control that basically gives equivalent results.

When this treatment is applied to lesions in the lung, the criteria that one uses to determine whether or not the treatment is effective is basically a one-month response rate, which is typically accepted for patients with chemotherapy.  Most definitions of response in medical oncology insist that a lesion be controlled for at least one month.  Now that's often not a sufficient criteria, and, in fact, the criteria that we're proposing for local tumor control is an evaluation at six months after treatment.

The tumor doubling time for metastatic tumors has been studied in dozens of papers.  For metastatic cancers it's on average a median doubling time anywhere from 40 to 136 days in multiple publications, and for primary tumors it varies in different publications from 67 to 134 days.

So with this doubling time, it's our feeling that for the cancer patient a tumor that either does not grow or has decreased in size over a six-month period represents a very clinically-meaningful response.  Obviously, if a tumor is growing over a six-month period, then that would not represent an effective treatment.  But stability or decrease over a six-month period would certainly be clinically beneficial for most patients, given the doubling time of tumors.

Now the overall experience supporting the use of this ablation technique in patients with lung tumors is actually quite extensive, as you can see.  Because the device is available to surgeons, many have just begun using it to treat lung tumors.  We will particularly talk about a treatment that RITA has had considerable involvement, and that is the treatment of Lencioni with 62 patients.

But, as you can see, there are many other investigators, including several that have published their work and papers that were made available to you, who have used this technique for lung tumor ablation, and I believe shown that it can be an effective treatment for patients with lung tumors that meet these criteria.  That is, non-surgical candidates or tumors that are less than 3.5 centimeters.

I believe that the information that's been obtained not only in our own RITA study, but in other studies that have been reported at meetings and in the literature, show that CT-guided radio frequency ablation is safe and can successfully destroy lung tumors, that there's minimum morbidity -- we'll look at some data in a moment -- that typically resolves within 24 hours, most often does not require any intervention.  In fact, this can often be performed as an outpatient procedure, and that RF ablation may prove to be a viable alternative for patients who cannot be treated by surgery or who are refractory to other forms of treatment.

RITA has sponsored a trial at three non-U.S. sites, and then an independent investigator basically used the exact same protocol for a clinical trial at UCLA.  Again, these were non-surgical candidates.  The endpoint for clinical effectiveness was a six-month tumor response on CT scan.  We'll talk about the results in 62 patients.  They had 79 different applications of 102 lesions, but 49 of the lesions are not evaluable for tumor response at six months.

As you'll see, we believe this to be a safe and effective method for tumor destruction and a kind of modality that patients that have no other alternatives available to them, who have a metastatic cancer or non-resectable primary lung cancers should have available to them.

The general patient characteristics, as you can see, the median age of the patients in this experience of 62 patients that have been carefully followed and carefully evaluated, is 64 years.  You see the number of lesions that were treated.  Again, the diameter of the lesions that we feel should be treated that can be completely ablated is 3.5 sonometers, given the fact that the probes that are available can cause ablatants that are, in fact, substantially larger than 3.5 centimeters.

We used the RECIST tumor criteria.  Again, this is a standard criteria that, as states here, is advocated by the World Health Organization.  A partial response is any response that causes more than a 30 percent decrease.  A stable disease is less than a 20 percent increase.  So we very rigidly used these World Health Organization criteria in evaluating these patients.

Now the results, in comparing these lesions, can be looked at in two ways.  Clearly, the most credible way to look at the lesions is to measure the lesions prior to treatment and then six months later measure the lesions again to see if they have grown.

You can see that 20 percent of the lesions had completely disappeared on x-ray; 22 percent had decreased, had had an objective partial response by classic oncologic RECIST criteria, and that is greater than a 30 percent decrease, and another 24 percent had stable disease; that is, did not grow.

So you can see that the number of patients that could be effectively treated and benefitted using this criteria of comparing a post-treatment lesion to a pre-treatment lesion was 67 percent.  Almost two-thirds of the patients showed clinical benefit.

Now another way to look at the data is to compare the six-month value at one month.  I don't believe this is quite as valid, but I do want to present this information because the lesion that is created by the radio frequency ablation, as you will see in one example, does cause an inflammatory lesion that appears to be maximum at one month.

So if you wanted to compare to the maximum inflammatory size, you could see that at six months -- that is, five months later -- 93 percent of the lesions had either disappeared, decreased, or were the same size.

This is an example of what one might typically see.  You can see a 2-sonometer lesion on the left that received radio frequency ablation.  Again, this can be done under conscious sedation, under CT guidance.  It can be minimum morbidity procedure.

You can see at one month the ablated area that has been created is larger than the original lesion, but you can see, as often happens in the lung, as lesions disappear, the lung collapses around them, the scar decreases, and you can see that lesion has substantially decreased, if not disappeared, by three months later.  So the general criteria we use are comparing to the pre-treatment lesion, not the one-month size, but you can see the reason for analyzing it both ways.

Complication rates from these procedures, again done in a conscious patient, can be performed as an outpatient, are quite minimal.  Pneumothorax, anytime one puts a needle in the chest, it's always an opportunity to obtain a pneumothorax.  There was a 19 percent incidence of pneumothorax that required a chest tube, generally for just one day; very rare pleural effusions.  One patient had pneumonia.  But these were all minimum morbidity complications in this series of the treatment of 79 lesions, and could be easily managed.

So, in summary, I do believe that the clinical evidence in the patients that have been treated thus far does support the safety and effectiveness of the treatment; that non-surgical patients who are refractory to treatment or aren't candidates for radiation could benefit from this, and that physicians could very readily benefit from access to radio frequency ablation tools, when appropriately instructed on the use of this device, just as they now use it as part of the standard care of patients that have cancers that have spread to the liver.  I think cancers that spread to the lung could be equally effectively treated.

So, very specifically, then, the FDA posed five questions that I would like to at least provide our suggested answers to.  The first question dealt with the request to discuss the ablation as a means of local control, to describe the specific patients for which this treatment might be appropriate.

So our feeling is that RFA would be appropriate for patients with primary or metastatic cancer to the lung that's less than 3.5 sonometers, because we can very reliably create an ablated lesion larger than that, but who are not candidates for surgical resection.  Surgery is, obviously, the standard treatment for patients that do have resectable, even metastatic, lung tumors and should be applied when the patient is a candidate for surgery.

The second question dealt with the appropriate endpoint for the study and the appropriate, in my view, as is the case throughout the field of oncology, is that clinical effectiveness should be judged by the stability or decrease in the treated lesions.

The third question dealt with the imaging modalities that would thought to be appropriate.  Of course, throughout oncology for cancers at any location, CAT scans are a standard way to evaluate these lesions, and we do believe that by evaluating the size of lung lesions, which can be easily measured quite accurately by CAT scan in the lung, that pre-treatment measurements and measurements at six months is the appropriate evaluatory mechanism to assess effectiveness of these lesions.

The fourth question dealt with the evaluation occurring at what duration, what duration would be suitable.  We do believe, given the doubling time of metastatic and primary lung tumors, an evaluation at six months is sufficient to assess safety and effectiveness, as we have shown in some of the data that I have presented, since the majority of untreated lesions will progress during this six-month period.

The final question which was posed, and that is the appropriate control group, endpoints, and duration of followup for a clinical study of the device to be used in place of surgical resection.  I would advocate that right now this should not be used in place of surgical resection.  If a patient is a surgical candidate, they should have surgery.  We do not plan to treat patients who are candidates for surgical resection and want to have available the use of this technique for patients who are not surgical candidates and do not have other effective options available to them.

Well, thank you for your very kind attention.

CHAIRMAN McCAULEY:  Are there any questions from the Panel for RITA Medical Systems, Incorporation?

DR. FERGUSON:  I have a question.  You have to forgive my ignorance, Steve, but the effectiveness of this device, as I understand it, is related to the temperature that is developed in the lesion.  My question to you is, given the difference in blood flow surrounding a lesion in the liver and the relative insulation that you have in the lung, have you developed parameters for different temperatures?

DR. ROSENBERG:  It's a good point.  We can measure the temperatures because there are probes that are present, because of temperature measurements that are contained in the probes, but this is the reason that we, in fact, ablate an area beyond the size of the lesion, which is why we would only treat lesions that are 3.5 sonometers.

In the studies that were performed in experimental animals and the clinical data do indicate that you can destroy tissue beyond a 3.5-centimeter diameter.  This, of course, was all done in experimental animals before these clinical efforts were undertaken.

CHAIRMAN McCAULEY:  Dr. Shure?

DR. SHURE:  The increasing size of the area surrounding the lesion at one month, have you looked at what that is?  Is that edema?  Is that necrosis?  What's going on there?

DR. ROSENBERG:  It mainly represents an inflammatory process, because, after all, you are destroying this tissue by heat.  And macrophages, other inflammatory cytokines, come into that location.  So it does represent inflammation, which is why it does resolve so quickly.

CHAIRMAN McCAULEY:  Dr. Miller?

DR. MILLER:  Does it matter where the location of the lesion is?  I mean, can you treat a lesion up against the hilum versus one that's peripheral?

DR. ROSENBERG:  One would have to be careful not to treat lesions that are near major vessels.  Now it turns out vessels are, of course, heat sinks.  So they do protect themselves to some extent.  But, clearly, we would not advocate the use of this for a lesion that was next to a major structure in the medial sternum or in the pulmonary hilum.  Most lung metastases, of course, are not in that location.

CHAIRMAN McCAULEY:  Dr. Blumenstein?

DR. BLUMENSTEIN:  If we were talking about a drug or a biologic agent or something like that, we would be talking in terms of patient benefit, but, yet, you're advocating here that you're just going to measure tumor response as a basis of a marketing approval for this device.  I am wondering why you would change.

DR. ROSENBERG:  Well, this is a local treatment.  Of course, we are talking about patients who had metastatic disease, and I would use the analogy of metastatic disease in the liver.

There's been virtually no study performed with controls that are untreated to show that virtually any surgical resection of a liver metastases prolongs survival.  That study has never been done.

The resection of lung metastases is standard procedure, and I am sure there are many tens of thousands of those procedures that are performed each year standing, that if one has a limited amount of metastatic cancer, taking it out or controlling it is better than leaving it in to grow.  But, again, despite the fact that many tens of thousands of those procedures have been done, there's never been a clinical trial that has compared patients who have otherwise resectable lesions be compared to patients -- that had them resected being compared to patients that don't have them resected.

So there's the general acceptance in the field of oncology that if you can control a tumor, make it shrink or keep it from growing, that you will benefit that patient.  It is exceedingly difficult in patients with metastatic or unresectable cancers to do survival studies which compare to no treatment.

The only patients that we plan to treat here that patients that do not have other options.  So I would say that is fairly standard.  In fact, it is absolutely standard in the field of oncology.  I would say 95 percent of all of the chemotherapy drugs that have been approved have never been shown in randomized trials to prolong survival compared to no treatment.

CHAIRMAN McCAULEY:  I would like to introduce Dr. Solomon.  Dr. Solomon, can you please tell us your specialty, title, institution, and your status with the Panel, prior to your question.

DR. SOLOMON:  My name is Steve Solomon.  I'm a Radiologist at Johns Hopkins, and I am a consultant to the Panel.

The question I just had for you is, have you looked at other imaging endpoints such as contrast enhancement or Houndsfield unit change to suggest that you're successful?

Additionally, another thing that we will talk about later is PET imaging and whether that has played a role in the analysis of success.

DR. ROSENBERG:  Yes, I'm delighted you mentioned that, because I think there are a lot of potential ways to measure effectiveness:  Houndsfield units certainly could be used.  PET scans can be used, although it's hard to quantitate the size of a lesion.

But there's no substitute for measuring the lesion and showing that it's not growing or it goes away, and the CAT scan is certainly the best way to do that.  You can accurately a lesion certainly much more accurately than you can on a PET scan, as you know.

If it goes away or it shrinks, I think that is evidence that will probably surpass many of these other measures, but they could presumably be used in conjunction with it.  I would urge, however, that if this gets approved, that CAT scan measurement of the size -- there's no substitute for seeing a cancer shrink or go away.

CHAIRMAN McCAULEY:  Any other comments from the Panel?

[No response.]

Thank you.

DR. ROSENBERG:  Thank you, Dr. McCauley.

CHAIRMAN McCAULEY:  We will now proceed with the open public comment sessions.  All persons addressing the Panel speak clearly into the microphone as the transcriptionist is dependent upon these means of providing an accurate record of this meeting.

We are requesting that all persons making statements during the open public hearing session of the meeting disclose whether they have financial interest in any medical device company before making your presentation to the Panel.  In addition to stating your name and affiliation, please state the nature of your financial interest, if any, and disclose if anyone besides yourself paid for your transportation or accommodations.

We will begin with those individuals who have notified the FDA of their request to present in the open forum.  The first presenter is Dr. Putnam.

DR. PUTNAM:  Mr. Chairman, thank you for the opportunity to present to the Panel.  I previously introduced myself and stated my affiliations and disclosures beforehand.  Should I repeat them?

CHAIRMAN McCAULEY:  No.

DR. PUTNAM:  Thank you.

I have, just for the purposes of review, I have received non-restricted grant funding from Radiotherapeutics Corporation, which has been purchased recently by Boston Scientific.  I have also received non-restricted grant funding from RITA Medical, also for studies.  These reflect both pre-clinical and clinical studies, respectively.  I have no financial affiliation with them, nor with any competing groups.

The use of radio frequency ablation for primary metastatic lung tumors is investigational, and its use outside of prospective clinical trials is not warranted.  Clinical trials, as the Panel has proposed, should be carefully designed to address the critical issues of patient safety and treatment efficacy with this technology.

Patient followup must be of sufficient duration to determine frequency and time before reoccurrence of subsequent disease, either primary or metastatic.

How do we distinguish between primary carcinoma of the lung and its treatment and patients with secondary tumors or metastasis to the lung?  These clinical trials to develop efficacy and safety, appropriate patient selection, indications and applications should be addressed.

The issues which should be addressed in the development of these protocols should include selection of patients.  I agree with Dr. Rosenberg who, by the way, was my Chief when I was a Surgical Fellow at the National Cancer Institute -- always happy to follow your professor.

The surgical evaluation must be performed by a thoracic surgeon.  Simply to say that a patient is not a surgical candidate is overly simplistic, and there is no single clinical criteria which excludes a patient from surgery.

Combinations of FED1 and DLCO together have shown to be of great significance.   However, no single criteria, nor combination of criteria, have completely identified a population who can be considered non-surgical candidates.

Surgery as the primary therapy should be considered, as Dr. Rosenberg mentioned.  However, radiation therapy, either external beam or three-dimensional conformal therapy, would be considered, either with or without chemotherapy, would be considered appropriate management as a second line in patients who are not deemed otherwise to be surgical candidates, in patients with primary lung cancers.

However, in patients with pulmonary metastatic disease, surgical management alone frequently fails because we are using a mechanical technique, surgical resection, to treat a biological problem.  As a thoracic surgeon who has spent the last 20 years trying to understand the surgical management of pulmonary metastasis, I realize the limitations of such a mechanical approach.

The peripheral location of pulmonary metastasis can be easily resected with thoracoscopic resection or with open thoracotomy resection.  Even so, there is a frequent incidence of patients having recurrence, even with the best technology.

The ability to state that this should only be used in patients who are not surgical candidates suggests that, once approved, any patient or physician who feels that this technology should be applied is certainly eligible to receive this treatment, potentially to the detriment of the patient.

The indications for application and reapplication of this technology must be considered.  Is a single application appropriate or are multiple applications of radio frequency ablation appropriate?  What about the variability and the completeness of the resection?  Even though the technology suggests that the entire area is ablated, is this area ablated completely?

It is easy to show that there is complete ablation which occurs in pulmonary parenchyma.  However, in patients with tumors in the lung, complete ablation may not occur, and there may be inconsistent killing of cancer in this area, suggesting that these patients may have rapid recurrence and that they may require additional management in a compromised pulmonary field.

The second question of the appropriate endpoints would be survival and also local control at two years.  Patients with limited-stage disease, such as stage 1 disease, and stage 1A and stage 1B lung cancer, for which this would be appropriate, patients with tumors 3 centimeters or less in diameter, or stage 1A lung cancer, patients with greater than 3 centimeters, stage 1B, they have a five-year survival when treated surgically of about 60 percent.

So a six-month followup would be too short.  Most recurrences, both local or metastatic, would occur within the first two years.  I would suggest that two years would be the minimum endpoint for both survival and for local control.

The required followup would also be important to identify secondary procedures that may be required for treatment, such as radiation therapy or other chemotherapy.

The injury to adjacent structures, either bronchi or vessels, certainly exists, as well as the potential exacerbation of the malignancy.  In my conduct of a clinical trial examining the role of radio frequency ablation in a study where patients who would otherwise undergo complete resection of their metastasis were initially treated with radio frequency ablation and then resected.  We found that the completeness of ablation is related to size, location, and consistency of the primary tumor.

In contrast to the mucous-producing colorectal carcinoma and hepatic metastasis, lung tumors, both primary and secondary, can have a wide range of tissue consistency, which may prevent accurate placement of the ablation catheter needle array.

A second major limitation which should be addressed is that micrometasasis to local and regional lymph nodes will remain untreated, and this micrometasasis cannot be diagnosed with certainty using conventional radiologic or radiopharmaceutical diagnostic procedures, and have variable location involving the bronchi, arteries, and pulmonary vasculature.

The limitations require sufficiently long-term followup to determine the efficacy of the procedure versus standard therapies.

CHAIRMAN McCAULEY:  Dr. Putnam, could you summarize your comments, please?

DR. PUTNAM:  I can.

The imaging modalities of CAT scan and PET scan are certainly appropriate.  Pathologic biopsy, however, may be necessary to confirm adequacy of treatment.

The appropriate control group is unclear at this time, given that many patients, that most patients can be treated with appropriate primary therapeutic modalities.

Thank you.

CHAIRMAN McCAULEY:  Thank you.

Are there any questions for Dr. Putnam?

[No response.]

Our next presenter is Dr. Sewell.

DR. SEWELL:  All right, good afternoon.  My name is Patrick Sewell.  I'm an Assistant Professor of Radiology and Surgery at the University of Mississippi Medical Center in Jackson, Mississippi.

I am here as an independent proponent of radio frequency ablation of lung tumor, and my travel is not compensated industry and I have no ownership/stock in any of the radio frequency companies nor the imaging companies that I'm going to discuss.

I have devoted the last four years to image-guided, minimally-invasive surgical thermal ablations of a variety of tumors in a variety of organs.  The encompass clinical partners that are equally involved from neurosurgery, orthopedics, GYN oncology, medical oncology, cardiothoracic surgery, and abdominal surgery or hepatic surgery.

They all involve some form of image guidance, ultrasound, CT, MRI, and the vast majority of them are done in the operating room.

Approximately 300 patients a year are admitted to or come to my University for these treatments.  I have eight ongoing IRB-approved, monitored clinical trials for these therapies in humans and some in animals as well, to investigate the uses of these technologies.

I don't know if I mentioned they involve laser cryo high-frequency ultrasound, focused ultrasound, and, of course, radio frequency ablation.  So what I am going to, of course, talk to you about today is the RFA.

I agree with many of the comments made previously by both Dr. Rosenberg and Dr. Putnam.  The clinical trial that Dr. Rosenberg described is essentially what I embarked upon about three-and-a-half years ago.  I began to treat primary and secondary malignancies with CT-guided radio frequency ablation.  I have to stress that the image guidance is probably one of the most important -- is of equal importance as the ablation technology because it is what allows the precision or the focused delivery of this destructive process.  There have been some clinical trials, as Dr. Putnam alluded to, that did not use CT guidance, and those, of course, are not as successful.

Dr. Rosenberg's description of primary cancer, non-surgical candidates are exactly the candidates treated by myself at the University.  I have a team of cardiothoracic surgery, pulmonology, and medical oncology that evaluate the patients if they are primary or non-small cell lung cancer, and if they are deemed unresectable by the cardiothoracic surgeon, then they enroll in this therapy.  The vast majority of them have exhausted chemotherapy and radiation or they have declined radiation and chemotherapy.

So I want to show you a couple of slides to illustrate the potential.  This is a non-small cell lung cancer pre-RFA, and this is the scar 36 months later.  Now I have approximately 16 patients that are a year-and-a-half out that are stage 1 or 2, and the data, survival data, on that group is the same as surgical, although it is only one year.  But keep in mind these are pulmonary cripples who are the sickest of the sick.  So I think that illustrates significant potential.

However, overall, what I wanted to do with this investigation over the last three-and-a-half years is to clearly define the safety because the efficacies can be very complicated, but the safety is pretty clear.  What are the complications?

When I looked at the data pooled on 180-something patients over that time period, in the peri-operative period, and I'm not talking about -- I'm talking about days in the hospital and shortly thereafter -- the major complications are COPD exacerbation, 2.2 percent with some of those very severe in these very sick patients, and intraparenchymal bleeding, 0.4 percent.

Now there are some minor complications such as pneumothorax.  When you treat a single lesion, it can be less than 5 percent.  For multiple lesions, obviously, every time you poke the pleura, the risk goes up, and it can be 50 percent or higher when you start going above five lesions in one treatment session.

Pleural effusion, which is more of a nuisance, except in the patients with significant COPD and already on oxygen, atelectasis, and loss of any functional lung can be a significant detriment.  Then an acceptable infection rate, and what one would expect in colonized pulmonary tissues with necrotic tissue.

I believe that the main benefit, the clearest benefit for the application of this technology is for pain management or pleural-based tumors, albeit primary or secondary.

The application of RF energy is well-documented.  It destroys tissue effectively and precisely.  The ability to target and enhancing primary pleural-based tumor growing into the chest wall and resulting in pain, and end up with a necrotic area that's slightly larger, and immediate pain relief upon the release the following day is a significant advance to patients when their options are surgical resection, which would be pretty much a nightmare in this patient; radiation therapy, which this patient had failed.

In addition, the thermal ablative processes are not cell-type-specific, unlike radiation or chemo.  It doesn't have to radiation-sensitive.  If you heat cell tumor or not to 60 degrees Celsius, it incurs reparable damage and dies within 24 to 48 hours.  It is easily done with local anesthesia, often on an outpatient basis.

Additionally, the complications of treating the painful pleural-based tumors are extremely low.  Pneumothorax is less than 1 percent.  It's almost zero because the vast majority of the time you can get to the tumor without traversing aerated lung.  Bleeding is very low.  Pleural fusion is not even significant, and abscess or pneumonia is, again, a very small percentage.

Now keep in mind I preselected very sick patients who are pretty knocked down from chemotherapy.  So their infection rate is probably going to be higher than if you did this to just a normal person and created a necrotic area.

CHAIRMAN McCAULEY:  Dr. Sewell, we're going to have to ask you to summarize your comments.

DR. SEWELL:  Okay.  The results are immediate and sustained pain relief and preservation of lung function by targeting and destroying a viable, enlarging lesion that would otherwise consume functioning lung.

So, in summary, I think RF ablation of lung tissue is an accepted treatment for soft tissue tumors with proven results, and it has shown very low morbidity and mortality, and it does preserve lung function and reduces pain in inoperable pulmonary malignancies, which results in a significant quality-of-life maintenance or improvement.

Thank you.

CHAIRMAN McCAULEY:  Thank you.

Are there any questions for Dr. Sewell?

[No response.]

Is there anyone else wishing to address the Panel?  Yes?

DR. COOPER:  I'm Joel Cooper.  I spoke this morning, representing, again, the ATS, and I have no financial interest or conflict or association in this situation.

We're all aware of how very vulnerable end-stage cancer patients are to any proposed new or innovation, how often they have flocked to an unsubstantiated treatment, and how high the standard must be if we are going to propose anything for this unfortunate group of patients.

Dr. Rosenberg, in his presentation, mentioned individuals who are not surgical candidates but another time he mentioned those who have no other option.  I just wanted to emphasize that there is a standard of care for solitary pulmonary lesions in patients who are not operable, and that is radiation therapy, with or without chemotherapy.  I would presume, therefore, that his comments related to "no other options" is in fact the operative one.

We're also all aware of the fact that off-label use of FDA-approved devices, often treatments promoted on the Internet by physicians, by companies, often patient-driven, sometimes physician-driven, unfiltered information which attracts patients to a particularly unproven therapy is, unfortunately, something which our Association views with great alarm, and it just puts an added burden, we believe, on the Panel and on the FDA to assure that these types of innovations for end-stage disease be viewed in a very scientific fashion and, in fact, be approved only when there is solid weight of evidence, because of the potential for the misuse.

I would echo Dr. Putnam, and I would say that he is more of an expert than I am because he has more personal experience than I have, but I would echo his comments, the importance of exhausting all other modalities, of making certain that traditional therapies are exhausted, and making certain that the patient has been offered those, and certainly in a trial there should be standards whereby the patient is assessed by a competent, experienced thoracic surgeon.

Dr. Sewell's comments in one of the particular cases was that surgery would be a nightmare for that particular patient.  I'm not sure that I would agree, but I don't know the details of that patient that he referred.  He showed a case of right upper lobe cancer.  I don't know whether that patient was a candidate resection or not, but I would argue that this be very carefully evaluated when pivotal studies are done, to make certain that safeguards for the patients are in place, that in fact there is no alternative therapy during the period of time that the patient is going to be considered for this particular therapy.

We would all be delighted to have an additional modality to treat patients with cancer for whom we now have no alternative therapies, but I would only urge caution in making certain that, indeed, high standards are maintained.

Thank you again for the opportunity of having our Association have input into your process.

CHAIRMAN McCAULEY:  Thank you.

Dr. Rosenberg?

DR. ROSENBERG:  You know surgeons love to operate, and we have at least three very distinguished surgeons on this Panel.  We're not advocating, and I would not advocate, the use of this modality now in place of surgery for patients who are surgical candidates.

Now I don't, frankly, see the logical consistency in some of the arguments that were made.  Dr. Putnam is an expert; I have enormous respect for him.  In fact, I author, as some of you know, a 3,000-page textbook of oncologic treatment which is the most widely-used text in the world.  In fact, I asked Dr. Putnam to actually write the chapter, which he did for the last several editions, on the treatment of lung metastases.

In fact, in that chapter he, as would any thoracic surgeon, advocates vigorously, and has advocated in that chapter, the resection of pulmonary nodules in patients with metastatic disease, because there are patients that have cancers that are metastatic to the lung that have gone on to survive 25 years.  It is hard to imagine that they would survive 25 years if you allowed the cancer to keep growing.

It is an absolute standard of treatment to resect pulmonary metastatic nodules.  If a patient is not a surgical candidate, radiation therapy is virtually never used for metastatic cancer to the lung.  It is used for primary lung cancers.  If a patient is not a surgical candidate, they have no other options.

Radio frequency ablation has been shown to be able to completely ablate those lesions.  So a surgical resection in a surgical candidate is reasonable.  You would certainly think RFA, RF ablation, would be reasonable in that same situation when patients are not surgical candidates.

These are the exact arguments that we heard before the approval of this modality for the treatment of metastatic cancer to the liver.  We know that patients that have metastatic cancer to the liver from colorectal cancer that have less than three lesions have a 25 percent five-year survival.  In the absence of that surgery, that survival is probably less than 5 percent.

So the control of isolated metastatic lesions, although never, to my knowledge, in oncology ever, not a single trial has been conducted in randomizing patients to that treatment versus no surgery have ever been conducted, it's absolutely standard of care.

I think for thoracic surgeons who are concerned that this would be used in place of surgical resection, I agree completely it should not be, and I agree with Dr. Putnam and I agree with the other comments that were made.  That is not what is being proposed.

What is being proposed is very simple, and it seems to me extraordinarily logical.  If surgical resection helps a patient that is a surgical candidate and you have an identical situation in a patient who is not a surgical candidate, we have a modality that can ablate that lesion.  This should not be considered in place of surgery; it should be considered as yet another tool.  It's just a scalpel that can be applied with minimum or as an outpatient and should not be considered as a replacement for surgery.

CHAIRMAN McCAULEY:  Thank you.

We will move now to the discussion of the FDA questions.  Do we have the --

DR. WITTEN:  Can I just make one comment as we move to the questions?  I just want to remind everybody just to focus the discussion on the questions as they would relate to any of these devices for ablation of lung tumor.  Although I think it is extremely helpful that a sponsor of a device presented their particular data, we're not here to discuss that particular dataset, but these questions as they relate to level of evidence for devices for ablation of lung tumors.

CHAIRMAN McCAULEY:  The first question.  I am going to ask Dr. Brenner to read the first question and start off the discussion.

DR. BRENNER:  "Please discuss where you believe the coagulation necrosis is a means of local control may fit into the current treatments for lung cancer.  Local control may mean reduction, stabilization, or a decrease in the growth of tumor size.  Please describe specific patients for which this treatment may be appropriate in terms of extent of disease, size of tumors, and evidence of failure of treatments."

I think that the discussion that we heard points out a number of major issues that the agency is going to have difficulty with.  They are essentially the heterogeneity of this disease that you are trying to treat, lung tumors or lung cancer.

Now, first of all, let's break this down a bit because I think from a regulatory point of view that is going to be important.  We are talking about primary tumors of the lung; in other words, tumors that are the result of growing out of epithelial sources usually, but not always, of the lung, bronchio epithelial.

That's a primary tumor, and there are multiple histologies.  Some histologies are generally more aggressive; some are less aggressive.  For example, a squamous cell tumor can have a very long natural history of five years.  So you can ablate a squamous cell primary or not ablate a squamous cell primary; survival is the same.

So, therefore, there's a concern right away about the concept of using decreasing size as a primary endpoint simply because of differences in natural history.  So there can be different histologies.

Now there's one histology that was not mentioned, and I need to mention it.  That's small cell lung cancer.  Small cell lung cancer is considered a metastatic systemic disease at one set and is, therefore, not considered for local ablation, local control, because of the systemic nature of the disease and the rapid mortality.

However, there are occasional instances where small cell tumors locally can be very destructive and, in fact, sometimes local and even systemic treatments might not be useful, particularly in the recurrent stage, and I will get to that in a few minutes.  So there is histologic variability.

The second issue is there is stage.  Stages are meaningful because they predict prognosis, but they also predict whether the disease is systemic or not systemic.  If the goal was local control and one has systemic diseases, the rate-limiting step, I'm not convinced that trying to control a local disease, a systemic disease with a local modality is going to be useful in the long run, and, in fact, most of the advocates here agree with that, since they focused more on the local disease.

However, there is likely to be a role for this modality in control symptoms of quality-of-life indications that were not mentioned here, except very briefly.  I'm a little surprised about that because I think that that would be a group of patients who might really benefit.

The question is, what would be the standard by which one would utilize to approve on the basis of a palliative indication or quality-of-life indication, rather than on the basis of a size or a complete response or an efficacy indication, a survival indication.

So, with those kinds of comments in hand, I'm going to break down the stages and the types and give you at least my own view in terms of addressing this question as to who would fit, so to speak.

Let's talk first about the curative intent, so to speak, or the complete ablation.  That would be generally what was described by the speakers as the 3.5-centimeter primary stage 1, non-small cell lung cancer.  The intent would be essentially a complete response or a complete ablation of the tumor.

I personally do not accept a partial response as a surrogate for survival unless there is a type of extenuating circumstance such as a prolonged time to recurrence, and we will get into that much later.  I'm poaching a little bit on two, but I think that it has to be in terms of describing who the cohort would be.  So that would be one group that has already been discussed.

I think that that is a legitimate group.  If there's no surgical availability, if there's no surgical option, I agree with that.  Surgery should be the primary modality, I agree with that.

In terms of stage 2, that's a real tough one.  The reason is that some stage 2's are resectable.  My own interpretation of the literature for stage 2 is that, since they generally have lymph nodes in the hilar regions, that for the most part, although some are resectable, 85 percent of the lymph nodes means that there is systemic dissemination.

Sometimes survival related to local controls is more related to histology and biology than it really is to the ablational procedure.  But one could argue that a stage 2, potentially a 2A, might also fit in a resectability group, which means that if they do not have a surgical option, they could also be eligible for this and might be a cohort.

Stage 3's, the primary care, primary ablational indication; these are generally almost all individuals who have disease that will be systemically disseminated on diagnosis.  However, the use of cytotoxic and radiation approaches, either in combination or in sequence, seem in a few of these groups to have ablated the tumor systemically.  And if there are many times a central recurrence, there could conceivably also be a local ablational indication for this.  I heard nothing of this here, although I could conceive of a trial that could actually be designed around this.

Stage 4 are generally systemic tumors and, therefore, for the purposes of ablational approaches for survival endpoints, I do not feel that this would be cohort that would be useful for regulatory purposes.  However, such patients do have palliative needs, and if there are things like chest wall tumors and pain from primaries, this cohort could certainly be used, but only on a palliative indication.

Then there is a cohort of patients who are either recurrent or obstructive, who either have received maximal radiation and do not have that available to them, who might be candidates for palliative maneuvers to either reduce the size of the tumor, to release the obstruction, or to have pain control as an endpoint.  That was alluded to by our final speaker, and I think that is a potential cohort.

So we break all these different cohorts.  You can see there are multiple small cohorts, and each one of these, if one is going to want to get an indication, we have to design a trial for that endpoint.  So to restrict it to just a CR-type endpoint or some type of reduction, tumor size endpoint, unresectable tumors, to my mind, are resectable tumors in a small size primary that is not surgically eligible seems a somewhat restrictive approach to the regulation, to the potential proof of whether or not this concept would be useful.

In terms of the concept of control of metastatic tumors, in other words, cancers that are metastatic to the lung that are not primaries from the lung, I have major concerns with that as an indication.  First of all, these, too, are a whole variety of different primaries.  One could have a renal cell primary that is metastatic to the lung that might have three or four mets, many times a surgical resection.  If they are not resected, some of these patients have prolonged survivals.  As Dr. Rosenberg notes, there is no data to show whether or not resection really is efficacious or not.

On the other hand, some other tumors may be colon cancers might grow, but usually the survival limitations in many of these patients are not the lung primaries themselves, but, rather, they are metastatic sites different, different sites other than the lung.  The same with breast primaries, for example.

So to say to lump all metastatic tumors together and argue that one could definitively or provide some type of evidence that there is real efficacy in terms of survival endpoints on the basis of reduction or ablation of metastatic lesions from a heterogeneous group of tumors without defining exactly what the primary source of this tumor, problematic, and I personally would be uncomfortable with that as an indication.

CHAIRMAN McCAULEY:  Dr. Brenner, we appreciate the complexity of your comments.   However, I'm not sure whether or not you defined the extent of disease in which this therapy may be utilized.

DR. BRENNER:  Stage.  In other words, stage.  Stage 1 would be a local regional tumor that meets the size definition.  Stage 2 would be a tumor that would be larger but would have a lymph node in the hilum.  Stage 3 would be one that is invasive in the chest wall or obstructive.  Stage 4 would be metastatic.

In terms of size indications, I think that is related as much to the placement of the needles, and I think that the sponsors define that as a 3.5-centimeter lesion, but not necessarily because other groups, and in fact the literature that was provided to us did not restrict the size for palliative indications.  So that is why there is complexity here.

There's multiple potential regulatory endpoints.

CHAIRMAN McCAULEY:  Are there any other comments from Panel members?

DR. LoCICERO:  Just one.  We're confusing here patients who are operable and those who are not operable.  Those who are operable fit in point No. 5, which we have not gotten to yet.  We're talking strictly about inoperable patients here.

CHAIRMAN McCAULEY:  Other comments?

DR. MILLER:  I guess I just to be clear that, when we say the patient is inoperable, it's because of medical contraindications?  It is not because of the nature of their tumor?  Am I correct in assuming that?

DR. BRENNER:  Stage issues.  In other words, generally, stage 1 would be operable tumors for the most part; in other words, on the size and location basis.

DR. SHURE:  Which would be operable versus resectable differences.

DR. BRENNER:  Operable/resectable.

DR. SHURE:  Could I just add that stage 2 disease might be one, if you're adding it to this, since there is nodal involvement, it would need another modality.  It couldn't be a sonomodality.

DR. BRENNER:  Yes, for the most part.

CHAIRMAN McCAULEY:  Other comments from Panel members?

Dr. Witten, it appears that the answer to Question No. 1 in terms of extent of disease would be basically patients who are inoperable with lesions less than 3.5 sonometers.  Is that the consensus of the Panel?

DR. BRENNER:  No.

CHAIRMAN McCAULEY:  Stage 1?

DR. BRENNER:  Oh, stage 1, okay.  But there are other -- to me, that's not the -- the question is broader than that.  I mean it seems to me you're asking a much broader question than that.

It says, in terms of the size of tumors and which tumors might be useful from a regulatory purpose.  For example, a palliative indication.

DR. WITTEN:  I think it is going to be hard to summarize Dr. Brenner's answer, but I think it's very helpful to us in terms of the Panel having answered our question.

DR. LoCICERO:  One point, and that is that stage 1 cancers are less than 3 centimeters in size.  T1 lesions are less than 3 centimeters.  So although you're talking about 3.5, that gets you in a stage 2.

DR. SOLOMON:  I guess the other comment is I'm not sure -- you seem to be very negative on the metastatic opportunity, and I think that if you look at what is current surgical treatment for metastatic, certain metastatic lesions are surgically operated on now.  If it's a patient who is unable to be operated on because of other co-morbidities, and where you would ordinarily decide that surgery is an option, I think that some of these ablations may be considered an option, too.

DR. BRENNER:  It gets to really what the standard of efficacy is here.  Dr. Rosenberg strongly argues that removal of lesions is efficacious in the long run, and that might be true, I believe, for some malignancies, but it might not be true for other malignancies.  How does one distinguish that in a regulatory issue?  That, to me, is the question.

In other words, if one is attempting to define efficacy as a regulatory endpoint and one is utilizing a very heterogeneous lesion -- in other words, metastatic from what?  How about a sarcoma versus a renal cell carcinoma?  Any metastatic lesion?

Are we allowed to ask people to respond, particularly Dr. Rosenberg?  Is he going to be given an opportunity to respond?

CHAIRMAN McCAULEY:  Dr. Witten?

DR. ROSENBERG:  Am I allowed to comment?  Am I allowed to comment?

DR. WITTEN:  It's up to the Panel, the Chair.

CHAIRMAN McCAULEY:  Yes.

DR. ROSENBERG:  I would make just two simple points.  One is we need to distinguish non-operable from a non-surgical candidate, because a patient might well have an operable lesion, but if they're not a surgical candidate because of co-morbidities, then, of course, they don't have an option available for therapy.

Now this question that you raise, Dr. Brenner, is one that comes up frequently in terms of the resection of metastatic disease to the lung.  But I do not think you could find one out of a hundred oncologists, be they medical or surgical, who would advocate that patients with solitary or two lung metastases who have refractory to chemotherapy treatments should not have those lesions resected.  It is absolute standard of care in this country, and any oncologic textbook would advocate the removal of those lesions, as Dr. Putnam has and as I'm sure an thoracic surgeon on this Panel would advocate.

Now there haven't been randomized trials to demonstrate its benefit, but the first time, the first patient actually that ever had resection of a lung metastasis with renal cell cancer went on to survive 30 years.  Patients with sarcomas and pulmonary metastases that are less than 3 have five-year survivals and ten-year survivals in excess of 20 percent.  It's absolute standard of care.

If we start questioning now whether or not it's reasonable to try to treat isolated metastatic disease to the liver or to the lung, we're flying in the face of what 99-plus percent of oncologists would advocate as standard of care in this country in the year 2003.

DR. BRENNER:  It has to be specified that it is a, quote, "isolated" metas.  In other words, one or two.  Generally, I think that there would be a fair amount of resistance if you go above two as a standard of care, and particularly for a regulatory purpose.

DR. ROSENBERG:  You know, again, it does vary a little bit for histology, but, in general, you're right.  I would think generally often in the liver one calls four as a cutoff.  Most thoracic surgeons might consider three or four, but I'm sure Dr. Putnam has probably resected twenty lesions in the lung, because very often you get into the lung and you find more than you thought were there.

Those patients have, with a definite incidence -- and I'm talking now about hundreds of papers, maybe 500 publications, that would show those patients have in a small incidence very long-term survivals.

CHAIRMAN McCAULEY:  Any comments from the Panel?

DR. MILLER:  Just one more question for you.  Where does radiotherapy fall into priority of options here?  As you mentioned, surgery, and if not a surgical candidate, then this is an option, but how about radiotherapy?

DR. BRENNER:  I think that really gets into the design of a pivotal trial, because that would be the alternate arm.

DR. SEWELL:  I would like to reiterate and comment that the issue of which metastatic lesion to treat, what type of met, is complicated.  The only clear indication, as I indicated on my last slide, would be pain reduction or pain elimination for maintenance or improvement of quality of life.

In the middle of that, and on the middle ground is an asymptomatic single met and two patients, one that can tolerate surgical resection, and that's what they're offered and that's what they get.  The other has co-morbidity, congestive heart failure, pulmonary disease from whatever cause.  Should we deny them radio frequency ablation of that metastasis, knowing that it's their only documented metastasis, knowing that radio frequency ablation is very extremely reliable and documented to destroy focal areas of tissue with the hope that both of those patients, that is their only metastatic disease, which we have no control over, but what do we do in that case?

And then the third point, primary disease, no one tonight has advocated treating any primary tumor that can be resected, as determined by the resection experts, cardiothoracic surgery.  That is universally agreed upon, and I think it will be for 20 more years, until we get long-term data as to what this does.

But what about those patients like the one I showed that lived three years after I destroyed his four lesions on his right side?  Ultimately, he died of congestive -- tumor-free, after he had failed chemotherapy and refused radiation, and at the age of 88.  I think that's a valid use of this technology, and I don't think it should be restricted.  Certainly it's subject to abuse, but can we really control that?

CHAIRMAN McCAULEY:  Dr. Witten, do you think this has been adequate discussion, so that we can move on to the second question?

DR. WITTEN:  Yes, thank you.

CHAIRMAN McCAULEY:  Thank you.

The second question, please.  I'm going to ask Dr. Shure to read the second question and lead off with her comments.

DR. SHURE:  The second questions is, "Please discuss what you believe to be the appropriate endpoints for study to demonstrate the effectiveness of ablation devices for the treatment of primary and secondary lung malignancies; i.e., increase in survival, reduction in tumor size, improvement in symptoms, improvement in quality of life."

I would like to combine -- some of this has been covered already in the previous discussion, but I would like to add one thing in here, which is safety as well as effectiveness, because I don't think that's covered throughout our discussions.

And I have some concerns with the safety of this technique in the lung, despite the study that was presented.  I think there are some things, based on what I've seen of lung ablation in my time at the University of Mississippi.

Patrick, I know that you presented your selected series, but I think it may have left out some of the patients that I have seen.  You did report the hemorrhage in an article, but I think we have to consider hemorrhage requiring intubation or transfusion.  We need to report death because there was that death and there have been other deaths.

Prolonged mechanical ventilation from respiratory failure in this group of patients is an important complication that has occurred and needs to be reported.  I think we need a large series to look at that.

I was aware of four patients while I was at the University of Mississippi who had significant  complications, including prolonged mechanical ventilation, infection, and sepsis.  I think we need to look at this within 30 days of the procedure, not just acutely, which may account for the differences that we see here.

So I think hemorrhage, mechanical ventilation, pneumothorax where they early degraded in seven days, death, and infection are important.

These patients, as has been pointed out, who are candidates for this have very poor respiratory status and very low FEV1s.  The claim has been made that this is safer than radiation.  I think we have to establish that.  We say that radiation decreases the FEV1.  I think we need to do a prospective analysis of these patients as we treat them that does show that they don't drop their FEV1 in response to this treatment, and that needs to be both short-term and long-term.

One might expect a short-term reduction, as we have seen with surgery, but they shouldn't lose function over a six-month to one-year period, we would hope, other than their natural disease, of course, but that can be looked at in a controlled population.  So I think those are important things related to safety.

In terms of effectiveness, I agree with what was said.  I think tumor size alone is not a significant parameter to look at.  I think it has to be looked at as part of the endpoint, but unless it is offering the patient some benefit, either in terms of quality of life, palliation of symptoms, or improvement in survival, it is not a meaningful parameter to follow.  It needs to be followed, but survival needs to be looked at, as it is in most studies of chemotherapy and radiation.

I think local recurrence is an important thing to look at as well.  I think those would be primary.

I think secondary parameters of the efficacy would be quality of life and, as I mentioned, FEV1 could be included in there.  So that would be my impressions of what needs to be looked at.

CHAIRMAN McCAULEY:  Any other comments from Panel members?  Yes?

DR. BRENNER:  I would like to amplify a number of things and also to make a few extra comments.

I concur with Dr. Putnam's points that an appropriate endpoint here is a survival surrogate.  In my opinion, a survival surrogate for a lung primary is a complete response, not a partial response.  Partial responses have been notoriously unuseful and unpredictive of survival as the endpoint.

So if one is choosing, I would concur a complete response.  This, however, is an interesting modality because the preliminary data that I have seen appear to create a conundrum where there is no complete response, but there's actually ablation of the primary.  Therefore, one sees a lesion that doesn't progress.

That is why a time period, time to failure, which Dr. Blumenstein mentioned initially, would also be considered a reasonable surrogate for a non-small cell lung cancer with an expected time of progressive.  Time to failure is the con-oncologic term.

In fact, this modality -- I think we should, as new targeted therapies are coming about, consider that as a potential surrogate as well, assuming that there is no other metastatic sites that go along, that come along during that period of time.

What is an appropriate time to failure in most lung cohorts?  It's six months.  So that a time-to-failure endpoint I think would actually be a valid endpoint, even if there is not a complete response because of the preliminary data that we're seeing.

So I think that, from a regulatory point of view, we have to understand that non-small cell lung cancers progress very rapidly, that time to failure in stage 2s and higher are common, and even stage 1s will usually fail within a year.  So that would be a design issue that I will get into later.

So, in fact, even though there are generally rules and cytotox ODAC has generally held survival endpoint, not a time-to-failure endpoint, for common solid tumors, with a modality like this, with a different potential mechanism of action, I think we have to be open-minded on this issue.

Quality-of-life surrogates I agree upon, but I do have some additional endpoints that I think are worthwhile that oncologists have used in the past for indications, and they are pain, pain endpoints.  And I alluded to the palliative issues in patients who have, for example, large chest masses that are in pain.

Most of these patients have pain, and we have used pill counts, total dose of narcotics over daily times as reasonable surrogates.  I think reduction usually at least of 50 percent were sometimes regulatory; it has been 100 percent, in other words, a complete cessation of narcotics for a period of time, usually a month, as a quality-of-life surrogate that you can really hit the nail on.

Another potential quality-of-life surrogate, particularly in these folks who might have obstruction or fevers, would be a loss of fever.  Then the real difficult one, but one that many have found to be reproducible, even though it's been difficult to bind, is the Karnofsky performance status.  As controversial as that is, it's surprising how it holds up in just about every analysis.

What we would be looking for is an improvement in the Karnofsky performance test that is real, not a 10 percent improvement where you power your study up with a thousand patients so that you prove a 10 percent improvement in Karnofsky performance test, but were somebody has a very low Karnofsky performance test, bed-to-chair existence, and is able to then get up to an 80 or 90 percent performance test, although I don't know that that has been used in a regulatory situation.

People have used the SF-36 for a partial regulatory indication as well.  So that has already been mentioned.

And then pulmonary function apropos to the previous discussion, to me that's a black box, but it could be defined by our experts on the Panel as a potential surrogate, particularly saturation in the palliative situation.

CHAIRMAN McCAULEY:  Other comments from the Panel?

[No response.]

Has there been enough discussion, Dr. Witten, to address the Question No. 2?

DR. WITTEN:  Yes, thank you.

CHAIRMAN McCAULEY:  We'll move on to Question 3.  Dr. Solomon, I'm going to ask you to read Question 3 and lead the discussion.

DR. SOLOMON:  "Please discuss the imaging modalities that you believe would be appropriate to assess tumor size during these studies."

I think if we look at the possible imaging modalities, you can divide them into sort of anatomic and physiologic.  I think anatomic measurements would be size, as Dr. Rosenberg described.

However, one of the problems that comes up with ablation technologies is that after a procedure, immediately afterwards, the lesion appearance could be larger than when -- you saw some of these in the initial images that Dr. Rosenberg showed us, that at one month the lesion looks actually bigger than it was initially.  So size can be somewhat confusing at different times.  So you probably need to follow that over time.

I think that if we look at the physiologic aspects of imaging, and that would include Houndsfield unit change and contrast enhancement, as well as glucose metabolism, as in PET imaging, you actually are going to get a much more realistic opportunity to evaluate these lesions.

So that I think on a practical sense, most people have access to CT scanning, and I think that's probably the way things should go.  We should use CT to measure size.

I would add using the Houndsfield units.  I would also add using contrast enhancement because I think in previous experience with other parts of the body, if there is an area of enhancement, it is suggestive of recurrence.  Then if you do see this area of enhancement, perhaps a biopsy would be a good endpoint to say, hey, this is actually an area of residual recurrent tumor.

PET imaging has shown to be very valuable in looking at post-ablation imaging, and recent literature in The Journal of Clinical Oncology suggested that it actually was very helpful in looking at success of treatment of liver lesions.  So I would suspect that in the lung we would appear to have similar results.  So I would advocate considering PET imaging with glucose metabolism, FDG tracer, as a second alternative to CAT scan.

So, again, to summarize, CAT scan, measuring size over time, but also use contrast enhancement; also use Houndsfield units.  Additionally, I think that PET imaging would be a valuable tool and using biopsy as a tool to confirm areas of increased enhancement or perhaps increased FDG uptake.

CHAIRMAN McCAULEY:  Any comments from the Panel?

Dr. Witten, does Dr. Solomon's comments -- I'm sorry, Dr. Brenner?

DR. BRENNER:  I wanted to state my concern about using PET as a regulatory tool to suggest that the mass that you're seeing is not viable unless there are data compelling to demonstrate that that is, indeed, the case.  In other words, linked biopsy data.

Since I don't know the literature in that field, I would like to bring that up and ask whether there is compelling information that supports that position because that would be an important regulatory point, I think.

DR. SOLOMON:  I guess I would say that I don't think the data is clearly -- it hasn't been shown at all in the lung.  There are a few case reports of using PET as a tool for looking at post-ablation in the lung.  There is definitely good data coming out as far as liver ablation, but I think that perhaps it would be a tool that could be used in conjunction with biopsy to confirm areas of residual.

DR. BRENNER:  Then my advise to FDA would be that, if PET is to be used, it has to be accompanied by a biopsy in order to verify that there is, indeed, stasis; in other words, ablation, and that that is not viable.

CHAIRMAN McCAULEY:  Comments?  Dr. Stoller?

DR. STOLLER:  I would just make the perhaps overarching comment that the challenge in this conversation is, as I think has been pointed out, the incredible heterogeneity of populations for whom this might be made available.  It seems to me that, perhaps in contrast to the earlier discussion, which had its own subtleties about different patient populations, this particular conversation is complicated by the fact that, if there were three or four patient populations -- heterogeneous, homogeneous -- in the prior discussion there are perhaps dozens, if not hundreds, of patient populations for whom this issue would be framed.

That said, I think the comments about primary outcome measures have to be focused completely on the patient population and focused question for which the issue is being channeled.  So, in my own mind, perhaps the simplest, and in the context of comments that have been made about trying to avoid vigilantly making this modality available as a primary curative -- as a surrogate for a primary curative modality when there is ample experience in, for example, bronchogenic carcinoma of resection or even radiation therapy, avoiding any opportunities, I think Dr. Cooper had pointed out, to allocate patients to this treatment when there are conventionally-available modalities is absolutely improper.  I agree with that comment completely.

That said, on the other end of the spectrum there are patients who have metastatic bronchogenic to the lung who have either a high degree of symptoms, whether it be pain related to some pleural disease or perhaps even hemophilus, et cetera, for whom a very clear palliative indication might exist that goes beyond radiation therapy or even local bronchoscopic opportunities.

So, in that context, the outcomes, to my mind, would have nothing to do with, in fact, size or PET FDG agreement.  It would have to do with control of the symptom that was the indication for the procedure.

So I'm aware that I may not be helping to drill down on the agency's questions about this, but I would submit that the complexity of answering this question has to do with the incredible heterogeneity of the patient populations for whom this modality might be made available.

To my way of thinking, the simplest answer would be confine it to a group in which a palliative indication is offered, which I think is largely the sense of comments made by many of the discussants before.

In that regard, one might focus perhaps less on the physiologic function or even survival, because survival is an unlikely outcome benefit in that particular context, but more around symptomatic control.  I personally think that perhaps functional status is unlikely to be dramatically improved, and I would discount the contribution of any pulmonary physiologic measurement, as I think that would be highly unlikely and difficult to understand in terms of its biologic contribution, even if demonstrated.

DR. SHURE:   I meant a lack of decrease, you know, not hurting them by the procedure.

DR. STOLLER:  I completely agree with that comment.  My comments are not directed at safety, and I agree with Dr. Shure's comments completely about concerns, you know, the temptation to enter this vicious cycle of technology about tumor growth, repeat invasion, repeat pneumothorax.  I mean this is a cascade on which these patients sometimes find themselves.

I think one of the important cautions, I think echoing comments by other discussants, is that the question ought to be vigorously framed to avoid a cycle of technologic interventions to these patients without demonstrable benefit.

CHAIRMAN McCAULEY:  I think many of those issues were addressed in the safety and efficacy question.  Right now we are focusing primarily on tumor imaging.

Dr. Witten, do Dr. Solomon's and Dr. Brenner's comments satisfy you at this point?

DR. WITTEN:  Yes, thank you.

CHAIRMAN McCAULEY:  We'll move on to the fourth question.  Dr. Brenner, can you take the lead and read the fourth question?

DR. BRENNER:  Now with all this heterogeneity, let me see if I can break it down and clean it up.

There are a number of proposed designs that might work here.  Let's take the group that was primarily discussed by our initial presenters.

They would be the group of stage 1, resectable but inoperable, non-small cell lung cancer.  Standard of care in those folks would be considered radiation alone.  So that a study design would be randomization.

Before I go forward, I would argue that any pivotal trial design, in my opinion, would need to be randomized and randomly allocated, the reason being, otherwise, there would induce selection bias into the design study, and Dr. Blumenstein could probably help with that some.

But for the cohort that was described by the industrial sponsors and two representatives, the group I believe they are really targeting are the resectable, inoperable group with medical contraindications.  The standard of care I would think would be radiation.

Therefore, an appropriate design would be radiation versus radiation plus heat.  Plus, this modality, I don't know how you could not provide radiation in either arm.  So one would have to evaluate the contribution of the heat.  That's problematic in terms of the local control of the lesion, because one might argue that the radiation might be successful in such a large number, particularly in stage 1, that the burden of proof would require a large cohort in order to get the power to deal with complete ablation.

I don't know what other panelists would think of a radiation versus radio frequency ablation alone as a standard design, with radiation available for failures, but I think that one has to realize that that -- and it was already mentioned by the Panel members this is not an entirely risk-free procedure with some early reports of some 70 percent pneumothoraces with chest tubes, some bleeding episodes.

So I agree.  I concur.  That kind of design would also allow good toxicity profile.

CHAIRMAN McCAULEY:  Dr. Shure, do you have any comments, as this partially relates to the question that you outlined and answered previously?

DR. SHURE:  Duration of followup with regard to safety and efficacy, well, I think with regard to safety, I think one has to look at least a 30-day followup.  I don't think -- it has to be looked at as an operative procedure and look at 30-day complications at least.

In terms of the FEV1, which does concern me that this hasn't been looked at in this population, I think that should be at least a three-to-six-month followup.

I think in terms of survival, and if we are talking about treating with stage 1 disease who are otherwise inoperable, but would be resectable, I think you need a minimum of one-year followup.  I think in the oncology world, it would be more like a three-to-five-year followup.

Given that these patients have bad lung disease, I think one should require in a controlled study that the original cohorts be followed until death.

CHAIRMAN McCAULEY:  Thank you.

Dr. Witten, has that been enough discussion to satisfy answers to Question 4?

DR. WITTEN:  Well, I think Dr. Brenner has something.

CHAIRMAN McCAULEY:  Dr. Brenner?

DR. BRENNER:  Yes, I was going to argue for a two-year endpoint, per Dr. Putnam's comments.  I agree that two years is appropriate for this cohort as a followup, as a survival endpoint, with a secondary endpoint of followup to death.

I think there are other designs that would be useful from a regulatory point of view that would be a palliative design, where one would use recurrent patients who have lesions that are symptomic in prior radiation fields who would benefit from a palliative endpoint, and that would be randomization between this modality and best supportive care.

I think there would be a large population for such a design.  Not only that, but I think that the risk/benefit ratio would be perhaps even more favorable because people, because of their distress, would be willing to accept more risk compared to, say, the stage 1 group.  So that is another design that I think is reasonable.

CHAIRMAN McCAULEY:  Are there other comments from Panel members?

DR. SOLOMON:  I mean this is a therapy that is a local therapy.  So I guess we have to keep in mind that what you're expecting out of this therapy is going to be a local response.  Therefore, sometimes you might consider looking at the imaging findings of the local therapy or the biopsy findings of that local therapy to prove that it actually worked the way you expected.  That is local control.

Obviously, the endpoint of why we are doctors and why we treat patients is for other issues, such as survival and quality of life, and things like that.  Obviously, quality of life will be another endpoint that is shorter-term.

I guess the question of whether or not this works for survival-length time periods, you are going to be looking at very long time periods.  Maybe the comparison should be made with surgical resection of mestastic lesions and see how it compares against that.  Maybe that would make it a more reasonable time course because a two-to-five-year time course may be too impractical.

DR. BRENNER:  A palliative endpoint would simply be pain; for example, number of pills.  That might be just as useful.  You might not even need, for example, a size.  As a matter of fact, a size estimator would be of less interest to me than a quality-of-life endpoint, because we're not treating, in essence, size; we're treating symptoms to a palliative endpoint.  So I would find that a standard that I could live on the basis of benefit, benefit here being quality-of-life benefit rather than survival benefit.

So I pose two proposed designs.  One is an efficacy survival endpoint for the primary -- for the resectable, inoperable group, but also then a palliative design which would not have a survival endpoint, but a quality-of-life endpoint, which I would find to be of benefit.

CHAIRMAN McCAULEY:  I think these were issues that Dr. Shure brought out in Question No. 2 as primary endpoints in evaluating efficacy of this particular treatment.

Has this discussion been beneficial for you, Dr. Witten, to move on to Question No. 5?

DR. WITTEN:  Yes, thank you.

CHAIRMAN McCAULEY:  Question No. 5, I'm going to ask Dr. LoCicero to read that question and lead the discussion.

DR. LoCICERO:  The answer's longer than the question.  "Please describe the appropriate control group endpoints and duration of followup for a clinical study of an ablative device that is to be used in place of surgical resection of lung tumors for patients who would be considered operable."

I think we have had a huge discussion already concerning this issue, that operable patients should receive surgical resection, and that at this time this device is not to be used in those patients.

CHAIRMAN McCAULEY:  Any other comments from Panel members?

DR. WITTEN:  Well, I am just wondering, as I had mentioned, that we're not talking just about the particular device of this particular sponsor.  So I suppose one question would be, does your answer mean that you don't consider that an ablation study for patients in place of an operative would ever be appropriate and, if so, why not?  Or if that's not the answer, then a discussion of the question, which, like I say, may not apply to this particular device, but ablation devices in general would still be helpful.

DR. LoCICERO:  There are a variety of ablation devices, photodynamic therapy, laser, ultrasound, radio frequency, focused IMRT, et cetera, all potential ablative devices for operable patients, but operation or surgical resection remains the standard.  No one here has questioned that, even those who have used other devices.

The only study that I'm aware of that looked at ablative devices was using photodynamic therapy in Japan, where they gave the photodynamic therapy followed by resection.  In that case, the tumor was removed, appropriate operation was performed, and then it was looked at under the microscope to see if there were any remaining viable cells.  That would be the only situation where this could be used or applied, as far as I'm concerned.

DR. WITTEN:  So what you're saying is that a first stage would be to look and see if there are viable cells at resection?  Suppose the sponsor performed that; they did that stuff, and they found that there weren't viable cells.  Then would there be a next stage, and what would that trial look like?

DR. LoCICERO:  I think that would not be a primary trial for approval of such a device, and that would have to be a preliminary step on the way to much further studies.  But as far as expanding indications at this point, this would really be not an indication.  Nobody at this point considers ablative therapy to replace surgery in operable patients.

DR. SHURE:  Can I make one quick comment to echo what Dr. LoCicero said.  I think one way to look at this is that there have been numerous studies that have shown that removing a nodule, removing a wedge, doing a wedge resection, or doing a segmentectomy do not have as good a survival as lobectomy.

All this can be, at best, is removing a nodule.  That's all.  It's local.  It's local, very, very local therapy.  So even at its best, it couldn't replace a lobectomy as we see the disease now.

CHAIRMAN McCAULEY:  Other comments from Panel members?

[No response.]

Dr. Witten, does that satisfy your answer to Question 5?

DR. WITTEN:  Yes.

CHAIRMAN McCAULEY:  The FDA would like to thank all of the Panel members for giving their time and efforts for the discussion of this product.

This meeting is now adjourned.

(Whereupon, the proceedings were concluded at 5:18 p.m.)