13_030922_CLN_01.doc

TO: DISTRIBUTION

FROM: C NEUMEYER

SUBJECT: TF JOINT INSTRUMENTATION

The fiber optic signal measurement system available from FISO Technologies, Inc., is recommended for use on the NSTX TF joint for temperature, strain, and displacement measurements which will be used to confirm the analysis and guide in the operation of the NSTX TF system.

The basis for selecting the FISO system is as follows:

- The fiber optic measurement technique is immune to magnetic fields and can be used in a high voltage environment.

- The FISO temperature probe can be supplied in a custom form (sheathless) such that the time lag of the measurement is minimized to less than 100mS, which is important for NSTX due to the short pulse length (600mS flat top time).

- Since the FISO signal conditioning system supports a variety of transducers (temperature, strain, displacement) it can satisfy all of the NSTX measurement requirements using the same modules. This avoids the use of multiple types of signal conditioners, and offers the flexibility that different transducer types can be plugged into the signal conditioning system depending on the type of testing underway.

- The FISO system is compatible with the NSTX pulse waveform since it samples at a fast rate (1kHz) and accepts an external trigger.

- The FISO system output is an analog voltage which can be easily interfaced to a digitizer and ported into the NSTX MDS data acquisition system.

- The signal conditioner module set-up and trigger is accomplished via an RS-232 port which can be interfaced an RS-232 driver module and the NSTX EPICS system.

- ORNL has used the FISO system for the SNS project and has commented favorably on their experience with same.

The following measurement points are planned:

Temperature

There are four unique locations on NSTX for which temperature measurements close to the flag joint are desirable, namely the inner and outer layer flags on top of the machine and the same on the bottom. Each of these locations is unique in terms of flag cooldown via conduction and convection. And of course the inner and outer layer flags are loaded differently in terms of current density and pressure across the joint. Temperature will be FISO type FOT-C. They will be installed in grooves on the flag edge closest to the midplane of NSTX, prior to application of flag insulation and potting in the flag boxes. Although detailed dimensional information still needs to be obtained from the manufacturer, it is anticipated that a 1/64” groove, 0.4” long, beginning 1/8” away from the joint, followed by a 1/16” groove extending to the end of the flag, will be suitable for these probes.  FISO will recommend a suitable epoxy adhesive. See figures 1 and 2.

Figure 1 – Edge View of Grooves in Flag to Accommodate Fiber Optic Probes

Figure 2 – Cross Section View of Grooves In Flag

(note that centers of 1/16” and 1/64” grooves are aligned)

Strain

It is desirable to measure radial strain near the flag edge closest to the midplane of NSTX, and axial (z direction) strain at the shear shoe. To confirm the analysis, both inner and outer flags need to be measured. For consistency with the temperature measurements these strain measurements should be made on the same flags as the temperature measurements, namely the inner and outer layer flags on top of the machine and the same on the bottom, total four flags.

The radial strain on the flag needs to be measured on both sides of the flag to properly capture out-of-plane effects. The FISO type FOS gauge shall be used. The strain range shall be 0 to 2500 strain (allowing up to 30ksi in copper with modulus of 18e6psi). The same groove style as the temperature probes can be used, except displaced 0.125” from the sides of the flag to capture the out-of-plane effects (figures 1 and 2).

The same probe type and groove style can be used in a vertical arrangement on the shear shoe as shown in figure 3. This will allow up to 50ksi in stainless steel with modulus of 28e6psi. The fiber optic cable can be routed out along the water cooling line for mechanical support.

Figure 3 – Grooves for Probe in Shear Shoe

Displacement

The FISO displacement transducer accommodates a 20mm (0.787”) linear stroke. One transducer will be positioned on the top of the machine to measure the relative axial displacement of the spline w.r.t. the umbrella cover, which is calculated to be 0.35” maximum.

Four additional displacement transducers will be use to measure the angular twist of the hub assembly and a flag end w.r.t. the umbrella, on the top and on the bottom of the machine. If the transducers are positioned on the umbrella at r=40” and actuated from extension arms mounted off the rotating elements, then allowing half of the stroke in either direction the measurable angular twist will be of order 0.55 degrees.

Measurements are summarized in the following table.

| Measurement | Location | Quantity | Range |  |
| --- | --- | --- | --- | --- |
| Temperature | Flag Edge | 4 x 1 = 4 | 0-150C |  |
| Strain | Flag Edge | 4 x 2 = 8 | 0-2500 Strain |  |
| Strain | Shear Shoe | 4 x 1 = 4 | 0-2500 Strain |  |
| Displacement | Spline | 1 | 0-20mm |  |
| Displacement | Hub | 2 x 1 = 2 | 0-20mm |  |
| Displacement | Flag | 2 x 1 = 2 | 0-20mm |  |

The following equipment will be required.

| Quantity | Part # | Description | Unit Cost | Total Cost |
| --- | --- | --- | --- | --- |
| 1 | BUS-CHA | 19” rackmount chassis for BUS system channel modules | 3300 | 3300 |
| 8 | BUS-MOD | BUS system channel module, 1kHz, +/-10V BNC output | 2750 | 22000 |
| 4 | FOT-C | Custom fiber optic temperature probe 0-150C | 500 | 2000 |
| 12 | FOT-S | Fiber optic strain gauge, temperature compensated, 0-2500 microStrain | 220 | 2640 |
| 5 | FOD | Fiber optic displacement transducer, 20mm linear stroke | 485 | 2425 |
| 16 | CFO-PO3-025 | 25 m fiber optic extension cable, ST connectors | 100 | 1600 |
| 1 | BUS-CAS | Carrying Case for BUS System | 270 | 270 |
|  |  |  |  | 34235 |

The above 25m extension lengths are based on input from F. Jones as to the distance required to connect to the instrumentation racks.

Cc:

J Chrzanowski 	J Dong 	T Egebo	P Heitzenroeder 	M Kalish

R Marsala 		G Oliaro 	M Ono		H Schneider		A Von Halle

M Williams		I Zatz