MET SERVICES

File STN/649/64

SYNOPTIC OR CLIMATE STATION INSPECTION REPORT

Name of Station:       Gobabeb	                  Station No:   68106

Observer’s name:  Calvin Sisamu		  Person interviewed: Calvin Sisamu

Site and enclosure (any changes in site, exposure, layout of instruments, general condition or alterations since last visit):

First visit, the instruments enclosure is situated a few meters in a met square on the north of Research center.

RAINGAUGES

MANUAL GAUGE

GAUGE A                                                                                                    GAUGE B

| Type | 5inch standard |
| --- | --- |
| General condition | Good |
| State of Rim | good |
| Is Rim Level? | Yes |

| Type | Plastic Rainguage |
| --- | --- |
| General condition | Good |
| State of Rim | Fine |
| Is Rim Level? | Not |

Diameter Test                                                                                        Diameter Test

| GAUGE | REPLACEMENT |
| --- | --- |
| 1. 128.67 | 1. |
| 2. 127.93 | 2. |
| Mean: 128.3 | Mean: |

| GAUGE | REPLACEMENT |
| --- | --- |
| 1. 126.42mm | 1. |
| 2. 126.71mm | 2. |
| Mean: 126.57mm | Mean: |
| Inner can tested for leaks? | Leaking | Not Leaking x |
| Outer can tested for leaks? |  | x |
| Funnel tested for leaks? |  | x |

| Inner can tested for leaks? Yes | Leaking | Not Leaking x |
| --- | --- | --- |
| Outer can tested for leaks? Yes |  | x |
| Funnel tested for leaks? Yes |  | x |

(For standard 5” gauge tolerances allowed are: mean of any two diameters +/- 0.010” or +/-0.3mm. For any single diameter +/- 0.020” or +/- 0.5mm)

Gauge Replaced:		No		Replacement Type: _no_____________________________

Measuring Glass Condition:  20mm broken		Replaced: 	no

Maintenance/Repairs

Leveling of the plastic gauge was not done but the observers were to look into the gauge in a way to level it.

## AUTOMATIC RAINGAUGE

| Type | Casella - London |
| --- | --- |
| Serial No | 04340 |
| General condition | Not working |
| State of Rim | Good |
| Is Rim Level? | Yes |
| Height of rim above ground | 130cm |

Is gauge calibrated correctly? (Provide details of measured quantities and attach copy chart used in calibration)

Calibration is not done because the automatic rain gauge is not working.

Maintenance/Repairs carried out:

None.

SCREENS

| Type | standard |
| --- | --- |
| Condition of Screen | The screen is old and may need replacement |
| Condition of Stand | good |

Repairs carried out and action taken:

None.

THERMOMETERS

| Thermometer | Serial No. /Make | Check Therm | Station Therm | Check Therm | Station Therm | Check Therm | Station Therm | Mean difference | Serial no. of replacement |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Inspector’s Therm | GTH 175 PT | 22.8 |  | 22.9 |  | 23.0 |  |  |  |
| Maximum | Lambrecht 2793/9 |  | 25.0 |  | 25.0 |  | 25.0 | -2.1 | No replacement |
| Minimum | Lambrecht 5287/04 |  | 22.8 |  | 22.8 |  | 23.0 | +0.03 | No replacement |
| Dry Bulb | C.W. Price 1607/90 |  | 22.8 |  | 22.8 |  | 23.0 | +0.03 | No replacement |
| Wet Bulb | Lambrecht 5908/74 |  | 22.5 |  | 22.9 |  | 22.9 | +0.13 | No replacement |
| Grass Minimum |  |  |  |  |  |  |  |  |  |

If the inspector’s thermometer is a Digitron 3204SR & PRT enter pre-visit and post visit checks below:        Mean Diff.

| Inspector’s thermometer |  | Date 29/09/06 | Dtron | Insp. | Dtron 22.9 | Insp. 23.1 | Dtron 22.95 | Insp. 22.3 | +0.28 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Inspector’s thermometer |  | Date 16/10/06 | Dtron 24.1 | Insp. 24.1 | Dtron 24.0 | Insp. 24.0 | Dtron 23.9 | Insp. 23.9 | +-0.0 |
| Inspector’s thermometer calibration due date | 27/02/07 |  |  |  |

Condition of Thermometers, wick and reservoir:

With the exception of the maximum thermometer which showed a high a margin of temperature difference from the traveling standard, all other thermometers are in good working condition. The wet bulb wick and reservoir are kept clean.

Condition of Earth Thermometers, Tubes and Plot:

All earth thermometers are available except the 100cm. The 120cm is broken and may be replaced.

Action taken and work carried out:

None.

BAROMETERS

| TYPE | SERIAL NUMBER | CONDITION |
| --- | --- | --- |
| No Borometer |  |  |

Comparisons made with inspector’s Barometer?

Correction cards held on Station?

Are Correction Cards correct for type of observation prepared by Station?

Action taken:

NON AUTOGRAPHIC INSTRUMENTS

### SUNSHINE RECORDER

Type:	Campbell Stokes			Serial No. not available

Condition of Recorder and Mast and any action taken:

The sunshine recorder is mounted on a steel base 127cm above the ground and installed in the met square. The instrument is working and out of obstruction.

### CUP COUNTER ANEMOMETER

Type: Cassela			Serial No. Not available

Condition of Anemometer and Mast and any action (note results of spin test):

This instrument is in good condition.

### WIND VANE

Type:	This is the Pressure Plate and wind vane           Orientation: magnetic north

Action taken:

No action was taken.

### EVAPORIMETER

Type: Class A evaporation Pan				Height above ground level: 18cm

Dipstick Number and Condition of instrument: A graduated ruler is being used. Advised to use standard mesh wire to cover the pan.

Action Taken, give details of measurements of any two random 0.1mm graduations on dipstick,(use calipers):

CLOUD SEARCHLIGHT:  This instrument is not available.

Type:					Maker &

No.

Location of instrument:

Condition of instrument and any action taken:

Additional instruments:

Dew collector. This is manufactured in Israel.

AUTOGRAPHIC INSTRUMENTS

THERMOGRAPH :

Type:							Serial Number:

Condition of instrument and maintenance carried out:

HYGROGRAPH

Type:							Serial Number:

Condition of instrument and maintenance carried out:

### THERMOHYGROGRAPH

Type: Weekly				 Serial Number:

Condition of instrument and maintenance carried out:

One of the pens is broken hence the instrument needs to be replaced.

### BAROGRAPH

Type:	Weekly						 Serial Number: 3.0812.10.000

Condition of instrument and maintenance carried out:  The instrument is in good working condition and is well maintained but the instrument pressure reading showed a significant difference when compared to the traveling standard barometer. The instrument was adjusted.

#### ANEMOGRAPH

Type:	not available.						 Serial Number:

Condition of instrument and maintenance carried out:

### SOLAR RADIATION

Sensor Type:	not available	                                                   Serial Number: 	_______________________

Sensor Type:	____________________________________	   Serial Number:	_______________________

Sensor Type:  	____________________________________           Serial Number:	_______________________

Recorder Type:	____________________________________           Serial Number:	_______________________

Condition of instrument and maintenance carried out:	_______________________________________________

_______________________________________________________________________________________________

Additional Instruments:

Automatic Weather Station and Dew collector are available at this station.

VISUAL OBSERVATIONS AND GENERAL

Wind direction (check correct use of the Beaufort Scale, bearings known?)

A pressure plate is used at this station and Beaufort scale is available and observers are conversant with it.

Visibility (marks known, table on station?)

The sand dunes lying 35km on the Southern side of the station and a hill on the Western side about 50km from the station are used as visibility marks.

Cloud amounts, heights and types: The table is available

(Heights of surrounding hills known, is there a cloud book and cloud chart on station?)

A cloud chart is available and heights of nearby hills are known.

Field books 301s and remarks:

Field books are available and are adequately stalked and observations are well recorded.

Synoptic/Metar Coding(Practical test, comments, action taken. Is more training required?)

This is a research station and observers are not employees of Namibian Meteorological Service.

### Supplies, are there adequate supplies on station?

1 year supply is provided in advance.

Instrument care, management and maintenance:

Regular maintenance is done three times in a year.

Sketch, photograph and horizon profile of site:

See attachment.

Are hydrogen and balloon filling procedures being correctly followed?

Not available at the station.

Other items and general remarks on impression and ability of observer(s):

Observers are mainly students. Three observations are done per day and send via email to Met Station in Windhoek. In general the station is well situated and instruments are well maintained

RECOMMENDATIONS                                                                               ACTION COMPLETED

| Pressure Plate to be replaced with a conventional wind system |  |
| --- | --- |
| 120 cm soil thermometer to be replaced. |  |
| Automatic rain gauge to be replaced. |  |

Inspected by: D. F. MOLOTSI and E. MAGAGWA			 Date: 06/10/2006