9.0	Measurement Traceability and Calibration

9.1	Policy

9.1.1	Standards and measuring and test equipment significantly affecting the measurement integrity of tests conducted by the laboratory are monitored for stability as part of the measurement control program.  Standards and equipment are calibrated and/or verified before use to ensure the recall or removal from service of any equipment or standards that are unreliable or that have exceeded the calibration interval if established.

9.2	Measurement Traceability

9.2.1	The established program of calibration and verification of measuring and test equipment is designed to ensure that the measurements made by the laboratory are traceable to national standards of measurement.  All laboratory reports contain a statement that the standards used are traceable to the National Institute of Standards and Technology.

9.2.2	To provide external evidence of traceability, the laboratory participates in measurement control programs, proficiency tests, and other interlaboratory and collaborative experiments (see Appendices L and K).

9.3	Calibration/Verification (See Procedure for Calibration Intervals, AP No. 13)

9.3.1	Standards (Calibration)

9.3.1.1	Primary standards are calibrated by the National Institute of Standards and Technology or by a NIST-accredited laboratory.

9.3.1.2	Primary standards are continuously monitored by the measurement control programs established in the laboratory and are under the custody of trained metrologists (Note: According to State requirements, if any, document here who maintains custody).  These standards are generally stable and not subject to change.  Any changes that occur are generally identified by the laboratory and the NIST OWM State audit program in ample time for corrective action.

9.3.1.3	Standards are recalibrated if there is damage to the standards or any significant change is observed in the measurement control program.

9.3.2	Standards (Verification)

9.3.2.1	Continuous verification of standards, through the measurement control program, ensures required measurement integrity of test and includes:

A.	Statistical data from check standards and/or control charts (see Appendix J); and

B.	Results from interlaboratory comparisons and/or proficiency tests (see Appendix K).

9.3.2.2	Measurement assurance procedures for verification of standards are maintained in the laboratory files (see Appendix H, AP No. 13)

9.3.3	Measuring and Test Equipment (M &TE) (Calibration)

9.3.3.1	Equipment used with nominal values and corrections is calibrated by NIST or by a calibration laboratory whose traceability to NIST has been validated through an accreditation process.  A calibration interval is established for the equipment (i.e., environmental equipment, balances).

9.3.3.2	When appropriate, each item of equipment is labeled, marked, or otherwise identified to indicate its calibration status.

A.	All equipment used with nominal values and corrections is labeled indicating the calibration status.  Examples of this equipment include thermometers, barometers, hygrometers, and balances.

B.	All equipment used to facilitate comparison between a standard artifact and an unknown is labeled "Used for Comparisons Only - No Calibration Required."

9.3.3.3	Procedures for setting and changing M&TE calibration intervals are maintained in the laboratory files (see Appendix H, AP No. 14).

9.3.3.4	Calibration of equipment is conducted at a frequency to ensure that the equipment remains in tolerance during its use in the laboratory.  Frequency of calibration is based on a review of calibration, maintenance, and repair history.  Reviews are conducted by the technical manager and records are maintained with internal audit records in the laboratory files.

9.4	Uncertainty of Measurement

    - Where applicable, the ISO Guide to the Expression of Uncertainty  in Measurement, 1993; the G. Harris paper, Implementation of the Guide to the Expression of Uncertainty in Measurement for Calibration Laboratories, 1994; and NIST HB 145, 1986, are used as the basis for expression of uncertainty in measurement (see Appendix I, Assessment of Uncertainties). Note: NIST HB 145 is being revised to meet the new ISO requirements for the expression of uncertainty in measurement.  The revised document will be adopted by the laboratory upon completion.

17025

5.4.6  Estimation of uncertainty of measurement

5.4.6.1  A calibration laboratory, or a testing laboratory performing its own calibrations, shall have and apply a procedure to estimate the uncertainty of measurement for all calibrations and types of calibrations.

5.4.6.2  Testing laboratories shall have and apply procedures for estimating uncertainty of measurement. In certain cases the nature of the test method may preclude rigorous, metrologically and statistically valid, calculation of uncertainty of measurement. In these cases the laboratory shall at least attempt to identify all the components of uncertainty and make a reasonable estimation, and ensure that the form of reporting of the result does not give a wrong impression of the uncertainty. Reasonable estimation shall be based on knowledge of the performance of the method and on the measurement scope and shall make use of, for example, previous experience and validation data.

NOTE 1  The degree of rigor needed in an estimation of uncertainty of measurement depends on factors as:

- requirement of the test method;

- requirement  by the client;

- there are narrow limits on which decisions on conformance to a specification are based.

NOTE 2  In those cases where a well-recognized test method specifies limits to the values of the major sources of uncertainty of measurement and specifies the form of presentation of calculated results, the laboratory is considered to have satisfied this clause by following the test method and reporting instructions (see 5.10).

5.4.6.3  When estimating the uncertainty of measurement, all uncertainty components which are of importance in the given situation, shall be taken into account using appropriate methods of analysis.

NOTE 1  Sources contributing to the uncertainty include, but are not necessarily limited to, the reference standards and reference materials used, methods and equipment used, the environmental conditions, the properties and the condition of the item being tested or calibrated and the operator.

NOTE 2  The predicted long-term behaviour of the tested and/or calibrated item is not normally taken into account when estimating the measurement uncertainty.

NOTE 3  For further information see ISO 5725 and the Guide to the Expression of Uncertainty in Measurement (see bibliography).

5.6  Measurement traceability

5.6.1  General

All equipment used for tests and/or calibrations, including equipment for subsidiary measurements (e.g., for environmental conditions) having a significant effect on the accuracy or validity of the result of the test, calibration or sampling shall be calibrated before being put into service. The laboratory shall have an established programme and procedure for the calibration of its equipment.

NOTE -  Such a programme should include a system for selecting, using, calibrating, checking, controlling and maintaining measurement standards, reference materials, reference standards used as measurement standards, and measuring and test equipment used to perform tests and calibrations.

5.6.2  Specific requirements

5.6.2.1  Calibration

5.6.2.1.1  For calibration laboratories, the programme for calibration of equipment shall be designed and operated so as to ensure that calibrations and measurements made by the laboratory are traceable to the SI (Système International) units of measurement.

A calibration laboratory establishes traceability of its own measurement standards and measuring instruments to the SI by means of an unbroken chain of calibrations or comparisons linking them to relevant primary standards of the SI units of measurement. The link to SI may be achieved by reference to national measurement standards. National measurement standards may be primary standards, which are primary realizations of the units of the SI or agreed representations of SI units based on fundamental physical constants or they may be secondary standards which are standards calibrated by another national metrology institute. When using external calibration services traceability of measurement shall be assured by the use of calibration services from laboratories that can demonstrate competence, measurement capability and traceability. The calibration certificates issued by these laboratories shall contain the measurement uncertainty and/or a statement of compliance with an identified metrological specification (see also 5.10.4.2).

NOTE 1 – Calibration laboratories fulfilling the requirements of this International Standard are considered to be competent. A calibration certificate bearing an accreditation body logo from a calibration laboratory accredited to this International Standard for the calibration concerned, is sufficient evidence of traceability of the calibration data reported.

NOTE 2  Traceability to SI units of measurement may be achieved by reference to an appropriate primary standard (see VIM:1993, 6.4) or by reference to a natural constant the value of which in terms of the relevant SI unit is known and recommended by the General Conference of Weights and Measures (CGPM) and the International Committee for Weights and Measures (CIPM).

# NOTE 3  Calibration laboratories that maintain their own primary standard or representation of SI units based on fundamental physical constants can claim traceability to the SI only after these standards have been compared, directly or indirectly, with other similar standards of a national metrology institute.

NOTE 4 The term "identified metrological specification" means that it must be clear from the calibration certificate against which specification the measurements have been compared by including the specification or by giving an unambiguous reference to the specification.

NOTE 5 When the terms "international standard" or "national standard" are used in connection with traceability, it is assumed that these standards fulfil the properties of primary standards for the realization of SI units.

NOTE 6 Traceability to national measurement standards does not necessarily require the use of the national metrology institute of the country in which the laboratory is located.

NOTE 7 If a calibration laboratory wishes or needs to obtain traceability from a national metrology institute other than in its own country, this laboratory should select a national metrology institute that actively participates in the activities of BIPM either directly or through regional groups.

NOTE 8 The unbroken chain of calibrations or comparisons may be achieved in several steps carried out by different laboratories that can demonstrate traceability.

5.6.2.1.2  . There are certain calibrations that currently cannot be strictly made in SI units. In these cases calibration shall provide confidence in measurements by establishing traceability to appropriate measurement standards such as:

- the use of certified reference materials provided by a competent supplier to give a reliable physical or chemical characterization of a material;

- the use of specified methods and/or consensus standards that are clearly described and agreed by all parties concerned.

Participation in a suitable programme of interlaboratory comparisons is required where possible.

5.6.2.2  Testing

5.6.2.2.1  For testing laboratories, the requirements given in 5.6.2.1 apply for measuring and test equipment with measuring functions used, unless it has been established that the associated contribution from the calibration contributes little to the total uncertainty of the test result. When this situation arises, the laboratory shall ensure that equipment used can provide the uncertainty of measurement needed.

NOTE  The extent to which the requirements in 5.6.2.1 should be followed depends on the relative contribution of calibration uncertainty to the total uncertainty. If calibration is the dominant factor, the requirements should be strictly followed.

5.6.2.2.2  Where traceability of measurements to SI units of measurement is not possible and/or not relevant, the same requirements for traceability to e.g. certified reference materials, agreed methods and/or consensus standards are required as for calibration laboratories (see 5.6.2.1.2).

5.6.3  Reference standards and reference materials

5.6.3.1  Reference standards

The laboratory shall have a programme and procedure for the calibration of its reference standards. Reference standards shall be calibrated by a body that can provide traceability as described in 5.6.2.1. Such reference standards of measurement held by the laboratory shall be used for calibration only and for no other purpose, unless it can be shown that their performance as reference standards would not be invalidated. Reference standards shall be calibrated before and after any adjustment.

5.6.3.2  Reference materials

Reference materials shall, where possible, be traceable to SI units of measurement, or to certified reference materials. Internal reference materials shall be checked as far as is technically and economically practicable.

5.6.3.3  Intermediate checks

Checks needed to maintain confidence in the calibration status of reference, primary, transfer or working standards and reference materials shall be carried out according to defined procedures and schedules.

5.6.3.4  Transport and storage

The laboratory shall have procedures for safe handling, transport, storage and use of reference standards and reference materials in order to prevent contamination or deterioration and in order to protect their integrity.

NOTE  Additional procedures may be necessary when reference standards and reference materials are used outside the permanent laboratory for tests, calibrations or sampling.