Online Calibration of VLPC chips

Jadwiga Warchol

Introduction

I evaluate the feasibility of calibrating of VLPC chips during the data taking run in Run II. Calibration event is an event when LED chips illuminate all fibers and all SVX II chips connected to VLPC chips are read out. Calibration data are useful only if at least 10 000 triggers can be collected. I assume that the calibration events should not exceed 0.1 % of the data taking time.

There are two possible sinks for calibration data: local processors in the crates containing VRB modules or Level 3 monitoring computers, they both are evaluated. It turns out that the serious obstacle to the feasibility of calibrating of VLPC chips is long time required to initialize SVX II chips while switching from data mode to calibration mode.

Data Flow

The signals from VLPC chips are fed into SVX II chips. Eight SVX II chips (a stream of data) are read out serially into a VRB (VME Readout Buffer) module, eight such streams (on four fibers since one fiber is 16 bits wide and SVX data eight bits wide) are input in parallel into one VRB module. Ten VRB modules can be placed in one VME create. All CFT and preshower VRB modules occupy three crates (all axial CFT fibers are in one crate). All three creates are connected to a single data cable transporting the data to Level 3 System.

Calibration Using Local Processors

The VME crate with VRB modules contains an embedded processor on VME board. The silicon group intends to monitor the silicon tracker with these embedded processors operating in VxWorks environment. In order to monitor the data without hampering DAQ system two components are built or considered: monitoring buffer in VRB module and the IEEE P-1394 high-speed bus. The monitoring buffer in VRB module is 1/8th of the data buffer, it accepts the data from 1 of total of 8 data streams input to VRB module, the logic on VRB module programs which data stream gets into data and monitoring buffer. The embedded processor accesses monitoring buffer via IEEE P-1394 bus; in such way the data for monitoring do not use VME bus, which is used entirely by DAQ.

The VLPC calibration event sinking in embedded processor is a Level 1 trigger. The first step is the initialization of SVX II chips. The initialization for calibration event sets to 1 Read All Channels bit in SVX , this is needed since in the data mode the SVX II chips suppress hits below the threshold between 1 pe. and about 1.7 pe. . To do that one needs fully initialize SVX chip, which is a serial operation involving clocking in of 190 bits, and since chips are daisy chained in groups of eight altogether 190·8=1520 bits have to be clocked in. The clocking in of the bits is done by 1553 with the speed of about 1.3 (s/bit, and it has to be done twice since after the calibration event the SVX chips have to be initialized again for data taking mode, total initialization time, tinit, is then:

tinit =1520 bits·1.3 (s/bit·2 = 4000 µs.

The time from the trigger to the beginning of readout, tSVX, is a few µs:

tSVX = 560ns(delay before digitization)+1300ns(after digitization)= 2 µs.

The time, ttransf, needed to transfer the data to VRB is very short:

ttransf =8chips(1 stream of data)·128 channels·19ns/byte·2bytes/channel(address,data)=40 µs.

Since monitoring buffer in VRB sees only 1/8th of data buffer one needs eight calibration triggers to obtain a calibration event with the data from all VLPC chips. To get one full calibration event one needs 32.104 ms; this means that the frequency of calibration events has to be 1 per 32 s or 112 events per hour to use 0.1% of total time.

The maximum speed with which one can clock in bits into SVX  is 53 MHz, leading to tinit of 58 µs (1520 bits·19ns/bit·2). In that case the frequency of calibration events is 4000 per hour.

Table 1. Time to read out calibration event with embedded processors.

|  | Initialization by 1553 | Initialization by 53 MHz clock |
| --- | --- | --- |
| Initialization of SVX, tinit | 4000 µs | 58 µs |
| Delay in SVX, tSVX | 2 µs | 2 µs |
| Transfer to VRB, ttransf | 40 µs | 40 µs |
| Pipeline reset, tpipereset | 6 µs | 6 µs |
| Total for 1/8th of VLPC chips | 4052 µs | 106 µs |
| Total for all VLPC chips | 32.416 ms | 0.85 ms |

Frequency of calibration events :                   112/hour                            4000/hour

Frequency of calibration events using Level 3 system: 300/hour

Calibration Using Level 3 System

Now the data buffers of VRB modules are used, the data are transferred on VME backplane with the speed of 20 Mbytes/s to the data cable of Level 3 system. The speed on the data cable is 48 Mbytes/s and this is the effective speed because three VRB crates are connected to the data cable. The size of calibration event is 400kbytes (100000 VLPC channels·2 bytes/channel·2 for unused SVX channels) and the time to read it is 10 ms (400kbytes/48 Mbytes/s). The total time for one calibration event is 12 ms (10ms for data transfer+2 ms for SXV initialization-the initialization for data mode can be done in parallel with data transfer), and the frequency of calibration events is 300 events per hour.

Sampling Calibration Event

This is calibration event when only a fraction of VLPC chips is calibrated. Here I present an example of such event. We calibrate eight VLPC chips reading axial fibers in one wedge and two SVX chips to which these chips are connected are permanently initialized for Real All Channels mode. This adds 2·128·2=0.5kbytes to each data event, and constitutes about 17% of the size of fiber tracker's data (74240fiber·0.02%occupancy·2=3kbytes).

With embedded processor total time for the fractional calibration event is 48 µs (tSVX + ttransf + tpipereset) which would allow us to take 75000 sampling calibration events per hour (3600sec/hour /48000µs). Since data size is small one could transfer the data to Embedded Processor via VME bus increasing somewhat the time.

Using of Level 3 would require illumination of only one wedge by LED. The size of calibration event would be 2kbytes (960 fibers in one wedge·2). Total time for transfer of the data of the calibration event is 50 µs (2kbytes/40 Mbytes/s) and the total time of the sampling calibration event is 58 µs, which is very close to 48µs in the case of embedded processor.

Conclusions

The present electronics does not allow for Online calibration of VLPC chips due to very slow initializing time of SVX chips. The calibration of eight chips is possible by permanently initializing two SVX II chips for Real All Channels mode, this increases the size of data event by 17 % of the size of fiber data.