| Measurement Quality Objectives - Parameter NO2 (Chemiluminescence) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Standard Reporting Units | All data | ppm | 40 CFR, Pt 50.11 |  |
| Shelter Temperature Temperature range Temperature control | Daily Daily | 20 to 30E C # " 2E C | 40 CFR, Pt. 53.20 Vol II, S 7.1 1/ Vol II, MS 2.3.2 | Instruments designated as reference or equivalent have been tested over this temperature range. Maintain shelter temperature above sample dewpoint. Shelter should have a 24- hour temperature recorder. Flag all data for which temperature range or fluctuations are outside acceptance criteria. |
| Equipment NO2 analyzer Air flow controllers Flowmeters | Purchase specification | Reference or equivalent method Flow rate regulated to " 2 % Accuracy " 2 % | 40 CFR, Pt 53.9 40 CFR, Pt 50, App F, S 2.2 EPA-600/4-75-003 |  |
| Detection Noise Lower detectable level | Purchase specification | 0.005 ppm 0.01 ppm | 40 CFR, Pt 53.20 & 23 “ | Instruments designated as reference or equivalent have been determined to meet these acceptance criteria |
| Completeness Hourly Data | Quarterly | 75 % | 40 CFR, Pt 50.11 |  |
| Compressed Gases Dilution gas (zero air) Gaseous standards | Purchase specification Purchase specification | Free of contaminants NIST Traceable (e.g., EPA Protocol Gas) | EPA-600/4-75-003 40 CFR, Pt 50, App F, S 1.3 EPA-600/R-97/121 | Return cylinder to supplier. Nitric oxide in nitrogen EPA Protocol Gases have a 24-month certification period and must be recertified to extend the certification. |
| Calibration Multipoint calibration (at least 5 points) Convertor efficiency Zero/span check- level 1 Flowmeters | > 1/6 months., after failure of QC check or after maintenance During multipoint calibrations 1/ 2 weeks 1/3 months | Residence time < 2 min Dynam. parameter > 2.75 ppm-min All points within " 2 % of full scale of best-fit straight line $ 96 % Zero drift # " 20 to 30 ppb Span drift # " 20 to 25 % Zero drift # " 10 to 15 ppb Span drift # " 15 % Accuracy " 2 % | 40 CFR, Pt 50, App F, S 1 Vol II, S 12.6 Vol II, MS 2.3.2 40 CFR, Pt. 50, App F Vol II, MS.2.3.2 Vol II, S 12.6 Vol II, MS 2.3.2 Vol II, S 12.6 Vol II, MS 2.3.2 Vol II, App 12 | Zero gas and at least four upscale calibration points. Points outside acceptance criterion are repeated. If still outside consult manufacturers manual and invalidate data to last acceptable multipoint calibration or zero/span check . Replace or service converter. If calibration factors are updated after each zero/span, invalidate data to last acceptable zero/span check, adjust analyzer, and perform multipoint calibration. If fixed calibration factors are used to calculate data, invalidate data to last acceptable zero/span check, adjust analyzer, and perform multipoint calibration. Flowmeter calibration should be traceable to NIST standards. |
| Performance Evaluation (NPAP) State audits | 1/year at selected sites 1/year | Mean absolute difference # 15 % State requirements | NPAP QAPP Vol II, App 15, S 3 | Use information to inform reporting agency for corrective action and technical systems audits. |
| Precision Single analyzer Reporting organization | 1/ 2 weeks 1/3 months | None 95 % Confidence Interval # " 15 % | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 6 | Concentration. = 0.08-0.10 ppm. |
| Accuracy Single analyzer Reporting organization | 25 % of sites quarterly (all sites yearly) | None 95% Confidence Interval # " 20% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 3 | Four concentration ranges. If failure, recalibrate analyzer and reanalyze samples. Repeated failure requires corrective action. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives - Parameter O3 (Ultraviolet Photometric) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Standard Reporting Units | All data | ppm | 40 CFR, Pt 50.9 |  |
| Shelter Temperature Temperature range Temperature control | Daily Daily | 20 to 30E C. # " 2E C | 40 CFR, Pt. 53.20 Vol II, S 7.1 1/ Determination of Ozone by Ultraviolet Analysis (draft) | Instruments designated as reference or equivalent have been tested over this temperature range. Maintain shelter temperature above sample dewpoint. Shelter should have a 24- hour temperature recorder. Flag all data for which temperature range or fluctuations are outside acceptance criteria. |
| Equipment O3 analyzer | Purchase specification | Reference or equivalent method | 40 CFR, Pt 53.9 EPA-600/4-79-057 | Air flow controllers must be capable of regulating air flows as necessary to meet the output stability and photometer precision requirements. The photometric measurement of absorption is not directly related to flow rate, but may be indirectly related due to thermal or other effects. |
| Detection Noise Lower detectable level | Purchase specification | 0.005 ppm 0.01 ppm | 40 CFR, Pt. 53.20 & 23 “ | Instruments designated as reference or equivalent have been determined to meet these acceptance criteria. |
| Completeness (seasonal) Maximum 1-hour concentration | Daily | 75% values from 9:01 AM to 9:00 PM (LST) | 40 CFR, Pt 50, App H, S 3 | A missing daily maximum ozone value may be assumed to be less than the standard if valid daily maxima on the preceding and following days do not exceed 75 percent of the standard. |
| Transfer standard Qualification and certification Recertification to local primary standard | Upon receipt of transfer standard 1/3 months (if at a fixed site) | "4% or "4 ppb (whichever greater) RSD of six slopes # 3.7% Std. dev. of six intercepts #1.5% New slope = "0.05 of previous | EPA-600/4-79-056 EPA-600/4-79-057 “ “ | 6 comparison runs that include, at minimum, 6 concentrations per comparison run including 0 and 90 + 5% of upper range. A single six-point comparison run. |
| Local primary standard Certification/recertification to Standard Photometer (if recertified via a transfer standard) | 1/year " | Difference # "5 % (preferably " 3%) Regression slopes = 1.00 " 0.03 and two intercepts are 0 " 3 ppb | Determination of Ozone by Ultraviolet Analysis (draft) " | The local primary standard is a standard in its own right, but it must be repaired and recertified if the acceptance criterion is exceeded. |
| EPA Standard Reference Photometer recertification | 1/year | Regression slope = 1.00 + 0.01 and intercept < 3 ppb | Protocol for Recertification of Standard Reference Photometers... (TRC Environmental Document) | 9 replicate analysis over 12 conc. ranges. Disagreement must be resolved. EPA Standard Reference Photometer rechecked with NIST. If OK Network STANDARD REFERENCE PHOTOMETER must be repaired. |
| Zero air | Purchase specification | Free of O3 or any substance that might react with O3 (e.g., NO, NO2, hydrocarbons, and particulates) | EPA-600/4-79-057 | Return cylinder to supplier |
| Ozone analyzer calibration Zero/span check -level 1 Multipoint calibration (at least 5 points) | 1/ 2 weeks Upon receipt, adjustment, or 1/ 6 months | Zero drift # " 20 to 30 ppb Span drift # " 20 to 25 % Zero drift # " 10 to 15 ppb Span drift # " 15% Linearity error <5% | Vol II, S 12.6 “ Vol II, S 12.6 “ 40 CFR, Pt 50, App D, S 5.2.3 EPA-600/4-79-057 S.5 Vol II, S 12.2 | If calibration updated at each zero/span, Invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration. If fixed calibration used to calculate data, Invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration. Zero gas and at least four upscale calibration points. Check verify accuracy of flow dilution. Redo analysis. If failure persists corrective action required. |
| Performance Evaluation (NPAP) State audits | 1/year at selected sites 1/year | Mean absolute difference # 15% State requirements | Vol II, S 16.3 Vol II, App 15, S 3 | Use information to inform reporting agency for corrective action and technical systems audits. |
| Precision Single analyzer Reporting organization | 1/ 2 weeks 1/3 months | None 95% CI < " 15% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 6 | Concentration = 0.08-0.10 ppm. |
| Accuracy Single analyzer Annual accuracy | 25 % of sites quarterly (all sites yearly) | None 95% CI # " 20% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 6 | Four concentration ranges. If failure, recalibrate and reanalyze. Repeated failure requires corrective action. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives - Parameter Lead (Atomic Absorption Spectroscopy) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Reporting Units | All data | :g/m3 | 40 CFR, Pt 50.12 |  |
| Filter Checks Visual defect check Filter Integrity Collection efficiency Integrity pH | All filters Purchase specification | See reference 99% 2.4 mg max weight loss 6 to 10 | Vol II, MS 2.2.4 40 CFR, Pt 50, App B, S 7.1 “ “ | Discard any defective filters Measure using DOP test (ASTM-2988). Reject shipment |
| Equipment Sampler Flow rate transfer standard | Purchase specification Purchase specification | Reference or equivalent method 0.02 std. m3/min | 40 CFR, Pt 53.9 40 CFR, Pt 50, App B, S 7 " |  |
| Detection Limit LDL | Not applicable | 0.07 Fg/m3 | 40 CFR, Pt 50, App G, S 2 | This value is based on a collaborative test of the method. Assumed air volume of 2,400 m3. |
| Completeness | Quarterly | 75% |  |  |
| Sampler calibration Orifice calibration unit (flow rate transfer standard) Elapsed time meter On/Off Timer Sampler flow rate | On receipt and yearly On receipt and 1/6 months On receipt and 1/3 months On receipt, if audit deviation > 7 %, after maintenance | Indicated flow rate within " 2 % of actual flow rate " 2 min/24 hours " 30 min/24 hour All points within " 5 % of full scale of best-fit straight line | Vol II, MS 2.8.1 Vol II, MS 2.2.2 “ Vol II, MS 2.2.2 " | Adopt a new calibration curve. A rotary-type, gas displacement meter is the recommended NIST-traceable reference standard. Adjust or replace meter Checked against elapsed time meter. Adjust or repair. Rerun points outside limits until acceptable. |
| Analytical calibration Reproducibility test Calibration stability | On receipt Before first sample, after every tenth sample, after last sample | # 5% # " 5 % deviation from calibration curve. | Vol II, MS 2.8.1 Vol II, MS 2.8.5 | Reproducibility = 100 ([high response-low response]/average response). Responses should be corrected for the blank level. If acceptance criterion is exceeded, instrument should be checked by a service rep or qualified operator. Alternate between two control standards with concentrations# 1 :g/mL or 1 to # 10 :g/mL. Take corrective action and repeat the previous ten analyses. |
| Performance Evaluation (NPAP) Sampler performance Audit (flow rate) | 1/year at selected sites 1/3 months | Mean absolute difference # 15% Percentage difference # "7% | Vol II, S 16.3 40 CFR, Pt 58, App A Vol II, MS 2.2.8 | Use information to inform reporting agency for corrective action and technical systems audits Recalibrate before any additional sampling |
| Precision Single analyzer Reporting organization | 1/6 days 1/ 3 months | None 95% CI < " 15% | 40 CFR, Pt 58, App A, S 5.3 40 CFR, Pt 58, App A, S 5.3 | Both lead values must be > 0.15 :g/m3 |
| Accuracy Single analyzer Reporting organization | 25 % of sites quarterly | Percentage difference # " 16% 95% CI # " 20% | Vol II, MS 2.8.8 40 CFR, Pt 58, App A, S 3.4 EPA-600/4-83-023 | Analyze three audit samples in each of the two concentration ranges. The audit samples shall be distributed as much as possible over the entire calendar quarter. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives - Parameter PM10 (Dichotomous Sampler) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Reporting Units | All data | :g/m3 | 40 CFR, Pt 50.7 |  |
| Filter Checks Visual defect check Filter Integrity Collection efficiency Integrity Alkalinity Filter Conditioning Equilibration time Temperature range Temperature control Humidity range Humidity control | All filters Purchase specification All Filters “ “ “ “ | See reference 99 % " 5 :g/m3 < 25.0 microequivalents/gram at least 24 hours 15 to 30E C " 3E C 20 to 45 % relative humidity " 5 % relative humidity | Vol II, MS 2.10.4 40 CFR, Pt 50, App M, S 7.2 “ “ 40 CFR, Pt 50, App M, S 9.3 40 CFR, Pt 50, App M, S 7.4 “ “ “ | Discard any defective filters As measure by DOP test (ASTM-2988). Reject shipment. Following 2 months storage at ambient temp and relative humidity. Reject filters Repeat equilibration Keep thermometer in balance room and record temperature daily. Keep hygrometer in the balance room and record humidity daily. |
| Equipment Sampler Flow rate transfer standard Analytical balance Mass reference standards | Purchase specification Purchase specification Purchase specification Purchase specification | Reference or equivalent method " 2 % accuracy (NIST traceable) Sensitivity = 0.1 mg NIST traceable (e.g., ANSI/ASTM Class 2) | 40 CFR, Pt 53.9 40 CFR, Pt 50, App M, S7.3 40 CFR, Pt 50, App M, S 7.5 Vol II, MS 2.10.4 Vol II, MS 2.10.4 | This acceptance criterion is inconsistent with other acceptance criteria for balance that are in the quality assurance handbook. |
| Detection Limit LDL | Not applicable | Not applicable | 40 CFR, Pt 50, App M, S 3.1 | The lower limit of the mass concentration is determined by the repeatability of filter tare weights, assuming the nominal air sample volume for the sampler. |
| Completeness | quarterly | 75% | 40 CFR, Pt 50, App K, S 2.3 |  |
| Sampler Calibration Flow control device Elapsed time meter Flow-rate transfer Standard | On installation, after repairs, after out-of-limits flow check On receipt and 1/6 months Periodically | <4% difference from manufacturers spec and actual " 15 min "2% over the expected range of ambient conditions | 40 CFR, Pt 50, App M, S 7.1 Vol II, MS 2.10.2 40 CFR, Pt 50, App M, S 7.1 Vol II, MS 2.10.1 40 CFR, Pt 50, App M, S 8.2 Vol II, MS 2.10.1 | Adopt new calibration curve if no evidence of damage, otherwise replace. Adjust or replace. Checked against NIST-traceable primary standard. |
| Balance Calibration | 1/year |  | Vol II, MS 2.10.4 | Calibrate and maintain according to the manufacturer's recommendations. |
| Performance Evaluation (NPAP) | 1/year at selected sites | Mean absolute difference # 15% | Vol II, S 16.3 | Use information to inform reporting agency for corrective action and technical systems audits |
| Precision Single analyzer Reporting organization | 1/6 days 1/ 3 months | # 5 Fg/m3 for conc. #80 :g/m3 7% for conc. >80 :g/m3 95% CI < " 15% | 40 CFR, Pt 50, App M, S 4.1 40 CFR, Pt 58, App A, S 5.3 EPA-600/4-83-023 | Both PM10 values must be > 20 :g/m3. |
| Accuracy Single analyzer Annual accuracy | 25 % of sites quarterly (all sites yearly) | None 95% CI # " 20% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 6 | Transfer standards different then those used in calibration. Recalibrate before any additional sampling. Invalidate data to last acceptable flow check if difference > 10%. |
| QC Checks Field calibration flow check "Standard" filter weighing Reweighing filters Balance zero and calibration check | 1/month at beginning of weighing day 5 exposed and 5 unexposed/day every fifth filter | Percentage difference # "7 % from sampler's indicated flow rate or # " 10 % from design condition flow rate " 20 Fg of original weight " 20 Fg of original weight " 4 Fg at zero " 2 Fg at 10 mg | 40 CFR, Pt 50, App M, S 8.2 Vol II, MS 2.10.3 Vol II, S 2.10.4 Vol II, S 2.10.4 Vol II, S 2.10.4 | Trouble shoot and recalibrate sampler. Trouble shoot and reweigh. Trouble shoot and reweigh. Trouble shoot and reweigh. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives - Parameter SO2 (Ultraviolet Fluorescence) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Standard Reporting Units | All data | ppm | 40 CFR, Pt 50.4 |  |
| Shelter Temperature Temperature range Temperature control | Daily Daily | 20 to 30E C # " 2E C | 40 CFR Pt. 53.20 Vol II, S 7.1 1/ Vol II, MS 2.9 | Instruments designated as reference or equivalent have been tested over this temperature range. Maintain temperature above sample dewpoint. Shelter should have a 24- hour temperature recorder. Flag all data for which temperature range or fluctuations are outside acceptance criteria. |
| Equipment SO2 analyzer Air flow controllers Flowmeters | Purchase specification | Reference or equivalent method Flow rate regulated to " 2 % Accuracy " 2 % | Vol II, MS 2.9 " " |  |
| Detection Noise Lower detectable level | Purchase specification | .005 ppm .01 ppm | 40 CFR, Pt 53.20 & 23 “ | Instruments designated as reference or equivalent have been determined to meet these acceptance criteria. |
| Completeness Annual standard 24-hour standard 3-hour standard | Quarterly 24 hours 3 hours | 75% 75% 75% | 40 CFR, Pt 50.43 “ “ |  |
| Compressed Gases Dilution gas (zero air) Gaseous standards | Purchase specification Purchase specification | SO2 free, 21 % O2/78 % N2, 300 to 400 ppm CO2, # 0.1 ppm aromatics NIST Traceable (e.g., permeation tube or EPA Protocol Gas | Vol II, MS 2.9.2 EPA-600/R97/121 | Return cylinder to supplier. It is recommended that a clean air system be used instead of compressed air cylinders. Sulfur dioxide in nitrogen EPA Protocol Gases have a 24-month certification period for concentrations between 40 and 499 ppm and a 36-month certification period for higher concentrations. |
| Calibration Multipoint calibration (at least 4 points) Zero/span check -level 1 Flowmeters | Upon receipt, adjustment, or 1/ 6 months 1/ 2 weeks 1/3 months | All points within + 2% of full scale of best-fit straight line Zero drift # " 20 to 30 ppb Span drift # " 20 to 25 % Zero drift # " 10 to 15 ppb Span drift # " 15% Accuracy " 2 % | Vol II, S 12.6 Vol II, MS 2.9.2 Vol II, S 12.6 “ Vol II, S 12.6 “ Vol II, App 12 | Zero gas and at least three upscale points. Note: two pages from Section 2.4 (Calibration Procedures) of Vol II, MS 2.9.2 are missing from the 1994 reprinting of the QA Handbook. If calibration updated at each zero/span- Invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration If fixed calibration used to calculate data. Invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration Flowmeter calibration should be traceable to NIST standards |
| Performance Evaluation (NPAP) State audits | 1/year at selected sites 1/year | Mean absolute difference # 15% State requirements | Vol II, S 16.3 Vol II, App 15, S 3 | Use information to inform reporting agency for corrective action and technical systems audits. |
| Precision Single analyzer Reporting organization | 1/2 weeks 1/3 months | None 95% CI < " 15% | 40 CFR, Pt 58, App EPA-600/4-83-023 Vol II, S 16, S2 | Concentration = 0.08-0.10 ppm. |
| Accuracy Annual accuracy check- Reporting organization | 25 % of sites quarterly (all sites yearly) | None 95% CI # " 20% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, S 16 | Four concentration ranges. If failure, recalibrate and reanalyze. Repeated failure requires corrective action. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives - Parameter CO (Nondispersive Infrared Photometry) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Standard Reporting Units | All data | ppm | 40 CFR, Pt 50.8 |  |
| Shelter Temperature Temperature range Temperature control | Daily Daily | 20 to 30E C. < " 2E C | 40 CFR, Pt. 53.20 Vol II, S 7.1 1/ | Instruments designated as reference or equivalent have been tested over this temperature range. Maintain shelter temperature above sample dewpoint. Shelter should have a 24- hour temperature recorder. Flag all data for which temperature range or fluctuations are outside acceptance criteria. |
| Equipment CO analyzer Flow controllers Flowmeters | Purchase specification | Reference or equivalent method Flow rate regulated to " 1% Accuracy " 2% | 40 CFR, Pt 50, App C " “ |  |
| Detection Limit Noise Lower detectable level | Purchase specification | 0.5 ppm 1.0 ppm | 40 CFR, Pt 53.20 & 23 “ | Instruments designated as reference or equivalent have been determined to meet these acceptance criteria. |
| Completeness 8-hour average | hourly | $75 % of hourly averages for the 8-hour period | 40 CFR, Pt 50.8 |  |
| Compressed Gases Dilution gas (zero air) Gaseous standards | Purchase specification Purchase specification | < 0.1 ppm CO NIST Traceable (e.g., EPA Protocol Gas) | 40 CFR, Pt 50, App C " EPA-600/R97/12 | Return cylinder to supplier. Carbon monoxide in nitrogen or air EPA Protocol Gases have a 36-month certification period and must be recertified to extend the certification. |
| Calibration Multipoint calibration (at least 5 points) Zero/span check-level 1 Flowmeters | Upon receipt, adjustment, or 1/ 6 months 1/ 2 weeks 1/3 months | All points within " 2% of full scale of best-fit straight line Zero drift # " 2 to 3 ppm Span drift # " 20 to 25 % Zero drift # " 1 to 1.5 ppm Span drift # " 15% Accuracy " 2 % | Vol II, S 12.6 Vol II, MS.2.6.1 Vol II, S 12.6 " Vol II, S 12.6 " Vol II, App 12 | Zero gas and at least four upscale calibration points. Points outside acceptance criterion are repeated. If still outside criterion, consult manufacturers manual and invalidate data to last acceptable calibration. If calibration updated at each zero/span, invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration. If fixed calibration used to calculate data, invalidate data to last acceptable check, adjust analyzer, perform multipoint calibration. Flowmeter calibration should be traceable to NIST standards. |
| Performance Evaluation (NPAP) State audits | 1/year at selected sites 1 /year | Mean absolute difference # 15% State requirements | Vol II, S 16.3 Vol II, pp 15, S 3 | Use information to inform reporting agency for corrective action and technical systems audits |
| Precision Single analyzer Reporting organization | 1/2 weeks 1/3 months | None 95% CI # " 15% | 40 CFR, Pt 58, App A EPA-600/4-83-023 Vol II, App 15, S 5 | Concentration = 8 to 10 ppm. Aggregation of a quarters measured precision values. |
| Accuracy Single analyzer Reporting organization | 25 % of sites quarterly (all sites yearly) | None 95% CI # " 20% | 40 CFR, Pt 58, App A | Four concentration ranges. If failure, recalibrate and reanalyze. Repeated failure requires corrective action. |

1/ - reference refers to the QA Handbook for Air Pollution Measurement Systems, Volume II . The use of “S” refers to sections within Part 1 of Volume II.  The use of “MS” refers to method-specific sections in Volume II.

| Measurement Quality Objectives- Parameter PM2.5 |
| --- |
| Requirement | Frequency | Acceptance Criteria | 40 CFR Reference | QA Guidance Document 2.12 Reference |
| Filter Holding Times Pre-sampling Post-sampling Weighing | all filters “ | < 30 days before sampling < 10 days at 25O C from sample end date < 30 days at 4OC from sample end date | Part 50, App.L Sec 8.3 “ “ | Sec. 7.9 Sec. 7.11 “ |
| Sampling Period | All data | 1380-1500 minutes or value if < 1380 and exceedance of NAAQS | Part 50, App.L Sec 3.3 |  |
| Reporting Units | All data | Fg/m3 | Part 50.3 | Sec. 11.1 |
| Detection Limit Lower DL Upper Conc. Limit | All data All data | 2 Fg/m3 200 Fg/m3 | Part 50, App.L Sec 3.1 Part 50, App.L Sec 3.2 |  |
| Sampling Instrument Flow Rate Filter Temp Sensor | every 24 hours of op “ “ “ | < 5% of 16.67 < 2% CV measured < 5% average for < 5 min. < 5o C of ambient for <30min | Part 50, App.L Sec 7.4 “ “ |  |
| Data Completeness | quarterly | 75% | Part 50, App. N, Sec. 2.1 |  |
| Filter Visual Defect Check Filter Conditioning Environment Equilibration Temp. Range Temp. Control Humidity Range Humidity Control Pre/post sampling RH Balance | All Filters All filters “ “ “ “ “ “ | See reference 24 hours minimum 20-23o C +2o C SD over 24 hr 30% - 40% RH or _+ 5% sampling RH but >20%RH + 5% SD over 24 hr. + 5% RH located in filter conditioning environment | Part 50, App.L Sec 6.0 Part 50, App.L Sec 8.2 “ “ “ “ Part 50, App.L Sec 8.3.3 “8.3.2 | Sec 7.5 Sec. 7.6 " " " " |
| Filter Checks Lot Blanks Exposure Lot Blanks | 3 filters per lot 3 filters per lot | less than 15 Fg change between weighings less than 15 Fg change between weighings | not described not described | Sec. 7.7 Sec. 7.7 |
| Lab QC Checks Field Filter Blank Lab Filter Blank Balance Check Duplicate Filter Weighing | 10% or 1 per weighing session 10% or 1 per weighing session beginning, every 10th sample, end 1 per weighing session | +30 Fg change between weighings +15 Fg change between weighings < 3 Fg +15 Fg change between weighings | Part 50, App.L Sec 8.3 Part 50, App.L Sec 8.3 not described not described | Sec. 7.7 “ Sec. 7.9 Sec 7.11 |
| Calibration/Verification Flow Rate (FR) Calibration FR multi-point verification One point FR verification External Leak Check Internal Leak Check Temperature Calibration Temp M-point Verification One-point temp Verification Pressure Calibration Pressure Verification Clock/timer Verification | If multi-point failure 1/yr 1/4 weeks every 5 sampling events every 5 sampling events If multi-point failure on installation, then 1/yr 1/4 weeks on installation, then 1/yr 1/4 weeks 1/ 4 weeks | + 2% of transfer standard + 2% of transfer standard + 4% of transfer standard 80 mL/min 80 mL/min + 2% of standard + 2EC of standard + 4EC of standard "10 mm Hg "10 mm Hg 1 min/mo | Part 50, App.L, Sec 9.2 Part 50, App.L, Sec 9.2.5 Part 50, App.L, Sec 9.2 Part 50, App.L, Sec 7.4 " Part 50, App.L, Sec 9.3 Part 50, App.L, Sec 9.3 " “ “ Part 50, App.L, Sec 7.4 | Sec 6.3 Sec 6.3 & 8.4 Sec 8.4 Sec. 6.6 & 8.4 Sec. 6.6 & 8.4 Sec. 6.4 Sec. 6.4 and 8.4 Sec. 6.4 and 8.4 Sec. 6.5 Sec. 8.2 not described |
| Accuracy FRM Performance Evaluation External Leak Check Internal Leak Check Temperature Audit Pressure Audit Balance Audit | 25% of sites 4/yr 4/yr 4/yr 4/yr 4/yr (?) 1/yr | + 10% < 80 mL/min < 80 mL/min + 2EC "10 mm Hg Manufacturers specs | Part 58, App A, Sec 3.5 not described not described not described not described not described | Sec 10.2 Sec. 10.2 " " " " |
| Accuracy Flow Rate Audit | 1/2wk (automated) 4/yr (manual) | + 4% of audit standard | Part 58, App A, Sec 3.5 | Sec. 10.2 |
| Precision Collocated samples Single analyzer Single Analyzer Reporting Org. | every 6 days for 25% of sites 1/3 mo. 1/ yr 1/ 3 mo. | CV < 10% CV < 10% CV < 10% CV < 10% | Part 58, App.A, Sec 3.5 and 5.5 not described not described not described | Sec. 10.2 not described not described not described |
| Calibration & Check Standards Flow Rate Transfer Std. Field Thermometer Field Barometer Working Mass Stds. Primary Mass Stds. | 1/yr 1/yr 1/yr 3-6 mo. 1/yr | +2% of NIST-traceable Std. + 0.1o C resolution + 0.5o C accuracy + 1 mm Hg resolution + 5 mm Hg accuracy 0.025 mg 0.025 mg | Part 50, App.L Sec 9.1 & 9.2 not described not described not described not described | Sec. 6.3 Sec 4.2 & 6.4 “ “ “ Sec 4.3 and 7.3 " |

| Measurement Quality Objectives - Parameter PAMS Volatile Organic Compounds (VOC) |
| --- |
| Requirement | Frequency | Acceptance Criteria | Reference | Information/Action |
| Standard Reporting Units | All data | ppbC | TAD, July 1997 |  |
| Shelter Temperature Temperature range | Daily | 20 to 30E C. | Vol II, S 7.1 1/ | Instruments designated as reference or equivalent have been tested over this temperature range. Maintain shelter temperature above sample dewpoint. Shelter should have a 24- hour temperature recorder. Flag all data for which temperature range or fluctuations are outside acceptance criteria. |
| Detection Limit System detection limit |  | 1 ppbC | TAD Sect 2.8 2.3 | Calculation based on multiple manual or automated analysis and 40 CFR recommendations |
| Completeness (sesonal) | annually | 85 % | TAD 2.8.1 |  |
| Calibration Multipoint retention time calibration standard | Start of analytical season | correlation coefficient > 0.995 | TAD 2.8.2.3 | Triplicate analysis of multiple level propane standards over the expected sample concentration range (a minimum of three levels) |
| Performance Evaluation NPAP | prior to start of sampling season and twice during monitoring season | In absence of specified objectives within 25% | TAD 2.8.2.3 | Useful for informing reporting agency for corrective actions and technical systems audits. |
| Precision Duplicate samples | once/2weeks automated 10% -manual | " 25% RSD or RPD | TAD 2.8.2.1.1 | Comparison of duplicate field samples, or replicate sample analysis using manual or automated field devices. |
| QC Checks Retention time (RT) calibration check Canister cleaning | Weekly | Response Factor within 10% RPD of calibration curve < 10 ppbC total | TAD 2.8.2.3 | Retention time checked versus annual PAMS retention time cylinder provided to each site in the program. Canister cleaning per approved methodology |
| Background/carryover | weekly and after calibration & RT | < 20 ppbC for both columns or <10 ppbC per column | TAD 2.8.2.3 | Background testing according to TAD |