## Sheet1
Lab Code (A-J) | Burner Type (P-Park, L-Lennox, C-Carlin, X-Other) | Igniter Position (1-12 O'clock) | Distance from Turbulator to Nozzle (inches) | Calorimeter Mfgr (V-Vatell, T-Thermoguage, M-Medtherm, X-Other) | Thermocouple Mfgr (T-Thermoelectric, X-Other) | Air Velocity Meter (O-Omega, X-Other) | Pre-test calibration date (mm/dd/yy) | Pre-test calibration time (hh:mm) | Fuel code (A-Jet A, D-Diesel, F-Fuel Oil, X-Other) | 30-second average Temperature of Thermocouple #1 (oF) | 30-second average Temperature of Thermocouple #2 (oF) | 30-second average Temperature of Thermocouple #3 (oF) | 30-second average Temperature of Thermocouple #4 (oF) | 30-second average Temperature of Thermocouple #5 (oF) | 30-second average Temperature of Thermocouple #6 (oF) | 30-second average Temperature of Thermocouple #7 (oF) | Average of Temperatures 1-7 (oF) | 30-second Heat Flux Average (Btu/ft2 -sec) | Intake Air Velocity (2150 +/- 50 ft/min) | Intake Airflow (ft3/min) | Fuel Pressure (lb/in2) | Fuel Flowrate (6.0 +/- 0.2 gal/hr) | Intake Air Temperature Avg (oF) | Fuel Temperature Avg (oF) | Test Cell Temperature Avg (oF) | Relative Humidity (%) | Barometric Pressure (inches Hg) | Material Identification (A-F) | Run number (1-6) | Test Date (mm/dd/yy) | Test Time (hh:mm) | Failure Time (seconds) | Maximum Heat Flux Calorimeter 1 (Btu/ft2 -sec) | Maximum Heat Flux Time Calorimeter 1 (seconds) | Maximum Heat Flux Calorimeter 2 (Btu/ft2 -sec) | Maximum Heat Flux Time Calorimeter 2 (seconds) | Intake Air Temperature Avg (oF) | Fuel Temperature Avg (oF) | Test Cell Temperature Avg (oF) | Relative Humidity (%) | Barometric Pressure (inches Hg) | Intake Air Temp Change (oF) | Fuel Temp Change (oF) | Test Cell Temp Change (oF) | Relative Humidity Change (%) | Barometric Pressure Change (inches Hg)
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | A | 1 | 2002-02-13 | 12:28 | 30 |  |  |  |  | 64 | 77 | 126 | 22 | 30.23 | 0 | 1 | 9 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | B | 1 | 2002-02-13 | 12:40 | 53 |  |  |  |  | 65 | 74 | 114 | 22 | 30.23 | 1 | -2 | -3 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | C | 1 | 2002-02-13 | 12:51 | 158 |  |  |  |  | 65 | 71 | 115 | 22 | 30.23 | 1 | -5 | -2 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | D | 1 | 2002-02-13 | 13:05 | 290 |  |  |  |  | 66 | 71 | 122 | 22 | 30.23 | 2 | -5 | 5 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | E | 1 | 2002-02-13 | 13:21 | 360 | 0.28 | 345 | 0.42 | 340 | 67 | 72 | 122 | 22 | 30.23 | 3 | -4 | 5 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 12:10 | A | 1890.1883 | 2009.9873 | 2026.363 | 2028.809 | 2029.9377 | 1990.071 | 1906.5263 | 1983.1261 | 15.7598 | 2165 | 81.3661 | 98 | 5.88 | 64 | 76 | 117 | 22 | 30.23 | F | 1 | 2002-02-13 | 13:43 | 360 | 0.19 | 345 | 0.28 | 345 | 67 | 75 | 122 | 22 | 30.23 | 3 | -1 | 5 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | A | 2 | 2002-02-13 | 15:21 | 23 |  |  |  |  | 70 | 75 | 127 | 22 | 30.23 | 1 | -1 | 2 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | B | 2 | 2002-02-13 | 14:08 | 79 |  |  |  |  | 68 | 78 | 118 | 22 | 30.23 | -1 | 2 | -7 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | C | 2 | 2002-02-13 | 14:45 | 142 |  |  |  |  | 66 | 73 | 113 | 22 | 30.23 | -3 | -3 | -12 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | D | 2
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | E | 2 | 2002-02-13 | 14:55 | 360 | eqpt fail | eqpt fail | eqpt fail | eqpt fail | 67 | 69 | 118 | 22 | 30.23 | -2 | -7 | -7 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 13:50 | A | 1839.8607 | 2023.5807 | 2022.394 | 2044.4363 | 2037.2857 | 1962.7643 | 1953.3673 | 1983.3841 | 15.7728 | 2161 | 81.2158 | 98 | 5.88 | 69 | 76 | 125 | 22 | 30.23 | F | 2 | 2002-02-13 | 15:13 | 360 | 0.24 | 353 | 0.56 | 346 | 70 | 70 | 128 | 22 | 30.23 | 1 | -6 | 3 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | A | 3 | 2002-02-13 | 16:27 | 24 |  |  |  |  | 72 | 80 | 132 | 22 | 30.23 | 2 | 3 | -3 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | B | 3 | 2002-02-13 | 16:34 | 91 |  |  |  |  | 72 | 78 | 133 | 22 | 30.23 | 2 | 1 | -2 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | C | 3 | 2002-02-13 | 15:55 | 144 |  |  |  |  | 71 | 77 | 133 | 22 | 30.23 | 1 | 0 | -2 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | D | 3
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | E | 3 | 2002-02-13 | 16:08 | 360 | 0.22 | 347 | 0.4 | 348 | 71 | 77 | 128 | 22 | 30.23 | 1 | 0 | -7 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-13 | 15:33 | A | 1831.077 | 1995.164 | 2023.3347 | 2046.048 | 2031.784 | 1973.0897 | 1900.6167 | 1971.5877 | 15.4469 | 2160 | 81.1782 | 98 | 5.88 | 70 | 77 | 135 | 22 | 30.23 | F | 3 | 2002-02-13 | 16:21 | 360 | 0.35 | 355 | 0.27 | 343 | 72 | 77 | 134 | 22 | 30.23 | 2 | 0 | -1 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | A | 4 | 2002-02-14 | 8:27 | 29 |  |  |  |  | 66 | 71 | 125 | 96 | 30.26 | 4 | 10 | 16 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | B | 4 | 2002-02-14 | 8:35 | 82 |  |  |  |  | 65 | 69 | 118 | 96 | 30.26 | 3 | 8 | 9 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | C | 4 | 2002-02-14 | 8:46 | 141 |  |  |  |  | 64 | 68 | 111 | 96 | 30.26 | 2 | 7 | 2 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | D | 4
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | E | 4 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:08 | A | 1820.162 | 1976.699 | 2006.106 | 2029.549 | 2008.0217 | 1927.7953 | 1881.384 | 1949.9596 | 15.6984 | 2155 | 80.9903 | 98 | 5.88 | 62 | 61 | 109 | 96 | 30.26 | F | 4 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | A | 5 | 2002-02-14 | 9:03 | 35 |  |  |  |  | 63 | 68 | 109 | 96 | 30.26 | -1 | 0 | -4 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | B | 5 | 2002-02-14 | 9:13 | 11 |  |  |  |  | 63 | 68 | 113 | 96 | 30.26 | -1 | 0 | 0 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | C | 5 | 2002-02-14 | 9:23 | 156 |  |  |  |  | 64 | 69 | 108 | 96 | 30.26 | 0 | 1 | -5 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | D | 5 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | E | 5
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 8:57 | A | 1877.6573 | 1992.557 | 2019.5767 | 2034.4487 | 2026.315 | 1965.8953 | 1902.6553 | 1974.1579 | 15.9749 | 2170 | 81.554 | 98 | 5.88 | 64 | 68 | 113 | 96 | 30.26 | F | 5 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | A | 6 | 2002-02-14 | 9:57 | 33 |  |  |  |  | 65 | 71 | 117 | 96 | 30.26 | 0 | 0 | -18 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | B | 6 | 2002-02-14 | 10:04 | 103 |  |  |  |  | 65 | 71 | 116 | 96 | 30.26 | 0 | 0 | -19 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | C | 6 | 2002-02-14 | 10:13 | 161 |  |  |  |  | 65 | 70 | 110 | 96 | 30.26 | 0 | -1 | -25 | 0 | 0
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | D | 6 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | E | 6 | 2002-02-14 | [9 empty] | 96 | 30.26
A | P | 11 | 3.75 | M | X | O | 2002-02-14 | 9:51 | A | 1832.9183 | 1970.5817 | 2014.9133 | 2040.2283 | 2026.5563 | 1961.0963 | 1859.0443 | 1957.9055 | 15.9749 | 2166 | 81.4037 | 98 | 5.88 | 65 | 71 | 135 | 96 | 30.26 | F | 6
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.3285 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | A | 1 | 2002-02-12 | 8:46 | 22 | 1.53 | 22 | 0.95 | 22 | 74.5 | 86.1 | 69.8 | 28 | 30.05 | 2.8 | 6.1 | 0.1 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.33 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | B | 1 | 2002-02-12 | 9:09 | 51 | 0.75 | 51 | 0.09 | 51 | 72.6 | 70.5 | 70.4 | 28 | 30.01 | 0.9 | -9.5 | 0.7 | 0 | -0.04
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.33 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | C | 1 | 2002-02-12 | 9:35 | 80 | 0.34 | 80 | 0.17 | 80 | 74.7 | 61.8 | 71.1 | 28 | 29.99 | 3 | -18.2 | 1.4 | 0 | -0.06
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.33 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | D | 1 | 2002-02-12 | 9:59 | 215 | 0.41 | 215 | 0.11 | 215 | 81.5 | 76 | 71.7 | 27 | 29.99 | 9.8 | -4 | 2 | -1 | -0.06
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.33 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | E | 1 | 2002-02-12 | 10:26 | 203 | 1.03 | 203 | 0.29 | 203 | 81.1 | 76.1 | 72.4 | 27 | 29.97 | 9.4 | -3.9 | 2.7 | -1 | -0.08
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 8:40 | A | 1938 | 2014 | 1959 | 1890 | 1911 | 1981 | 1835 | 1932 | 14.72 | 2164 | 81.33 | 100 | 5.97 | 71.7 | 80 | 69.7 | 28 | 30.05 | F | 1 | 2002-02-12 | 10:59 | 135 | 0.84 | 135 | 0.21 | 135 | 77.2 | 68.9 | 73.2 | 27 | 29.95 | 5.5 | -11.1 | 3.5 | -1 | -0.1
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.103 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | B | 2 | 2002-02-12 | 12:38 | 54 | 1.22 | 54 | 0.22 | 54 | 79.9 | 78.8 | 74.9 | 27 | 29.95 | 7.2 | 18.2 | 0.3 | 0 | 0.06
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.1 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | C | 2 | 2002-02-12 | 13:01 | 74 | 0.17 | 74 | 0.24 | 74 | 76.6 | 72.5 | 75.4 | 27 | 29.87 | 3.9 | 11.9 | 0.8 | 0 | -0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.1 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | D | 2 | 2002-02-12 | 13:22 | 279 | 0.37 | 279 | 0.21 | 279 | 86.9 | 85.4 | 75.8 | 27 | 30.42 | 14.2 | 24.8 | 1.2 | 0 | 0.53
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.1 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | E | 2 | 2002-02-12 | 13:51 | 219 | 0.94 | 219 | 0.29 | 219 | 86.4 | 87.5 | 76.2 | 27 | 30.4 | 13.7 | 26.9 | 1.6 | 0 | 0.51
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.1 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | F | 2 | 2002-02-12 | 14:18 | 242 | 0.57 | 242 | 0.2 | 242 | 86 | 85.4 | 76.7 | 27 | 30.38 | 13.3 | 24.8 | 2.1 | 0 | 0.49
B | P | 3 | 3.75 | V | T | O | 2002-02-12 | 11:04 | A | 1965 | 1992 | 1930 | 1869 | 1898 | 1979 | 1845 | 1925 | 14.45 | 2158 | 81.1 | 100 | 5.97 | 72.7 | 60.6 | 74.6 | 27 | 29.89 | A | 2 | 2002-02-12 | 14:39 | 16 | 0.11 | 16 | 0.89 | 16 | 74.5 | 76.1 | 76.9 | 27 | 30.38 | 1.8 | 15.5 | 2.3 | 0 | 0.49
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.2909 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | C | 3 | 2002-02-13 | 7:54 | 59 | 0.24 | 59 | 0.13 | 59 | 77 | 73.6 | 71.4 | 27 | 30.28 | 10.2 | 23.8 | 0.1 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.29 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | D | 3
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.29 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | E | 3 | 2002-02-13 | 8:16 | 215 | 0.31 | 215 | 0.6 | 215 | 80.9 | 74.4 | 71.7 | 27 | 30.3 | 14.1 | 24.6 | 0.4 | 0 | 0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.29 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | F | 3 | 2002-02-13 | 8:40 | 232 | 0.99 | 232 | 0.45 | 232 | 80.2 | 75.4 | 72 | 27 | 30.3 | 13.4 | 25.6 | 0.7 | 0 | 0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.29 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | A | 3 | 2002-02-13 | 8:58 | 14 | 0.3 | 14 | 1.51 | 14 | 71.2 | 68.7 | 72.4 | 27 | 30.3 | 4.4 | 18.9 | 1.1 | 0 | 0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 7:38 | A | 1892 | 1978 | 1901 | 1817 | 1829 | 1934 | 1843 | 1885 | 14.16 | 2163 | 81.29 | 100 | 5.97 | 66.8 | 49.8 | 71.3 | 27 | 30.28 | B | 3 | 2002-02-13 | 9:14 | 69 | 0.43 | 69 | 0.96 | 69 | 74.4 | 66.5 | 72.7 | 28 | 30.32 | 7.6 | 16.7 | 1.4 | 1 | 0
B | P | 3 |  | V | T | O |  |  | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.28 | D | 4
B | P | 3 |  | V | T | O |  |  | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.28 | E | 4
B | P | 3 |  | V | T | O |  |  | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.28 | F | 4
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 9:28 | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.32 | A | 4 | 2002-02-13 | 9:47 | 19 | 0.66 | 19 | 0.71 | 19 | 73.8 | 70.7 | 73.4 | 28 | 30.3 | 1.9 | -0.4 | 0.4 | 0 | -0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 9:28 | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.32 | B | 4 | 2002-02-13 | 10:05 | 44 | 0.57 | 44 | 0.05 | 44 | 74.5 | 70.2 | 73.8 | 28 | 30.32 | 2.6 | -0.9 | 0.8 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 9:28 | A | 1920 | 2007 | 1947 | 1889 | 1901 | 1973 | 1838 | 1925 | 15.03 | 2155 | 80.99 | 100 | 5.97 | 71.9 | 71.1 | 73 | 28 | 30.32 | C | 4 | 2002-02-13 | 10:24 | 73 | 0.16 | 73 | 0.29 | 73 | 76.1 | 71 | 73.9 | 28 | 30.32 | 4.2 | -0.1 | 0.9 | 0 | 0
B | P | 3 |  | V | T | O |  |  | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 |  | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | E | 5
B | P | 3 |  | V | T | O |  |  | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 |  | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | F | 5
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 10:41 | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 | 1935 | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | A | 5 | 2002-02-13 | 10:54 | 26 | 0.48 | 26 | 0.5 | 26 | 75.6 | 80.7 | 73.9 | 28 | 30.3 | 1.2 | 2.7 | 0.1 | 0 | -0.02
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 10:41 | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 | 1935 | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | B | 5 | 2002-02-13 | 11:09 | 49 | 0.51 | 49 | 0.13 | 49 | 74.4 | 76.6 | 74 | 28 | 30.32 | 0 | -1.4 | 0.2 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 10:41 | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 | 1935 | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | C | 5 | 2002-02-13 | 11:28 | 84 | 0.36 | 84 | 0.17 | 84 | 77.9 | 75.6 | 74.1 | 28 | 30.32 | 3.5 | -2.4 | 0.3 | 0 | 0
B | P | 3 |  | V | T | O |  |  | A | 1973 | 2008 | 1954 | 1897 | 1924 | 1978 | 1816 |  | 15.1 | 2152 | 80.88 | 100 | 5.97 | 74.4 | 78 | 73.8 | 28 | 30.32 | D | 5
B | P |  |  | V | T | O |  |  | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.73 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | F | 6
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 12:28 | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.7272 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | A | 6 | 2002-02-13 | 12:43 | 15 | 0.16 | 15 | 0.34 | 15 | 74.5 | 74 | 74.8 | 28 | 30.3 | 1.6 | 9.5 | 0.2 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 12:28 | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.73 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | B | 6 | 2002-02-13 | 12:54 | 84 | 0.46 | 84 | 0.38 | 84 | 78.6 | 75.9 | 75 | 28 | 30.3 | 5.7 | 11.4 | 0.4 | 0 | 0
B | P | 3 | 3.75 | V | T | O | 2002-02-13 | 12:28 | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.73 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | C | 6 | 2002-02-13 | 13:13 | 71 | 0.29 | 71 | 0.14 | 71 | 77.6 | 73.4 | 75.3 | 28 | 30.3 | 4.7 | 8.9 | 0.7 | 0 | 0
B | P | 3 |  | V | T | O |  |  | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.73 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | D | 6
B | P | 3 |  | V | T | O |  |  | A | 1882 | 2010 | 1944 | 1871 | 1872 | 1956 | 1856 | 1913 | 14.52 | 2148 | 80.73 | 100 | 5.97 | 72.9 | 64.5 | 74.6 | 28 | 30.3 | E | 6
C | P | 10 | 3.75 | V | T | X | 02/20/2002 | 11:00 | A | 1807.2413 | 1972.5452 | 1993.9187 | 1970.3794 | 1948.7039 | 1946.3116 | 1885.2277 | 1932.0468 | 15.7 | 2162 | 81.2534 | 100 | 5.9 | 75 | 62 | 62 | 49 | 29.53 | A | 1 | 02/20/2002 | 13:45 | 28 |  |  |  |  | 72 | 64.9 | 69.8 | 48 | 29.53 | -3 | 2.9 | 7.8 | -1 | 0
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | B | 1 | 02/20/2002 | 14:05 | 135 |  |  |  |  | 71 | 64 | 71 | 55 | 29.53
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | C | 1 | 02/20/2002 | 14:40 | 201 |  |  |  |  | 70.9 | 63.9 | 68 | 55 | 29.53
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | D | 1 | 02/20/2002 | 15:00 | 360 | 0.24 | 360 | 0.21 | 360 | 71.2 | 65.7 | 72 | 55 | 29.53
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | E | 1 | 02/20/2002 | 15:25 | 360 | 0.3 | 360 | 0.28 | 360 | 71.2 | 66.6 | 74 | 53 | 29.53
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | F | 1 | 02/20/2002 | 15:50 | 214 |  |  |  |  | 71 | 65.1 | 72 | 53 | 29.53
C | P | 10 | 3.75 | V | T | X | 02/21/2002 | 9:30 | A | 1822 | 1932 | 1969 | 1957 | 1938 | 1901 | 1803 | 1903 | 15.3 | 2106 | 79.1487 | 100 | 5.9 | 69.8 | 65.5 | 69.8 | 45 | 29.57 | B | 2 | 02/21/2002 | 9:40 | 66 |  |  |  |  | 72 | 69.1 | 71.6 | 43 | 29.57
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | C | 2 | 02/21/2002 | 10:00 | 232 |  |  |  |  | 77.9 | 65.1 | 79.6 | 45 | 29.57
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | D | 2 | X | X | X
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | E | 2 | 02/21/2002 | 10:30 | 360 | 0.24 | 360 | 0.22 | 360 | 78.4 | 64.6 | 79.4 | 43 | 29.57
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | F | 2 | 02/21/2002 | 11:20 | 360 | 0.24 | 360 | 0.22 | 360 | 78.3 | 64.3 | 78.9 | 44 | 29.57
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | A | 2 | 02/21/2002 | 13:45 | 48 |  |  |  |  | 72.3 | 65 | 73.5 | 43 | 29.57
C | P | 10 | 3.75 | V | T | X | 02/22/2002 | 8:55 | A | 1839 | 1949 | 1978 | 1964 | 1948 | 1920 | 1813 | 1916 | 15.9 | 2204 | 82.8318 | 100 | 5.9 | 71 | 64 | 79 | 33 | 29.92 | C | 3 | 02/22/2002 | 8:55 | 179 |  |  |  |  | 79 | 65 | 82 | 33 | 29.92
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | D | 3 | X | X | X | [8 empty] | 29.92
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | E | 3 | 02/22/2002 | 9:30 | 360 | 0.16 | 360 | 0.17 | 360 | 76.5 | 65 | 76.5 | 33 | 29.92
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | F | 3 | 02/22/2002 | 9:55 | 360 | 0.16 | 360 | 0.14 | 360 | 76.8 | 61.9 | 77 | 36 | 29.92
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | A | 3 | 02/22/2002 | 10:25 | 60 |  |  |  |  | 72.4 | 63.7 | 74 | 31 | 29.92
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | B | 3 | 02/22/2002 | 10:40 | 200 |  |  |  |  | 75.3 | 62.8 | 76 | 33 | 29.92
C | P | 10 | 3.75 | V | T | X | 02/22/2002 | 13:45 | A | 1839 | 1953 | 1982 | 1966 | 1948 | 1926 | 1838 | 1922 | 15.9 | 2146 | 80.652 | 100 | 5.9 | 72 | 63 | 79 | 37 | 29.7 | D | 4 | X | X | X
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | E | 4
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | F | 4
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | A | 4 | 02/22/2002 | 13:55 | 1 |  |  |  |  | 77.3 | 66.4 | 79.7 | 32 | 29.7
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | B | 4 | 02/22/2002 | 14:15 | 90 |  |  |  |  | 77.8 | 65.7 | 80 | 36 | 29.7
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | C | 4 | 02/22/2002 | 14:45 | 234 |  |  |  |  | 79.8 | 63.9 | 78.3 | 39 | 29.7
C | P | 10 | 3.75 | V | T | X |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.2 | 2150 |  | 100 | 5.9 | 68 | 63 | 72 | 30 | 30.08 | E | 5
C | P | 10 | 3.75 | V | T | X | 2002-02-06 |  | A | [13 empty] | 68 | 63 | 72 | 30 | 30.08 | F | 5
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | A | 5 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | B | 5 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | C | 5 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | D | 5 | X | X | X
C | P | 10 | 3.75 | V | T | X |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.2 | 2150 |  | 100 | 5.9 | 74 | 69 | 75 | 30 | 29.99 | F | 6
C | P | 10 | 3.75 | V | T | X | 2002-02-06 |  | A | [13 empty] | 74 | 69 | 75 | 30 | 29.99 | A | 6 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | B | 6 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | C | 6 |  |  | 1
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | D | 6 | X | X | X
C | P | 10 | 3.75 | V | T | X |  |  | A | [18 empty] | E | 6
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 1 |  |  | 1 |  |  |  |  | 74.5 | 86.1 | 69.8 | 28 | 30.05 | 74.5 | 86.1 | 69.8 | 28 | 30.05
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 1 |  |  | 1 |  |  |  |  | 72.6 | 70.5 | 70.4 | 28 | 30.01 | 72.6 | 70.5 | 70.4 | 28 | 30.01
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 1 |  |  | 1 |  |  |  |  | 74.7 | 61.8 | 71.1 | 28 | 29.99 | 74.7 | 61.8 | 71.1 | 28 | 29.99
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | D | 1 |  |  | 1 |  |  |  |  | 81.5 | 76 | 71.7 | 27 | 29.99 | 81.5 | 76 | 71.7 | 27 | 29.99
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | E | 1 |  |  | 1 |  |  |  |  | 81.1 | 76.1 | 72.4 | 27 | 29.97 | 81.1 | 76.1 | 72.4 | 27 | 29.97
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | F | 1 |  |  | 1 |  |  |  |  | 77.2 | 68.9 | 73.2 | 27 | 29.95 | 77.2 | 68.9 | 73.2 | 27 | 29.95
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 2 |  |  | 1 |  |  |  |  | 79.9 | 78.8 | 74.9 | 27 | 29.95 | 79.9 | 78.8 | 74.9 | 27 | 29.95
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 2 |  |  | 1 |  |  |  |  | 76.6 | 72.5 | 75.4 | 27 | 29.87 | 76.6 | 72.5 | 75.4 | 27 | 29.87
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | D | 2 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | E | 2 |  |  | 1 |  |  |  |  | 86.4 | 87.5 | 76.2 | 27 | 30.4 | 86.4 | 87.5 | 76.2 | 27 | 30.4
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | F | 2 |  |  | 1 |  |  |  |  | 86 | 85.4 | 76.7 | 27 | 30.38 | 86 | 85.4 | 76.7 | 27 | 30.38
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 2 |  |  | 1 |  |  |  |  | 74.5 | 76.1 | 76.9 | 27 | 30.38 | 74.5 | 76.1 | 76.9 | 27 | 30.38
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 3 |  |  | 1 |  |  |  |  | 77 | 73.6 | 71.4 | 27 | 30.28 | 77 | 73.6 | 71.4 | 27 | 30.28
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | D | 3 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | E | 3 |  |  | 1 |  |  |  |  | 80.9 | 74.4 | 71.7 | 27 | 30.3 | 80.9 | 74.4 | 71.7 | 27 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | F | 3 |  |  | 1 |  |  |  |  | 80.2 | 75.4 | 72 | 27 | 30.3 | 80.2 | 75.4 | 72 | 27 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 3 |  |  | 1 |  |  |  |  | 71.2 | 68.7 | 72.4 | 27 | 30.3 | 71.2 | 68.7 | 72.4 | 27 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 3 |  |  | 1 |  |  |  |  | 74.4 | 66.5 | 72.7 | 28 | 30.32 | 74.4 | 66.5 | 72.7 | 28 | 30.32
E | P |  | 3.75 | T | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | D | 4 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | E | 4 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | F | 4 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 4 |  |  | 1 |  |  |  |  | 73.8 | 70.7 | 73.4 | 28 | 30.3 | 73.8 | 70.7 | 73.4 | 28 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 4 |  |  | 1 |  |  |  |  | 74.5 | 70.2 | 73.8 | 28 | 30.32 | 74.5 | 70.2 | 73.8 | 28 | 30.32
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 4 |  |  | 1 |  |  |  |  | 76.1 | 71 | 73.9 | 28 | 30.32 | 76.1 | 71 | 73.9 | 28 | 30.32
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | E | 5 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | F | 5 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 5 |  |  | 1 |  |  |  |  | 75.6 | 80.7 | 73.9 | 28 | 30.3 | 75.6 | 80.7 | 73.9 | 28 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 5 |  |  | 1 |  |  |  |  | 74.4 | 76.6 | 74 | 28 | 30.32 | 74.4 | 76.6 | 74 | 28 | 30.32
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 5 |  |  | 1 |  |  |  |  | 77.9 | 75.6 | 74.1 | 28 | 30.32 | 77.9 | 75.6 | 74.1 | 28 | 30.32
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | D | 5 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | F | 6 | [12 empty] | 0 | 0 | 0 | 0 | 0
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | A | 6 |  |  | 1 |  |  |  |  | 74.5 | 74 | 74.8 | 28 | 30.3 | 74.5 | 74 | 74.8 | 28 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | B | 6 |  |  | 1 |  |  |  |  | 78.6 | 75.9 | 75 | 28 | 30.3 | 78.6 | 75.9 | 75 | 28 | 30.3
E | P |  | 3.75 | V | T | O |  |  | A | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 |  |  | 100 | 5.97 | [5 empty] | C | 6 |  |  | 1 |  |  |  |  | 77.6 | 73.4 | 75.3 | 28 | 30.3 | 77.6 | 73.4 | 75.3 | 28 | 30.3
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | D | 6
E | P |  | 3.75 | V | T |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | [9 empty] | E | 6
F | P | 11 | 3.75 | V | T | O | 2002-02-12 | 9:30 | A | 1930 | 1916 | 1907 | 1905 | 1925 | 1950 | 1935 | 1924 | 15.317 | 2150 | 80.8024 | 92 | 5.96 | 53.24 | 56.26 | 49.53 | 28 | 30.08 | A | 1 | 2002-02-12 | 10:30 | 25 |  |  |  |  | 63 | 63 | 56 | 27 | 30.08 | 9.76 | 6.74 | 6.47 | -1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | 1930 | [6 empty] | 1924 | 15.32 | 2150 |  | 92 | 5.96 | 53.2 | 56.3 | 49.5 | 28 | 30.08 | B | 1 | 2002-02-12 | 11:00 | 159 |  |  |  |  | 65 | 66 | 57 | 27 | 30.08 | 11.8 | 9.7 | 7.5 | -1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | 1930 | [6 empty] | 1924 | 15.32 | 2150 |  | 92 | 5.96 | 53.2 | 56.3 | 49.5 | 28 | 30.08 | C | 1 | 2002-02-12 | 11:30 | 120 |  |  |  |  | 68 | 72 | 57 | 25 | 30.03 | 14.8 | 15.7 | 7.5 | -3 | -0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | 1930 | [6 empty] | 1924 | 15.32 | 2150 |  | 92 | 5.96 | 53.2 | 56.3 | 49.5 | 28 | 30.08 | D | 1 | 2002-02-12 | 12:00 | 360 |  |  |  |  | 66 | 74 | 57 | 25 | 30.03 | 12.8 | 17.7 | 7.5 | -3 | -0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | 1930 | [6 empty] | 1924 | 15.32 | 2150 |  | 92 | 5.96 | 53.2 | 56.3 | 49.5 | 28 | 30.08 | E | 1 | 2002-02-12 | 12:25 | 360 |  |  |  |  | 72 | 79 | 57 | 27 | 30.03 | 18.8 | 22.7 | 7.5 | -1 | -0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | 1930 | [6 empty] | 1924 | 15.32 | 2150 |  | 92 | 5.96 | 53.2 | 56.3 | 49.5 | 28 | 30.08 | F | 1 | 2002-02-12 | 12:50 | 360 |  |  |  |  | 74 | 82 | 57 | 27 | 30.03 | 20.8 | 25.7 | 7.5 | -1 | -0.05
F | P | 11 | 3.75 | V | T | O | 2002-02-12 | 14:00 | A | 1933 | 1929 | 1924 | 1921 | 1936 | 1958 | 1958 | 1937 | 15.4996 | 2150 | 80.8024 | 92 | 5.96 | 62.87 | 70.85 | 57.54 | 27 | 29.85 | B | 2 | 2002-02-12 | 14:50 | 140 |  |  |  |  | 69 | 69 | 60 | 27 | 29.85 | 6.13 | -1.85 | 2.46 | 0 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1937 | 15.5 | 2150 |  | 92 | 5.96 | 62.9 | 70.9 | 57.5 | 27 | 29.85 | C | 2 | 2002-02-12 | 15:20 | 110 |  |  |  |  | 72 | 75 | 61 | 27 | 29.85 | 9.1 | 4.1 | 3.5 | 0 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1937 | 15.5 | 2150 |  | 92 | 5.96 | 62.9 | 70.9 | 57.5 | 27 | 29.85 | D | 2 | 2002-02-12 | 15:45 | 360 |  |  |  |  | 72 | 77 | 60 | 28 | 29.85 | 9.1 | 6.1 | 2.5 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1937 | 15.5 | 2150 |  | 92 | 5.96 | 62.9 | 70.9 | 57.5 | 27 | 29.85 | E | 2 | 2002-02-12 | 16:10 | 360 |  |  |  |  | 77 | 82 | 62 | 28 | 29.85 | 14.1 | 11.1 | 4.5 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1937 | 15.5 | 2150 |  | 92 | 5.96 | 62.9 | 70.9 | 57.5 | 27 | 29.85 | F | 2 | 2002-02-12 | 16:40 | 360 |  |  |  |  | 75 | 83 | 59 | 28 | 29.85 | 12.1 | 12.1 | 1.5 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1937 | 15.5 | 2150 |  | 92 | 5.96 | 62.9 | 70.9 | 57.5 | 27 | 29.85 | A | 2 | 2002-02-12 | 17:00 | 38 |  |  |  |  | 72 | 88 | 60 | 29 | 29.85 | 9.1 | 17.1 | 2.5 | 2 | 0
F | P | 11 | 3.75 | V | T | O | 2002-02-13 | 9:00 | A | 1929 | 1919 | 1910 | 1909 | 1925 | 1950 | 1948 | 1927 | 15.227 | 2150 | 80.8024 | 92 | 5.96 | 55.36 | 56.98 | 50.36 | 33 | 30.07 | C | 3 | 2002-02-13 | 10:00 | 135 |  |  |  |  | 56 | 55 | 54 | 33 | 30.07 | 0.64 | -1.98 | 3.64 | 0 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1927 | 15.23 | 2150 |  | 92 | 5.96 |  |  |  | 33 |  | D | 3 | X | X | X
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1927 | 15.23 | 2150 |  | 92 | 5.96 | 55.4 | 57 | 50.4 | 33 | 30.07 | E | 3 | 2002-02-13 | 11:00 | 360 |  |  |  |  | 65 | 68 | 57 | 33 | 30.07 | 9.6 | 11 | 6.6 | 0 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1927 | 15.23 | 2150 |  | 92 | 5.96 | 55.4 | 57 | 50.4 | 33 | 30.07 | F | 3 | 2002-02-13 | 11:20 | 360 |  |  |  |  | 67 | 74 | 55 | 31 | 30.12 | 11.6 | 17 | 4.6 | -2 | 0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1927 | 15.23 | 2150 |  | 92 | 5.96 | 55.4 | 57 | 50.4 | 33 | 30.07 | A | 3 | 2002-02-13 | 11:45 | 31 |  |  |  |  | 67 | 80 | 57 | 31 | 30.12 | 11.6 | 23 | 6.6 | -2 | 0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1927 | 15.23 | 2150 |  | 92 | 5.96 | 55.4 | 57 | 50.4 | 33 | 30.07 | B | 3 | 2002-02-13 | 12:00 | 145 |  |  |  |  | 67 | 73 | 56 | 31 | 30.12 | 11.6 | 16 | 5.6 | -2 | 0.05
F | P | 11 | 3.75 | V | T | O |  |  | A | 1932 | 1928 | 1921 | 1917 | 1929 | 1952 | 1957 | 1934 | 15.47 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | D | 4 | X | X | X
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1934 | 15.47 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | E | 4
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1934 | 15.47 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | F | 4
F | P | 11 | 3.75 | V | T | O | 2002-02-13 | 13:00 | A | 1932 | 1928 | 1921 | 1917 | 1929 | 1952 | 1957 | 1934 | 15.468 | 2150 | 80.8024 | 92 | 5.96 | 60.58 | 65.47 | 56.64 | 20 | 30.13 | A | 4 | 2002-02-13 | 13:40 | 25 |  |  |  |  | 67 | 67 | 57 | 21 | 30.13 | 6.42 | 1.53 | 0.36 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1934 | 15.47 | 2150 |  | 92 | 5.96 | 60.6 | 65.5 | 56.6 | 20 | 30.13 | B | 4 | 2002-02-13 | 13:55 | 96 |  |  |  |  | 68 | 69 | 57 | 21 | 30.13 | 7.4 | 3.5 | 0.4 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1934 | 15.47 | 2150 |  | 92 | 5.96 | 60.6 | 65.5 | 56.6 | 20 | 30.13 | C | 4 | 2002-02-13 | 14:15 | 149 |  |  |  |  | 68 | 73 | 57 | 21 | 30.13 | 7.4 | 7.5 | 0.4 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | 1946 | 1949 | 1944 | 1937 | 1947 | 1969 | 1969 | 1952 | 15.86 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | E | 5
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1952 | 15.86 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | F | 5
F | P | 11 | 3.75 | V | T | O | 2002-02-13 | 14:30 | A | 1946 | 1949 | 1944 | 1937 | 1947 | 1969 | 1969 | 1952 | 15.86 | 2150 | 80.8024 | 92 | 5.96 | 63.63 | 70.66 | 55.68 | 20 | 30.13 | A | 5 | 2002-02-13 | 15:10 | 30 |  |  |  |  | 65 | 68 | 56 | 18 | 30.13 | 1.37 | -2.66 | 0.32 | -2 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1952 | 15.86 | 2150 |  | 92 | 5.96 | 63.6 | 70.7 | 55.7 | 20 | 30.13 | B | 5 | 2002-02-13 | 15:30 | 225 |  |  |  |  | 67 | 70 | 57 | 18 | 30.16 | 3.4 | -0.7 | 1.3 | -2 | 0.03
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1952 | 15.86 | 2150 |  | 92 | 5.96 | 63.6 | 70.7 | 55.7 | 20 | 30.13 | C | 5 | 2002-02-13 | 15:50 | 153 |  |  |  |  | 68 | 77 | 57 | 17 | 30.16 | 4.4 | 6.3 | 1.3 | -3 | 0.03
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1952 | 15.86 | 2150 |  | 92 | 5.96 |  |  |  | 20 |  | D | 5 | X | X | X
F | P | 11 | 3.75 | V | T | O |  |  | A | 1950 | 1942 | 1933 | 1929 | 1942 | 1967 | 1967 | 1947 | 15.76 | 2150 |  | 92 | 5.96 |  |  |  | 17 |  | F | 6
F | P | 11 | 3.75 | V | T | O | 2002-02-13 | 16:20 | A | 1950 | 1942 | 1933 | 1929 | 1942 | 1967 | 1967 | 1947 | 15.756 | 2150 | 80.8024 | 92 | 5.96 | 61.519 | 64.598 | 56.37 | 17 | 30.19 | A | 6 | 2002-02-13 | 16:45 | 26 |  |  |  |  | 64 | 68 | 55 | 18 | 30.19 | 2.481 | 3.402 | -1.37 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1947 | 15.76 | 2150 |  | 92 | 5.96 | 61.5 | 64.6 | 56.4 | 17 | 30.19 | B | 6 | 2002-02-13 | 17:00 | 148 |  |  |  |  | 66 | 70 | 55 | 18 | 30.19 | 4.5 | 5.4 | -1.4 | 1 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1947 | 15.76 | 2150 |  | 92 | 5.96 | 61.5 | 64.6 | 56.4 | 17 | 30.19 | C | 6 | 2002-02-13 | 17:25 | 321 |  |  |  |  | 67 | 75 | 56 | 17 | 30.19 | 5.5 | 10.4 | -0.4 | 0 | 0
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1947 | 15.76 | 2150 |  | 92 | 5.96 |  |  |  | 17 |  | D | 6 | X | X | X
F | P | 11 | 3.75 | V | T | O |  |  | A | [7 empty] | 1947 | 15.76 | 2150 |  | 92 | 5.96 |  |  |  | 17 |  | E | 6
G | P | 11 | 3.75 | M | T | O | 2002-02-14 | 8:36 | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.9298 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | A | 1 | 2002-02-14 | 9:35 | 143 |  |  |  |  | 63 | 67 | 66 | 53 | 29.84 | -4 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.93 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | B | 1 | 2002-02-14 | 9:57 | 210 |  |  |  |  | 63 | 67 | 66 | 53 | 29.84 | -4 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.93 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | C | 1 | 2002-02-14 | 10:13 | 360 | 0.76 | 359 | 0.4 | 356 | 66 | 67 | 66 | 53 | 29.84 | -1 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.93 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | D | 1 | 2002-02-14 | 10:33 | 360 | 0.16 | 351 | 0.12 | 346 | 66 | 67 | 66 | 53 | 29.84 | -1 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.93 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | E | 1 | 2002-02-14 | 11:07 | 360 | 0.47 | 360 | 0.39 | 360 | 66 | 67 | 66 | 53 | 29.84 | -1 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1751 | 1800 | 1892 | 1977 | 2015 | 1987 | 1871 | 1902 | 15.56 | 2180 | 81.93 | 104 | 6.09 | 67 | 67 | 67 | 54 | 29.84 | F | 1 | 2002-02-14 | 11:33 | 360 | 0.34 | 349 | 0.36 | 349 | 66 | 67 | 66 | 53 | 29.84 | -1 | 0 | -1 | -1 | 0
G | P | 11 | 3.75 | M | T | O | 2002-02-14 | 12:18 | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | B | 2 | 2002-02-14 | 12:56 | 200 |  |  |  |  | 66 | 67 | 66 | 48 | 29.84 | -1 | 0 | -1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | C | 2 | 2002-02-14 | 13:09 | 360 | 0.64 | 358 | 0.55 | 359 | 66 | 67 | 66 | 48 | 29.84 | -1 | 0 | -1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | D | 2 | X | X | X
G | P | 11 | 3.75 | M | T | O |  |  | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | E | 2 | 2002-02-14 | 13:31 | 360 | 0.37 | 353 | 0.37 | 350 | 66 | 67 | 66 | 48 | 29.84 | -1 | 0 | -1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | F | 2 | 2002-02-14 | 13:54 | 360 | 0.59 | 355 | 0.4 | 355 | 66 | 67 | 66 | 48 | 29.84 | -1 | 0 | -1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1994 | 1992 | 1889 | 1871 | 1922 | 1944 | 1857 | 1924 | 15.2 | 2200 | 82.6815 | 104 | 6.18 | 67 | 67 | 67 | 48 | 29.84 | A | 2 | 2002-02-14 | 14:17 | 100 |  |  |  |  | 66 | 67 | 66 | 48 | 29.84 | -1 | 0 | -1 | 0 | 0
G | P | 11 | 3.75 | M | T | O | 2002-02-15 | 7:18 | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | C | 3 | 2002-02-15 | 7:51 | 230 |  |  |  |  | 68 | 67 | 68 | 54 | 29.89 | 0 | 0 | 1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | D | 3 | X | X | X
G | P | 11 | 3.75 | M | T | O |  |  | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | E | 3 | 2002-02-15 | 8:11 | 340 |  |  |  |  | 68 | 67 | 68 | 54 | 29.89 | 0 | 0 | 1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | F | 3 | 2002-02-15 | [11 empty] | -68 | -67 | -67 | -54 | -29.89
G | P | 11 | 3.75 | M | T | O |  |  | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | A | 3 | 2002-02-15 | 8:29 | 80 |  |  |  |  | 68 | 67 | 68 | 54 | 29.89 | 0 | 0 | 1 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1837 | 1878 | 1995 | 2043 | 2016 | 1968 | 1799 | 1934 | 15.05 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 67 | 54 | 29.89 | B | 3 | 2002-02-15 | 8:40 | 230 |  |  |  |  | 68 | 67 | 68 | 54 | 29.89 | 0 | 0 | 1 | 0 | 0
G | P | 11 | 3.75 | M | T | O | 2002-02-15 | 8:56 | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | D | 4 | X | X | X
G | P | 11 | 3.75 | M | T | O |  |  | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | E | 4
G | P | 11 | 3.75 | M | T | O |  |  | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | F | 4
G | P | 11 | 3.75 | M | T | O |  |  | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | A | 4 | 2002-02-15 | 9:28 | 75 |  |  |  |  | 68 | 67 | 68 | 48 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | B | 4 | 2002-02-15 | 9:40 | 150 |  |  |  |  | 68 | 67 | 68 | 48 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1911 | 1920 | 1910 | 1915 | 1936 | 1956 | 1937 | 1927 | 15.57 | 2100 | 78.9232 | 104 | 6.14 | 68 | 67 | 68 | 48 | 29.89 | C | 4 | 2002-02-15 | 10:06 | 113 |  |  |  |  | 68 | 67 | 68 | 48 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O | 2002-02-15 | 10:30 | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | E | 5
G | P | 11 | 3.75 | M | T | O |  |  | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | F | 5
G | P | 11 | 3.75 | M | T | O |  |  | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | A | 5 | 2002-02-15 | 10:48 | 24 |  |  |  |  | 68 | 67 | 68 | 52 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | B | 5 | 2002-02-15 | 11:01 | 112 |  |  |  |  | 68 | 67 | 68 | 52 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | C | 5 | 2002-02-15 | 11:14 | 85 |  |  |  |  | 68 | 67 | 68 | 52 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1941 | 1974 | 1887 | 1845 | 1873 | 1919 | 1858 | 1900 | 15.49 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 52 | 29.89 | D | 5 | X | X | X
G | P | 11 | 3.75 | M | T | O | 2002-02-15 | 11:24 | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | F | 6
G | P | 11 | 3.75 | M | T | O |  |  | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | A | 6 | 2002-02-15 | 11:36 | 25 |  |  |  |  | 68 | 67 | 68 | 46 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | B | 6 | 2002-02-15 | 11:44 | 120 |  |  |  |  | 68 | 67 | 68 | 46 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | C | 6 | 2002-02-15 | 12:09 | 105 |  |  |  |  | 68 | 67 | 68 | 46 | 29.89 | 0 | 0 | 0 | 0 | 0
G | P | 11 | 3.75 | M | T | O |  |  | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | D | 6 | X | X | X
G | P | 11 | 3.75 | M | T | O |  |  | D | 1869 | 1962 | 1903 | 1839 | 1835 | 1889 | 1890 | 1884 | 15.83 | 2100 | 78.9232 | 104 | 6.1 | 68 | 67 | 68 | 46 | 29.89 | E | 6
H | P | 11 | 3.75 | V | T | O | 2002-02-19 | 9:16 | A | 1834 | 1909 | 1907 | 1890 | 1910 | 1936 | 1863 | 1923 | 15.5 | 2160 | 81.1782 | 104 |  | 66 | 87 | 326 | 45 |  | A | 1 | 2002-02-19 | 9:00 | 32 | 0 | ---------- | 0 | ------------- | 72 | 52 | 288 |  |  | 6 | -35 | -38 | #VALUE! | #VALUE!
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | B | 1 |  | 9:47 | 176 | 0.8 | 176 | 0.43 | 176 | 95 | 58 | 321 | 34 | 29.6
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | C | 1 |  | 10:13 | 154 | 0.28538 | 154 | 0.39 | 154 | 105 | 54 | 317
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | D | 1 |  | 10:32 | 360 | 0.29 | 360 | 0.29 | 360 | 155 | 58 | 347
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | E | 1 |  | 10:57 | 360 | 0.25 | 360 | 0.45 | 360 | 153 | 57 | 346
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | F | 1 |  | 11:29 | 360 | 0.26 | 360 | 0.76 | 360 | 157 | 57 | 346
H | P | 11 | 3.75 | V | T | O | 2002-02-19 | 11:49 | A | 1832.482 | 1907.0473 | 1905.4879 | 1891.312 | 1917.3957 | 1934.6903 | 1868.2053 | 1924 | 15.59 | 2160 | 81.1782 | 104 |  | 82 | 58 | 250 |  |  | B | 2 | 2002-02-19 | 12:00 | 67 | 0.1 | 67 | 0.1 | 67 | 84 | 54 | 318
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | C | 2 |  | 13:30 | 90 | 0.21 | 90 | 0.38 | 90 | 88 | 56 | 301
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | D | 2
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | E | 2 |  | 13:46 | 360 | 0.3 | 360 | 0.34 | 360 | 124 | 58 | 335
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | F | 2 |  | 14:16 | 360 | 0.3 | 360 | 0.57 | 360 | 145 | 57 | 350
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | A | 2 |  | 14:45 | 26 | 0 | ---------- | 0 | ------------- | 83 | 55 | 315
H | P | 11 | 3.75 | V | T | O | 2002-02-19 | 14:42 | A | 1830.9226 | 1906.3385 | 1903.3615 | 1882.8065 | 1908.7484 | 1938.5178 | 1895.1395 | 1925 | 15.8 | 2160 | 81.1782 | 104 |  | 77 | 54 | 210 |  |  | C | 3 | 2002-02-19 | 15:00 | 125 | 0.28 | 125 | 0.34 | 125 | 100 | 56 | 344
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | D | 3 | X | X | X
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | E | 3 |  | 15:25 | 360 | 0.28 | 360 | 0.4 | 360 | 149 | 60 | 357
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | F | 3 |  | 16:00 | 360 | 0.23 | 360 | 0.32 | 360 | 158 | 58 | 365
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | A | 3 |  | 16:45 | 20 | 0 | ------- | 0 | ----- | 90 | 55 | 323
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | B | 3 |  | 17:35 | 194 | 0.44 | 194 | 0.66 | 194 | 108 | 57 | 345
H | P | 11 | 3.75 | V | T | O |  |  | A | 1844 | 1912 | 1903 | 1889 | 1921 | 1956 | 1908 | 1927 | 15.9 | 2160 | 81.18 | 104 | 5.96 | 81 | 60 | 259 |  |  | D | 4 | X | X | X
H | P | 11 | 3.75 | V | T | O | 2002-02-19 | 18:00 | A | 1843.6809 | 1911.7253 | 1903.078 | 1889.0439 | 1920.6561 | 1955.9541 | 1908.3231 | 1927 | 15.9 | 2160 | 81.1782 | 104 | 5.96 | 81 | 60 | 259 |  |  | E | 4 | 2002-02-19
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | F | 4
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | A | 4 |  | 18:15 | 22 | 0 | ----- | 0 | ------- | 72 | 54 | 311
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | B | 4 |  |  | 83 | 0.12 | 83 | 0.23 | 83 | 85 | 54 | 323
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | C | 4 |  |  | 89 | 0.25 | 89 | 0.31 | 89 | 86 | 54 | 324 | 31.1
H | P | 11 | 3.75 | V | T | O |  |  | A | 1833 | 1910 | 1899 | 1896 | 1925 | 1939 | 1853 | 1925 | 15.2 | 2150 | 80.8 | 104 | 5.96 | 90 | 61 | 249 | 49 | 31 | E | 5
H | P | 11 | 3.75 | V | T | O | 2002-02-20 | 8:15 | A | 1832.8531 | 1909.6281 | 1899.0295 | 1896.4445 | 1924.6212 | 1939.4851 | 1853.0162 | 1925 | 15.2 | 2150 | 80.8024 | 104 | 5.96 | 90 | 61 | 249 | 49 | 31 | F | 5 | 2002-02-20
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | A | 5 |  | 8:30 | 16 | 0 | --- | 0 | ------ | 76 | 56 | 285 | 54 | 29.9
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | B | 5 |  | 8:45 | 133 | 0.21 | 133 | 0.65 | 133 | 98 | 57 | 312
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | C | 5 |  | 9:15 | 103 | 0.19 | 103 | 0.29 | 103 | 96 | 58 | 308
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | D | 5 | X | X | X
H | P | 11 | 3.75 | V | T | O |  |  | A | 1827 | 1900 | 1908 | 1889 | 1913 | 1940 | 1882 | 1927 | 15.57 | 2150 | 80.8 | 104 | 5.96 | 74 | 69 | 75 |  |  | F | 6
H | P | 11 | 3.75 | V | T | O | 2002-02-20 | 9:30 | A | 1827.1066 | 1900.3884 | 1907.9898 | 1889.3427 | 1913.2157 | 1939.7016 | 1882.454 | 1927 | 15.57 | 2150 | 80.8024 | 104 | 5.96 | 74 | 69 | 75 |  |  | A | 6 | 2002-02-20 | 9:40 | 18 | 0 | ----- | 0 | ------ | 79 | 56 | 296
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | B | 6 |  | 10:00 | 74 | 0.13 | 74 | 0.23 | 74 | 84 | 56 | 305
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | C | 6 |  | 10:30 | 183 | 0.17 | 183 | 0.25 | 183 | 83 | 57 | 305 | 49.8
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | D | 6 | X | X | X
H | P | 11 | 3.75 | V | T | O |  |  | A | [18 empty] | E | 6
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | A | 1 | 2002-02-08 | 9:00 | 25 | 0.041 | 28 | 0.969 | 28 | 62.6 | 60 | 62.1927 | 80 | 30.2 | 2.1675 | 0 | 4.5227 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | B | 1 | 2002-02-08 | 9:14 | 37 | 2.698 | 38 | 0.47 | 38 | 62.87 | 60 | 62.1927 | 80 | 30.2 | 2.4375 | 0 | 4.5227 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | C | 1 | 2002-02-08 | 9:32 | 87 | 0.017 | 92 | 0.128 | 88 | 62.8 | 60 | 62.1927 | 80 | 30.2 | 2.3675 | 0 | 4.5227 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | D | 1 | 2002-02-08 | 10:05 | 260 | 0.038 | 260 | 0.094 | 260 | 62.7713 | 60 | 62.1927 | 80 | 30.2 | 2.3388 | 0 | 4.5227 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | E | 4 | 2002-02-09 | 9:10 | 229 | 0.134 | 224 | X | 229 | 59.4634 | 60 | 58.3 | 77 | 30.25 | -0.96914 | 0 | 0.63 | -3 | 0.05
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 8:00 | A | 1837 | 1923 | 1859 | 1854 | 1916 | 1972 | 1882 | 1892 | 15.88 | 2120 | 79.6749 | 96 | 5.96 | 60.4325 | 60 | 57.67 | 80 | 30.2 | F | 1 | 2002-02-08 | 11:10 | 207 | 0.037 | 66 | 0.043 | 206 | 61.226 | 60 | 62.1927 | 80 | 30.2 | 0.7935 | 0 | 4.5227 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | B | 2 | 2002-02-08 | 12:12 | 66 | 0.377 | 66 | 0 | 66 | 61.5528 | 60 | 65.4 | 75 | 30.18 | 2.8171 | 0 | 9.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | C | 2 | 2002-02-08 | 12:23 | 115 | 0.083 | 109 | 0.073 | 115 | 61.1853 | 60 | 65.4 | 75 | 30.18 | 2.4496 | 0 | 9.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | D | 2 | X | X | X | X | X | X | X | X | X | X | X | X
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | E | 5 | 2002-02-09 | 12:13 | 224 | 0.048 | 224 | 0.29 | 224 | 59.8376 | 60 | 59.7 | 69 | 30.3 | 1.102 | 0 | 3.7 | -6 | 0.12
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | F | 2 | 2002-02-08 | 12:57 |  | 0.082 | 245 | 0.038 | 252 | 60.01 | 60 | 65.4 | 75 | 30.18 | 1.2744 | 0 | 9.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 12:00 | A | 1824 | 1912 | 1846 | 1852 | 1924 | 1982 | 1898 | 1891 | 15.78 | 2138 | 80.3514 | 96 | 5.96 | 58.7356 | 60 | 56 | 75 | 30.18 | A | 2 | 2002-02-08 | 13:22 | 36 | 1.947 | 36 | 0.287 | 36 | 60.57 | 60 | 65.4 | 75 | 30.18 | 1.8344 | 0 | 9.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | C | 3 | 2002-02-08 | 13:51 | 133 | 0.091 | 132 | 0.195 | 132 | 61.1624 | 60 | 61.1 | 78 | 30.18 | 0.062407 | 0 | 6.1 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | D | 3 | X | X | X | X | X | X | X | X | X | X | X | X
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | E | 6 | 2002-02-09 | 13:54 | 225 | 0.229 | 225 | 0.315 | 225 | 59.0709 | 60 | 59.9 | 67 | 30.28 | -2.0291 | 0 | 4.9 | -11 | 0.1
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | F | 3 | 2002-02-08 | 14:33 | 241 | 0.103 | 241 | 0.265 | 241 | 60.8818 | 60 | 61.1 | 78 | 30.18 | -0.21819 | 0 | 6.1 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | A | 3 | 2002-02-08 | 15:00 | 24 | 0.002 | 24 | 0.247 | 24 | 60.1466 | 60 | 61.1 | 78 | 30.18 | -0.95336 | 0 | 6.1 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-08 | 13:30 | A | 1850 | 1941 | 1898 | 1902 | 1962 | 1997 | 1920 | 1924 | 15.98 | 2137 | 80.3138 | 96 | 5.96 | 61.1 | 60 | 55 | 78 | 30.18 | B | 3 | 2002-02-08 | 15:13 | 67 | 0 | 67 | 0.251 | 67 | 59.6181 | 60 | 61.1 | 78 | 30.18 | -1.4819 | 0 | 6.1 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | D | 4 | X | X | X | X | X | X | X | X | X | X | X | X
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | E | 1 | 2002-02-08 | 10:37 |  | 0.121 | 260 | 0 | 260 | 61.8 | 60 | 62.1927 | 80 | 30.2 | 61.8 | 60 | 62.1927 | 80 | 30.2
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | F | 4 | 2002-02-09 | 9:23 | 233 | 0.003 | 233 | 0.119 | 233 | 59.7836 | 60 | 58.3 | 77 | 30.25 | 1.5527 | 0 | 3.8 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | A | 4 | 2002-02-09 | 9:45 | 34 | 0.074 | 34 | 0.209 | 34 | 59.739 | 60 | 58.3 | 77 | 30.25 | 1.5081 | 0 | 3.8 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | B | 4 | 2002-02-09 | 9:54 | 55 | 0.102 | 55 | 0.042 | 55 | 60.2275 | 60 | 58.3 | 77 | 30.25 | 1.9965 | 0 | 3.8 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 8:45 | A | 1845 | 1934 | 1860 | 1846 | 1912 | 1974 | 1927 | 1900 | 16.05 | 2119 | 79.6373 | 96 | 5.96 | 58.231 | 60 | 54.5 | 77 | 30.25 | C | 4 | 2002-02-09 | 10:06 | 107 | 0.289 | 75 | 0.349 | 107 | 59.6833 | 60 | 58.3 | 77 | 30.25 | 1.4523 | 0 | 3.8 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | E | 2 | 2002-02-08 | 12:38 |  | 0.073 | 253 | 0.043 | 252 | 60.76 | 60 | 65.4 | 75 | 30.18 | 60.76 | 60 | 65.4 | 75 | 30.18
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | F | 5 | 2002-02-09 | 10:37 |  | 0.125 | 254 | 0.037 | 236 | 60.2651 | 60 | 59.7 | 69 | 30.3 | 0.57468 | 0 | 3.7 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | A | 5 | 2002-02-09 | 11:05 | 25 | 0.142 | 25 | 0.68 | 25 | 59.9878 | 60 | 59.7 | 69 | 30.3 | 0.29736 | 0 | 3.7 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | B | 5 | 2002-02-09 | 11:17 | 47 | 0.203 | 47 | 0.002 | 47 | 59.8409 | 60 | 59.7 | 69 | 30.3 | 0.1505 | 0 | 3.7 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | C | 5 | 2002-02-09 | 11:38 | 140 | 0.111 | 140 | 0.271 | 140 | 59.5011 | 60 | 59.7 | 69 | 30.3 | -0.18931 | 0 | 3.7 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 10:15 | A | 1832 | 1938 | 1875 | 1855 | 1908 | 1967 | 1924 | 1900 | 15.62 | 2133 | 80.1635 | 96 | 5.96 | 59.6904 | 60 | 56 | 69 | 30.3 | D | 5 | X | X | X | X | X | X | X | X | X | X | X | X
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.939 | 60 | 56.5 | 67 | 30.28 | F | 6 | 2002-02-09 | 14:27 | 238 | 0.039 | 237 | 0.344 | 237 | 59.4325 | 60 | 59.9 | 67 | 30.28 | -0.50647 | 0 | 3.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.9 | 60 | 56.5 | 67 | 30.28 | A | 6 | 2002-02-09 | 12:48 | 22 | 0.003 | 22 | 0.118 | 22 | 59.3812 | 60 | 59.9 | 67 | 30.28 | -0.51876 | 0 | 3.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.939 | 60 | 56.5 | 67 | 30.28 | B | 6 | 2002-02-09 | 13:03 | 49 | 0.005 | 49 | 0.5 | 48 | 59.1715 | 60 | 59.9 | 67 | 30.28 | -0.76749 | 0 | 3.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.939 | 60 | 56.5 | 67 | 30.28 | C | 6 | 2002-02-09 | 13:25 | 114 | 0.192 | 114 | 0.47 | 113 | 59.3257 | 60 | 59.9 | 67 | 30.28 | -0.61329 | 0 | 3.4 | 0 | 0
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.939 | 60 | 56.5 | 67 | 30.28 | D | 6 | X | X | X | X | X | X | X | X | X | X | X | X
I | P | 11 | 3.75 | V | T | O | 2002-02-09 | 12:25 | A | 1846 | 1951 | 1889 | 1871 | 1930 | 1992 | 1945 | 1918 | 15.89 | 2120 | 79.6749 | 96 | 5.96 | 59.939 | 60 | 56.5 | 67 | 30.28 | E | 3 | 2002-02-08 | 14:08 |  | 0.093 | 245 | 0.463 | 245 | 60.9931 | 60 | 61.1 | 78 | 30.18 | 60.9931 | 60 | 61.1 | 78 | 30.18
J | P | 11 | 3.75 | V | X | O | 2002-02-14 | 9:30 | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 | 2130-2180 |  | 80.05-81.92 | 6.1 | 65 | 65 | 63 | 50 | 30.02 | A | 1 | 2002-02-14 | 10:00 | 29 | 0.26 | 29 | 0.24 | 30 | 70 | 68 | 75 | 45 | 30.02 | 5 | 3 | 12 | -5 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 |  |  |  | 6.1 | 65 | 65 | 63 | 50 | 30.02 | B | 1 |  |  | 80 | 0.3 | 80 | 0.25 | 80 | 68 | 68 | 75 | 46 | 30.02 | 3 | 3 | 12 | -4 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 |  |  |  | 6.1 | 65 | 65 | 63 | 50 | 30.02 | C | 1 |  |  | 137 | 0.21 | 137 | 0.16 | 141 | 68 | 68 | 76 | 45 | 30.02 | 3 | 3 | 13 | -5 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 |  |  |  | 6.1 | 65 | 65 | 63 | 50 | 30.02 | D | 1 |  |  | 349 | 0.2 | 349 | 0.25 | 349 | 68 | 68 | 75 | 45 | 30.02 | 3 | 3 | 12 | -5 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 |  |  |  | 6.1 | 65 | 65 | 63 | 50 | 30.02 | E | 1 |  |  | 350 | 0.27 | 350 | 0.23 | 240 | 69 | 69 | 75 | 45 | 30.02 | 4 | 4 | 12 | -5 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1882 | 1871 | 1817 | 1801 | 1828 | 1880 | 1895 | 1853 | 12.4 |  |  |  | 6.1 | 65 | 65 | 63 | 50 | 30.02 | F | 1 |  |  | 360 | 0.25 | 360 | 0.23 | 360 | 70 | 70 | 75 | 45 | 30.02 | 5 | 5 | 12 | -5 | 0
J | P | 11 | 3.75 | V | X | O | 2002-02-14 | 13:15 | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 | 2130-2180 |  | 80.05-81.92 | 6.1 | 70 | 68 | 70 | 51 | 30.01 | B | 2 | 2002-02-14 | 13:30 | 120 | 0.19 | 120 | 0.25 | 129 | 70 | 70 | 75 | 42 | 30.01 | 0 | 2 | 5 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 |  |  |  | 6.1 | 70 | 68 | 70 | 51 | 30.01 | C | 2 |  |  | 144 | 0.24 | 144 | 0.21 | 145 | 72 | 71 | 75 | 42 | 30.01 | 2 | 3 | 5 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 |  |  |  | 6.1 | 70 | 68 | 70 | 51 | 30.01 | D | 2 |  |  | X | X | X | X | X | X | X | X | X | X
J | P | 11 | 3.75 | V | X | O |  |  | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 |  |  |  | 6.1 | 70 | 68 | 70 | 51 | 30.01 | E | 2 |  |  | 300 | 0.25 | 300 | 0.15 | 300 | 72 | 72 | 75 | 42 | 30.01 | 2 | 4 | 5 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 |  |  |  | 6.1 | 70 | 68 | 70 | 51 | 30.01 | F | 2 |  |  | 348 | 0.46 | 348 | 0.29 | 348 | 72 | 72 | 75 | 44 | 30.01 | 2 | 4 | 5 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1887 | 1853 | 1818 | 1829 | 1889 | 1917 | 1871 | 1866 | 12.5 |  |  |  | 6.1 | 70 | 68 | 70 | 51 | 30.01 | A | 2 |  |  | 27 | 0.33 | 27 | 0.19 | 29 | 72 | 72 | 75 | 44 | 30.01 | 2 | 4 | 5 | -7 | 0
J | P | 11 | 3.75 | V | X | O | 2002-02-14 | 8:15 | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 | 2130-2180 |  | 80.05-81.92 | 6.1 | 65 | 65 | 65 | 57 | 30.02 | C | 3 | 2002-02-15 | 8:30 | 150 | 0.25 | 150 | 0.19 | 152 | 68 | 68 | 71 | 49 | 30.02 | 3 | 3 | 6 | -8 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 |  |  |  | 6.1 | 65 | 65 | 65 | 57 | 30.02 | D | 3 |  |  | X | X | X | X | X | X | X | X
J | P | 11 | 3.75 | V | X | O |  |  | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 |  |  |  | 6.1 | 65 | 65 | 65 | 57 | 30.02 | E | 3 |  |  | 360 | 0.23 | 360 | 0.23 | 360 | 70 | 70 | 73 | 48 | 30.02 | 5 | 5 | 8 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 |  |  |  | 6.1 | 65 | 65 | 65 | 57 | 30.02 | F | 3 |  |  | 360 | 0.16 | 240 | 0.23 | 360 | 70 | 70 | 75 | 48 | 30.02 | 5 | 5 | 10 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 |  |  |  | 6.1 | 65 | 65 | 65 | 57 | 30.02 | A | 3 |  |  | 27 | 0.26 | 28 | 0.26 | 27 | 71 | 71 | 75 | 48 | 30.02 | 6 | 6 | 10 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1881 | 1855 | 1818 | 1821 | 1875 | 1901 | 1880 | 1862 | 12.5 |  |  |  | 6.1 | 65 | 65 | 65 | 57 | 30.02 | B | 3 |  |  | 80 | 0.24 | 80 | 0.21 | 84 | 71 | 71 | 74 | 48 | 30.02 | 6 | 6 | 9 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 |  |  |  | 6.1 | 71 | 71 | 70 | 55 | 30.02 | D | 4 |  |  | X | X | X | X | X | X | X | X
J | P | 11 | 3.75 | V | X | O |  |  | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 |  |  |  | 6.1 | 71 | 71 | 70 | 55 | 30.02 | E | 4
J | P | 11 | 3.75 | V | X | O |  |  | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 |  |  |  | 6.1 | 71 | 71 | 70 | 55 | 30.02 | F | 4
J | P | 11 | 3.75 | V | X | O | 2002-02-15 | 10:45 | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 | 2130-2180 |  | 80.05-81.92 | 6.1 | 71 | 71 | 70 | 55 | 30.02 | A | 4 | 2002-02-15 | 11:00 | 27 | 0.3 | 27 | 0.19 | 29 | 72 | 72 | 75 | 48 | 30.02 | 1 | 1 | 5 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 |  |  |  | 6.1 | 71 | 71 | 70 | 55 | 30.02 | B | 4 |  |  | 102 | 0.21 | 102 | 0.19 | 106 | 72 | 72 | 75 | 48 | 30.02 | 1 | 1 | 5 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1879 | 1860 | 1815 | 1801 | 1867 | 1887 | 1882 | 1856 | 12.4 |  |  |  | 6.1 | 71 | 71 | 70 | 55 | 30.02 | C | 4 |  |  | 134 | 0.19 | 134 | 0.19 | 135 | 72 | 72 | 75 | 48 | 30.02 | 1 | 1 | 5 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 56 | 30.01 | E | 5
J | P | 11 | 3.75 | V | X | O |  |  | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 56 | 30.01 | F | 5
J | P | 11 | 3.75 | V | X | O | 2002-02-15 | 1:45 | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 | 2130-2180 |  | 80.05-81.92 | 6.1 | 71 | 71 | 72 | 56 | 30.01 | A | 5 | 2002-02-15 | 2:00 | 29 | 0.29 | 29 | 0.15 | 30 | 71 | 71 | 75 | 47 | 30.01 | 0 | 0 | 3 | -9 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 56 | 30.01 | B | 5 |  |  | 120 | 0.35 | 120 | 0.28 | 120 | 72 | 72 | 75 | 46 | 30.01 | 1 | 1 | 3 | -10 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 56 | 30.01 | C | 5 |  |  | 148 | 0.21 | 148 | 0.28 | 148 | 72 | 72 | 75 | 46 | 30.01 | 1 | 1 | 3 | -10 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1880 | 1870 | 1815 | 1804 | 1840 | 1901 | 1890 | 1857 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 56 | 30.01 | D | 5 |  |  | X | X | X | X | X | X | X | X
J | P | 11 | 3.75 | V | X | O |  |  | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 51 | 30.01 | F | 6
J | P | 11 | 3.75 | V | X | O | 2002-02-15 | 2:45 | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 | 2130-2180 |  | 80.05-81.92 | 6.1 | 71 | 71 | 72 | 51 | 30.01 | A | 6 | 2002-02-15 | 3:00 | 27 | 0.19 | 27 | 0.26 | 28 | 71 | 71 | 75 | 44 | 30.01 | 0 | 0 | 3 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 51 | 30.01 | B | 6 |  |  | 83 | 0.26 | 83 | 0.23 | 87 | 71 | 71 | 76 | 44 | 30.01 | 0 | 0 | 4 | -7 | 0
J | P | 11 | 3.75 | V | X | O |  |  | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 51 | 30.01 | C | 6 |  |  | 157 | 0.21 | 157 | 0.2 | 168 | 70 | 70 | 75 | 44 | 30.01 | -1 | -1 | 3 | -7 | 0
J |  |  | 3.75 | V | X | O |  |  | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 51 | 30.01 | D | 6 |  |  | X | X | X | X | X | X | X | X
J |  |  | 3.75 | V | X | O |  |  | A | 1870 | 1860 | 1814 | 1810 | 1834 | 1890 | 1876 | 1851 | 12.4 |  |  |  | 6.1 | 71 | 71 | 72 | 51 | 30.01 | E | 6