Question: Is medical outcomes audit data collected during facility inspections used in a national database?

Answer: No.  While medical outcomes audit data is reviewed during the inspection to assure the facility is complying with the regulations, it is not routinely collected by FDA.  FDA does not maintain nor does it supply information to such a national database.

Question:  What type of lay summary should be sent to a patient who has a normal mammogram, but an abnormal physical exam?

Answer:  The facility has a great deal of flexibility in the format and substance of the lay summary.  The lay summary must inform the patient of the results of the mammographic examination and follow-up actions, if needed.  In the example given, the lay summary should say that the mammogram was normal, but the patient should consult with her/his physician because of the abnormal physical exam.

Question: Must the compression paddle be placed in the x-ray beam during half value layer (HVL) measurements?

Answer: If the HVL measurement is being used to calculate patient dose, the compression paddle must be placed in the x-ray beam during the measurement.  If the HVL measurement is being used for purposes other than patient dose, the position of the paddle is not specifically stated in the regulations.  However, FDA recommends inclusion of the compression paddle in the x-ray beam because most manufacturers have designed their equipment to be tested in this configuration.

Question: I have a Masters/Bachelor’s degree specifically in physics.  Do I still need to document my number of semester hours in physics?

Answer: If the degree is stated to be specifically in physics (or any of its sub-specialties), it will routinely be accepted during inspections as adequate documentation of the required number of semester hours in physics.  Therefore, during routine inspections, you will not need to provide detailed transcripts to all the facilities where you provide physics services.

If the degree was stated to be in the more general term, physical science, or in one of the three non-physics categories of physical science (chemistry, radiation science, or engineering), then you will have to provide additional documentation demonstrating compliance with the appropriate requirement for semester hours in physics.

Question: What are the minimum tests and/or review that the medical physicist must perform for a facility survey, survey of a mammography unit, equipment evaluation of a unit or processor that has been installed or disassembled and reassembled, and an equipment evaluation of a unit or processor that has undergone a major repair?

Answer:

| Facility Survey | All tests as described in 900.12(e)(2), (e)(5), and if applicable (e)(6). Evaluate adequacy of the results of all tests conducted by the facility in accordance with 900.12(e)(1) through (e)(7) |
| --- | --- |
| Survey of a mammography unit | All tests as described in 900.12(e)(2), (e)(5), and if applicable (e)(6) |
| Equipment evaluation of a unit or processor that has been installed or disassembled and reassembled | All applicable tests and equipment requirements described in 900.12(b) and (e) |
| Equipment evaluation of a unit or processor that has undergone a major repair | Only those tests and equipment requirements described in 900.12(b) and (e) that are applicable to the repaired component of the unit or processor. The decision as to what constitutes an applicable test for the repaired component should be made by the medical physicist. |

Question: While performing a physics survey, unit survey, or equipment evaluation, the medical physicist determines that the unit fails a required test(s) and needs to be adjusted or repaired.  After the adjustment or repair has been completed, does the test(s) have to be repeated and must the medical physicist perform the test(s)?

Answer: After the adjustment or repair has been completed, the test(s) have to be repeated and must show that the unit is within the appropriate action limit before it can be used on patients.  As to who must perform the test(s), if the adjustment or repair constitutes a “major repair,” then the medical physicist must repeat the test(s) as part of the required follow-up equipment evaluation.  If the adjustment or repair does not constitute a “major repair,” then the test(s) can be performed by any qualified person (e.g. radiologic technologist or repairperson with the appropriate training/experience).  However, in this circumstance, FDA recommends, at a minimum, that the medical physicist be consulted during this process.

Question: Our facility has permanently glued the acrylic contrast disk to the phantom at the center of the cover plate.  The regulation states that the OD be measured at the center of the phantom which would require our repositioning the disk (which may leave permanent artifacts) or replacement of the phantom cover.  Is there another way to leave the disk in place and still fulfill the requirement?

Answer: For the purpose of making the density measurements, FDA interprets the “center” of the phantom to be equivalent to measurements taken on the center line as described later in this paragraph.  As an alternative to repositioning the disk or replacement of the phantom cover, you may choose to keep the disk in place.  Facilities may measure the background density at a different location along the center line of the phantom in the direction parallel to the chest wall, as long this location does not intercept any of the phantom objects.  In addition, you must measure the background density at the same location each time the test is conducted.

In the process of removing the disk, permanent defects may be produced in the cover plate.  Such defects may produce permanent “artifacts” in the phantom image.  As long as they do not obscure the test objects, facilities can continue to use the phantom.  The facility should not subtract these permanent “artifacts” from the phantom score.

Question:  We use the same locum tenens interpreting physician on a recurring basis throughout the year.  Do we need to handle the personnel records for this person differently from our “permanent” interpreting physicians?  Which of his/her records will be reviewed at the time of our next inspection?

Answer:  If the locum tenens interpreting physician reads for the facility on a recurring basis throughout the year, the facility can treat this person as “permanent” for purposes of the annual inspection.  In that event, the facility has to show that this person met all the requirements (including the continuing requirements, when appropriate) on the date of the inspection.

If the facility considered this person “temporary” and not working at the facility at the time of the inspection, the facility has to show that this person met all the requirements (including the continuing requirements, when appropriate) on each of the multiple occasions this person started reading at the facility.  For example, if the person worked for a week at a time starting the 15th day of each month, the facility is responsible for documenting that the person was qualified on each of those 12 occasions.  This is not as daunting a task as it first appears.  The facility has to document that the person was fully qualified the first time he/she started working at the facility.  In addition, in each of the calendar quarters the person worked, the facility has to document (before letting the person provide mammography services) that the person met the continuing requirements (when appropriate) and maintained a valid medical license.

Question:  We use the same “temporary” radiologic technologist on a recurring basis throughout the year.  Do we need to handle the personnel records for this person differently from our “permanent” radiologic technologists?  Which of his/her records will be reviewed at the time of our next inspection?

Answer:  If the “temporary” radiologic technologist performs mammograms for the facility on a recurring basis throughout the year, the facility can treat this person as “permanent” for purposes of the annual inspection.  In that event, the facility has to show that this person met all the requirements (including the continuing requirements, when appropriate) on the date of the inspection.

If the facility considered this person “temporary” and not working at the facility at the time of the inspection, the facility has to show that this person met all the requirements (including the continuing requirements, when appropriate) on each of the multiple occasions this person started reading at the facility.  For example, if the person worked for a week at a time starting the 15th day of each month, the facility is responsible for documenting that the person was qualified on each of those 12 occasions.  This is not as daunting a task as it first appears.  The facility has to document that the person was fully qualified the first time he/she started working at the facility.  In addition, in each of the calendar quarters the person worked, the facility has to document (before letting the person provide mammography services) that the person met the continuing requirements (when appropriate) and maintained a valid State license or certification.

Question:  Under what situations should facilities establish new processor operating levels?

Answer:  The most warranted and common situations for a facility to establish processor operating levels are when processor QC testing is initiated for a new processor or when a significant change is made in the processing system, i.e., different film brand/type, different chemical brand/type, or a change in processing conditions (standard vs. extended).

Facilities should not use the establishment of new operating levels to correct problems in the processing system, but should troubleshoot and solve the problem with appropriate corrective action.

Question: During the time a facility is establishing new operating levels (typically done by performing a five-day data plot average): A) does the facility continue to plot the data on the processor chart?  B) Is the facility exempt from having to stay within the old processor action limits during the five-day averaging period?

Answer:  While establishing new operating levels (during which time the facility can continue to process mammograms), the facility must continue to perform the daily processor QC tests and should plot the data in the same manner it usually does.  This may be done on the same graph as the previous data or on a different graph.  In either event, this new data should be clearly identified as being derived during the establishment of the new operating levels, so that both the facility and the inspector are aware of the origins of this data.  Because no operating level has yet been established, the facility is exempt from having to stay within any processor action limits during this five-day averaging period.

Question:  What constitutes an equipment evaluation (what tests must the medical physicist perform) for a processor that has undergone major repairs or is a new processor to the facility?

Answer:  At a minimum, the following tests must be done for a processor that has been replaced, undergone major repairs or is a new processor to the facility: processor testing as described in 900.12(e)(1), phantom testing as described in 900.12(e)(2), and system artifact evaluation as described in 900.12(e)(5)(ix).  If the major repairs or the use of the new processor necessitates a change in clinical technique factors (for the standard breast), that could significantly increase patient dose, a determination of dose as described in 900.12(e)(5)(vi) must be done.

Question:  Must the equipment evaluation report be sent to the facility within 30 days?

Answer: The regulations do not specify when the equipment evaluation report must be sent to the facility.  However, the facility cannot use the equipment until it has documentation (equipment evaluation report) showing that the equipment passes all the appropriate tests.

Question:  What assessment category should be used for post lumpectomy patients, whose mammograms are otherwise negative?

Answer: In the case described, the results could be categorized as; “benign.”  In addition, qualifying statements such as “post surgical changes noted” may be added.  The regulations require that each mammography report contain one of the six overall assessment categories.  The decision as to which category to assign is left to the interpreting physician.

Question: Our group practice interprets mammograms sent to us by other facilities under a contractual arrangement. This is the only service that we provide in the mammography area. Does my group practice need an FDA certificate to interpret mammograms?

Answer: No. The basic requirement established by MQSA is that any facility that produces, processes, or interprets mammograms after October 1, 1994, must have an MQSA certificate to continue to operate lawfully. Generally, however, where procedures such as processing and interpretation of the films are performed in a location different from where the mammography is performed, the facility performing the mammography will be responsible for obtaining a certificate.

There is no mechanism at the present time for accrediting and certifying an organization such as yours. FDA's policy is that partial providers, that is, groups such as yours that provide only part of the services required for mammography, will be certified as part of a "system" for producing, processing, and interpreting mammograms.

The provider of some component of that system that performs mammography must take the lead in obtaining a certificate. FDA expects that the owner of an X-ray unit or units will apply for accreditation and receive a certificate. The application for accreditation must show that all components of the system used to produce, process, and interpret mammograms meet the MQSA requirements. If one or more of the facilities for which you interpret mammograms is applying for accreditation and certification, your responsibility will be to provide them with the information necessary to prove that the physicians in your group meet the MQSA quality standards requirements for interpreting physicians.

When the facilities performing mammography receive their certificate, you will be included for the purpose of interpreting mammograms for that facility. If you interpret mammograms for several facilities, your group will be included under several certificates. Conversely, if one of the facilities for which you interpret mammograms is not certified, then it would be unlawful for them to continue to produce mammograms. Although, as mentioned, the owner of the X-ray unit(s) will probably take the lead in obtaining the certificate in most cases, any partial provider in a "system" can seek the required certificate.

Your group practice could apply for accreditation and receive the certificate, as a part of a system. In that case, your practice would be responsible for assuring not only that your practice meets the MQSA quality standards for interpreters, but also that the facilities that produce and process the mammograms for your interpretation meet the quality standards that apply to them. Your practice would also be responsible for passing the review of clinical images from each facility that produces images for your interpretation, and for meeting the other requirements for accreditation. Finally, as a certificate holder, you would be responsible for paying an annual FDA MQSA inspection fee.

Question:  The regulations at 900.12(e)(5)(i) require that an x-ray unit pass an annual test for AEC performance over a range of 2 to 6 cm thick absorbers.  If a unit is used clinically at combinations of kVp and filtration that include tissue thickness outside the 2 to 6 cm range, must it meet the AEC performance requirements at the thickness’ where it operates and must it be tested at those technique factors under the annual quality control requirements?

Answer:  No, the unit is not required to meet the AEC performance specification outside the 2 to 6 cm range and the physicist is not required to test the AEC performance requirements for thickness' outside this range during the annual survey.  However, we strongly recommend that in addition to the required testing in the 2 to 6 cm range, the unit also be tested at all clinically used thickness' outside this range and that the action limits specified in the regulations be applied to the extended test.

You should note that under the Equipment Evaluation outlined at 900.12(e)(10), an evaluation of the AEC under all conditions of use is required, not merely recommended.  This is because 900.12(e)(10) mandates conformance with all pertinent aspects of 900.12(b) and (e).  Under 900.12(b)((10) the AEC is required to be “operable” under “configurations provided.”   In this context, “operable” means that it must meet the density and reproducibility requirements of (e)(5)(i) under all clinical use situations occurring at the facility.

Question:  The Bucky assembly is being replaced on our x-ray unit.  Is it necessary that the medical physicist test and pass this prior to its use on patients?

Answer:  If the installation does not involve the replacement of the AEC detector(s) on the system, the exchange of Bucky assemblies would not be considered a major repair and would therefore not require the medical physicist to evaluate the assembly prior to use on patients.  If however, the replacement does include the AEC detector, the medical physicist must evaluate the assembly prior to use on patients.  There are tests that should be performed when the bucky is replaced (table artifacts, grid artifacts, and x-ray field/image receptor alignment), but these can be done under the direction of the physicist and do not require his/her evaluation prior to the next annual survey.

Question:  We are a mobile facility with a van that does not have on-board processing.  We have a film-changing room on the van where we load and unload cassettes with film during the day.  When the van returns to our main office, we batch process the films.  Do we have to perform a semiannual test for darkroom fog in this film-changing room?

Answer:  Yes.  Since you use a room to load and unload cassettes with film, you must test this room for darkroom fog as well as the darkroom for your processor.  Both rooms have the potential to fog films and degrade image quality.

Question: I have my ARRT (M) certificate.  Will this certificate be sufficient documentation to show that I have received adequate training in breast anatomy and physiology, positioning and compression, quality assurance/quality control techniques and imaging of patients with breast implants as specifically spelled out in 900.12(a)(2)(ii)(A)?

Answer: Yes.  The ARRT (M) certificate will be accepted as documenting that the areas of training in breast anatomy and physiology, positioning and compression, quality assurance/quality control techniques and imaging of patients with breast implants have been covered.

# Required QC Tests* for Facilities Using Multiple Units & Screen-Film Combinations

(11/22/99)

| Test | Units Tested | S-F Combinations Tested With Each Unit |
| --- | --- | --- |
| Focal Spot Condition a) System Resolution OR b) Focal Spot Dimensions (up to 10/28/2002) | All units** | For a), all S-F combinations clinically used with the unit in question. For b), one S-F combination or direct exposure film. |
| Phantom Image | All units used to image the standard*** breast | All S-F combinations clinically used for the standard breast |
| Phantom Image | All units that are used only for non-standard breasts, and or magnification work | One S-F combination, using clinical techniques that would be used for the standard breast |
| Dose | All units used to image the standard breast | All S-F combinations clinically used for the standard breast with their corresponding techniques |
| Dose | All units that are used only for non-standard breasts, and or magnification work | One S-F combination, using clinical techniques that would be used for the standard breast |
| Darkroom Fog | Any unit (one only) | All film types clinically used |
| Uniformity of Screen Speed Screen-Film Contact | Any unit (one only) | All clinically used screens (cassettes), one film type |
| AEC Performance - Reproducibility | All units | One S-F combination (typically used with the unit) |
| AEC Performance – kVp & Thickness Tracking | All units used clinically in the 2-6 cm thickness range | One S-F combination (typically used with the unit) |
| Collimation System Artifacts | All units | One S-F combination (typically used with the unit) |
| KVp Accuracy & Reproducibility Decompression Radiation Output | All units | None |

* This table does not cover other required QC tests that are independent of the x-ray unit or screen-film combination, e.g., the daily processor QC, the fixer retention, and repeat analysis.

** In this table, “All units” refers to those that are used clinically.

*** The standard breast referenced in this table is defined in the regulations as “a 4.2 centimeter (cm) thick compressed breast consisting of 50 percent glandular and 50 percent adipose tissue.”

Verification Testing for Items that Fail QC Tests

Question: For repairs made in response to a failure of one > of the QC tests, which of the following items on the list below would be considered a > “major repair”, and hence must be evaluated by the medical physicist following the repair and prior to clinical use of the equipment?

Answer: The extent to which the medical physicist must be involved in the following repairs is summarized in the table below.

#### Item						Extent of Involvement by MP

> Compression device adjustment:

> a. pressure							Not required*

b. thickness scale inaccuracy				Not required

> c. repair of auto decompression aspect.			Not required

>

> Compression paddle adjustment or replacement:

> a. for deflection problem					Not required

b. for extension beyond allowable limit			Oversight**

> >

> Collimation adjustment:					>

> extension beyond allowable limit				Oversight

>

> AEC Adjustment:

a. thickness compensation					Conduct evaluation in person

b. image mode tracking					Conduct evaluation in person

> c. reproducibility						Conduct evaluation in person

> d. bring down dose or decrease output			Conduct evaluation in person

>

> Artifacts caused by filter:>

> a. filter replacement					Conduct evaluation in person			>

> Processor artifacts:>

> a. roller replacement					Not required

b. replenishment adjustment				Not required

> >

> KvP adjustment						Conduct evaluation in person >

> Addition or subtraction of filtration (HVL problem)	Conduct evaluation in person

*not a major repair

**the facility must, at a minimum, consult with the medical physicist to determine whether this represents a major repair for the facility’s specific unit, and how the evaluation is to be performed