7.0	RELIABILITY OF THE 3T3 AND NHK NRU TEST METHODS	7-3

7.1	Substances Used to Determine the Reliability of the 3T3 and NHK 
NRU Test Methods	7-4

7.2	Reproducibility Analyses for the 3T3 and NHK NRU Test Methods	7-6

7.2.1	ANOVA Results for the 3T3 and NHK NRU Test Methods	7-6

7.2.2	CV Results for the 3T3 and NHK NRU Test Methods	7-17

7.2.3	Comparison of Laboratory-Specific Linear Regression Analyses for the Prediction of In Vivo Rodent LD50 Values from In Vitro NRU IC50 Values		7-24

7.2.4	Laboratory Concordance for the Prediction of GHS Acute Oral Toxicity Category	7-24

7.3	Historical Positive Control Data	7-27

7.3.1	ANOVA and Linear Regression Results for the 3T3 NRU 
Test Method	7-30

7.3.2	ANOVA and Linear Regression Results for the NHK NRU 
Test Method	7-32

7.4	Laboratory Concordance for Solvent Selection	7-34

7.5	Summary	7-37

[This Page Intentionally Left Blank]

7.0	RELIABILITY OF THE 3T3 AND NHK NRU TEST METHODS

This section discusses the reliability of the 3T3 and NHK NRU test methods.  Reliability is the degree to which a test method can be performed reproducibly within and among laboratories over time (ICCVAM 2003).  It is assessed by calculating intra- and inter-laboratory reproducibility and repeatability.  Reproducibility is the consistency of individual test results obtained in a single laboratory (intralaboratory reproducibility) or in different laboratories (interlaboratory reproducibility) using the same protocol and test samples.  Repeatability, usually applied to results within a laboratory, is the closeness of agreement between test results obtained within a single laboratory when the procedure is performed on the same substance under identical conditions within a given time.  The NICEATM/ECVAM study was not designed to assess intralaboratory repeatability.

For the NICEATM/ECVAM validation study, reliability was assessed by determining both intra- and inter-laboratory reproducibility.  Intralaboratory reproducibility is the agreement of results produced when qualified people within the same laboratory perform the test method using the same test protocol at different times (ICCVAM 2003).  Interlaboratory reproducibility is the agreement of results from different qualified laboratories using the same protocol and reference substances.  Interlaboratory reproducibility indicates the extent to which a test method can be successfully transferred among laboratories.

Intra- and inter-laboratory reproducibility of the 3T3 and NHK NRU test methods were determined using ANOVA and CV analysis as discussed in Section 5.3.3 (see Sections 7.2.1 and 7.2.2).  Interlaboratory reproducibility of the 3T3 and NHK NRU test methods was also assessed by comparing the laboratory-specific IC50-LD50 regressions (from Table 6-1) to one another for each test method (see Section 7.2.3) and by evaluating laboratory concordance for the GHS acute oral toxicity category predictions reported in Sections 6.3.1 through 6.3.3 (see Section 7.2.4).  Laboratory concordance for the solvent selection process using the solubility protocol (described in Section 2.9) is provided in Section 7.4.

# 7.1	Substances Used to Determine the Reliability of the 3T3 and NHK NRU Test Methods

The SMT intended to use the IC50 results of all 72 reference substances identified for testing in Table 3-2 to determine the reliability of the 3T3 and NHK NRU test methods.  Unfortunately, IC50 results for all substances could not be obtained in all the laboratories.  Table 7-1 shows the substances that failed to yield sufficient cytotoxicity for the calculation of an IC50 and the number of substances left to determine intralaboratory reproducibility.  The laboratories failed to obtain IC50 results for three to five substances in the 3T3 NRU test method and two to three substances with the NHK NRU test method.

For the 3T3 NRU test method, no laboratory achieved sufficient cytotoxicity to obtain IC50 values for carbon tetrachloride or methanol and only one laboratory obtained IC50 results for lithium carbonate and xylene.  Thus, interlaboratory reproducibility for the 3T3 NRU test method was assessed using the remaining 68 reference substances.  For the NHK NRU test method, no laboratory obtained IC50 values for carbon tetrachloride and only one laboratory achieved IC50 results for xylene and 1,1,1-trichloroethane.  Interlaboratory reproducibility for the NHK NRU test method was assessed using the IC50 results for the remaining 69 reference substances.

Despite the fact that IC50 values were not obtained by all the laboratories for all reference substances, Table 7-2 shows that the complete range of LD50 responses, as defined by the GHS classification for acute oral toxicity in Table 3-1, was covered by the remaining substances.  The IC50 values also covered a wide range of responses (see Table 7-3).  IC50 values for the 3T3 NRU test method ranged from 0.005 to 38,878 (g/mL.  IC50 values for the NHK NRU test method covered a larger range, from 0.00005 to 49,800 (g/mL.

Table 7-1	Reference Substances That Failed to Yield IC50 Values1 And Number of Reference Substances Available for Intralaboratory Reproducibility Analyses

| Laboratory | 3T3 NRU Test Method | NHK NRU Test Method |
| --- | --- | --- |
|  | Reference Substances Lacking IC50 Results | N2 | Reference Substances Lacking IC50 Results | N2 |
| ECBC | Carbon tetrachloride Methanol Xylene | 69 | Carbon tetrachloride Methanol Xylene | 69 |
| FAL | Carbon tetrachloride Gibberellic acid Lithium carbonate Methanol Xylene | 67 | 1,1,1-Trichloroethane Carbon tetrachloride Xylene | 69 |
| IIVS | Carbon tetrachloride Lithium carbonate Methanol | 69 | 1,1,1-Trichloroethane Carbon tetrachloride | 70 |

1Due to insufficient cytotoxicity.

2Number of substances available for intralaboratory reproducibility analyses.

Table 7-2	Number of Reference Substances Tested vs Number of Reference Substances Yielding IC50 Values in Each GHS Toxicity Category1 for Two Sets of LD50 Values

| GHS Category1 (LD50 in mg/kg) | Initial Oral LD502 | Reference Oral LD503 | Results from 3T3 NRU Test Method | Results from NHK NRU Test Method |
| --- | --- | --- | --- | --- |
|  |  |  | Initial Oral LD502 | Reference Oral LD503 | Initial Oral LD502 | Reference Oral LD503 |
| LD50 ≤ 5 | 12 | 7 | 12 | 7 | 12 | 7 |
| 5 < LD50 ≤ 50 | 12 | 12 | 12 | 12 | 12 | 12 |
| 50 < LD50 ≤ 300 | 12 | 12 | 12 | 12 | 12 | 12 |
| 300 < LD50 ≤ 2000 | 12 | 16 | 11 | 15 | 12 | 16 |
| 2000 < LD50 ≤ 5000 | 12 | 12 | 10 | 10 | 10 | 10 |
| LD50 > 5000 | 12 | 13 | 11 | 12 | 11 | 12 |

1GHS-Globally Harmonized System of Classification and Labelling of Chemicals (UN 2005).

2Number of reference substances that yielded an IC50 value in at least one laboratory based on initial oral LD50 in Table 3-2.  Initial oral LD50 values, used during the reference substance selection process, were those used by the Registry of Cytotoxicity (RC) (from 1983/84 RTECS®) when applicable.  The RC is a database of acute oral LD50 values for rats and mice obtained from RTECS® and IC50 values from in vitro cytotoxicity assays using multiple cell lines and cytotoxicity endpoints for chemicals with known molecular weights (Halle 1998).

Values for reference substances not included in the RC came from HSDB or RTECS®.

3Number of reference substances that yielded an IC50 value in at least one laboratory based on reference oral LD50 in Table 4-2.  Reference oral LD50 values from rats and mice were derived after evaluating LD50 values located through literature searches and references from toxicity databases such as RTECS®.

7.2	Reproducibility Analyses for the 3T3 and NHK NRU Test Methods

Reproducibility of the 3T3 and NHK NRU test methods were performed using ANOVA and CV as described in Section 5.3.3.  Table 7-3 reports the results of these analyses for each reference substance and test method.

# 7.2.1	ANOVA Results for the 3T3 and NHK NRU Test Methods

ANOVA was performed as discussed in Section 5.3.3.  Since the sample sizes from this study were small, usually three observations per laboratory, the ANOVA results may be misleading.  There may be some differences that are statistically significant only because there are too few observations within the laboratories to adequately characterize the variability, and/or the within-laboratory variability estimate is small.

Differences Among the Laboratories for the 3T3 NRU Test Method

The ANOVA results in Table 7-3 indicate that there were statistically significant (p < 0.01) differences among the laboratories for 26 reference substances.  These chemicals are listed in Table 7-4 along with columns showing the laboratory statistically significantly differing from the other two laboratories (as indicated by the contrast results).  Since significant laboratory differences may be produced by insolubility or volatility, Table 7-4 also indicates whether any laboratory reported insolubility or volatility during conduct of the test.  Insolubility was suggested by the presence of precipitates in either the stock solutions or in cell culture.  Volatility was identified by the use of plate sealers to contain volatile contamination of lower concentration wells by higher concentrations.  Insolubility and volatility were reported for only nine of the 26 chemicals.

| Table 7-3 Reproducibility Results for the 3T3 and NHK NRU Test Methods |
| --- |
| Reference Substance/Laboratory | 3T3 NRU Test Method | NHK NRU Test Method |
|  | Arithmetic Mean IC50 (g/mL)1 | Arithmetic IntraLab %CV | Arithmetic InterLab %CV | Log Arithmetic Mean IC50 (g/mL)1 | ANOVA P2 | Contrast P3 | Arithmetic Mean IC50 (g/mL)1 | Arithmetic IntraLab %CV | Arithmetic InterLab %CV | Log Arithmetic Mean IC50 (g/mL)1 | ANOVA P2 | Contrast P3 |
| Acetaminophen | 50.1 |  | 28 | 1.70 | 0.171 |  | 526 |  | 13 | 2.72 | 0.181 |  |
| ECBC | 40.8 | 22 |  | 1.61 |  | NA | 558 | 15 |  | 2.75 |  | NA |
| FAL | 66.2 | 35 |  | 1.82 |  | NA | 447 | 19 |  | 2.65 |  | NA |
| IIVS | 43.4 | 26 |  | 1.64 |  | NA | 571 | 14 |  | 2.76 |  | NA |
| Acetonitrile | 8484 |  | 21 | 3.93 | 0.553 |  | 10104 |  | 8 | 4.00 | 0.9641 |  |
| ECBC | 6433 | 2 |  | 3.81 |  | NA | 10868 | 72 |  | 4.04 |  | NA |
| FAL | 9690 | 58 |  | 3.99 |  | NA | 10153 | 19 |  | 4.01 |  | NA |
| IIVS | 9330 | 13 |  | 3.97 |  | NA | 9290 | 4 |  | 3.97 |  | NA |
| Acetylsalicylic acid | 760 |  | 56 | 2.88 | <0.001 |  | 613 |  | 15 | 2.79 | 0.060 |  |
| ECBC | 646 | 10 |  | 2.81 |  | 0.581 | 631 | 3 |  | 2.80 |  | NA |
| FAL | 1234 | 24 |  | 3.09 |  | <0.001 | 694 | 14 |  | 2.84 |  | NA |
| IIVS | 401 | 16 |  | 2.60 |  | <0.001 | 514 | 15 |  | 2.71 |  | NA |
| 5-Aminosalicylic acid | 1698 |  | 19 | 3.23 | 0.054 |  | 52.3 |  | 47 | 1.72 | 0.044 |  |
| ECBC | 1467 | 14 |  | 3.17 |  | 0.092 | 29.9 | 22 |  | 1.48 |  | 0.025 |
| FAL | 2070 | 16 |  | 3.32 |  | 0.021 | 78.2 | 54 |  | 1.89 |  | 0.033 |
| IIVS | 1557 | 12 |  | 3.19 |  | 0.312 | 48.8 | 16 |  | 1.69 |  | 0.832 |
| Aminopterin | 0.007 |  | 54 | -2.14 | 0.036 |  | 682 |  | 27 | 2.83 | 0.0250 |  |
| ECBC | 0.005 | 20 |  | -2.28 |  | 0.216 | 889 | 20 |  | 2.95 |  | 0.017 |
| FAL | 0.012 | 46 |  | -1.93 |  | 0.013 | 545 | 8 |  | 2.74 |  | 0.041 |
| IIVS | 0.005 | 23 |  | -2.33 |  | 0.079 | 611 | 12 |  | 2.79 |  | 0.345 |
| Amitriptyline HCl | 7.23 |  | 14 | 0.86 | 0.348 |  | 9.76 |  | 19 | 0.99 | 0.365 |  |
| ECBC | 6.03 | 23 |  | 0.78 |  | 0.163 | 10.8 | 31 |  | 1.03 |  | NA |
| FAL | 7.86 | 28 |  | 0.90 |  | 0.469 | 7.57 | 72 |  | 0.88 |  | NA |
| IIVS | 7.81 | 18 |  | 0.89 |  | 0.445 | 10.9 | 10 |  | 1.04 |  | NA |
| Arsenic trioxide | 2.51 |  | 61 | 0.40 | 0.004 |  | 10.4 |  | 91 | 1.02 | <0.001 |  |
| ECBC | 2.41 | 33 |  | 0.38 |  | 0.527 | 7.77 | 33 |  | 0.89 |  | 0.694 |
| FAL | 1.04 | 7 |  | 0.02 |  | 0.002 | 2.55 | 75 |  | 0.41 |  | <0.001 |
| IIVS | 4.09 | 52 |  | 0.61 |  | 0.006 | 20.9 | 31 |  | 1.32 |  | 0.0006 |
| Atropine sulfate | 85.6 |  | 49 | 1.93 | 0.049 |  | 91.9 |  | 13 | 1.96 | 0.9881 |  |
| ECBC | 54.1 | 55 |  | 1.73 |  | 0.046 | 85.4 | 12 |  | 1.93 |  | 0.8903 |
| FAL | 133 | 31 |  | 2.12 |  | 0.024 | 104 | 85 |  | 2.02 |  | 0.9069 |
| IIVS | 70.0 | 8 |  | 1.85 |  | 0.641 | 83.2 | 25 |  | 1.92 |  | 0.9832 |
| Boric acid | 2228 |  | 69 | 3.35 | 0.010 |  | 473 |  | 8 | 2.67 | 0.9306 |  |
| ECBC | 1497 | 32 |  | 3.18 |  | 0.189 | 440 | 31 |  | 2.64 |  | 0.9692 |
| FAL | 3987 | 17 |  | 3.60 |  | 0.004 | 517 | 73 |  | 2.71 |  | 0.7391 |
| IIVS | 1202 | 48 |  | 3.08 |  | 0.021 | 464 | 2 |  | 2.67 |  | 0.7680 |
| Busulfan | 135 |  | 119 | 2.13 | 0.002 |  | 278 |  | 11 | 2.44 | 0.659 |  |
| ECBC | 40.0 | 48 |  | 1.60 |  | 0.012 | 253 | 27 |  | 2.40 |  | NA |
| FAL | 321 | 56 |  | 2.51 |  | < 0.001 | 268 | 72 |  | 2.43 |  | NA |
| IIVS | 43.7 | 4 |  | 1.64 |  | 0.033 | 313 | 12 |  | 2.50 |  | NA |
| Cadmium chloride | 0.565 |  | 39 | -0.25 | 0.124 |  | 1.98 |  | 10 | 0.30 | 0.733 |  |
| ECBC | 0.480 | 14 |  | -0.32 |  | NA | 2.20 | 37 |  | 0.34 |  | NA |
| FAL | 0.400 | 32 |  | -0.40 |  | NA | 1.88 | 65 |  | 0.27 |  | NA |
| IIVS | 0.817 | 53 |  | -0.09 |  | NA | 1.86 | 8 |  | 0.27 |  | NA |
| Caffeine | 161 |  | 18 | 2.21 | 0.481 |  | 661 |  | 21 | 2.82 | 0.296 |  |
| ECBC | 133 | 10 |  | 2.12 |  | NA | 817 | 31 |  | 2.91 |  | NA |
| FAL | 157 | 52 |  | 2.20 |  | NA | 591 | 32 |  | 2.77 |  | NA |
| IIVS | 191 | 7.5 |  | 2.28 |  | NA | 574 | 1 |  | 2.76 |  | NA |
| Carbamazepine | 109 |  | 35 | 2.04 | 0.049 |  | 128 |  | 85 | 2.11 | 0.432 |  |
| ECBC | 83.0 | 14 |  | 1.92 |  | NA | 66.1 | 13 |  | 1.82 |  | NA |
| FAL | 152 | 37 |  | 2.18 |  | NA | 253 | 129 |  | 2.40 |  | NA |
| IIVS | 91.8 | 12 |  | 1.96 |  | NA | 63.9 | 8 |  | 1.81 |  | NA |
| Carbon tetrachloride | NA |  | NA | NA | NA |  | NA |  | NA | NA | NA |  |
| ECBC | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| FAL | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| IIVS | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| Chloral hydrate | 187 |  | 25 | 2.27 | 0.004 |  | 137 |  | 17 | 2.14 | 0.302 |  |
| ECBC | 151 | 10 |  | 2.18 |  | 0.008 | 140 | 24 |  | 2.15 |  | NA |
| FAL | 241 | 10 |  | 2.38 |  | 0.002 | 159 | 32 |  | 2.20 |  | NA |
| IIVS | 170 | 12 |  | 2.23 |  | 0.181 | 112 | 2 |  | 2.05 |  | NA |
| Chloramphenicol | 161 |  | 67 | 2.21 | <0.001 |  | 366 |  | 13 | 2.56 | 0.750 |  |
| ECBC | 55.3 | 22 |  | 1.74 |  | <0.001 | 318 | 45 |  | 2.50 |  | NA |
| FAL | 273 | 30 |  | 2.44 |  | 0.001 | 414 | 44 |  | 2.62 |  | NA |
| IIVS | 156 | 18 |  | 2.19 |  | 0.165 | 367 | 22 |  | 2.56 |  | NA |
| Citric acid | 829 |  | 41 | 2.92 | 0.002 |  | 424 |  | 25 | 2.63 | 0.006 |  |
| ECBC | 473 | 29 |  | 2.68 |  | 0.001 | 526 | 16 |  | 2.72 |  | 0.009 |
| FAL | 1148 | 13 |  | 3.06 |  | 0.003 | 312 | 17 |  | 2.49 |  | 0.002 |
| IIVS | 865 | 19 |  | 2.94 |  | 0.298 | 433 | 5 |  | 2.64 |  | 0.483 |
| Colchicine | 0.047 |  | 85 | -1.33 | 0.001 |  | 0.007 |  | 22 | -2.16 | 0.174 |  |
| ECBC | 0.020 | 11 |  | -1.70 |  | 0.0028 | 0.005 | 46 |  | -2.28 |  | NA |
| FAL | 0.093 | 45 |  | -1.03 |  | 0.0005 | 0.008 | 10 |  | -2.12 |  | NA |
| IIVS | 0.028 | 1 |  | -1.55 |  | 0.0914 | 0.008 | 21 |  | -2.09 |  | NA |
| Cupric sulfate pentahydrate | 70.6 |  | 85 | 1.85 | <0.001 |  | 197 |  | 4 | 2.29 | 0.374 |  |
| ECBC | 82.7 | 4 |  | 1.92 |  | 0.001 | 190 | 10 |  | 2.28 |  | NA |
| FAL | 123 | 44 |  | 2.09 |  | <0.001 | 195 | 6 |  | 2.29 |  | NA |
| IIVS | 5.70 | 31 |  | 0.76 |  | <0.001 | 207 | 3 |  | 2.32 |  | NA |
| Cycloheximide | 0.293 |  | 104 | -0.53 | 0.021 |  | 0.082 |  | 43 | -1.09 | 0.302 |  |
| ECBC | 0.125 | 45 |  | -0.90 |  | 0.118 | 0.053 | 22 |  | -1.28 |  | NA |
| FAL | 0.647 | 70 |  | -0.19 |  | 0.007 | 0.120 | 78 |  | -0.92 |  | NA |
| IIVS | 0.109 | 23 |  | -0.96 |  | 0.076 | 0.071 | 19 |  | -1.15 |  | NA |
| Dibutyl phthalate | 78.3 |  | 124 | 1.89 | < 0.001 |  | 32.6 |  | 41 | 1.51 | 0.408 |  |
| ECBC | 23.5 | 17 |  | 1.37 |  | 0.012 | 28.3 | 27 |  | 1.45 |  | NA |
| FAL | 191 | 50 |  | 2.28 |  | <0.001 | 47.4 | 73 |  | 1.68 |  | NA |
| IIVS | 20.7 | 7 |  | 1.32 |  | 0.005 | 22.0 | 6 |  | 1.34 |  | NA |
| Dichlorvos | 20.3 |  | 57 | 1.31 | 0.002 |  | 11.1 |  | 20 | 1.05 | 0.181 |  |
| ECBC | 9.80 | 35 |  | 0.99 |  | 0.001 | 8.56 | 27 |  | 0.93 |  | NA |
| FAL | 32.8 | 6 |  | 1.52 |  | 0.002 | 12.4 | 30 |  | 1.09 |  | NA |
| IIVS | 18.3 | 11 |  | 1.26 |  | 0.823 | 12.2 | 3 |  | 1.09 |  | NA |
| Diethyl phthalate | 113 |  | 28 | 2.05 | 0.127 |  | 145 |  | 44 | 2.16 | 0.049 |  |
| ECBC | 85.5 | 34 |  | 1.93 |  | 0.092 | 174 | 8 |  | 2.24 |  | 0.196 |
| FAL | 147 | 26 |  | 2.17 |  | 0.070 | 71.5 | 94 |  | 1.85 |  | 0.018 |
| IIVS | 106 | 24 |  | 2.03 |  | 0.846 | 189 | 18 |  | 2.28 |  | 0.127 |
| Digoxin | 520 |  | 62 | 2.72 | 0.043 |  | 0.00314 |  | 88 | -2.50 | <0.001 |  |
| ECBC | 351 | 39 |  | 2.54 |  | 0.167 | 0.00538 | 13 |  | -2.27 |  | <0.001 |
| FAL | 892 | 36 |  | 2.95 |  | 0.017 | 0.00005 | 36 |  | -4.29 |  | <0.001 |
| IIVS | 317 | 21 |  | 2.50 |  | 0.144 | 0.00398 | 7 |  | -2.40 |  | <0.001 |
| Dimethylformamide | 5242 |  | 6 | 3.72 | 0.296 |  | 7856 |  | 19 | 3.90 | <0.001 |  |
| ECBC | 5343 | 10 |  | 3.73 |  | NA | 9353 | 2 |  | 3.97 |  | <0.001 |
| FAL | 5483 | 9 |  | 3.74 |  | NA | 7817 | 1 |  | 3.89 |  | 0.508 |
| IIVS | 4900 | 4 |  | 3.69 |  | NA | 6397 | 3 |  | 3.81 |  | <0.001 |
| Diquat dibromide monohydrate | 15.1 |  | 120 | 1.18 | 0.017 |  | 4.73 |  | 37 | 0.67 | 0.217 |  |
| ECBC | 3.90 | 23 |  | 0.59 |  | 0.040 | 3.59 | 23 |  | 0.56 |  | NA |
| FAL | 36.1 | 98 |  | 1.56 |  | 0.006 | 6.77 | 55 |  | 0.83 |  | NA |
| IIVS | 5.40 | 25 |  | 0.73 |  | 0.190 | 3.84 | 8 |  | 0.58 |  | NA |
| Disulfoton | 98.6 |  | 55 | 1.99 | 0.003 |  | 378 |  | 99 | 2.58 | <0.001 |  |
| ECBC | 137 | 55 |  | 2.14 |  | NA | 140 | 19 |  | 2.15 |  | 0.002 |
| FAL | NA | NA |  | NA |  | NA | 808 | 26 |  | 2.91 |  | <0.001 |
| IIVS | 60.4 | 87 |  | 1.78 |  | NA | 186 | 32 |  | 2.27 |  | 0.018 |
| Endosulfan | 8.02 |  | 78 | 0.90 | 0.046 |  | 2.35 |  | 43 | 0.37 | 0.029 |  |
| ECBC | 5.30 | 57 |  | 0.72 |  | 0.447 | 3.44 | 17 |  | 0.54 |  | 0.020 |
| FAL | 15.2 | 78 |  | 1.18 |  | 0.018 | 1.42 | 50 |  | 0.15 |  | 0.018 |
| IIVS | 3.60 | 42 |  | 0.56 |  | 0.080 | 2.19 | 20 |  | 0.34 |  | 0.927 |
| Epinephrine bitartrate | 59.4 |  | 12 | 1.77 | 0.048 |  | 90.6 |  | 24 | 1.96 | 0.119 |  |
| ECBC | 51.5 | 12 |  | 1.71 |  | 0.018 | 115 | 9 |  | 2.06 |  | NA |
| FAL | 63.4 | 11 |  | 1.80 |  | 0.165 | 81.7 | 35 |  | 1.91 |  | NA |
| IIVS | 63.4 | 3 |  | 1.80 |  | 0.149 | 75.0 | 16 |  | 1.88 |  | NA |
| Ethanol | 6731 |  | 23 | 3.83 | 0.075 |  | 10184 |  | 18 | 4.01 | 0.035 |  |
| ECBC | 5360 | 33 |  | 3.73 |  | NA | 8290 | 5 |  | 3.92 |  | 0.019 |
| FAL | 8420 | 14 |  | 3.93 |  | NA | 12013 | 19 |  | 4.08 |  | 0.029 |
| IIVS | 6413 | 5 |  | 3.81 |  | NA | 10250 | 9 |  | 4.01 |  | 0.752 |
| Ethylene glycol | 25292 |  | 26 | 4.40 | 0.007 |  | 42600 |  | 15 | 4.63 | 0.063 |  |
| ECBC | 18325 | 9 |  | 4.26 |  | 0.004 | 38000 | 12 |  | 4.58 |  | NA |
| FAL | 31650 | 24 |  | 4.50 |  | 0.010 | 49800 | 9 |  | 4.70 |  | NA |
| IIVS | 25900 | 12 |  | 4.41 |  | 0.505 | 40000 | 13 |  | 4.60 |  | NA |
| Fenpropathrin | 27.2 |  | 49 | 1.43 | 0.301 |  | 2.60 |  | 39 | 0.41 | 0.031 |  |
| ECBC | 22.6 | 11 |  | 1.35 |  | NA | 3.73 | 27 |  | 0.57 |  | 0.013 |
| FAL | 42.4 | 63 |  | 1.63 |  | NA | 2.23 | 28 |  | 0.35 |  | 0.375 |
| IIVS | 16.7 | 12 |  | 1.22 |  | NA | 1.82 | 17 |  | 0.26 |  | 0.044 |
| Gibberellic Acid | 7842 |  | 3 | 3.89 | 0.621 |  | 2866 |  | 2 | 3.46 | 0.862 |  |
| ECBC | 8027 | 11 |  | 3.90 |  | NA | 2850 | 14 |  | 3.45 |  | NA |
| FAL | NA | NA |  | NA |  | NA | 2940 | 9 |  | 3.47 |  | NA |
| IIVS | 7657 | 10 |  | 3.88 |  | NA | 2807 | 4 |  | 3.45 |  | NA |
| Glutethimide | 192 |  | 43 | 2.28 | < 0.001 |  | 177 |  | 5 | 2.25 | 0.968 |  |
| ECBC | 167 | 4 |  | 2.22 |  | 0.029 | 187 | 34 |  | 2.27 |  | NA |
| FAL | 284.3 | 7 |  | 2.45 |  | <0.001 | 170 | 14 |  | 2.23 |  | NA |
| IIVS | 125.3 | 7 |  | 2.10 |  | <0.001 | 176 | 16 |  | 2.24 |  | NA |
| Glycerol | 28904 |  | 33 | 4.46 | 0.846 |  | 27108 |  | 31 | 4.43 | 0.200 |  |
| ECBC | 20000 | 15 |  | 4.30 |  | NA | 34267 | 45 |  | 4.53 |  | NA |
| FAL | 38878 | 73 |  | 4.59 |  | NA | 18023 | 46 |  | 4.26 |  | NA |
| IIVS | 27833 | 39 |  | 4.44 |  | NA | 29033 | 16 |  | 4.46 |  | NA |
| Haloperidol | 6.26 |  | 24 | 0.80 | 0.006 |  | 3.57 |  | 7 | 0.55 | 0.935 |  |
| ECBC | 5.30 | 12 |  | 0.72 |  | 0.030 | 3.69 | 27 |  | 0.57 |  | NA |
| FAL | 8.00 | 8 |  | 0.90 |  | 0.002 | 3.72 | 49 |  | 0.57 |  | NA |
| IIVS | 5.50 | 12 |  | 0.74 |  | 0.061 | 3.29 | 35 |  | 0.52 |  | NA |
| Hexachlorophene | 4.48 |  | 27 | 0.65 | 0.174 |  | 0.031 |  | 41 | -1.50 | 0.097 |  |
| ECBC | 5.00 | 48 |  | 0.70 |  | NA | 0.027 | 16 |  | -1.57 |  | NA |
| FAL | 5.30 | 33 |  | 0.72 |  | NA | 0.046 | 44 |  | -1.34 |  | NA |
| IIVS | 3.10 | 9 |  | 0.49 |  | NA | 0.021 | 11 |  | -1.67 |  | NA |
| Lactic acid | 3073 |  | 12 | 3.49 | 0.160 |  | 1308 |  | 1 | 3.12 | 0.904 |  |
| ECBC | 2943 | 11 |  | 3.47 |  | NA | 1290 | 4 |  | 3.11 |  | NA |
| FAL | 3487 | 16 |  | 3.54 |  | NA | 1320 | 5 |  | 3.12 |  | NA |
| IIVS | 2790 | 9 |  | 3.45 |  | NA | 1313 | 11 |  | 3.12 |  | NA |
| Lindane | 161 |  | 58 | 2.21 | 0.066 |  | 19.3 |  | 20 | 1.29 | 0.203 |  |
| ECBC | 125 | 95 |  | 2.10 |  | NA | 19.1 | 17 |  | 1.28 |  | NA |
| FAL | 266 | 36 |  | 2.43 |  | NA | 23.2 | 31 |  | 1.37 |  | NA |
| IIVS | 90.4 | 122 |  | 1.96 |  | NA | 15.6 | 15 |  | 1.19 |  | NA |
| Lithium carbonate | NA |  | NA | NA | NA | NA | 477 |  | 13 | 2.68 | 0.295 |  |
| ECBC | 564 | 12 |  | 2.75 |  | NA | 411 | 29 |  | 2.61 |  | NA |
| FAL | NA | NA |  | NA |  | NA | 486 | 20 |  | 2.69 |  | NA |
| IIVS | NA | NA |  | NA |  | NA | 535 | 6 |  | 2.73 |  | NA |
| Meprobamate | 539 |  | 54 | 2.73 | <0.001 |  | 516 |  | 61 | 2.71 | 0.027 |  |
| ECBC | 353 | 14 |  | 2.55 |  | 0.001 | 761 | 15 |  | 2.88 |  | 0.0758 |
| FAL | 877 | 15 |  | 2.94 |  | <0.001 | 163 | 116 |  | 2.21 |  | 0.0098 |
| IIVS | 386 | 2 |  | 2.59 |  | 0.005 | 624 | 14 |  | 2.80 |  | 0.1648 |
| Mercury chloride | 4.32 |  | 33 | 0.64 | 0.021 |  | 5.87 |  | 15 | 0.77 | 0.120 |  |
| ECBC | 3.50 | 5 |  | 0.54 |  | 0.083 | 6.87 | 15 |  | 0.84 |  | NA |
| FAL | 6.00 | 31 |  | 0.78 |  | 0.008 | 5.40 | 19 |  | 0.73 |  | NA |
| IIVS | 3.50 | 3 |  | 0.54 |  | 0.110 | 5.35 | 2 |  | 0.73 |  | NA |
| Methanol | NA |  | NA | NA | NA | NA | 1616 |  | 42 | 3.21 | 0.007 |  |
| ECBC | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| FAL | NA |  |  | NA |  | NA | 1133 | 19 |  | 3.05 |  | NA |
| IIVS | NA |  |  | NA |  | NA | 2100 | 11 |  | 3.32 |  | NA |
| Nicotine | 378 |  | 25 | 2.58 | 0.128 |  | 113 |  | 17 | 2.05 | 0.700 |  |
| ECBC | 272 | 24 |  | 2.43 |  | NA | 94.3 | 26 |  | 1.97 |  | NA |
| FAL | 412 | 33 |  | 2.61 |  | NA | 134 | 59 |  | 2.13 |  | NA |
| IIVS | 450 | 12 |  | 2.65 |  | NA | 112 | 25 |  | 2.05 |  | NA |
| Paraquat | 23.3 |  | 8 | 1.37 | 1.000 |  | 66.1 |  | 40 | 1.82 | 0.047 |  |
| ECBC | 21.3 | 34 |  | 1.33 |  | NA | 48.3 | 13 |  | 1.68 |  | 0.089 |
| FAL | 24.9 | 67 |  | 1.40 |  | NA | 96.6 | 39 |  | 1.98 |  | 0.018 |
| IIVS | 23.7 | 64 |  | 1.37 |  | NA | 53.4 | 10 |  | 1.73 |  | 0.279 |
| Parathion | 61.8 |  | 111 | 1.79 | 0.014 |  | 31.4 |  | 8 | 1.50 | 0.845 |  |
| ECBC | 22.7 | 53 |  | 1.36 |  | 0.064 | 34.0 | 30 |  | 1.53 |  | NA |
| FAL | 141 | 70 |  | 2.15 |  | 0.005 | 31.2 | 38 |  | 1.49 |  | NA |
| IIVS | 22 | 22 |  | 1.34 |  | 0.081 | 29.0 | 29 |  | 1.46 |  | NA |
| Phenobarbital | 612 |  | 21 | 2.79 | 0.232 |  | 478 |  | 39 | 2.68 | 0.027 |  |
| ECBC | 634 | 21 |  | 2.80 |  | NA | 693 | 26 |  | 2.84 |  | 0.010 |
| FAL | 726 | 35 |  | 2.86 |  | NA | 360 | 27 |  | 2.56 |  | 0.072 |
| IIVS | 476 | 23 |  | 2.68 |  | NA | 381 | 18 |  | 2.58 |  | 0.173 |
| Phenol | 70.9 |  | 41 |  | 0.011 |  | 77.7 |  | 22 | 1.89 | 0.094 |  |
| ECBC | 50.2 | 22 |  | 1.70 |  | 0.022 | 59.1 | 36 |  | 1.77 |  | NA |
| FAL | 104 | 24 |  | 2.02 |  | 0.004 | 93.2 | 6 |  | 1.97 |  | NA |
| IIVS | 58.1 | 12 |  | 1.76 |  | 0.206 | 80.8 | 6 |  | 1.91 |  | NA |
| Phenylthiourea | 119 |  | 90 | 2.08 | 0.007 |  | 346 |  | 19 | 2.54 | 0.133 |  |
| ECBC | 30.1 | 66 |  | 1.48 |  | 0.004 | 363 | 16 |  | 2.56 |  | NA |
| FAL | 239 | 28 |  | 2.38 |  | 0.006 | 401 | 21 |  | 2.60 |  | NA |
| IIVS | 89 | 25 |  | 1.95 |  | 0.718 | 272 | 26 |  | 2.44 |  | NA |
| Physostigmine | 28.8 |  | 30 | 1.46 | 0.149 |  | 172 |  | 22 | 2.24 | 0.623 |  |
| ECBC | 28.2 | 53 |  | 1.45 |  | NA | 164 | 3 |  | 2.21 |  | NA |
| FAL | 37.8 | 5 |  | 1.58 |  | NA | 213 | 112 |  | 2.33 |  | NA |
| IIVS | 20.4 | 33 |  | 1.31 |  | NA | 139 | 6 |  | 2.14 |  | NA |
| Potassium chloride | 3635 |  | 7 | 3.56 | 0.846 |  | 2279 |  | 13 | 3.36 | 0.396 |  |
| ECBC | 3352 | 14 |  | 3.53 |  | NA | 2560 | 17 |  | 3.41 |  | NA |
| FAL | 3842 | 31 |  | 3.58 |  | NA | 2287 | 28 |  | 3.36 |  | NA |
| IIVS | 3710 | 11 |  | 3.57 |  | NA | 1990 | 8 |  | 3.30 |  | NA |
| Potassium cyanide | 64.3 |  | 127 | 1.81 | <0.001 |  | 45.1 |  | 86 | 1.65 | 0.340 |  |
| ECBC | 15.3 | 25 |  | 1.18 |  | 0.001 | 29.3 | 24 |  | 1.47 |  | NA |
| FAL | 159 | 52 |  | 2.20 |  | <0.001 | 89.0 | 112 |  | 1.95 |  | NA |
| IIVS | 18.9 | 5 |  | 1.28 |  | 0.006 | 16.9 | 13 |  | 1.23 |  | NA |
| Procainamide HCl | 443 |  | 11 | 2.65 | 0.007 |  | 1764 |  | 16 | 3.25 | 0.053 |  |
| ECBC | 400 | 4 |  | 2.60 |  | 0.008 | 1480 | 14 |  | 3.17 |  | NA |
| FAL | 431 | 1 |  | 2.63 |  | 0.396 | 1787 | 12 |  | 3.25 |  | NA |
| IIVS | 497 | 8 |  | 2.70 |  | 0.003 | 2027 | 11 |  | 3.31 |  | NA |
| 2-Propanol | 3563 |  | 23 | 3.55 | 0.001 |  | 5541 |  | 26 | 3.74 | 0.033 |  |
| ECBC | 2610 | 9 |  | 3.42 |  | < 0.001 | 5263 | 11 |  | 3.72 |  | 0.797 |
| FAL | 3970 | 4 |  | 3.60 |  | 0.004 | 4273 | 27 |  | 3.63 |  | 0.026 |
| IIVS | 4110 | 4 |  | 3.61 |  | 0.002 | 7087 | 7 |  | 3.85 |  | 0.018 |
| Propranolol HCl | 14.9 |  | 16 | 1.17 | 0.488 |  | 36.9 |  | 21 | 1.57 | 0.003 |  |
| ECBC | 13.6 | 32 |  | 1.13 |  | NA | 38.27 | 12 |  | 1.58 |  | 0.325 |
| FAL | 13.5 | 51 |  | 1.13 |  | NA | 43.8 | 6 |  | 1.64 |  | 0.006 |
| IIVS | 17.6 | 21 |  | 1.25 |  | NA | 28.6 | 11 |  | 1.46 |  | 0.001 |
| Propylparaben | 29.9 |  | 64 | 1.48 | 0.001 |  | 16.8 |  | 16 | 1.23 | 0.066 |  |
| ECBC | 20.9 | 16 |  | 1.32 |  | 0.045 | 18.1 | 13 |  | 1.26 |  | NA |
| FAL | 51.8 | 29 |  | 1.71 |  | < 0.001 | 18.6 | 15 |  | 1.27 |  | NA |
| IIVS | 17.1 | 12 |  | 1.23 |  | 0.003 | 13.8 | 9 |  | 1.14 |  | NA |
| Sodium arsenite | 0.873 |  | 55 | -0.06 | 0.028 |  | 0.532 |  | 44 | -0.27 | 0.061 |  |
| ECBC | 0.500 | 6 |  | -0.30 |  | 0.032 | 0.790 | 32 |  | -0.10 |  | NA |
| FAL | 1.40 | 57 |  | 0.15 |  | 0.012 | 0.336 | 56 |  | -0.47 |  | NA |
| IIVS | 0.700 | 17 |  | -0.15 |  | 0.478 | 0.470 | 14 |  | -0.33 |  | NA |
| Sodium chloride | 4764 |  | 3 | 3.68 | 0.759 |  | 2724 |  | 51 | 3.44 | 0.045 |  |
| ECBC | 4790 | 5 |  | 3.68 |  | NA | 3583 | 7 |  | 3.55 |  | 0.141 |
| FAL | 4625 | 13 |  | 3.67 |  | NA | 1118 | 124 |  | 3.05 |  | 0.017 |
| IIVS | 4877 | 9 |  | 3.69 |  | NA | 3470 | 9 |  | 3.54 |  | 0.161 |
| Sodium dichromate dihydrate | 0.602 |  | 9 | -0.22 | 0.822 |  | 0.737 |  | 19 | -0.13 | 0.258 |  |
| ECBC | 0.603 | 14 |  | -0.22 |  | NA | 0.784 | 14 |  | -0.11 |  | NA |
| FAL | 0.657 | 37 |  | -0.18 |  | NA | 0.851 | 36 |  | -0.07 |  | NA |
| IIVS | 0.547 | 17 |  | -0.26 |  | NA | 0.576 | 17 |  | -0.24 |  | NA |
| Sodium fluoride | 79.8 |  | 22 | 1.90 | 0.016 |  | 47.4 |  | 15 | 1.68 | 0.313 |  |
| ECBC | 61.3 | 9 |  | 1.79 |  | 0.007 | 48.7 | 14 |  | 1.69 |  | NA |
| FAL | 96.1 | 18 |  | 1.98 |  | 0.019 | 39.7 | 24 |  | 1.60 |  | NA |
| IIVS | 82.0 | 7 |  | 1.91 |  | 0.463 | 53.7 | 13 |  | 1.73 |  | NA |
| Sodium hypochlorite | 1211 |  | 57 | 3.08 | 0.040 |  | 1580 |  | 20 | 3.20 | 0.313 |  |
| ECBC | 823 | 13 |  | 2.92 |  | 0.257 | 1863 | 31 |  | 3.27 |  | NA |
| FAL | 805 | 46 |  | 2.91 |  | 0.119 | 1243 | 46 |  | 3.09 |  | NA |
| IIVS | 2005 | 44 |  | 3.30 |  | 0.015 | 1633 | 11 |  | 3.21 |  | NA |
| Sodium oxalate | 40.8 |  | 23 | 1.61 | 0.643 |  | 355 |  | 1 | 2.55 | 0.926 |  |
| ECBC | 42.0 | 41 |  | 1.62 |  | NA | 355 | 15 |  | 2.55 |  | NA |
| FAL | 31.0 | 28 |  | 1.49 |  | NA | 350 | 42 |  | 2.54 |  | NA |
| IIVS | 49.5 | 53 |  | 1.69 |  | NA | 360 | 26 |  | 2.56 |  | NA |
| Sodium selenate | 34.5 |  | 60 | 1.54 | <0.001 |  | 11.2 |  | 40 | 1.05 | 0.134 |  |
| ECBC | 12.7 | 13 |  | 1.10 |  | <0.001 | 7.47 | 12 |  | 0.87 |  | NA |
| FAL | 54.2 | 19 |  | 1.73 |  | < 0.001 | 16.1 | 59 |  | 1.21 |  | NA |
| IIVS | 36.5 | 14 |  | 1.56 |  | 0.026 | 10.0 | 13 |  | 1.00 |  | NA |
| Strychnine | 199 |  | 83 | 2.30 | <0.001 |  | 69.3 |  | 39 | 1.84 | 0.364 |  |
| ECBC | 389 | 21 |  | 2.59 |  | <0.001 | 100 | 76 |  | 2.00 |  | NA |
| FAL | 124 | 16 |  | 2.09 |  | 0.018 | 52.5 | 53 |  | 1.72 |  | NA |
| IIVS | 83.5 | 6 |  | 1.92 |  | <0.001 | 55.1 | 6 |  | 1.74 |  | NA |
| Thallium Sulfate | 7.50 |  | 72 | 0.88 | 0.165 |  | 0.16 |  | 23 | -0.80 | 0.405 |  |
| ECBC | 2.80 | 24 |  | 0.45 |  | NA | 0.198 | 51 |  | -0.70 |  | NA |
| FAL | 13.4 | 78 |  | 1.13 |  | NA | 0.153 | 20 |  | -0.82 |  | NA |
| IIVS | 6.30 | 28 |  | 0.80 |  | NA | 0.127 | 16 |  | -0.90 |  | NA |
| Trichloroacetic acid | 928 |  | 27 | 2.97 | 0.005 |  | 427 |  | 24 | 2.63 | 0.134 |  |
| ECBC | 762 | 13 |  | 2.88 |  | 0.022 | 348 | 18 |  | 2.54 |  | NA |
| FAL | 1220 | 6 |  | 3.09 |  | 0.002 | 541 | 28 |  | 2.73 |  | NA |
| IIVS | 801 | 14 |  | 2.90 |  | 0.069 | 394 | 13 |  | 2.60 |  | NA |
| 1,1,1-Trichloroethane | 15538 |  | 52 | 4.19 | <0.001 |  | NA |  | NA | NA | NA |  |
| ECBC | NA | NA |  | NA |  | NA | 8137 | 7 |  | 3.91 |  | NA |
| FAL | 21250 | 11 |  | 4.33 |  | NA | NA | NA |  | NA |  | NA |
| IIVS | 9827 | 2 |  | 3.99 |  | NA | NA | NA |  | NA |  | NA |
| Triethylenemelamine | 0.568 |  | 135 | -0.25 | <0.001 |  | 1.95 |  | 12 | 0.29 | 0.562 |  |
| ECBC | 0.086 | 11 |  | -1.07 |  | <0.001 | 1.69 | 57 |  | 0.23 |  | NA |
| FAL | 1.45 | 18 |  | 0.16 |  | <0.001 | 2.03 | 23 |  | 0.31 |  | NA |
| IIVS | 0.169 | 29 |  | -0.77 |  | 0.002 | 2.13 | 23 |  | 0.33 |  | NA |
| Triphenyltin hydroxide | 0.022 |  | 29 | -1.66 | 0.688 |  | 0.013 |  | 55 | -1.89 | 0.088 |  |
| ECBC | 0.026 | 17 |  | -1.59 |  | NA | 0.021 | 32 |  | -1.68 |  | NA |
| FAL | 0.026 | 81 |  | -1.59 |  | NA | 0.007 | 106 |  | -2.15 |  | NA |
| IIVS | 0.015 | 55 |  | -1.83 |  | NA | 0.011 | 32 |  | -1.96 |  | NA |
| Valproic acid | 1177 |  | 76 | 3.07 | < 0.001 |  | 533 |  | 28 | 2.73 | 0.081 |  |
| ECBC | 547 | 12 |  | 2.74 |  | NA | 468 | 25 |  | 2.67 |  | 0.331 |
| FAL | 1807 | 10 |  | 3.26 |  | NA | 702 | 23 |  | 2.85 |  | 0.032 |
| IIVS | NA | NA |  | NA |  | NA | 430 | 17 |  | 2.63 |  | 0.135 |
| Verapamil HCl | 35.2 |  | 10 | 1.55 | 0.230 |  | 68.7 |  | 14 | 1.84 | 0.624 |  |
| ECBC | 32.0 | 18 |  | 1.51 |  | NA | 60.5 | 22 |  | 1.78 |  | NA |
| FAL | 34.6 | 5 |  | 1.54 |  | NA | 79.4 | 42 |  | 1.90 |  | NA |
| IIVS | 38.9 | 11 |  | 1.59 |  | NA | 66.2 | 8 |  | 1.82 |  | NA |
| Xylene | NA |  | NA | NA | NA | NA | NA |  | NA | NA | NA |  |
| ECBC | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| FAL | NA | NA |  | NA |  | NA | NA | NA |  | NA |  | NA |
| IIVS | 724 | 12 |  | 2.86 |  | NA | 486 | 38 |  | 2.69 |  | NA |

1Results reported on the same row with chemical names are the means of all the laboratories.  Results reported on the same row as laboratories are the laboratory means.

2p < 0.01 indicated statistical significance.

3Contrasts were performed if ANOVA was significant (p < 0.01) to determine which laboratory was different from the other two laboratories.  Significant contrasts were denoted by p < 0.01.  If only two laboratories reported results, no contrast tests were necessary.

Abbreviations: Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.  NA - no acceptable IC50 results reported or calculation was not performed (e.g., for contrast results).

Table 7-4	Reference Substances with Significant Differences between Laboratories for 3T3 NRU Test Method Results

| Reference Substance | Significant Contrast Results1 | Insoluble/ Volatile2 |
| --- | --- | --- |
|  | ECBC | FAL | IIVS |  |
| Acetylsalicylic acid |  | H | L |  |
| Arsenic trioxide |  | L | H | Precipitate |
| Busulfan |  | H |  |  |
| Chloral hydrate | L | H |  |  |
| Chloramphenicol | L | H |  |  |
| Citric acid | L | H |  |  |
| Colchicine | L | H |  |  |
| Cupric sulfate pentahydrate | X | H | L |  |
| Dibutyl phthalate |  | H | L | Precipitate |
| Dichlorvos | L | H |  | Precipitate |
| Disulfoton3 |  |  |  | Precipitate |
| Ethylene glycol | L |  |  |  |
| Glutethimide |  | H | L |  |
| Haloperidol |  | H |  |  |
| Meprobamate | L | H | X |  |
| Phenylthiourea | L | H |  |  |
| Potassium cyanide | L | H | X | Precipitate /Volatility |
| Procainamide HCl | L |  | H |  |
| 2-Propanol | L | X | H | Volatility |
| Propylparaben |  | H | L |  |
| Sodium selenate | L | H |  |  |
| Strychnine | H |  | L | Precipitate |
| Trichloroacetic acid |  | H |  |  |
| 1,1,1-Trichloroethane4 |  |  |  | Precipitate |
| Triethylenemelamine | L | H |  |  |
| Valproic acid5 |  |  |  | Precipitate |

1Laboratories significantly different from the other two at p < 0.01.  H – Laboratory reported the highest mean IC50.  L – Laboratory reported the lowest mean IC50.  X – Laboratory reported a mean IC50 between the values of the other two laboratories.

2From Table 5-8.  Precipitate reported by at least one laboratory is indicated by “Precipitate”.  Use of plate sealers by at least one laboratory to prevent volatile contamination of control wells indicated by “Volatility”.

3Significant ANOVA (p < 0.01), but no contrast analysis since only two laboratories (ECBC and IIVS) reported IC50 values.

4Significant ANOVA (p < 0.01), but no contrast results since only two laboratories (FAL and IIVS) reported IC50 values.

5Significant ANOVA (p < 0.01), but no contrast results since only two laboratories (ECBC and FAL) reported IC50 values.

Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

For the 26 substances that yielded significantly different results among the laboratories, contrast analyses indicated that ECBC and FAL were frequently different from the other laboratories.  ECBC tended to report the lowest IC50 values among the laboratories while FAL tended to report the highest values of the three laboratories.  ECBC reported significantly different results from the other two laboratories for 15 of the 26 substances.  For 13 of the 15 substances, ECBC reported the lowest mean IC50 value among the three laboratories.  FAL reported significantly different results from the other two laboratories for 20 of the 26 substances.  For 18 of the 20 substances, FAL reported the highest mean IC50 value among the three laboratories.  IIVS reported significantly different results for 11 of the 26 substances, with no great majority of highest or lowest IC50 values.

# Differences Among the Laboratories for the NHK NRU Test Method

The ANOVA results in Table 7-3 indicate that there were statistically significant (p < 0.01) laboratory differences for seven substances.  These substances are listed in Table 7-5 along with columns showing the laboratory statistically significantly differing from the other two laboratories (as indicated by the contrast results), and indications of whether any laboratory reported insolubility or volatility during conduct of the assay.  Insolubility was reported for three of the seven substances.

Table 7-5	Reference Substances with Significant Differences between Laboratories for NHK NRU Test Method Results

| Reference Substance | Significant Contrast Results1 | Solubility/ Volatility2 |
| --- | --- | --- |
|  | ECBC | FAL | IIVS |  |
| Arsenic trioxide |  | L | H | Precipitate |
| Citric acid | H | L |  | Precipitate |
| Digoxin | H | L |  |  |
| Dimethylformamide | H |  | L |  |
| Disulfoton | L | H |  | Precipitate |
| Methanol3 |  |  |  |  |
| Propranolol HCl |  | H | L |  |

1Laboratories significantly different from the other two at p < 0.01.  H – Laboratory reported the highest mean IC50.  L – Laboratory reported the lowest mean IC50.  X – Laboratory reported a mean IC50 between the values of the other two laboratories.

2From Table 5-8.  Precipitate reported by at least one laboratory is indicated by “Precipitate”.  Use of plate sealers by at least one laboratory to prevent volatile contamination of control wells indicated by “Volatility”.

3Significant ANOVA (p < 0.01), but no contrast results since only two laboratories (FAL and IIVS) reported IC50 values.

Laboratories: ECBC – U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

For the seven substances that yielded significantly different results among the laboratories, ECBC and FAL were frequently different from the other laboratories.  ECBC tended to report the highest IC50 value among the laboratories (4/7 substances) while FAL tended to report the lowest values among the three laboratories (3/7 substances).

7.2.2	CV Results for the 3T3 and NHK NRU Test Methods

CV was calculated as described in Section 5.3.3.  Table 7-3 provides the intra- and inter-laboratory CV values for individual substances.  Table 7-6 summarizes the CV results for each test method.  Table 7-6 shows that median and mean CV values were often similar.  Median CV values appeared always lower than the corresponding means, which indicated that large individual CV values skewed the CV distributions somewhat to the right.

Intralaboratory CV

Table 7-6 shows that both test methods had similar ranges for the intralaboratory CV.  The mean intralaboratory CV values were the same, 26%.  The median intralaboratory CVs were also similar: 23% for the 3T3 NRU test method and 24% for the NHK NRU test method.  Of the three laboratories, FAL had the highest mean and median CV for both test methods and IIVS had the lowest mean and median CV for both test methods.

# Interlaboratory CV

The mean and median interlaboratory CV for the reference substances was lower for the NHK NRU test method (mean = 28%; median = 21%) than for the 3T3 NRU test method (mean = 46%; median = 40%) (see Table 7-6).

Table 7-6	Summary of CV Results for the 3T3 and NHK NRU Test Methods

| CV | 3T3 NRU Test Method | NHK NRU Test Method |
| --- | --- | --- |
|  | N | Mean | Median | Range | N | Mean | Median | Range |
| Intralaboratory CV | 202 | 26% | 23% | 1-122% | 208 | 26% | 24% | 1-129% |
| ECBC | 68 | 23% | 17% | 2-95% | 69 | 23% | 19% | 2-76% |
| FAL | 66 | 33% | 30% | 1-98% | 69 | 42% | 32% | 1-129% |
| IIVS | 68 | 21% | 13% | 1-122% | 70 | 14% | 13% | 1-38% |
| Interlaboratory CV | 68 | 46% | 40% | 2-135% | 68 | 28% | 21% | 1-99% |

Abbreviations: N- number of values.  Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

Note: For the 3T3 NRU test method, the following laboratories/substances did not obtain sufficient IC50 data for the calculation of an intralaboratory CV: carbon tetrachloride at any laboratory; disulfoton at FAL; gibberellic acid at FAL; lithium carbonate at FAL and IIVS; methanol at any laboratory; 1,1,1-trichloroethane at ECBC; valproic acid at IIVS; and xylene at ECBC and FAL.  For the NHK assay, the following laboratories/substances did not obtain sufficient IC50 data for the calculation of an intralaboratory CV: carbon tetrachloride at any laboratory; methanol at ECBC; 1,1,1-trichloroethane at FAL and IIVS; and xylene at ECBC and FAL.  For the 3T3 NRU test method, the following substances did not obtain sufficient IC50 data for the calculation of an interlaboratory CV: carbon tetrachloride, lithium carbonate; methanol; and xylene.  For the NHK assay, the following substances did not yield sufficient IC50 data for the calculation of an interlaboratory CV: carbon tetrachloride; 1,1,1-trichloroethane; and xylene.

# Variation of CV with Chemical Property

To identify the chemical characteristics that may yield high or low CV values, CV values were analyzed to determine their association with the following chemical attributes: physical state (i.e., solid or liquid), solubility, volatility, chemical class, GHS acute oral toxicity class (UN 2005), molecular weight, log Kow, IC50, and boiling point.  For categorical characteristics such as physical form, solubility (i.e., precipitate/no precipitate), volatile/not volatile, and chemical class, the mean CV values and CV ranges for the groups were compared to one another and to the overall mean CV and CV range for each test method.  No statistical analyses were performed.  For chemical characteristics measured by continuous variables, such as molecular weight, log Kow, and IC50, and boiling point, Spearman correlation analyses were performed.

# Results of Intralaboratory CV Analysis

Table 7-7 shows the analysis of intralaboratory CV.  The analysis of intralaboratory CV uses one mean intralaboratory CV for each reference substances that was calculated from the intralaboratory CV values from each laboratory.  With the exception of the amides, which had relatively low intralaboratory CV values (for both 3T3 and NHK NRU test methods), and organophosphates and halogenated hydrocarbons (for the 3T3 NRU test method only), which had relatively high intralaboratory CV values, there seemed to be little difference in CV values for the categorical physical/chemical/toxicological attributes.  The mean intralaboratory CV values for solids and liquids were similar (26 vs. 24% for the 3T3 NRU test method; 27 vs. 23% for the NHK NRU test method).  The mean intralaboratory CV values for reference substances for which precipitates were observed were similar to the mean intralaboratory CV values for substances for which no precipitates were observed (29 vs. 23% for the 3T3 NRU test method; 24 vs. 27% for the NHK NRU test method).  The mean intralaboratory CV values for substances that exhibited volatility (i.e., indicated by laboratory use of film plate sealers to prevent contamination of control wells) were relatively similar to those that did not (31 vs. 24% for the 3T3 NRU test method; 27 vs. 26% for the NHK NRU test method).  Similarly, the substances grouped by GHS toxicity category (UN 2005) had mean intralaboratory CV values that were similar (19-33% for the 3T3 NRU test method; 18-31% for the NHK NRU test method) to the overall mean CV values (26% for both the 3T3 and NHK NRU test methods).

Reference substances in the amide chemical class had unusually low mean intralaboratory CV values for both the 3T3 NRU test method (13%) and NHK NRU test method (10%) compared with the overall mean CV (26% for both test methods), but there were only three substances in the class (acetaminophen, dimethylformamide, and procainamide HCl).  Reference substances in the organophosphate chemical class had unusually high mean intralaboratory CV values for the 3T3 NRU test method (46%), but not for the NHK NRU test method (26%) compared with the overall mean CV (26% for the 3T3 and NHK NRU test methods).  There were only three substances in the class (dichlorvos, disulfoton, and parathion), but two of the three substances had relatively high mean intralaboratory CV values (17, 48 and 71%).  Halogenated hydrocarbons also had high mean intralaboratory CV for the 3T3 NRU test method (46%), but not for the NHK NRU test method (14%) compared with the overall mean intralaboratory CV for each test method (26%).  However, the mean intralaboratory CV for the 3T3 NRU test method was calculated from only two values; 7% for 1,1,1-trichloroethane and 84% for lindane.  No laboratory obtained sufficient toxicity for the calculation of an IC50 for the carbon tetrachloride, the third halogenated hydrocarbon.

Table 7-7	Intralaboratory CV by Chemical Characteristics for the 3T3 and NHK NRU Test Methods

| Class/Attribute | 3T3 NRU Test Method | NHK NRU Test Method |
| --- | --- | --- |
|  | Na | Range | Mean | Nb | Range | Mean |
| All chemicals | 70 | 1-122% | 26% | 71 | 1-129% | 26% |
| Chemical form |  |  |  |  |  |  |
| Solid | 53 | 4-84 | 26 | 53 | 6-50 | 27 |
| Liquid | 17 | 6-71 | 24 | 18 | 2-40 | 23 |
| Solubility |  |  |  |  |  |  |
| Precipitatec | 24 | 7-84 | 29 | 2a | 2-47 | 24 |
| No precipitate | 46 | 4-55 | 23 | 50 | 7-57 | 27 |
| Volatilityd |  |  |  |  |  |  |
| Volatile | 10 | 6-84 | 31 | 9 | 11-50 | 27 |
| Nonvolatile | 62 | 4-71 | 24 | 63b | 2-57 | 26 |
| Chemical Class |  |  |  |  |  |  |
| Alcohols | 9 | 6-42 | 22 | 10 | 10-37 | 21 |
| Carboxylic acids | 12 | 10-41 | 20 | 12 | 7-48 | 26 |
| Heterocyclics | 14 | 6-59 | 30 | 14 | 13-50 | 31 |
| Organophosphorous | 3 | 17-71 | 46 | 3 | 20-32 | 26 |
| Amides | 3 | 4-28 | 13 | 3 | 2-16 | 10 |
| Halogenated hydrocarbons | 2 | 7-84 | 46 | 2 | 7-21 | 14 |
| Inorganics | 15 | 9-43 | 24 | 15 | 6-50 | 29 |
| Toxicity Class |  |  |  |  |  |  |
| ≤ 5 mg/kg | 7 | 9-71 | 33 | 7 | 20-40 | 30 |
| > 5 - ≤ 50 | 12 | 13-59 | 32 | 12 | 12-50 | 31 |
| > 50 - ≤ 300 | 12 | 11-84 | 33 | 12 | 17-37 | 25 |
| > 300 - ≤ 2000 | 16 | 4-51 | 21 | 16 | 6-57 | 25 |
| > 2000 - ≤ 5000 | 10a | 9-32 | 19 | 10a | 7-50 | 31 |
| > 5000 | 13b | 6-42 | 19 | 14 | 2-40 | 18 |
| Correlations | N | rs | P value | N | rs | P value |
| Molecular weight | 70a,b | 0.323 | 0.006 | 71b | 0.199 | 0.097 |
| Log Kow | 50e | 0.117 | 0.421 | 51e | 0.311 | 0.026 |
| IC50 | 70a,b | -0.436 | 0.0002 | 71b | -0.362 | 0.002 |
| Boiling point | 27 | 0.576 | 0.002 | 28 | 0.277 | 0.154 |

aOne intralaboratory CV for each chemical was calculated by averaging the CV values for the laboratories that reported sufficient data for the calculation of a CV.  No CV was calculable for carbon tetrachloride or methanol.

bOne intralaboratory CV for each chemical was calculated by averaging the CV values for the laboratories that reported sufficient data for the calculation of a CV.  No CV was calculable for carbon tetrachloride.

cDenoted by laboratory reports of precipitate in the stock reference substance solutions or in cell culture (see Table 5-8).

dDenoted by laboratory reports of using plate sealers to avoid contamination of the VC wells (see Table 5-8).

eNumber of reference substances with CV values and log Kow data.

fNumber of reference substances with CV values and boiling point data.

For the characteristics amenable to correlation analysis, none of the correlation coefficients were large (absolute value of rs < 0.6), but several were statistically significantly different from zero for the 3T3 NRU test method.  Molecular weight (p = 0.006), IC50 (p = 0.0002), and boiling point (p = 0.002) exhibited statistically significant correlations (p < 0.05) to intralaboratory CV for the 3T3 NRU test method.  For molecular weight, the higher molecular weight substances had higher intralaboratory CV values.  For IC50, however, the substances with lower IC50 values had higher CV values.  The inverse correlation between intralaboratory CV values and IC50 is consistent with the common observation that measurements with very low values tend to have high CV values.  The fact that substances with higher boiling points had higher CV values was consistent with the categorical analysis of volatility.  The substances that exhibited volatile characteristics (i.e., high reference substance concentration wells contaminated the VC wells) in the 3T3 NRU test method had higher mean intralaboratory CV values (31%) than the substances that did not exhibit volatile characteristics (24%), but the difference did not seem large.

Likewise, for the NHK NRU test method, two of the characteristics amenable to correlation analysis were statistically significantly different from zero, but the correlation coefficients did not have large magnitudes (absolute value of rs < 0.4).  Log Kow (p = 0.026) and IC50         (p = 0.002) exhibited statistically significant correlations (p < 0.05) to intralaboratory CV for the NHK NRU test method.  Log Kow was positively correlated to the mean intralaboratory CV for each substance, but IC50, was negatively correlated to the mean IC50 for each substance.

# Results of Interlaboratory CV Analysis

Table 7-8 shows the analysis of interlaboratory CV.  With the exception of chemical class, there seemed to be little difference in interlaboratory CV values for most of the categorical physical/chemical characteristics.  The mean interlaboratory CV values for solids and liquids were similar (48 vs. 46% for the 3T3 NRU test method and 28 vs. 27% for the NHK NRU test method).  The mean interlaboratory CV values for substances for which precipitates were observed was similar to the mean interlaboratory CV values for substances for which no precipitates were observed (56 vs. 43% for the 3T3 NRU test method and 29 vs. 28% for the NHK NRU test method).  The mean interlaboratory CV values for substances that exhibited volatile characteristics appeared similar to those that did not (51 vs. 46% for the 3T3 NRU test method and 32 vs. 28% for the NHK NRU test method).

Reference substances in the amide chemical class had unusually low mean interlaboratory CV values for both the 3T3 NRU test method (15%) and NHK NRU test method (16%) compared with the overall mean interlaboratory CV (46% for the 3T3 NRU test method and 28% for the NHK NRU test method).  Chemicals in the organophosphate chemical class had unusually high mean interlaboratory CV values for the 3T3 NRU test method (74%) and moderately higher mean interlaboratory CV values for the NHK NRU test method (42%) compared with the overall mean interlaboratory CV (46% for the 3T3 NRU test method and 28% for the NHK NRU test method).  The high mean interlaboratory CV value for organophosphates in the NHK NRU test method, however, was produced largely by the high interlaboratory CV of 99% for disulfoton.  The interlaboratory CV values for dichlorvos and parathion were 20% and 8%, respectively.  Heterocyclic compounds also had higher mean interlaboratory CV values for the 3T3 NRU test method but not for the NHK NRU test method.  As a group, the 14 heterocyclic compounds had a mean interlaboratory CV of 61% while the overall mean interlaboratory CV for the 3T3 NRU test method was 46%.  Although there were a few low CV values (e.g., 8, 18) in the heterocyclic group, there were seven values greater than the overall mean CV of 46%.  The median interlaboratory CV for the heterocyclic group was 52%.

Table 7-8	Interlaboratory CV by Chemical Characteristics for the 3T3 and NHK NRU Test Methods

| Class/Attribute | 3T3 NRU Test Method | NHK NRU Test Method |
| --- | --- | --- |
|  | N | Range | Mean | N | Range | Mean |
| All chemicals | 68a | 2-135% | 46% | 69b | 1-99% | 28% |
| Chemical Form |  |  |  |  |  |  |
| Solids | 52 | 3-135 | 48 | 53 | 1-91 | 28 |
| Liquids | 16 | 6-124 | 46 | 16 | 1-99 | 27 |
| Solubility |  |  |  |  |  |  |
| Precipitatec | 22 | 3-127 | 56 | 19 | 1-99 | 29 |
| No precipitate | 47 | 3-135 | 43 | 50 | 1-88 | 28 |
| Volatility |  |  |  |  |  |  |
| Volatiled | 10 | 21-127 | 51 | 9 | 8-86 | 32 |
| Nonvolatile | 58 | 3-135 | 46 | 60 | 1-99 | 28 |
| Chemical Class |  |  |  |  |  |  |
| Alcohols | 9 | 12-119 | 38 | 10 | 11-42 | 22 |
| Carboxylic acids | 12 | 12-124 | 46 | 12 | 1-61 | 27 |
| Heterocyclics | 14 | 8-135 | 61 | 14 | 5-85 | 32 |
| Organophosphorous | 3 | 57-111 | 74 | 3 | 8-99 | 42 |
| Amides | 3 | 6-28 | 15 | 3 | 13-19 | 16 |
| Halogenated hydrocarbons | 2 | 52-58 | 55 | 1 | 20 | 20 |
| Inorganics | 14 | 3-127 | 48 | 15 | 4-91 | 29 |
| Toxicity Class |  |  |  |  |  |  |
| ≤ 5 mg/kg | 7 | 12-135 | 69 | 7 | 12-99 | 37 |
| > 5 - ≤ 50 | 12 | 33-127 | 78 | 12 | 8-91 | 41 |
| > 50 - ≤ 300 | 12 | 8-120 | 37 | 12 | 10-41 | 26 |
| > 300 - ≤ 2000 | 15 | 11-85 | 38 | 15 | 1-61 | 20 |
| > 2000 - ≤ 5000 | 9 | 3-69 | 29 | 9 | 1-85 | 27 |
| > 5000 | 13 | 3-124 | 39 | 13 | 2-44 | 25 |
| Correlations |  | rs | P value |  | rs | P value |
| Molecular weight | 68 | 0.193 | 0.115 | 69 | 0.136 | 0.265 |
| Log Kow | 49e | 0.194 | 0.182 | 49 | 0.170 | 0.244 |
| IC50 | 68 | -0.295 | 0.015 | 69 | -0.271 | 0.024 |
| Boiling point | 24f | 0.467 | 0.021 | 26 | -0.131 | 0.525 |

aThe following chemicals did not have sufficient IC50 data for the calculation of an interlaboratory CV: carbon tetrachloride, lithium carbonate; methanol; and xylene.

bThe following substances did not yield sufficient IC50 data for the calculation of an interlaboratory CV: carbon tetrachloride; 1,1,1-trichloroethane; and xylene.

cDenoted by laboratory reports of precipitate in the stock reference substance solutions or in cell culture (see Table 5-8).

dDenoted by laboratory reports of using plate sealers to avoid contamination of the VC wells (see Table 5-8).

eNumber of reference substances with CV values and log Kow data.

fNumber of reference substances with CV values and boiling point data.

Mean interlaboratory CV values tended to be large for chemicals in the most toxic GHS acute categories, especially for the 3T3 NRU test method.  For the 3T3 NRU test method, the mean interlaboratory CV for chemicals in the classes for LD50 ≤ 5 mg/kg (69%) and 5 < LD50 ≤ 50 mg/kg (78%) were much larger than the mean overall interlaboratory CV (46%,).  For the NHK NRU test method, the mean interlaboratory CV for chemicals in the classes for 5 < LD50 ≤ 5 mg/kg (37%) and 5 < LD50 ≤ 50 mg/kg (41%) were much larger than the mean overall interlaboratory CV (28%).

For the characteristics amenable to correlation analysis, none of the correlation coefficients were large (absolute value of rs < 0.5), but IC50 (p = 0.015) and boiling point (p = 0.021) exhibited statistically significant correlations (p < 0.05) to interlaboratory CV for the 3T3 NRU test method.  There was a negative correlation between interlaboratory CV and IC50, but the correlation between boiling point and interlaboratory CV was positive.  The positive correlation of CV with boiling point was largely consistent with the categorical analysis of volatility.  The substances that exhibited volatile characteristics in the 3T3 NRU test method had slightly higher mean CV than for the substances that did not exhibit volatile characteristics (51 vs. 46%).  For the NHK NRU test method, only IC50 was significantly correlated (p = 0.024) to interlaboratory CV with a negative correlation (rs = -0.271).

# 7.2.3	Comparison of Laboratory-Specific Linear Regression Analyses for the Prediction of In Vivo Rodent LD50 Values from In Vitro NRU IC50 Values

The laboratory-specific regressions presented in Table 6-1 of Section 6.1.1 were compared to one another (for each test method) with a goodness of fit F-test as described in Section 5.3.3.  The comparisons indicated that the laboratory-specific regressions for both test methods were not significantly different (p < 0.05) from one another.  The comparison of the laboratory-specific 3T3 NRU regressions to one another yielded p = 0.796.  The comparison of the laboratory-specific NHK NRU regressions to one another yielded p = 0.985.  Because the laboratory-specific regressions were not statistically different, data were combined into a single regression for each test method using a geometric mean of the laboratory-specific IC50 values for each substance (see Section 6.1.1).

7.2.4	Laboratory Concordance for the Prediction of GHS Acute Oral Toxicity Category

# This section provides the percentage of substances for which the laboratory-specific IC50 data yielded the same (for all three laboratories) GHS toxicity categorization when used with the regressions evaluated in Sections 6.3.1 through 6.3.3.  Data for the same reference substances for each test method were evaluated to determine the laboratory concordance for each regression.  Forty-three substances were evaluated for the 3T3 NRU test method and 44 substances were evaluated for the NHK NRU test method.  Of the original 72 substances tested, epinephrine bitartrate, colchicine, and propylparaben were excluded from all analyses because they were removed from the calculation of the RC rat-only weight regressions due to the lack of rat oral reference LD50 data.  The 21 substances with specific mechanisms of toxicity in Table 6-3 were excluded from all analyses to be consistent with those removed from the RC rat-only weight regression excluding substances with specific mechanisms of toxicity.  These substances have known mechanisms of toxicity that are not expected to be active in the 3T3 or NHK cell cultures.  Carbon tetrachloride, methanol, gibberellic acid, lithium carbonate, and xylene were excluded from the 3T3 NRU evaluations because at least one laboratory failed to attain sufficient toxicity in any test for the calculation of an IC50.  Carbon tetrachloride, methanol, 1,1,1-trichloroethane, and xylene were excluded from the NHK NRU analyses because at least one laboratory failed to attain sufficient toxicity in any test for the calculation of an IC50.

# Laboratory Concordance for the 3T3 and NHK NRU Test Methods with the RC Millimole Regression

Appendix J (Table J-1 for the 3T3 NRU test method and Table J-3 for the NHK NRU test method) shows the laboratory concordance of the observed (i.e., in vivo categories for the initial LD50 values in Table 3-2) and predicted GHS toxicity categories for each substance determined in each in vitro NRU cytotoxicity test method using the laboratory-specific geometric mean IC50 values and the RC millimole regression, log LD50 (mmol/kg) = 0.435 x log IC50 (mM) + 0.625.  The observed LD50 values are the rodent LD50 values from Table 3-2.

For the 43 substances that yielded IC50 results in all laboratories using the 3T3 NRU test method, the laboratories agreed on the GHS toxicity category for 31 substances (72%).  The 12 substances that produced discordant results among the laboratories were cupric sulfate pentahydrate, cycloheximide, dimethylformamide, diquat dibromide, phenol, phenylthiourea, sodium arsenite, sodium oxalate, sodium selenate, thallium sulfate, triethylenemelamine, and 1,1,1-trichloroethane.  The laboratory predictions for these substances disagreed by one GHS toxicity category.

For the 44 substances that yielded IC50 results in all laboratories using the NHK NRU test method, the laboratories agreed on toxicity category for 39 substances (89%).  The five substances that produced discordant results among the laboratories were arsenic trioxide, digoxin, ethanol, 2-propanol, and sodium arsenite.  The laboratory predictions for these substances disagreed by one toxicity category.  Laboratory concordance was greater for the NHK assay than for the 3T3 assay (89% vs 72%).

Laboratory Concordance of the 3T3 and NHK NRU Test Methods with the RC Rat-Only Weight Regression

Appendix J (Table J-5 for the 3T3 NRU test method and Table J-6 for the NHK NRU test method) shows the laboratory concordance of the observed (i.e., in vivo reference categories for LD50 values in Table 4-2) and predicted GHS toxicity categories for each substance as determined for each test method using the laboratory-specific geometric mean IC50 in the RC rat-only weight regression, log LD50 (mg/kg) = log IC50 ((g/mL) x 0.372 + 2.024, from Table 6-2.

For the 43 substances that yielded IC50 results in all laboratories using the 3T3 NRU test method, the laboratories agreed on the GHS toxicity category for 34 substances (79%).  The nine substances that produced discordant results among the laboratories were boric acid, cupric sulfate pentahydrate, cycloheximide, 2-propanol, propranolol HCl, sodium selenate, thallium sulfate, triethylenemelamine, and 1,1,1-trichloroethane.  The laboratory predictions for these substances disagreed by one GHS toxicity category.

For the 44 substances that yielded IC50 results in all laboratories using the NHK NRU test method, the laboratories agreed on toxicity category for 39 substances (89%).  The five substances that produced discordant results among the laboratories were arsenic trioxide, digoxin, glycerol, sodium chloride, and thallium sulfate.  The laboratory predictions for these substances disagreed by one toxicity category.  Laboratory concordance was greater for the NHK assay than for the 3T3 assay (89% vs 79%).

Laboratory Concordance of the 3T3 and NHK NRU Test Methods with the RC Rat-Only Weight Regression Excluding Substances with Specific Mechanisms of Toxicity

Appendix J (Table J-7 for the 3T3 NRU test method and Table J-8 for the NHK NRU test method) shows the laboratory concordance of the observed (i.e., in vivo) and predicted GHS toxicity categories for each substance as determined for each test method using the laboratory-specific geometric mean IC50 values in the RC rat-only weight regression after exclusion of substances with specific mechanisms of toxicity, log LD50 (mg/kg) = log IC50 ((g/mL) x 0.357 + 2.194 (Table 6-2).

For the 43 substances considered in the analysis of the 3T3 NRU test method, the three laboratories agreed on the toxicity category for 36 (84%) of the substances.  The seven substances that produced discordant results among the laboratories were boric acid, cupric sulfate pentahydrate, diquat dibromide, sodium hypochlorite, thallium sulfate, 1,1,1-trichloroethane, and valproic acid.  The laboratory predictions for these substances disagreed by one GHS toxicity category.

The extent of laboratory concordance for the RC rat-only weight regression after excluding substances with specific mechanisms of toxicity was the same for the NHK NRU test method (i.e., 84%, 37/44).  The seven substances that produced discordant results among the laboratories were arsenic trioxide, digoxin, glycerol, hexachlorophene, mercury chloride, sodium chloride, and sodium hypochlorite.  The laboratory predictions for these substances disagreed by one GHS toxicity category.

# 7.3	Historical Positive Control Data

The reproducibility of the positive control (SLS) data was assessed by CV analysis, ANOVA, and linear regression over time as described in Section 5.3.4.  The SLS data analyzed for variability are slightly different from those used to determine the PC acceptance limits shown in Table 5-2.  To get an assessment of the true variation of SLS IC50 values, the reproducibility analyses included IC50 values from SLS tests that failed the test acceptance criterion for the IC50 acceptance limits determined for each study phase.  These additional SLS tests, however, passed all other test acceptance criteria.  If more than one SLS test was performed in a single day (for each test method and laboratory), the IC50 values were averaged to determine a single IC50 for the day so that multiple results from a single day would not overly influence the average for each phase.

Figure 7-1 shows the average SLS IC50 values for each test method, laboratory, and study phase.  Graphically, it appears that the SLS IC50 for the 3T3 NRU test method was relatively consistent over the entire period of the study (approximately 2.5 years).  The intralaboratory CV values (shown in Figure 7-1) for the individual study phases ranged from 5% to 24%.  With the exception of the Phase Ib CV at FAL, the CV values for each laboratory and phase were less than 20%.  The interlaboratory CV values were even smaller: 6% for Phases Ia and Ib; 10% for Phase II; and 2% for Phase III.

Figure 7-1 shows that the SLS IC50 for the NHK NRU test method tended to vary with time, but, with the exception of the SLS IC50 results from FAL, there appeared to be no consistent trend.  The IC50 values from FAL, which changed NHK cell culture methods after Phase Ib (see Section 5.1.3), tended to decrease over time.  Although the change in cell culture methods reduced the magnitude of the IC50, the variability (as evidenced by the intralaboratory CV values shown in Figure 7-1) remained relatively high (CV ≥ 34% for all FAL study phases).  The CV values for all the laboratories and study phases indicated that the SLS IC50 values for the NHK NRU test method was more variable within laboratories than the SLS IC50 for 3T3 NRU test method.  CV values for the SLS IC50 for the NHK NRU test method ranged from 11 to 51%, with nine of the 12 values greater than 20%.  The interlaboratory CV values, which were also greater than those for the 3T3 NRU test method, were: 39% for Phase Ia; 21% for Phase Ib; and 31% for Phase II; and 8% for Phase III.

Figure 7-1	SLS IC50 for Each Laboratory and Study Phase

a	3T3 NRU Test Method

b	NHK NRU Test Method

Bars show mean IC50 values.  Error bars show standard deviation.  Percent values above error bars are intralaboratory CVs.

Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

7.3.1	ANOVA and Linear Regression Results for the 3T3 NRU Test Method

# SLS IC50 Variation with Time

Table 7-9 shows the ANOVA results for SLS from the 3T3 NRU test method.  When the IC50 values within each laboratory were compared by study phase (i.e., the ANOVA factor was study phase), there were no statistically significant differences (p < 0.01) between study phases for any laboratory.  Table 7-10 shows that the slopes of the linear regressions of the IC50 values over time (expressed as index values) were statistically different from zero for ECBC and FAL (p = 0.001 and 0.012, respectively).  Since the slopes were so small (0.000204 and -0.000324), they were considered to be unimportant.  The slope of the IIVS regression of SLS IC50 over time was not statistically different from zero (p = 0.651; Table 7-10), which was entirely consistent with the ANOVA (Table 7-9) indicating that SLS IC50 from IIVS did not vary with study phase (p = 0.854).  The ANOVA with study phase as the factor (with laboratories combined) indicated that the 3T3 NRU IC50 values from all the laboratories were consistent over time since data from the various study phases were not statistically significantly different (p = 0.304).

# Comparison of SLS IC50 Among the Laboratories

When all study phases from each laboratory were combined, ANOVA, with laboratory as the factor, indicated that the SLS IC50 for the 3T3 NRU test method differed in some statistically significant fashion among the laboratories (p < 0.006).  However, the differences between laboratories look rather small in Figure 7-1 since the SDs for the laboratories clearly overlap one another.

Table 7-9	ANOVA Results for SLS IC50 from the 3T3 NRU Test Method

| Study Phase/ Laboratory | ECBC | FAL | IIVS |
| --- | --- | --- | --- |
|  | Log Mean IC50 (mM) | SD | N | P1 | Log Mean IC50 (mM) | SD | N | P1 | Log Mean IC50 (mM) | SD | N | P1 |
| Test for differences between phases within each laboratory |  |  |  |  |  |  |  |  |
| Phase Ia | -0.876 | 0.042 | 6 | 0.031 | -0.811 | 0.046 | 9 | 0.015 | -0.850 | 0.034 | 7 | 0.854 |
| Phase Ib | -0.864 | 0.066 | 6 |  | -0.846 | 0.065 | 8 |  | -0.838 | 0.025 | 5 |  |
| Phase II | -0.848 | 0.027 | 16 |  | -0.796 | 0.057 | 19 |  | -0.854 | 0.025 | 8 |  |
| Phase III | -0.842 | 0.036 | 36 |  | -0.851 | 0.066 | 27 |  | -0.844 | 0.041 | 23 |  |
| Test for differences between laboratories (phases combined) |  |  |  |  |  |  |  |  |
| All Phases | -0.849 | 0.039 | 64 | 0.006 | -0.826 | 0.062 | 63 |  | -0.847 | 0.035 | 44 |  |
| Test for differences between phases (laboratories combined) |  |  |  |  |  |  |  |  |
| Phase Ia | -0.839 | 0.049 | 22 | 0.304 |  |  |  |  |  |  |  |  |
| Phase Ib | -0.850 | 0.056 | 19 |  |  |  |  |  |  |  |  |  |
| Phase II | -0.831 | 0.047 | 34 |  |  |  |  |  |  |  |  |  |
| Phase III | 0.845 | 0.045 | 86 |  |  |  |  |  |  |  |  |  |

1Statistically significant at p < 0.01.

Abbreviations: N- number of values; SD – standard deviation.  Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

Table 7-10	Linear Regression Analysis of SLS IC50 Over Time1

| Test Method/ Laboratory | Slope | P-value (Slope)2 | Intercept |
| --- | --- | --- | --- |
| 3T3 NRU |
| ECBC | 0.000204 | 0.001 | -0.874 |
| FAL | -0.000324 | 0.012 | -0.796 |
| IIVS | 0.0000304 | 0.651 | -0.850 |
| NHK NRU |
| ECBC | -0.000559 | 0.002 | -1.901 |
| FAL | -0.00112 | < 0.001 | -1.737 |
| IIVS | -0.000445 | 0.002 | -1.885 |

1Time was expressed as index values. The index value of each test reflected the order of testing without respect to the time lapsing between tests.

2Statistically significant from zero at p < 0.05.

Laboratories: ECBC- U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

7.3.2	ANOVA and Linear Regression Results for the NHK NRU Test Method

SLS IC50 Variation with Time

Table 7-11 shows the ANOVA results for the NHK NRU test method.  When the IC50 values within each laboratory were compared by study phase (i.e., the ANOVA factor was phase), the phases were statistically different (p < 0.01) at each laboratory.  The IC50 values from the various study phases were also significantly different from one another when the laboratory data were combined (p < 0.001).  Linear regression analyses showed that the slopes for IC50 over time (expressed as an index values) were statistically significantly greater than zero for each laboratory (see Table 7-10).  Since the slopes were so small (-0.000559, -0.00112, and -0.000445), they were considered to be unimportant.

Table 7-11	ANOVA Results for SLS IC50 from the NHK NRU Test Method

| Study Phase/ Laboratory | ECBC | FAL | IIVS |
| --- | --- | --- | --- |
|  | Log Mean IC50 (mM) | SD | N | P1 | Log Mean IC50 (mM) | SD | N | P1 | Log Mean IC50 (mM) | SD | N | P1 |
| Test for differences between phases within each laboratory |  |  |  |  |  |  |  |  |
| Phase Ia | -1.867 | 0.135 | 5 | 0.001 | -1.656 | 0.125 | 5 | < 0.001 | -1.904 | 0.060 | 12 | < 0.001 |
| Phase Ib | -1.936 | 0.092 | 6 |  | -1.829 | 0.141 | 10 |  | -1.965 | 0.046 | 5 |  |
| Phase II | -2.007 | 0.109 | 11 |  | -1.982 | 0.173 | 15 |  | -1.863 | 0.058 | 12 |  |
| Phase III | -1.990 | 0.098 | 31 |  | -1.941 | 0.113 | 34 |  | -1.972 | 0.070 | 19 |  |
| Test for differences between laboratories (phases combined) |  |  |  |  |  |  |  |  |
| All Phases | -1.971 | 0.113 | 53 | < 0.001 | -1.879 | 0.175 | 64 |  | -1.924 | 0.073 | 48 |  |
| Test for differences between phases (laboratories combined) |  |  |  |  |  |  |  |  |
| Phase Ia | -1.833 | 0.143 | 22 | < 0.001 |  |  |  |  |  |  |  |  |
| Phase Ib | -1.891 | 0.125 | 21 |  |  |  |  |  |  |  |  |  |
| Phase II | -1.964 | 0.139 | 38 |  |  |  |  |  |  |  |  |  |
| Phase III | -1.971 | 0.100 | 84 |  |  |  |  |  |  |  |  |  |

1Statistically significant at p < 0.01.

Abbreviations: N- number of values; SD – standard deviation.  Laboratories: ECBC – U.S. Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences.

Comparison of SLS IC50 Among the Laboratories

The ANOVA results, with laboratory as a factor (Table 7-11) indicated that the SLS IC50 was statistically different among the laboratories when the data from the study phases were pooled (p < 0.001).  Figure 7-1 shows that the SLS data from ECBC and IIVS were rather similar for Phases Ia, Ib, and III.  The SLS IC50 data from FAL looks different from the other two laboratories for Phases Ia, Ib, and II, but the bars and SDs for Phase III show that the data from all laboratories were similar.

# 7.4 	Laboratory Concordance for Solvent Selection

The solvents used to dissolve the reference substances are shown in Table 7-12.  For Phases Ib and II, the SMT selected the solvents to use for cytotoxicity testing based on the solubility results provided by BioReliance (see Table 5-7) using the solubility protocol in Appendix G2.  Despite the fact that the solubility of an individual substance in 3T3 medium and NHK medium might be different, the SMT chose the same solvent for both test methods, rather than choosing one for the 3T3 assay and one for the NHK assay.  For example, if solubility in the 3T3 medium was ≥ 2 mg/mL and solubility in the NHK medium was < 2 mg/mL, and the substance was soluble in DMSO at 200 mg/mL, then the SMT selected DMSO as the solvent for cytotoxicity testing.

During Phases Ib and II, the SMT noted that BioReliance sometimes achieved higher solubility than the cytotoxicity laboratories (e.g., see the results for arsenic trioxide, aminopterin, and chloramphenicol in Table 5-7).  In an attempt to avoid the selection of a solvent for which one or more laboratories could not achieve the desired solubility, the SMT used the solubility data from all the laboratories to determine solvent selections for cytotoxicity testing in Phase III.  The SMT viewed BioReliance’s NHK and 3T3 media solubility results for each substance in Phases Ib and II to be one result for media and took a similar approach in Phase III when considering all the laboratory results to determine the solvent to use for cytotoxicity testing.  For example, if one laboratory had achieved solubility at 2 mg/mL in medium, but the other laboratories had not, and the substance was soluble in DMSO at 200 mg/mL, then the SMT selected DMSO as the solvent.  Table 7-12 shows that cell culture medium was used to test as the solvent for 38 substances and DMSO was used as the solvent for 34 substances.

The solubility protocol used by the cytotoxicity laboratories failed to guide the selection of a solvent for five substances because they were insoluble at all concentrations tested in at least one laboratory.  Arsenic trioxide was insoluble at all the cytotoxicity laboratories.  IIVS also found sodium oxalate, strychnine, and triethylenemelamine insoluble in any solvent, and FAL found thallium sulfate insoluble in any solvent.  To select a solvent for cytotoxicity testing of these substances, the SMT used the solubility results from the laboratories that did achieve solubility.

| Table 7-12 Solvent Determinations by Laboratory |
| --- |
| Reference Substance | Solvent for Testing1 | ECBC | FAL | IIVS |
| Acetaminophen | DMSO | Medium | Medium | DMSO |
| Acetonitrile | Medium | Medium | Medium | Medium |
| Acetylsalicylic acid | DMSO | Medium | DMSO | Medium |
| Aminopterin | DMSO | DMSO | DMSO | DMSO |
| 5-Aminosalicylic acid | Medium | Medium | Medium | Medium |
| Amitriptyline HCl | DMSO | DMSO | DMSO | DMSO |
| Arsenic III trioxide | Medium | ID | ID | ID |
| Atropine sulfate | Medium | Medium | Medium | Medium |
| Boric aid | Medium | Medium | Medium | Medium |
| Busulfan | DMSO | DMSO | DMSO | DMSO |
| Cadmium II chloride | DMSO | DMSO | DMSO | DMSO |
| Caffeine | Medium | Medium | Medium | Medium |
| Carbamazepine | DMSO | Medium | DMSO | DMSO |
| Carbon tetrachloride | DMSO | Medium | DMSO | Medium |
| Chloral hydrate | Medium | Medium | Medium | Medium |
| Chloramphenicol | DMSO | DMSO | DMSO | Medium |
| Citric acid | Medium | Medium | Medium | Medium |
| Colchicine | Medium | Medium | Medium | Medium |
| Cupric sulfate pentahydrate | Medium | Medium | Medium | Medium |
| Cycloheximide | Medium | Medium | Medium | Medium |
| Dibutyl phthalate | DMSO | DMSO | DMSO | DMSO |
| Dichlorvos (DDVP) | DMSO | Medium | DMSO | Medium |
| Diethyl phthalate | DMSO | DMSO | DMSO | DMSO |
| Digoxin | DMSO | DMSO | DMSO | DMSO |
| Dimethylformamide | Medium | Medium | Medium | Medium |
| Diquat dibromide monohydrate | Medium | Medium | Medium | Medium |
| Disulfoton | DMSO | DMSO | DMSO | DMSO |
| Endosulfan | DMSO | DMSO | DMSO | DMSO |
| Epinephrine bitartrate | Medium | Medium | Medium | Medium |
| Ethanol | Medium | Medium | Medium | Medium |
| Ethylene glycol | Medium | Medium | Medium | Medium |
| Fenpropathrin | DMSO | DMSO | DMSO | DMSO |
| Gibberellic acid | Medium | Medium | Medium | Medium |
| Glutethimide | DMSO | DMSO | DMSO | DMSO |
| Glycerol | Medium | Medium | Medium | Medium |
| Haloperidol | DMSO | DMSO | DMSO | DMSO |
| Hexachlorophene | DMSO | DMSO | DMSO | DMSO |
| Lactic acid | Medium | Medium | Medium | Medium |
| Lindane | DMSO | DMSO | DMSO | DMSO |
| Lithium I carbonate | Medium | Medium | Medium | Medium |
| Meprobamate | DMSO | Medium | Medium | DMSO |
| Mercury II chloride | DMSO | DMSO | DMSO | DMSO |
| Methanol | DMSO | Medium | Medium | DMSO |
| Nicotine | Medium | Medium | Medium | Medium |
| Paraquat | Medium | Medium | Medium | Medium |
| Parathion | DMSO | DMSO | DMSO | DMSO |
| Phenobarbital | DMSO | Medium | DMSO | DMSO |
| Phenol | Medium | Medium | Medium | Medium |
| Phenylthiourea | DMSO | DMSO | Medium | DMSO |
| Physostigmine | DMSO | Medium | DMSO | DMSO |
| Potassium I chloride | Medium | Medium | Medium | Medium |
| Potassium cyanide | Medium | Medium | Medium | Medium |
| Procainamide HCl | Medium | Medium | Medium | Medium |
| 2-Propanol | Medium | Medium | Medium | Medium |
| Propranolol HCl | DMSO | Medium | Medium | Medium |
| Propylparaben | DMSO | DMSO | DMSO | DMSO |
| Sodium arsenite | Medium | Medium | Medium | Medium |
| Sodium chloride | Medium | Medium | Medium | Medium |
| Sodium dichromate dihydrate | Medium | Medium | Medium | Medium |
| Sodium fluoride | Medium | Medium | Medium | Medium |
| Sodium hypochlorite | Medium | Medium | Medium | Medium |
| Sodium oxalate | Medium | Medium | Medium | ID |
| Sodium selenate | Medium | Medium | Medium | Medium |
| Strychnine | Medium | Medium | Medium | ID |
| Thallium I sulfate | Medium | Medium | ID | Medium |
| Trichloroacetic acid | Medium | Medium | Medium | Medium |
| 1,1,1-Trichloroethane | Medium | Medium | Medium | Medium |
| Triethylenemelamine | DMSO | Medium | DMSO | ID |
| Triphenyltin hydroxide | DMSO | DMSO | DMSO | DMSO |
| Valproic acid | DMSO | Medium | DMSO | DMSO |
| Verapamil HCl | DMSO | DMSO | DMSO | DMSO |
| Xylene | DMSO | DMSO | DMSO | DMSO |
| DMSO Total | 34 | 22 | 29 | 28 |
| Medium Total | 38 | 49 | 41 | 40 |

ID-insufficient data to select solvent.

1Solvents for testing as determined by the SMT and used in the study by each laboratory: Medium = cell culture medium; DMSO = dimethyl sulfoxide

ECBC – US Army Edgewood Chemical Biological Center; FAL – FRAME Alternatives Laboratory; IIVS – Institute for In Vitro Sciences

The cytotoxicity laboratories selected the same solvent for 55 of the 72 reference substances (76%).  Excluding the five substances that were found to be insoluble in any solvent by at least one laboratory, there were 12 substances for which the cytotoxicity laboratories disagreed: acetaminophen, acetylsalicylic acid, carbamazepine, carbon tetrachloride, chloramphenicol, dichlorvos, meprobamate, methanol, phenobarbital, phenylthiourea, physostigmine, and valproic acid.  Every laboratory reported relatively low solubility, ≤ 2 mg/mL, in medium for these substances.  Since 2 mg/mL in medium is the departure point for the selection medium or DMSO, a small variation in results causes the laboratories to select different solvents.  The solubility of acetaminophen, for example was reported as 2 mg/mL in culture media by ECBC and FAL, but < 2 mg/mL by IIVS.  IIVS found it soluble in 200 mg/mL DMSO and selected DMSO as the solvent.  ECBC and FAL selected the culture media as the solvent.  The SMT selected DMSO as the solvent for acetaminophen to be used by all laboratories.

  - Summary

Intra- and inter-laboratory reproducibility were assessed using ANOVA, CV analysis, comparison of the laboratory-specific IC50-LD50 regressions to one another (for each test method) and laboratory concordance for the GHS acute oral toxicity category predictions.  ANOVA permits statistical comparisons of laboratories and experimental averages, while controlling for other factors.  CV analysis is an empirical way of expressing the relative magnitudes of variability on a standardized scale.  ANOVA results for the reference substances showed significant laboratory differences for 26 substances for the 3T3 NRU test method and seven substances for the NHK test method.  Intralaboratory CV values were 1-122% for the 3T3 NRU test method and 1-129% for the NHK NRU test method.  Mean intralaboratory CV values were 26% for both test methods, but the NHK NRU test method had a lower interlaboratory CV (28% vs 46%).  Interlaboratory CV values were 2-135% for the 3T3 NRU test method and 1-99% for the NHK NRU test method.  FAL had the highest mean intralaboratory CV for both test methods (33% for the 3T3 NRU test method and 42% for the NHK NRU test method).

An analysis to determine the relationship between the chemical attributes and interlaboratory CV indicated that physical form, solubility, and volatility had little effect on CV.  CV seemed to be related, however, to chemical class, GHS acute toxicity category, IC50, and boiling point.  Reference substances in the amide class had unusually low mean interlaboratory CV values for both the 3T3 NRU test method (15%) and NHK NRU test method (16%) compared with the overall mean interlaboratory CV values (46% for the 3T3 NRU test method and 28% for the NHK NRU test method).  Reference substances in the organophosphate and heterocyclic classes had unusually high mean interlaboratory CV values for the 3T3 NRU test method (74% and 71%, respectively), but not for the NHK NRU test method.  Mean interlaboratory CV values were large for substances in the most toxic GHS acute categories, especially for the 3T3 NRU test method.  The mean interlaboratory CV for substances in the classes for LD50 ≤ 5 mg/kg (69%) and 5 < LD50 ≤ 50 mg/kg (78%) were larger than the mean overall interlaboratory CV (46%,) for the 3T3 NRU test method.  For the NHK NRU test method, the mean interlaboratory CV was 37% for substances with LD50 ≤ 5 mg/kg and 41% for substances with 5 < LD50 ≤ 50 mg/kg while the mean overall interlaboratory CV was 28%.  A Spearman correlation analysis indicated that IC50 was negatively correlated to interlaboratory CV for both 3T3 (p = 0.015) and NHK (p = 0.024) NRU test methods and that boiling point was positively correlated to interlaboratory CV (p = 0.021) for the 3T3 NRU test method.

The analysis of interlaboratory reproducibility by evaluating the similarity of the laboratory specific IC50-LD50 regressions indicated that the laboratory regressions for both test methods were not significantly different (p < 0.05) from one another (p = 0.796 for the 3T3 NRU and p = 0.985 for the NHK NRU).  The evaluation of laboratory concordance for the prediction of GHS acute oral toxicity category when the laboratory-specific IC50 data were applied to the same regression yielded the following proportions of substances for which all laboratories agreed on the GHS acute oral toxicity categorization:

- 78% (52/67) for the 3T3 NRU and 87% (59/68) for the NHK NRU with the RC regression

- 81% (52/64) for the 3T3 NRU and 91% (59/65) for the NHK NRU with the RC rat only weight regression

- 84% for the both test methods (36/43 for the 3T3 NRU and 37/44 for the NHK NRU) with the RC rat only weight regression excluding substances with specific mechanisms of action

ANOVA results for the positive control, SLS, IC50 in the 3T3 NRU test method indicated that there were significant differences among laboratories (p = 0.006) and but not between study phases within laboratories (p > 0.01).  However, interlaboratory CV values, which ranged from 2% to 10% for the study phases, indicated that the laboratories were similar.  Intralaboratory CV values for the study phases ranged from 5% to 24%.  SLS IC50 values for the NHK NRU test method were more variable than those for the 3T3 NRU test method.  ANOVA results for SLS in the NHK NRU test method indicated that there were significant differences between laboratories (p < 0.001) and between study phases within laboratories (p ≤ 0.001).  A change in cell culture methods at FAL decreased the SLS IC50 from Phase Ib to Phase II.  Intralaboratory CV values for the NHK NRU SLS IC50 during the various study phases ranged from 11% to 51%.  Interlaboratory CV values for SLS in the NHK NRU test method ranged from 8% to 39%.

Cell culture medium was used as the solvent for testing 38 substances and DMSO was used for 34 substances.  The laboratory concordance in selecting solvent for the reference substances using the solubility protocol was 76% (55/72).

[This Page Intentionally Left Blank]