SVX II Silicon Strip Detector Upgrade Project

TEST FIBER INTERFACE BOARD

--PRELIMINARY--

August 18, 1995

S. Zimmermann, J. Andresen, J. Chramowicz, H. Gonzalez,

K. Treptow, T.Zmuda

Document # ESE-SVX-950817

1. GENERAL INFORMATION

1.1 System Introduction

1.2 Description of Component & How it Fits into the System

1.3 List Of Component Requirements

2. THEORY OF OPERATION AND OPERATING MODES

2.1 Basic Features & Operation

2.2 Configuration of The TPC and SVX-II Chips

2.3 High Level Command Emulation

2.4 Data FIFO AB and C

2.5 G-Link Boards

2.6 Test Port Card Analog to Digital Converters

2.7 Calibration Inject

2.8 Number of Clocks for the SVX II Analog to Digital Converter

3. EMBEDDED & DIAGNOSTIC/DEVELOPMENT SOFTWARE

4. INTERFACE SPECIFICATIONS

4.1 VMEbus Interface

4.1.1 Addressing Modes

4.1.2 Data Cycles Types

4.1.3 VME Register Descriptions

4.2 STAR/SRC Interface

4.2.1 J3 Backplane Control Connector

4.2.2 Front Panel Control Connector

4.2.3 Command Protocol

4.3 Test Port Card Interface

4.4 G-Link Interface

4.5 Front Panel

5. RESET, POWER UP RESET & INITIALIZATIONS,

6. DIAGNOSIS

6.1 Diagnosis of the TFIB and TPC

6.1.1 Reset

6.1.2 Cables and Drivers/Receivers between TFIB and TPC

6.1.3 Configuration of the SVX-II Chips or FPGA

6.1.4 HDI Cables and Partial Test of SVX-II Chips

6.2 Diagnosis of the Command Lines between STAR/SRC and TFIB

6.3 Diagnosis of the G-Links

6.4 Status Latch

7. ELECTRICAL & MECHANICAL SPECIFICATIONS

7.1 Connectors and Dip Switches

7.2 Packaging & Physical Size

7.3 PC Board Construction

7.4 Power Requirements

7.5 Cooling Requirements

8. SAFETY FEATURES & QUALITY ASSURANCE PROCEDURES

8.1 Module Fusing & Transient Suppression

8.2 Other Safety & Quality Assurance Subsections

9. APPENDICES

9.1 List Of Component Documentation

9.2 Schematics

9.3 PAL, FPGA Equations

9.4 Timing Diagrams

9.5 Parts List

9.6 Additional Appendices

GENERAL INFORMATION

System Introduction

This document describes the “TEST FIBER INTERFACE BOARD”, hereafter referred to as the TFIB. The TFIB is a 9Ux400 VME card which interfaces to a customized J3 backplane and the J1 VMEbus backplane. It is one of the modules of a test stand designed to control and readout the SVX-II chips [ref. 1]. Figure 1 illustrates the basic components of the test stand which are the SVX Test Acquisition and Readout (STAR) board [ref. 2], the TFIB, the Test Port Card (TPC) board [ref. 3], and the MVME162/7 motherboard. The TFIB board can also be interface with the Silicon Readout Controller (SRC) [ref. 4] board through the J3 backplane, allowing further testing.

Description of Component & How it Fits into the System

The TFIB will be used to fully test the SVX-II chips and silicon detectors. This testing will be used in the development of the next generation of the CDF SVX detector. The control of the SVX-II chips is divided into two boards which are the TFIB and the TPC. The TFIB executes commands from three different sources:

Commands delivered by the STAR board or by the SRC board through the J3 backplane or front panel.

Commands previously stored by the VME motherboard inside the Configuration/Command FIFO.

Commands executed via selectable bits of the TFIB’s Control Register.

When the TFIB executes commands from its Configuration/Command FIFO or Control register, the TFIB works independently of the STAR/SRC. These commands are executed synchronously with the TPC. The TFIB interprets these commands and delivers microcommands to the TPC. The word “microcommands” refers specifically to the sequence of bits that the TFIB-CT sends to the Test Port Card controller (TPC-CT). The TPC generates the logic levels to control the internal features of the SVX-II chips. One TPC can control up to three chains of SVX-II chips which are mounted on a hybrid. These hybrids are connected to the TPC through the High Density Interconnect (HDI) cable.

The following is the flow of commands and clocks from the STAR/SRC to the SVX-II chips:

The STAR or the SRC sends a 53 MHz clock and commands to the TFIB through the J3 backplane or front panel.

The TFIB interprets the commands from the STAR/SRC and sends the SVX-II clock, the serial clock, and microcommands to the TPC using differential PECL levels.

The TPC controller interprets the microcommands and delivers the proper CMOS logical level sequences to the SVX-II chips through the HDI.

The clock sequences for the SVX-II chips are generated directly by the TFIB. These clocks are fanned out by the TPC to the three HDIs.

The data from the SVX-II chips travels through the following path:

SVX-II chips transmit the data to the TPC, through the HDI, using CMOS levels.

The TPC converts these signals to differential PECL levels and transmits the data to the TFIB through three 8 bit parallel wire links.

The TFIB board accepts data from three of these HDI links.

The STAR board accepts data from three of these HDI links.

The STAR/SRC board has a large eight-bit buffer to hold the data from the SVX-II chips. The TFIB has three 2K x 8 buffers which corresponds to one buffer for the data from each HDI.

The following sections give detailed descriptions of the TFIB.

List Of Component Requirements

The major components of the TFIB design are as follows:

The TFIB is a 9 U x 400 mm VME card.

The TFIB is connected to the J1 and J3 backplanes of the VME crate. Connections to J2 are not implemented.

The J1 backplane follows the VMEbus specification (ANSI/IEEE STD1014) [ref. 5].

The J3 backplane is a custom made backplane [ref. 6]. The STAR board sends commands and clocks to the TFIB through this backplane. Pins on the J3 connector provide -5.2 V for the ECL logic on the TFIB.

The STAR supplies one command every 132 ns.

The TFIB controls one TPC. The TFIB is capable of reading from all three HDIs connected to the TPC. The readout data rate is 26.5 MBytes/sec in each HDI.

The digital electrical interface between TFIB and TPC is differential PECL using AT&T 41 Series chips [ref. 7].

Two G-Link daughter boards can be connected to the TFIB. The G-Link daughter boards are used to send data that has been read by the TFIB to the SAR board [ref. 8].

The TFIB supplies the TPC with the Calibration, Ramp-Reference and Ramp-Pedestal analog voltages.

Cooling for the TFIB will be supplied by fans mounted on the VME crate.

The TFIB requires +5.0 V and -5.2 V power supplies.

THEORY OF OPERATION AND OPERATING MODES

Basic Features & Operation

Figure 2 is a block diagram of the TFIB. The main features of the TFIB are the following:

The execution of commands sent by the STAR/SRC or from the Configuration/Command FIFO.

Saves the sequence of commands into a Silo FIFO.

Synchronizes and shapes the appropriate clock for the SVX-II chips.

Translates these commands into microcommands to the TPC.

Receives and stores into the Data FIFOs the parallel data from the TPC.

Appends HDI identification to the data from the TPC.

Allows diagnostic testing of the SVX-II chips and TPC.

Delivers the Calibration, Ramp-Pedestal and Ramp-Reference bit stream to the DACs on the TPC.

Controls the configuration of the TPC controller which is a FPGA.

Reads back the configuration of the TPC controller.

Control the configuration of the SVX-II chips.

Read back the configuration of the SVX-II chips.

For future testing, allows the installation of boards with G-Links and laser diodes

All units of the TFIB operate synchronously with a 53 MHz clock. This 53 MHz clock is supplied by the STAR/SRC through the J3 backplane or a connector on the front panel. The TFIB also has the capability to operate independently of the STAR/SRC as the TFIB has its own 53 MHz clock generator. The source of the clock for the TFIB is selected by one bit of the Control Word and DIP Switch S1.

The functions of the TFIB are controlled by the TFIB controller (TFIB-CT). The TFIB-CT has several VME registers. One of the functions of these VME registers is to configure the TFIB-CT to execute commands. In this document, the word “commands” refers specifically to commands that the following units pass to the TFIB-CT:

The STAR/SRC through the J3 backplane or front panel.

The list of commands previously downloaded into the Configuration/Command FIFO.

The VME interface through the Control Register.

Depending on which unit supplied the command, the TFIB-CT has three modes of operation which are named Real Mode, Emulation Mode, and Immediate Mode. When the TFIB-CT executes commands received from the STAR/SRC, it is working in the Real Mode. When the TFIB-CT executes the list of commands from the Configuration/Command FIFO, it is in Emulation Mode. When the TFIB-CT executes commands from the Control Register, it is in Immediate Mode.

The main function of the Real Mode and Emulation Mode are for the data acquisition of the SVX-II chips. However, there are commands for calibration of the SVX-II chips, for the resetting of the TFIB, TPC, and SVX-II chips, for the regaining of lock of the G-Links, and for diagnostic testing. These commands are listed in Section 4.2.3 of this document. The Emulation Mode is implemented to allow the TFIB-CT to test all commands that the STAR/SRC can request independently of these STAR/SRC modules. This Emulation Mode is a powerful tool for diagnostic testing as well as for the development of the TFIB. The VME interface writes to the Configuration/Command FIFO the list of commands to be executed when in the Emulation Mode. The VME interface then uses the Immediate Command code 1 (Execute Emulation Mode) of the Control Register to trigger the Emulation Mode.

The Immediate Mode is used for configuration of the SVX-II chips and the FPGA of the TPC. This mode is also used for diagnostic testing.

The VME motherboard can monitor the Executing Real/Emulate Command and Executing Immediate Command bits to check when the FIB is executing commands. When executing commands, the VME motherboard should access solely the Status and Control registers of the TFIB. For example, consider that the TFIB is executing real commands. To access other registers inside the TFIB, the VME motherboards first has to disable real commands (set to “one” bit 5 of the Control Low Register). Then, it has to monitor if the TFIB already finished the execution of the present real command. Finally, it can access the registers of the TFIB.

When the TFIB-CT receives a command, it executes the appropriate operation. If the command is associated with the TPC, the TFIB-CT sends the necessary serial microcommands (through the line SR-CMD) to the TPC. The TFIB receives commands, interprets these commands, and when necessary sends microcommands to the TPC. The bits of the SR-CMD are validated by the Serial Clock (SR-CLK). After the microcommand is delivered to the TPC-CT, the TFIB and the TPC work synchronously. The TFIB-CT sends SR-CLK to the TPC-CT to advance the state machine of the TPC-CT. The TFIB-CT delivers these clocks with the appropriate timing required by the SVX-II chips and TPC-CT.

In Real Mode, the STAR/SRC delivers commands (CMD [4:0]), SYNC pulse, advance pipeline (AD-PIPE), and the 53 MHz master clock (MCLK) through the J3 backplane or front panel. A set of 2_inputs/1_output multiplexers, controlled by a DIP switch S1, is used to select the source of these signals between the J3 backplane or front panel. Section 4 of this document describes the timing of this interface. When the TFIB is in this mode, the source of clock is the J3 backplane or front panel. When the TFIB-CT receives the commands, TFIB-CT executes the requested actions.

The SVX-II chip clock has different shapes and frequencies [ref. 1]. The maximum and minimum delay of the SVX-II chip acquisition clock has to be controlled within strict boundaries in order to guarantee that all SVX-II chips controlled by different TFIBs and TPCs will be acquiring data synchronously. When the TFIB-CT is in Real Mode with commands being supplied by the J3 backplane and the SVX-II chips are in acquisition mode, the SVX-II Clock Shaper formats the SVX-II chip acquisition clock using the internal advance pipeline signal (INT-AD-PIPE) and the J-SYNC pulse delivered by the J3 backplane. When the same is happening, but the commands being supplied by the front panel, the TFIB uses the INT-SYNC pulse to generate the acquisition clock. The INT-SYNC pulse is synchronous with the master clock (the D flip-flop synchronizes it). For all other operations, the clock is supplied directly by the TFIB-CT, by the SVX_GEN_CLK line. The signal SVX_CLK_SEL, supplied by the TFIB controller, informs the SVX-II Clock Shaper when to use the SYNC pulse of the J3 backplane. AT&T differential PECL drivers are used to send the SVX-II Clock, Serial Clock, and microcommands to the TPC.

One of the commands executed in Real and Emulation modes is the Calibration Inject. This command forces the SVX-II chip to inject charge into its own pre-amplifiers. The TFIB-CT uses a delay line, programmable by the VMEbus, to time this command in relation to the acquisition clock. So, when the TFIB-CT recognizes this command, it sends the proper microcommands to the TPC-CT, then, it times the Serial Clocks to advance the state machine of the TPC-CT using this delay line. With this feature, one can control the timing relation between the acquisition clock and Calibration Inject operation.

The Silo FIFO is a 2K x 9 FIFO which stores the code of all commands delivered through the J3 backplane or front panel. The VME motherboard can access this FIFO and read out the sequence of these commands. When the FIFO gets full, the TFIB-CT removes the old commands from the FIFO in order to be sure that  the most recent 2K commands are saved. The store can be disabled by a bit in the Control Register or by the execution of the Latch Status command.

The bottom portion of the TFIB block diagram has a set of components to receive the data read out from the TPC. The TPC can read out data from a total of three High Density Interconnects (HDI). The data is transmitted to the TFIB using the same AT&T PECL protocol. The AT&T receivers translate this data to TTL electrical levels. The main components of this section is the set of three Data FIFOs (A, B and C). These Data FIFOs are used to synchronize the data from different HDIs and for diagnostic testing. The different modes of operation of these Data FIFOs are set by the Control Register located inside the TFIB-CT. The TFIB-CT communicates with the FIFO controller logic (DFIFO-CT) which then generates the proper clocking for the various components of this data readout section of the logic.

The HDI IDs are a set of latches accessible through the VME interface. The VME motherboard downloads the HDI identification number into this latches. The HDI ID is appended to the top of the SVX-II data.

The TFIB reads out data from the SVX-II chips and transmits this data to the SAR through the G-Links. Further details on the format of this data are in the Data FIFO AB and C section of this document. In this Real mode of operation, the FIFOs are used to synchronize data from different HDIs. Before the TFIB-CT actually starts the readout of the data, it requests to the DFIFO-CT to store the HDI ID into the FIFOs. Then, the TFIB-CT sends the readout clock for the SVX-II chips and informs the DFIFO-CT to enable the data write operation. The SVX-II clock (A-CLK, B-CLK or C-CLK clock signals in Figure 2) arrives in parallel with the data through the AT&T receivers. Figure 3 shows the relation between this data and clock. New data is available at each transition of the SVX-II clock which operates at 26.5 MHz. The delay between the transition of the clock and valid data on the output of the SVX-II chip has not been determined, therefore there are delay lines to set the proper relationship between data and clocks.

The DFIFO-CT logic delays the clock in order to synchronize the clock with valid data. Then, it shapes the clock to a duty cycle of approximately 50% (Latch Clock in Figure 3). This The Latch Clock stores the data into the input latches. After another delay, the data is transferred to the Data FIFOs. As soon as the two FIFOs connected to one G-Link have three or more words, the DFIFO-CT reads the data out of the FIFOs and sends the information through the G-Links.

The HDI Enable register allows the VME motherboard to enable the input circuitry associated with the HDI that are properly connected and operational. The TFIB-CT synchronizes itself with the SVX-II chips, by means of special data patterns End Of Readout (EOR). The EOR informs to rest of the data acquisition system that all data from that HDI was read out. When the DFIFO-CT detects this pattern, it informs the TFIB-CT. The TFIB-CT sends at least four more SVX-II readout clocks (STAR requirement) and then suspends the SVX-II readout clock to conclude the readout operation. Of course, the DFIFO-CT should not look for EOR in HDI that are not operational.

As previously mentioned, this section of the TFIB has other operating modes. The G-Links may be disabled resulting in all SVX-II data staying inside the Data FIFOs. The VME motherboard can then transfer this data to its own memory through the use of the VME interface. Another mode of operation allows the VME motherboard to actually write different patterns into the Data FIFOs. The VME motherboard can read back the Data FIFOs, or the DFIFO-CT can transmit the pattern in the Data FIFOs through the G-Links.

For configuration of the SVX-II chips and the TPC-CT, the TFIB-CT uses the Configuration/Command FIFO and specific microcommands. When the VME motherboard requests the download of the configuration of the SVX-II chips, the TPC-CT sends microcommands to the TPC to set the SVX-II chips into configuration mode. The TPC-CT then reads the configuration bytes from the Configuration/Command FIFO, serializes the configuration bytes, and sends them to the TPC, using the Serial Microcommand line. A similar operation occurs when the TFIB-CT configures the TPC-CT.

The Configuration/Command FIFO is not connected directly to the VME interface. The data flows through the TFIB-CT because the TFIB-CT also has to write to this FIFO when it is uploading the configurations of the SVX-II chips or the TPC-CT. When the TFIB-CT is uploading the configuration, the bit stream arrives as one bit of the data path connected to the TPC in parallel with the clock. This is similar to Figure 3 except that it is one data bit instead of DATA[7:0]. The bit is latched but not stored inside the Data FIFOs. The DFIFO-CT intercepts this data stream and forwards the bits to the TFIB-CT. The TFIB-CT assembles a byte and stores it inside the Configuration/Command FIFO. The VME motherboard can then check the consistency of the information.

The VME motherboard also has to program serial digital to analog converters (DAC) on the TPC. They set analog voltages used for calibration, ramp-pedestal and ramp-reference. The TFIB has several VME registers that are used for this task. When the motherboard writes to these registers, a specific controller serializes the data, and sends it to the TPC. There, it is used to program the DAC. Soon after dowloanding a DAC, the DAC Controller reads back the data stored inside the DAC Readback Data register.

Configuration of The TPC and SVX-II Chips

The TFIB allows the VME motherboard to download and upload the configuration of the SVX-II chips connected to one TPC and the FPGA mounted on the TPC. These operations are performed through the following VME registers:

Configuration/Command FIFO Register.

HDI address Register.

Number of chips/HDI Register.

Control Register.

Status Register.

The VME motherboard sees the download and upload of the configuration of the SVX as one operation. The TFIB actually performs these download/upload sequence automatically. The VME motherboard writes the configuration bits to the Configuration/Command FIFO. Then, the VME motherboard programs the HDI Address Register and the number of SVX-II chips connected to the HDI into the Number of Chips/HDI Register. Next, it requests the immediate command Download/Upload Configuration of SVX-II chips of the Control Register which triggers the sequence. The CFIFO controller reads the Configuration/Command FIFO and sends the appropriate clocks, microcommands, and serial data string to the TPC. The TPC controller programs the SVX-II chips. Then, the TFIB controller sends the appropriate clocks and microcommands to read back the configuration of the SVX chips. The CFIFO controller reads the data string and stores it into the Configuration/Command FIFO. The VME motherboard may poll the Executing Immediate Command bit of the Status Register to recognize when the operation is completed. Finally, the motherboard can check if the uploaded data is correct.

To download the FPGA of the TPC, the VME motherboard writes the configuration to the same Configuration/Command FIFO. Then, the VME motherboard request the immediate command Download Configuration of the FPGA in the Control Register which triggers the download. The CFIFO controller reads the FIFO and sends the appropriate clocks and data string to the FPGA. The VME motherboard may poll the Executing Immediate Command bit of the Status Register to recognize when the operation is completed.

To upload the configuration of the FPGA, the VME motherboard uses the same Configuration/Command FIFO as before. The VME motherboard request the immediate command Upload Configuration of the FPGA in the Control Register which triggers the upload. The TFIB controller sends the appropriate clocks and microcommands. The CFIFO controller reads the data string and stores the data string into the Configuration/Command FIFO. The VME motherboard may poll the Executing Immediate Command bit of the Status Register to recognize when the operation is completed. The VME motherboard reads the Configuration/Command FIFO to check the configuration.

We have also implemented the immediate command Reset Configuration of the FPGA, which allows the VME motherboard to clear the configuration of the FPGA if the FPGA is not correctly downloaded or if the configuration is corrupted. Then, the motherboard can re-configure the part. It is good practice to reset the configuration of the FPGA before downloading its configuration. Finally, before downloading or uploading any configuration one should reset the CFIFO, to guarantee that it is empty. To perform this operation, use the Reset Configuration/Command FIFO bit of the Configuration/Command FIFO Control/Status register.

The motherboard has to store the configuration bits of the SVX-II chips into the Command/Configuration FIFO as indicated in Table 1 (C0 is configuration bit 0, etc.).

Data Lines

D7 - - - - - - D0

C7 - - - - - - C0

C13 - - - - - - C8

:

:

:

X X C181 -- C176

C7 - - - - - - C0

etc.

Table 1. Configuration data format SVX-II chips.

Note that the number of bits to download is not a multiple of eight. Therefore, the last byte have two bits with “don't care” information (X). Normally, when configuring the SVX-II chips, it is necessary to download several chips which are connected to the same HDI. Table 1 shows the bits for the first chip, followed by the bits of the second chip, etc. The configuration bits of the next SVX-II chip start at bit 0 of the next byte. When uploading the configuration, the bits are saved in the same format.

The configuration bit stream for the FPGA are available in a Motorola S3 format file. The user has to remove the control characters of this format before storing then into the Configuration/Command FIFO. Three ZERO bytes (0x0) have to be append in the end of the configuration. The comparison of the readback bit stream of the configuration of the FPGA is not straight forward, because the data does not have the same format that was downloaded. See the Xilinx manual [ref. 12] for more details or the software drivers developed for this task (see Ref. 12).

High Level Command Emulation

The TFIB can emulate the high-level commands that it receives from the STAR/SRC. The VME motherboard stores the commands, one or several commands, into the Configuration/Command FIFO. The motherboard starts the emulation by asserting the immediate command Execute Emulation Mode in the Control Register. The TFIB controller will perform the operations in the same way it operates when it receives commands from the STAR/SRC. The VME motherboard may poll the Executing Immediate Command bit of the Status Register to recognize when all emulated commands were executed. Note that the bit Executing Real/Emulate Command of the Status Register will also be asserted.

One of the commands that the TFIB can emulate is the readout of the SVX-II chips. The Disable G-Links bit in the Control Register disables the G-Links and forces the data into the Data FIFOs. Normally, the Data FIFOs are used to synchronize different data paths from the TPC. The VME motherboard can read this data from the Data FIFOs and check data consistency. This allows the VME motherboard to read the data from the SVX chips without the STAR board.

Data FIFO AB and C

There are a set of three 2K x 8 FIFOs with each FIFO associated with one HDI. The Data FIFOs are shown in Figure 2. These FIFOs receive data from the VMEbus or the AT&T receivers and are read by the VMEbus.

The Data FIFOs have two main functions:

The primarily function is the synchronization of the data coming from different HDIs. The data comes from the TPC in parallel with a clock and is stored inside the FIFOs. When all FIFOs have three or more words, the data is transmitted through the G-Links.

Other functions are associated with the following diagnostic operations:

a) the VME motherboard can download a known data pattern into these FIFOs and use the     G-Links to ship the data to the SAR.

b) the G-Links may be disabled with the data that comes from the TPC being read by the VME motherboard instead of being transmitted to the SAR.

The data from the TPC has the format shown in Table 3 (a). End Of Readout (EOR) is differentiated from CHIP ID and Channel Number by the two most significant bits (see Table 2). Further details are provided within the SVX-II and TPC specifications.

| Control Byte | Bit Assignment |
| --- | --- |
| Chip ID | 10dddddd |
| Channel # | 0ddddddd |
| EOR | 11dddddd |

Table 2. Control Byte and Bit Assignment

The TFIB board appends the HDI identification number to the beginning of the data coming from the TPC. This is shown in Table 3 (b).

(a)			(b)

CHIP ID

HDI MSB ID

STATUS

HDI LSB ID

CH #

CHIP ID

DATA

STATUS

CH #

CH #

:

DATA

CHIP ID

CH #

STATUS

:

CH #

CHIP ID

:

STATUS

EOR

CH #

:

EOR

Table 3. Data format (a) from Test Port Card, (b) with HDI ID appended

G-Link Boards

The present test stand does not require the G-Links. However, in order to provide a way to further enhance and test the concept of the final data acquisition system, the TFIB has an interface that allows connection to G-Link and Finisar transmitters. These G-Link and Finisar transmitters are assembled on daughter boards which are plugged into the top of the TFIB [ref. 8]. Two of these daughter boards can be connected to the TFIB. Figure 4 is a diagram of this daughter card.

The implementation of this high speed data link as a self-contained board makes for greater flexibility of the overall data acquisition system in view of future upgrades. The daughter card concept permits convenient independent testing of the high speed communication system. Maintenance is simplified due to the boards’ interchangeability. The approximate dimensions of this daughter board is 4.5" x 2". The SMA connector allows for the bypassing of the Finisar Laser Diode. For more information about this interface, refer to the G-Link/Finisar Transmitter and Receiver Boards specifications [ref. 9 and 10].

Test Port Card Analog to Digital Converters

The TPC has a total of 18 12 bits serial Digital to Analog Converters (DAC) that control the various reference voltages of the three HDIs. The TFIB has the DAC Controller which remotely controls the DACs on the TPC.

The VME motherboard writes the Select Code into the DAC Select Register and the 12 bit value into the DAC Data Register. As soon as the DAC Controller recognizes that the VME motherboard finished the write to the DAC Data Register, it initiates the transfer to the TPC. It assembles a serial stream compose by part of the select code and the data and transmit to the TPC, in addition with several control signals required. After the data is properly transferred to the select DAC, the controller serially uploads the data and stores into the DAC Readback Data Register. The VME motherboard has to wait for 8 ms before the readback data is available. Presently, there is no circuit inside the TFIB to indicate when the readback is done and therefore the VME motherboard has to use its own resources to count this time.

Calibration Inject

The Calibration Inject command forces the SVX-II chip to inject charge into its own pre-amplifiers. The TFIB-CT uses a delay line, programmable by the VMEbus, to time this command in relation to the acquisition clock. Figure 5 shows the relationship among the several signals associated with the calibration inject and the SVX-II data sampling. Note that, presently, the input of the delay line is timed by the second raising edge of the master clock (MCLK) after the SYNC pulse.

Number of Clocks for the SVX II Analog to Digital Converter

The SVX-II chips have an analog-to-digital Gray Code counter, which counts until it reaches the Counter Module value programmed inside the configuration parameter bits [174:167] of the SVX-II chip. The analog to digital conversion is completed when the Gray counter reaches the Counter Module value, and the TFIB can proceed with the readout. Presently, the TFIB does not have a way to program the number of digitization clocks to deliver to the Test Port Card. It always delivers enough clocks to convert the maximum possible value programmed in the Counter Module register of the SVX-II chip.

EMBEDDED & DIAGNOSTIC/DEVELOPMENT SOFTWARE

The control and diagnosis of the TFIB and TPC are done through a MVME162/7 VME motherboard. The TFIB is a slave of this motherboard. This VME motherboard controls and executes the diagnostic features of the TFIB. The software is being currently designed. The document “TFIB Diagnostic Software” describes this software [ref. 11].

INTERFACE SPECIFICATIONS

The TFIB board has the following interfaces, which we will now describe:

With the TPC.

With the STAR/SRC through the specially designed J3 backplane.

With VMEbus through the J1 backplane.

With G-Link transmitters to connect with the SAR board during future testing.

VMEbus Interface

The TFIB is a VMEbus slave which implements A24:D16 addressing. Block transfers are not implemented. Presently, the TFIB does not generate VMEbus errors neither VMEbus interrupts. The VME interface has the following characteristics:

The TFIB board is always a VMEbus slave. The TFIB board has no VME master capability.

The VME interface is implemented through the J1 connector.

Only a 24 bit addressing scheme is allowed.

The VME interface only executes single data transfers.

The VME interface allows 16 or 8 bit data transfers.

The VME interface does not generate interruptions.

The VME interface generates VMEbus errors.

Addressing Modes

The TFIB is addressed using a 24 bit address mode and the VME Address Modifiers either 0x29 or 0x39. The base address has a fixed pattern, encoded inside a programmable logical device (PAL U86 on the schematics), and a variable one, which depends if the board is connected to the J3 backplane. We will call Addressing Mode 1 when the board is connected to the J3 backplane, and Addressing Mode 2 when it is not.

The fixed part of the address encoded inside the PAL is compared with address lines A[23:16] of the VME interface. Presently it is programmed for pattern 0x10 (we may change if necessary).

If the board is connected to the J3 backplane (Addressing Mode 1), the address lines A[15:12] are compared with the geographic address supplied by the backplane. Table 4 shows the relationship between A[15:12] and GA[3:0]. When the higher order 12 bits of the VME address, A[23:12], compare with the eight PAL bits and with the four bits GA[3:0], the lower address lines A[11:0] select one register inside the TFIB. The TFIB recognizes that it is being addressed by the VME motherboard and performs the requested action.

| GEOGRAPHIC ADDRESS | ADDRESS LINE |
| --- | --- |
| GA3 (msb) | A15 |
| GA2 | A14 |
| GA1 | A13 |
| GA0 | A12 |

Table 4. Geographic Address And Address Line Relationships

If the board is NOT connected to the J3 backplane (Addressing Mode 2), the address lines A[15:12] are compared with the setting of DIP switch S2 (see Section 7.1 for position and assignment of this DIP switches). Table 40 shows the relationship between A[15:12] and the position of the DIPs. Bits are set to zero by having the switch in the on position. Bits are set to one by having the switch in the off position. When the higher order 12 bits of the VME address, A[23:12], compare with the eight PAL bits and A[15:12] with the four DIP switch positions, the lower address lines A[11:0] select one register inside the TFIB. The TFIB recognizes that it is being addressed by the VME motherboard and performs the requested action.

Data Cycles Types

The TFIB only implements 8 and 16 bit single data transfers. The TFIB does not use the signal lines on the J2 connector and does not have block transfer capability. The following contains a description of each TFIB register.

VME Register Descriptions

Table 5 shows the VME registers and associated information.

| REGISTER | TYPE | OPERATION | ADDRESS |
| --- | --- | --- | --- |
| Status | byte | read only | Base Address + 0x00 |
| Status Latch | byte | read only | Base Address + 0x02 |
| Control Low | byte | read/write | Base Address + 0x04 |
| Control High | byte | read/write | Base Address + 0x06 |
| Spare |
|  | HDI Address | byte | read/write | Base Address + 0x0A |
| Conf/Cmd FIFO Control/Status | byte | read/write | Base Address + 0x0C |
| Number of Chips/HDI | byte | read/write | Base Address + 0x0E |
| Conf/Cmd FIFO | 9 bits | read/write | Base Address + 0x10 |
| G-Link Control/Status | byte | write/read | Base Address + 0x12 |
| Data FIFOs Control/Status | byte | write/read | Base Address + 0x14 |
| HDI_AB Identification LSB | word | read/write | Base Address + 0x16 |
| HDI_AB Identification MSB | word | read/write | Base Address + 0x18 |
| HDI_C Identification LSB | byte | read/write | Base Address + 0x1A |
| HDI_C Identification MSB | byte | read/write | Base Address + 0x1C |
| HDI_AB Contents | word | read only | Base Address + 0x1E |
| HDI_C Contents | byte | read only | Base Address + 0x20 |
| HDI Enabled | byte | read/write | Base Address + 0x22 |
| Data FIFO AB | word | read/write | Base Address + 0x24 |
| Data FIFO C | byte | read/write | Base Address + 0x26 |
| Silo FIFO | 9 bits | read only | Base Address + 0x28 |
| Silo FIFO Control/Status | byte | read/write | Base Address + 0x2A |
| SVX-II Pipeline Reset Time | byte | write only | Base Address + 0x2C |
| Calibration Inject Delay Line | byte | write only | Base Address + 0x2E |
| Spare |
|  | Base Address + 0x30 |
| VME Control/Status (N.I.) | byte | read/write | Base Address + 0x32 |
| DAC Data | word | read/write | Base Address + 0x34 |
| DAC Select | byte | write/write | Base Address + 0x36 |
| DAC Readback Data | word | read only | Base Address + 0x38 |

Table 5. VME Registers

Status Register

Table 6 provides the Status Register bit assignment.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | Real/Emulate State Machine [3] |
| 5 | Real/Emulate State Machine [2] |
| 4 | Real/Emulate State Machine [1] |
| 3 | Real/Emulate State Machine [0] |
| 2 | Executing Real/Emulate Command |
| 1 | Executing Immediate Command |
| 0 | Real command source: J3 => 0, Front Panel => 1 |

Table 6. Status Register

Real/Emulate State Machine [3:0]: Indicates the state of the state machine which controls the execution of real/emulate commands. Table 7 describes the state assignment. Note that read pipeline/digitize/readout is controlled by another state machine, not accessible through the Status register.

| STATE NUMBER | MEANING |
| --- | --- |
| 10 | G-Link send fill frames |
| 9 | Port Card controller reset |
| 8 | TFIB reset |
| 7 | SVX chip reset |
| 6 | Start preamplifier reset |
| 5 | Enter in data sampling mode phase 1 |
| 4 | Enter in data sampling mode phase 0 |
| 3 | Stop preamplifier reset |
| 2 | Calibration inject phase 1 |
| 1 | Calibration inject phase 0 |
| 0 | Idle |

Table 7. State assignment for the Real/Emulate State Machine

Executing Real/Emulate Command: when set to ONE, the TFIB-CT is executing a Real command.

Executing Immediate Command: when set to ONE, the TFIB-CT is execution an Immediate Command.

Real command source: This bit allows the readout of the position of the first switch of DIP switch S1 which selects the source of Real commands (see Section 7.1). When this bit is ZERO, the TFIB receives Real commands through the J3 backplane, when it is ONE, it receives Real commands through the front panel.

Status Latch

The Status latch is a byte wide latch which is used for diagnostic purposes. When the TFIB controller executes the Latch Status command (see Section 4.2.3), the Status Latch saves the contents of the Status Register. The VME motherboard can read this register. The Status latch is not resetable and have the same bit assignment of the Status Register (see Table 6, Status Register).

Control Low Register

Table 8 gives the bit assignment for Control Low register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | TFIB reset |
| 6 | 53 MHz clock source: 0 => external, 1 => internal |
| 5 | Disable Real Commands |
| 4 | Execute immediate command |
| 3 | Immediate command [3] |
| 2 | Immediate command [2] |
| 1 | Immediate command [1] |
| 0 | Immediate command [0] |

Table 8. Control Low Register

TFIB reset: When this bit is set to one, it resets the TFIB. It is a volatile bit, nad has no meaning when readout out by VME motherboard.

53 MHz clock source: This bit selects the source of the 53 MHz clock for the TFIB. In order to change the 53 MHz source and guarantee that the TFIB will work properly, the VME motherboard has to execute the following sequence:

The TFIB should not be executing any Immediate command.

Disable Real commands.

Check if all Real commands are done (bit Executing Real/Emulate Commands of the Status Register has to be zero).

Change the clock source

DIP Switch S1 selects the source of external clock: the front panel or the J3 backplane (see Section 7.1 for position and assignment of this DIP switch).

Disable Real Commands: When set to ONE, this bit disables the external commands and, therefore, the TFIB ignores all commands from the STAR/SRC. It may happen that, when the Real commands are disabled, the TFIB is still executing a Real command whose execution started prior the disabling. Therefore, the VME motherboard should check the Executing Real command bit of the Status Register to know when the TFIB is idle.

Execute immediate command: When this bit is set to one, the TFIB executes the immediate command. It is a volatile bit, and has no meaning when readout by VME motherboard. To execute Immediate commands, the Real commands have to be disabled and not in execution.

Immediate command [0:3]: The Immediate commands are encoded into bits 0 to 3 of the Control Register. Table 9 shows the Immediate command codes ICMD[0:3]. Codes 9 to 0xF are not presently used.

| ICMD[0:3] | MEANING |
| --- | --- |
| 8 | Reset SVX-II chips |
| 7 | Erase configuration of the FPGA |
| 6 | Reset Port Card controller |
| 5 | Download/upload configuration of the SVX-II chips |
| 4 | Upload configuration of the FPGA |
| 3 | Download configuration of the FPGA |
| 2 | G-Link transmit Fill Frames |
| 1 | Execute Emulation mode |
| 0 | No operation |

Table 9. Immediate Command Codes

The TFIB-CT uses the Executing Immediate Command of the Status register to inform the VME motherboard when the operation is completed.

Control High Register

Table 10 gives the bit assignment for Control High register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | not used |
| 3 | not used |
| 2 | not used |
| 1 | not used |
| 0 | Silo FIFO store disable |

Table 10. Control High Register

Silo FIFO store disable: When set to one, the store of Real commands inside the Silo FIFO is disabled. The Silo FIFO also can be automatically disable (and this bit is automatically set to ONE) when the TFIB executes the Real command Latch Status.

High Density Interconnect Address Register

Table 11 shows the HDI Address Register bit assignments.

| Bit 1 | Bit 0 | HDI selected |
| --- | --- | --- |
| 0 | 0 | all |
| 0 | 1 | A |
| 1 | 0 | B |
| 1 | 1 | C |

Table 11. HDI Address Register

This register is used to select the proper hybrid which is connected to one HDI. There can be three HDIs connected to one TPC. Certain operations are associated with specific HDIs as in the case of the specific HDI used to configure the SVX-II chips. This register is only used in immediate mode. Bits [2:7] are not used.

Configuration/Command FIFO Control/Status Register

Table 12 shows the bit assignment of the Configuration/Command FIFO Control/Status Register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | not used |
| 3 | not used |
| 2 | Configuration/command FIFO Flag 2 |
| 1 | Configuration/command FIFO Flag 1 |
| 0 | Reset Configuration/command FIFO |

Table 12. Conf./Com. FIFO Control/Status

Configuration/Command FIFO Flag 1 and 2: See Table 13. Note that this flags may change during a VME read, if the Configuration/Command FIFO is in use.

| FLAG 1 | FLAG 2 | STATE |
| --- | --- | --- |
| 0 | 0 | Empty |
| 1 | 0 | Almost empty: 1 to 16 bytes |
| 1 | 1 | Intermediate range: 17 to 2031 bytes |
| 0 | 1 | Almost full or full: 2032 to 2048 bytes |

Table 13. FIFO Flag True Table

Reset Configuration/Command FIFO: When set to one, the Configuration/command FIFO is reseted. In order to guarantee the reset, the TFIB can not be executing commands associated with the Configuration/command FIFO, i.e., executing emulating commands or configuring the SVX-II chips or FPGA. It is a volatile bit and is not readable.

Number of SVX-II Chips Connected to HDI Register

The number of SVX-II chips per HDI is programmed in this register. This information is used to download/upload the configuration of the SVX-II chips. Bits No.Chips/HDI[0:4] is a binary code, programmed from 0 to 31 chips. Zero means one chip, while 31 means 32 chips. Therefore, this register should be loaded with the number of chips minus one. Note that this register should be programmed before any download/upload operation can be executed. Table 14 shows the bit assignment.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | No.Chips/HDI [4] MSB |
| 3 | No.Chips/HDI [3] |
| 2 | No.Chips/HDI [2] |
| 1 | No.Chips/HDI [1] |
| 0 | No.Chips/HDI [0] |

Table 14. Number of Chips/HDI Register

Configuration/Command FIFO

The Configuration/Command FIFO is a 2K x 9 FIFO which is used for configuration and the emulation of commands. Table 15  shows the bit assignment of this FIFO when it is used for emulation commands.

| BIT NUMBER | MEANING |
| --- | --- |
| 8 | spare 2 |
| 7 | spare 1 |
| 6 | BXING |
| 5 | AD_PIPE |
| 4 | CMD[4] |
| 3 | CMD[3] |
| 2 | CMD[2] |
| 1 | CMD[1] |
| 0 | CMD[0] |

Table 15. Conf./Com. FIFO assignment for Emulation commands

Section 2.2 provides further information about the format of the data when this FIFO is used for configuration of the TPC controller (a FPGA) or the SVX-II chips.

Due to the simplicity of the VME interface being used, the following procedure is recommended to read the FIFOs:

a) Check if the FIFO is empty by reading the Status/Control Register.

b) If the FIFO is empty, the readout is done.

c) If the FIFO is not empty, read one more byte.

d) Go to step (a).

G-Link Control/Status Register

The G-Link transmitter supplies status information. This information is shown in Table 16 and is readable through the G-Link Status Register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | G-Links transmit DFIFOs DONE |
| 3 | G-Links transmit DFIFOs |
| 2 | Disable G-Links |
| 1 | G-Link_1, Locked |
| 0 | G-Link_0, Locked |

Table 16. G-Link Status Register

G-Link_0 and _1 Locked: This is a read only bit. It indicates the lock status of the transmitter’s PLL. A ONE indicates lock. This signal has to be debounced by the VME motherboard through several reads.

Disable G-Links: This is a read/write bit. When set to one, disable the G-Links, forcing the SVX-II data to stay into the Data FIFOs.

G-Links Transmit DFIFOs: When set to one, the DFIFO-CT transmit the contents of the DFIFOs through the G-links. This is a volatile bit, and is not readable.

G-Links Transmit DFIFOs DONE: This bit is set to ONE when the transmission of the content of the DFIFOs by the G-links is done. It is a read only bit, it does not change when the VME motherboard writes to this register.

Data FIFOs Control/Status Register

This register allows the VME motherboard to read the status and reset the three Data FIFOs. Table 17 shows the bit assignment of this register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | Data FIFO_C Flag 2 |
| 5 | Data FIFO_C Flag 1 |
| 4 | Data FIFO_B Flag 2 |
| 3 | Data FIFO_B Flag 1 |
| 2 | Data FIFO_A Flag 2 |
| 1 | Data FIFO_A Flag 1 |
| 0 | Reset Data FIFOs |

Table 17. Data FIFOs Status Register

Reset Silo FIFO: When set to one, the Data FIFOs are reseted. In order to guarantee the reset, the Data FIFOs should be idle. It is a volatile bit and is not readable.

Silo FIFO Flag 1 and 2: See Table 13 for the meaning of flags 1 and 2 associated with the FIFOs. Note that this flags may change during a VME read, if the Data FIFOs are in use.

HDI_AB Identification LSB Register

The value of this 16 bit register is appended to the beginning of the data read out from the TPC. They are the least significant bytes (LSB) of the HDI identification number for the A and B HDI and are programmable by the VME motherboard. The bit assignment is shown in Table 18.

| HDI Ident. | Data Lines |
| --- | --- |
| B | D[15:8] |
| A | D[7:0] |

Table 18. HDI_AB Identification LSB Register

HDI_AB Identification MSB Register

The value of this 16 bit register is appended to the beginning of the data read out from the TPC. They are the most significant bytes (MSB) of the HDI identification number for the A and B HDI and are programmable by the VME motherboard. The bit assignment is shown in Table 19.

| HDI Ident. | Data Lines |
| --- | --- |
| B | D[15:8] |
| A | D[7:0] |

Table 19. HDI_AB Identification MSB Register

HDI_C Identification LSB Register

The value of this 8 bit register is appended to the beginning of the data read out from the TPC. They are the least significant byte (LSB) of the HDI identification number for the C HDI and are programmable by the VME motherboard. The bit assignment is shown in Table 20.

| HDI Ident. | Data Lines |
| --- | --- |
| C | D[7:0] |

Table 20. HDI_C Identification LSB Register

HDI_C Identification MSB Register

The value of this 8 bit register is appended to the beginning of the data read out from the TPC. They are the least significant byte (MSB) of the HDI identification number for the C HDI and are programmable by the VME motherboard. The bit assignment is shown in Table 21.

| HDI Ident. | Data Lines |
| --- | --- |
| C | D[7:0] |

Table 21. HDI_C Identification MSB Register

HDI_AB Contents Register

When the VME motherboard reads this register, the TFIB sets on the VME data bus the actual bit pattern of HDIs A and B. Table 22 shows the bit assignment of this register.

| HDI Contents | Data Lines |
| --- | --- |
| B | D[15:8] |
| A | D[7:0] |

Table 22. HDI_AB Content Register

HDI_C Contents Register

When the VME motherboard reads this register, the TFIB sets on the VME data bus the actual bit pattern of HDI C. Table 23 shows the bit assignment of this register.

| HDI | Data Lines |
| --- | --- |
| C | D[7:0] |

Table 23. HDI_C Content Register

HDI Enable Register

The TFIB can be connected up to three HDIs. This register allows the VME motherboard to enable the input circuitry associated with the HDI that are connected by setting the appropriate bit to ONE. Also, note that in case of failure of one HDI, this register can also be used to disable that specific part of the circuit. Table 24 shows the bit assignment of this register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | not used |
| 3 | not used |
| 2 | Enable HDI C |
| 1 | Enable HDI B |
| 0 | Enable HDI A |

Table 24. HDI Enable Register

Data FIFO_AB

These are two 2K x 8 FIFOs with each FIFO associated with HDI A and HDI B. Their primary use is the synchronization of the data coming from different HDIs. These two FIFOs are connected to the Data FIFO_AB Register as a 16 bit word. Table 25 provides further details on these two FIFOs.

| HDI Ident. | Data Lines |
| --- | --- |
| B | D[15:8] |
| A | D[7:0] |

Table 25. Data FIFO AB

The G-Links can be disabled with the Disable G-Link bit of the Control Register. In this case, all data from the TPC are stored inside these data FIFOs. The format of the data is the same as the data readout from the SVX-II chips but with HDI identification appended to the beginning of the data as shown in Table 2.

Due to the simplicity of the VME interface being used, the following procedure is recommended to read the FIFOs:

a) Check if the FIFO is empty by reading the Data FIFO Status/Control Register.

b) If the FIFO is empty, the readout is done.

c) If the FIFO is not empty, read one more word.

d) Go to step (a).

Data FIFO_C

This is a 2K x 8 FIFO associated to HDI C. Its primary use is the synchronization of the data coming from different HDIs. This third FIFO is connected to Data FIFO C Register as a 8 bit byte. Table 26 provides further details on this FIFO.

| HDI Ident. | Data Lines |
| --- | --- |
| C | D(0:7) |

Table 26. Data FIFO C

Due to the simplicity of the VME interface being used, the following procedure is recommended to read the FIFOs:

a) Check if the FIFO is empty by reading the Status/Control Register.

b) If the FIFO is empty, the readout is done.

c) If the FIFO is not empty, read one more byte.

d) Go to step (a).

Silo FIFO

The Command Silo FIFO is a 9 bits wide 2K FIFO used for diagnostic purposes. The bit assignment of this FIFO is the same as the Configuration/Command FIFO when it is used to emulate commands  (see Table 15).

All REAL MODE commands executed by the TFIB are stored in this FIFO. When it gets full, the older commands are removed from the FIFO. The VME motherboard can read this register for diagnostic purposes. The Silo FIFO is not resetable by the TFIB Reset command or the TFIB Reset bit in the Control register. When the TFIB executes the Latch Status command, it disable any further stores inside the Silo FIFO. The VME motherboard has to reset the Silo FIFO Store Disable bit in the Control register in order to enable again the store.

Due to the simplicity of the VME interface being used, the following procedure is recommended to read the FIFOs:

a) Check if the FIFO is empty by reading the Status/Control Register.

b) If the FIFO is empty, the readout is done.

c) If the FIFO is not empty, read one more byte.

d) Go to step (a).

Silo FIFO Control/Status Register

Table 27 shows the bit assignment of this register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | not used |
| 3 | not used |
| 2 | Silo FIFO Flag 2 |
| 1 | Silo FIFO Flag 1 |
| 0 | Reset Silo FIFO |

Table 27. Silo FIFO Control/Status

Reset Silo FIFO: When set to one, the Silo FIFO is reseted. In order to guarantee the reset, the Silo FIFO has to be disabled. It is a volatile bit and is not readable.

Silo FIFO Flag 1 and 2: See Table 13.

SVX-II Pipeline Reset Time Register

This register programs the pipeline reset pulse width of the acquisition clock of the SVX-II chip. When it is set to ZERO, the reset pulse width is 19 ns, when it is set to 15, the reset pulse width is 34 ns. This register's bits are shown in Table 28.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | not used |
| 3 | Pipeline reset time [3] MSB |
| 2 | Pipeline reset time [2] |
| 1 | Pipeline reset time [1] |
| 0 | Pipeline reset time [0] |

Table 28. SVX-II Preamp Speed Register

Calibration Inject Delay Line Register

This register allows the VME motherboard to define the time relationship between the SVX-II acquisition clock and the calibration inject pulse.

Figure 5 shows the relationship between clock, SYNC pulse, SVX-II Acquisition clock and the trigger of the delay line TFIB controller triggers this delay line. When the delay expires, the TFIB controller sends a calibration inject pulse to the Port Card Controller. The delay is programmed from 5 ns to 315 ns, with increments of 5 ns. Table 29 shows the bit assignment of this register.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | Calibration Inject Delay [5] |
| 4 | Calibration Inject Delay [4] |
| 3 | Calibration Inject Delay [3] |
| 2 | Calibration Inject Delay [2] |
| 1 | Calibration Inject Delay [1] |
| 0 | Calibration Inject Delay [0] |

Table 29 Calibration Delay Line

VME Control/Status Register (Not Implemented)

This register indicates which unit generated the VME bus error. Table 30 shows its bit assignment.

| BIT NUMBER | MEANING |
| --- | --- |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | DACs controller bus error |
| 3 | Conf/Cmd FIFO controller bus error |
| 2 | Data FIFO controller bus error |
| 1 | TFIB controller bus error |
| 0 | Reset bus errors status |

Table 30. VME Status/Control Register

Reset bus error status: When set to ONE, the bus error bits of this register are reseted. This is a volatile bit, and is not readable.

DAC Data Register

This register is used to specify the data to be downloaded to the DAC selected by the DAC Select register. The TPC uses 12 bit DACs and they correspond to VME data bits 11 to 0. The programming of a selected DAC is initiated as soon as the VME write operation to the DAC Data register is completed.

DAC Select Register

The DAC Select register is used to specify the DAC that will be programmed by the logic on the TFIB. The DAC assignment is shown on Table 31.

| DAC Select CODE | Data Selected |
| --- | --- |
| 0x00 | HDI-A Calibration f |
| 0x01 | HDI-A Calibration z |
| 0x02 | HDI-A Ramp Ref. f |
| 0x03 | HDI-A Ramp Ref. z |
| 0x04 | HDI-A Ramp Ped. f |
| 0x05 | HDI-A Ramp Ped. z |
| 0x08 | HDI-B Calibration f |
| 0x09 | HDI-B Calibration z |
| 0x0a | HDI-B Ramp Ref. f |
| 0x0b | HDI-B Ramp Ref. z |
| 0x0c | HDI-B Ramp Ped. f |
| 0x0d | HDI-B Ramp Ped. z |
| 0x10 | HDI-C Calibration f |
| 0x11 | HDI-C Calibration z |
| 0x12 | HDI-C Ramp Ref. f |
| 0x13 | HDI-C Ramp Ref. z |
| 0x14 | HDI-C Ramp Ped. f |
| 0x15 | HDI-C Ramp Ped. z |

Table 31. DAC Select Code

DAC Readback Data Register.

This register is used to store the DAC data transmitted by the TPC everytime that a DAC is downloaded. Under proper operation, the contents of this register must match the data stored on the DAC Data Register. Note that the transmission of data from the TFIB to the TPC and from the TPC to the TFIB are serial and enough time (more than 8 ms) should be allowed for the Readback register to have to correct data.

STAR/SRC Interface

The STAR/SRC communication is done through the J3 backplane of the VME crate or a connector on the front panel of the TFIB. The J3 backplane was designed by Fermilab’s D0 and has more lines available than the TFIB requires.

The next sections describe the electrical and logical protocol of this interface.

J3 Backplane Control Connector

The following is a short description of the J3 backplane STAR/SRC interface signals that the TFIB will use. For further details, refer to specific documentation [ref. 4].

•	Fan out the 53 MHz differential ECL clock signal named MCLK.

•	Fan out a 18.9 ns wide, 132 ns period, "SYNC" pulse in differential ECL.

•	Command(4:0) named CMD[4:0] which are bussed TTL level signals.

•	Advance-Pipeline named AD-PIPE and Beam Crossing named BXING (no presently used) which are TTL level signals.

•	Provide for geographic addressing of the TFIB Modules. These lines are connected to ground and require pull-up resistors.

•	Provide +5.0 Vdc, -5.2 Vdc, and Ground.

•	Provide a "MAC 2" fiber optic backplane connector for future use.

Table 32 shows the pin assignment of the J3 VME backplane connector. Table 33 gives the correspondence between Geographic Address and VME slot number.

| PIN NUMBER | ROW A | ROW B | ROW C |
| --- | --- | --- | --- |
| 1 | GND | GND | J-MCLK |
| 2 | GND | GND | J-MCLK* |
| 3 | GND | GND | J-SYNC |
| 4 | GND | -5.2V | J-SYNC* |
| 5 | GND | NC | -5.2V |
| 6 | GND | NC | -5.2V |
| 7 | +5V | NC | -5.2V |
| 8 | +5V | J-AD-PIPE | -5.2V |
| 9 | +5V | J-BXING | -5.2V |
| 10 | +5V | NC | -5.2V |
| 11 | +5V | NC | J-CMD0 |
| 12 | +5V | NC | J-CMD1 |
| 13 | GA3 | NC | J-CMD2 |
| 14 | GA2 | NC | J-CMD3 |
| 15 | GA1 | NC | J-CMD4 |
| 16 | GA0 | NC | NC |

Table 32. J3 backplane pin assignment

| SLOT NUMBER | GA[0:3] |
| --- | --- |
| 6 | 0 |
| 7 | 1 |
| : | : |
| 20 | 0xE |
| 21 | 0xF |

Table 33. Geographic address and VME slot number

Front Panel Control Connector

The front panel Control connector uses a 20 pin 100 mils spacing connector. The electrical protocol is differential positive ECL provided by AT&T receivers. It is differentially terminated by 100 ohm resistors. Table 34 shows the pin assignment. The TFIB employs a jumper (J13) to open the ground, in case is necessary to avoid ground loops.

| Pin number | signal |
| --- | --- |
| 1 | F-CMD0* |
| 2 | F-CMD0 |
| 3 | F-CMD1* |
| 4 | F-CMD1 |
| 5 | F-CMD2* |
| 6 | F-CMD2 |
| 7 | F-CMD3* |
| 8 | F-CMD3 |
| 9 | F-CMD4* |
| 10 | F-CMD4 |
| 11 | F-AD-PIPE* |
| 12 | F-AD-PIPE |
| 13 | F-BXING* |
| 14 | F-BXING |
| 15 | F-SYNC* |
| 16 | F-SYNC |
| 17 | F-MCLK* |
| 18 | F-MCLK |
| 19 | GND |
| 20 | GND |

Table 34. Front Panel Command Connector Assignment

Command Protocol

The commands that the STAR/SRC sends to the TFIB controller are associated with the modes of operation of the SVX-II chip and the data acquisition. These commands are sent in parallel with the SYNC line. A zero-to-one transition in the SYNC line validates the commands. Table 35 lists the code assignment of these commands.

| CODE | MEANING |
| --- | --- |
| 0x15 | G-Links send Fill Frames |
| 0x14 | Reset TPC Controller |
| 0x13 | Reset TFIB |
| 0x12 | Reset SVX-II chip |
| 0x11 | Latch status |
| 9 | Calibration inject |
| 8 | Stop data sampling clocks |
| 5 | Read-pipeline/Digitize/Readout |
| 4 | Preamplifier start reset |
| 3 | Enter in data sampling mode |
| 2 | Preamplifier stop reset |
| 0 | No operation |

Table 35. Real and Emulation Mode Code Assignment

Commands 0x12, 0x13, and 0x14 set the SVX-II chip, the TFIB, and the TPC into known states. Command 0x12 does not re-configure the SVX-II chips but sets the SVX-II chips to a known state (specifically, to the data sampling mode, but no SVX-II acquisition clock is delivered). Command 0x15 forces the G-Link transmitter to send fill frames to allow the G-Link receiver on the SAR to regain lock. Command 0x11 is used for diagnostic purposes (see Section 4.1.3.2). The commands that are not listed have the same behavior as No Operation but may be used in future upgrades and should be considered RESERVED.

The STAR/SRC forwards these commands and signals every 132 ns. CMD[4:0] and AD-PIPE are validated by the SYNC pulse and change half way between SYNC pulses as is shown in Figure 6.

The TFIB-CT can NOT execute more than one command at a time as the TFIB-CT does not have pipeline capability. After the STAR/SRC sends a command, the STAR/SRC has to wait until the end of the execution of this command, which can take several multiples of 132 ns. However, this rule has exceptions that we will now describe:

the Latch Status and Stop Data Sampling Clocks command can be sent when another command is in execution.

the execution of Preamplifier Start Reset and Preamplifier Stop Reset commands takes more than 132 ns (see Table 36), and is normally issue during data sampling mode. However, during data sampling, after the STAR/SRC issued one of these commands, a trigger may arrive. In this case, the STAR/SRC does not need to wait until the completition of the Preamplifier Start or Stop Reset. The STAR/SRC can issue the Read-Pipeline/Digitize/Readout command right after the Preamplifier Start or Stop Reset commands.

The number of SYNC pulses to execute different SVX-II operations, including the one used to validate the command, are listed in Table 36. The operation Read-pipeline/digitize /readout is not listed because the TPC supplies the EOR control byte to inform when the operation is completed.

| OPERATION | NUMBER OF SYNC PULSES |
| --- | --- |
| Reset TPC Controller | 10 |
| Reset TFIB | 300 ms |
| Reset SVX-II chip | 6 |
| Latch status | 1 |
| Calibration inject (minimum) | 6 |
| Calibration inject (maximum) | 8 |
| Stop data sampling clocks | 1 |
| Preamplifier start reset | 4 |
| Enter in data sampling mode | 7 |
| Preamplifier stop reset | 2 |

Table 36. Number of SYNC pulse to execute different SVX-II operations

Test Port Card Interface

The TPC receives microcommands from the TFIB through the Serial-Clock (SR-CLK) and Serial Microommand (SR-CMD) lines. These microcommands when executed perform specific operations on the SVX-II chips.

Sequencer  microcommands are executed by sending the microcommand and then a specific number of clocks to cycle the sequencer through its states. The state of the SVXII control lines is determined by the sequencer state and the timing between states is determined by the time between clocks.  The TPC will not except another microcommand until the previous command is completed.

The electrical communication is accomplished by several sets of differential Positive ECL (PECL) lines which are driven by AT&T chips. There are several spare AT&T transmitters connected between the TFIB and TPC. These lines are implemented for future use and are not currently used. Data is received from the TPC from eight data lines in parallel with the SVX-II clock. Further information about the TPC interface is contained in the Test Port Card documentation [ref. 3].

G-Link Interface

The electrical interface between the G-Link and the TFIB is described in G-Link/Finisar Transmitter and Receiver Boards specification [ref. 9 and 10]. The data being read out from two hybrids are combined into one G-Link transfer as is shown in the example of Table 37.

| HDI # x | HDI # y |
| --- | --- |
| HDI MSB ID | HDI MSB ID |
| HDI LSB ID | HDI LSB ID |
| CHIP ID | CHIP ID |
| STATUS | STATUS |
| CH # | CH # |
| DATA | DATA |
| CH # | CH # |
| : | : |
| CHIP ID | CH # |
| STATUS | DATA |
| CH # | CH # |
| : | DATA |
| CH # | EOR |
| DATA | XXX |
| EOR | XXX |

Table 37. G-Link or Data FIFO data format

The data of two HDIs are transferred at a rate of 53 MHz per word with a word being 16 bits. The data transmitted has appended a HDI identification word, the MSB ID, and LSB ID. These two bytes correspond to two bytes used to identify the source of the data packet. EOR is flagged independently on each HDI data packet by asserting bits 6 and 7 of the data byte (see Table 2). The Data FIFO AB and C Section provides further explanation on this topic. The EOR is inserted by the TPC controller as soon as it recognizes that the last SVX-II chip of the HDI is done transmitting its data. The readout data from one HDI can finish before the other, and then we have unknown data (X’s on the Table) being asserted in that HDI. The HDI data packet is formed by an even number of bytes.

The outputs of the Finisar LEDs are connected to MAC-II connectors through short fiber optic links which are assembled as part of the J3 Backplane. This is shown in Figure 8.

Front Panel

Figure 7 shows the front panel of the TFIB.

RESET, POWER UP RESET & INITIALIZATIONS,

The TFIB is reseted as the result of four different actions:

Power up

Press the reset button on the front panel

The execution of the Real/Emulate command Reset TFIB

The VME motherboard asserts the TFIB Reset bit of the Control Register.

All registers and FIFOs are reseted as a result of these actions, with three exceptions: the Status and Status Latch registers, as well as the Silo FIFO. These units are associated with diagnostic features of the TFIB, which we keep the information intact for future examination. Therefore, if for some reason the TFIB is not operating properly, the user can execute some of these resets, and then examine the registers. Of course, on power up, the content of  the Silo FIFO is unknown, and the VME motherboard has to specifically reset it to guarantee that it is empty.

Below, we describe some procedures to assist the setup of the TFIB:

If  the TFIB is not connect to an external Master Clock (MCLK) through the J3 backplane or front panel, it will not operate until the internal MCLK is enabled. For this, the user has to set to ONE the 53 MHz Clock Source bit of the Control Register. Note that after reset this bit is logical level ZERO.

To execute Immediate Commands, the Real Commands have to be disabled. Therefore, bit Disable Real Commands of the Control Register has to be set to ONE. After reset, this bit is logical level ZERO, and the Real commands are enabled.

The Silo FIFO is enabled after reset.

The FIFOs should be reset before they are used. This guaranties that they are empty. However, there is no need to do this when the TFIB reads out data from the SVX chips. The Data FIFOs are automatically reset by the DFIFO Controller before it stores the data into the Data FIFOs.

The VME motherboard should employ a timeout, when it is reading out data from the SVX chips. The TFIB has no internal timeout to detect the absence of End Of Readout (EOR), and it will continue sending readout clocks until the VME motherboard aborts the operation.

DIAGNOSIS

Diagnostic features are an important part of this system. There are several methods to test and check the proper behavior of the TFIB. As an example, the TFIB can operate in Emulation Mode independent of the STAR/SRC. Also, the TFIB can read back the configuration of the SVX-II chips. The use of the test stand will point out new procedures to diagnose the TFIB and associated electronics.

The calibration mode is a powerful diagnostic method of the test stand. The calibration mode produces a known set of data which can be used to test the various inputs of the SVX-II chips as well as the overall test stand at maximum frequency. It is necessary to provide means to separately test the various facilities of the test stand. The following is a description of these tests.

Diagnosis of the TFIB and TPC

The focus of the proposed diagnostic method is to test the various units associated with the TFIB and TPC. These associated units are the following:

The TPC controller.

The SVX-II chips.

The cables.

The AT&T drivers and receivers.

The HDI cables.

The following is a description of these tests.

Reset

The TFIB, TPC, and SVX-II chips can be set into a known state through Reset. After Reset, the state of the data/control lines of the SVX-II chips is known. Specifically on the case of the reset of the SVX-II chips, a VME read of the HDI_AB and HDI_C Contents registers allows the VME motherboard to read the current contents of the HDI.

Cables and Drivers/Receivers between TFIB and TPC

(NOT CURRENTLY 	IMPLEMENTED) The Byte Pattern Register is used to test the cables and the drivers/receivers.

Configuration of the SVX-II Chips or FPGA

The configuration of the SVX-II chips can be read back to determine if the appropriate bits in the configuration of the SVX-II has been set. When read back, these bits should be equal to the value that was downloaded. The same with the FPGA used for the TPC controller. However, the verification of the configuration of the FPGA is more involving (see ref. 12).

HDI Cables and Partial Test of SVX-II Chips

There are two ways of testing the HDI cables and partially test the SVX-II chips. One method is to perform a readout of the SVX-II chips without actually performing a digitization. The other method is to perform a full read pipeline/digitize/readout with the configuration READOUT ALL CHANNEL bit of the SVX-II chips asserted.

Readout Without Digitization

(NOT CURRENTLY IMPLEMENTED) To perform this test, the READOUT ALL CHANNELS configuration must be unasserted, which means that the SVX-II chip must be in the sparsifier mode. The SVX-II chip ID is programmed during configuration. This test allows the readout of SVX-II chip ID and the status of all of the SVX-II chips in a HDI that do not contain data. The SVX-II chip ID can be configured to fully test the HDI lines. Table 38 shows the format of the data.

NOTE: RAY YAREMA DID NOT TEST THIS MODE OF OPERATION. HE BELIEVES IT WILL WORK AS PRESENTED.

HDI

MSB ID

LSB ID

CHIP ID

STATUS

CHIP ID

STATUS

:

:

CHIP ID

STATUS

EOR

Table 38. Data FIFO format for readout without digitize

Readout with Digitization

To perform this test, the READOUT ALL CHANNELS configuration of the SVX-II chips must be asserted. The TFIB executes the read-pipeline/digitize/readout command. Meaningful information is not connected to the input of the SVX-II chips. The data will be invalid, but the Chip ID and the Channel Number are known. The full readout will set the Channel Number from ZERO to 127. This will test all of the HDI lines as well as the Priority In/Priority Out scheme of readout. This mode can also be used to test part of the configuration of the SVX-II chips such as the Chip ID.

Diagnosis of the Command Lines between STAR/SRC and TFIB

The STAR/SRC sends commands and clock pulses to the TFIB. In order to test this connection, the motherboard can read the Command Silo FIFO. This FIFO saves all REAL MODE commands.

Diagnosis of the G-Links

For further testing of the TFIB, two G-Links can be installed. The G-Links are assembled in a daughter board which is plugged into the top of the TFIB. The VME motherboard writes a data pattern to the Data FIFOs and set the G-Links Transmit DFIFOs bit of the G-Link Control/Status register. The TFIB transmits, through the G-Links, the data stored inside the Data FIFO. The G-Links Transmit DFIFOs DONE bit of the G-Link Control/Status register indicates when the operation is completed.

The G-Link Status register and the immediate command G-Link Send Fill Frame of the Control Register are facilities available to help diagnose the G-Links as well as regain synchronization with the receiver.

Status Latch

When the TFIB-CT executes the Latch Status command, it saves the Status register into the Status Latch register. The VME motherboard can later check the contend of the Status Latch register. The Status Latch is a non-resetable register.

ELECTRICAL & MECHANICAL SPECIFICATIONS

Connectors and Dip Switches

Figure 8 shows the location of limo connectors and DIP switches.

Both the SVX-CLK and 53 MHz-CLK limos are single-ended ECL output designed to connect to a 50 ohms coaxial cable, terminated to ground by a 50 ohms load. The SVX-CLK limo shows the clocks sent to the SVX chips, while the 53 MHz-CLK is the master clock of the board. The 53 MHz-CLK has three sources: front panel commands connector, the J3 backplane command connector or the internal 53 MHz clock generator (for selection of this clock, see DIP switch S1 and Control Low Register).

DIP switch S1 is used to select the external source of the 53 MHz clock (Master Clock) for the TFIB. This DIP switch is enabled when the 53 MHz Clock Source bit of the Control register is set to ZERO. Table 39 show the DIP switch S1 assignment.

| DIP SWITCH NUMBER | FUNCTION |
| --- | --- |
| 4 | Not used |
| 3 | Not used |
| 2 | Not used |
| 1 | External Clock select: Open => Front panel, Close => J3 backplane. |

Table 39. DIP Switch S1 Assignment

DIP switch S2 is used to set the base address to the TFIB (see Section 4.1.1). Table 40 shows the assignment of DIP switch S2. Bits are set to zero by having the switch in the on position. Bits are set to one by having the switch in the off position.

| DIP SWITCH OPEN = 1, CLOSE = 0 | ADDRESS LINE |
| --- | --- |
| 8 | not used |
| 7 | not used |
| 6 | not used |
| 5 | not used |
| 4 | A15 |
| 3 | A14 |
| 2 | A13 |
| 1 | A12 |

Table 40. DIP Switch S2 Assignment

Packaging & Physical Size

The TFIB is a 9Ux400 card.

PC Board Construction

Power Requirements

The TFIB requires +5.0V and -5.2V power supplies.

Cooling Requirements

Cooling for the TFIB will be supplied by fans mounted on the VME crate.

SAFETY FEATURES & QUALITY ASSURANCE PROCEDURES

Module Fusing & Transient Suppression

The TFIB card is protected with fuses and transorbs.

Other Safety & Quality Assurance Subsections

APPENDICES

List Of Component Documentation

Schematics

PAL, FPGA Equations

Timing Diagrams

Parts List

Additional Appendices

REFERENCES

[1] R. Yarema et. al., "A Beginners Guide to the SVX-II", Fermilab document.

[2] J. Oliver, "SVX Test Acquisition and Readout Board".

[3] H. Gonzalez, K. Treptow, S. Zimmermann, J. Andresen "Test Port Card", Fermilab document, May 20, 1994.

[4] C. Gay and J. Oliver, "Silicon Readout Controller".

[5] VMEbus Specifications, ANSI/IEEE STD1014

[6] M. Baert, "D0 Silicon Acquisition and Readout Electronics: Custom J3 Backplane Specifications".

[7]  AT&T, "The 41 Series High-Performance Line Drivers, Receivers and Transceivers".

[8] “Low Cost Gigabit Rate Transmit/Receive”, Hewlett Packard, November 1993.

[9] C. Bonacorso, "G-Link/Finisar Transmitter and Receiver Boards", Fermilab document, SVX-II Silicon Strip Detector Upgrade Project.

[10] “Finisar Low Cost Gigabit Optics + G-Link Chips”, Finisar Corp.

[11] D. Walsh and S. Zimmermann, "TFIB Diagnostic Software", Fermilab document.

[12] “The Programmable Logic Data Book,” Xilinx Corp., 1994.

# Footnotes

The FPGA becomes the TPC Controller after it is programmed.

X = Don't care

0x means hexadecimal.

The two HDI ID bytes and the first byte (Chip ID) of the first SVX-II chip.

d is data bit.

The term "byte" is 8 bits. The term “9 bits” is a 9 bits. The term "word" is 16 bits.

Both Control Low and High are reference as Control Registers.

N.I. = Not Implemented

Bits that are marked as "not used" have no meaning, and they are readable or writable as ZERO or ONE. When the VME motherboard reads these bits, it should mask them off.

The signal BXING is not presently used by the TFIB logic.

NC = Not Connected

Not presently used.

This parameter does not use “number of SYNC pulses”. This time is given by the reset circuitry of the TFIB, formed by a MAX707 manufactured Maxim. It maximum reset pulse is 280 ms.

# Text boxes

Figure 1. Test Stand

MVME162/7

SVX Chips

6U to 9U

Adapter

HDI

HYBRID

TPC

J3 BACKPLANE

TFIB

VME Crate

STAR

DFIFO

Controller

M

U

X

FRONT-PANEL

J3 BACKPLANE

Figure 2. TFIB Block Diagram

Delays

&

Logic

TFIB

Controller

CFIFO

Controller

Com/Conf

FIFO

MUX

F-CMD[0:4]

J-CMD[0:4]

J-SYNC

J-AD-PIPE

J-MCLK

F-AD-PIPE

F-SYNC

F-MCLK

53 MHz

Clock

Delay

D

Q

Silo

FIFO

RE-CMD

CLOCK

EXT-CMD

EXT-AD-PIPE

EXT-SYNC

PC_INT

Controller

VME

Interface

VME (J1)

GA[3:0]

SVX-CLK

SR-CLK

SR-CMD

PROG*

SVX CLK

Shaper

PECL

-

I

NT

INT-SYNC

SVX-CLK-SEL

SVX-GEN-CLK

INT-AD-PIPE

DAC

Controller

PECL

-

I

NT

PECL

-

RECEIVERS

LATCH

TRI

BUF

HDI

ID

DATA

FIFO_A

TRI

BUF

A-CLK

A-DATA[7:0]

B-CLK

B-DATA[7:0]

C-CLK

C-DATA[7:0]

G-LINK

AB

G-LINK

C

SAR

SAR

DAC-SR-CLK

DAC-SR-DATA

DAC-SEL

DAC-SR-RB

DATA FIFO_B

& LOGIC

DATA FIFO_C

& LOGIC

PORT CARD

PORT CARD

A-CLK, B-CLK

or C-CLK

18.9 ns

DATA[7:0]

LATCH CLOCK

Figure 3. Relationship between SVX-II clock and data

Connector

G-Link

Xmiter

Figure 4. G-Link Board

Finisar

LED

DELAY LINE INPUT

Adjustable: 19ns to 34 ns

(1)

(1), (2): Gate delay, to be determined.

18.9 ns

SVX-II ACQ CLK

MCLK

SYNC

(2)

Figure 5. Calibration inject timing

Figure 6. Command Signal Timming

18.9 ns

CMD[4:0]

AD-PIPE

SYNC

MCLK

1

7

6

5

1

4

3

2

HDI - C

HDI - B

HDI - A

FRONT PANEL

COMMANDS

RESET

VME

SVX CLK

- 2V

TERM IN

EXEC

PC CLK

+ 5V

- 5V

Figure 7. TFIB Front Panel

TEST PORT

CARD CONTROL

TPC CONTROL

DIP SWITCH

S2

53MHz-CLK LIMO

SVX-CLK LIMO

Figure 8. Connectors and DIP switches

RESET

DIP SWITCH

S1

pin 1

pin 1

J1 VME

CONNECTOR

MAC-II

CONNECTOR

J3 BACKPLANE

COMMANDS

FRONT PANEL

COMMANDS

HDI A

HDI B

HDI C