Safety  Assessment  Document

FOR  THE   CMS  DETECTOR

MUON   ALIGNMENT  SYSTEMS

Prepared by:

D. Eartly     (Endcap Muon )

(LINK System)

(Barrel System)

November 9, 2001

### TABLE  OF  CONTENTS

1.	Introduction

  - Purpose

  - CMS Muon Alignment Project

2.  Summary  and  Conclusions

3.  Description of the Alignment Subsystems

3.1	 Endcap  Muon  Subsystem

3.1.1	Equipment Description

3.1.2	System Critical Elements

3.1.3	System Operations

3.1.4.	System Hazards

3.2  Barrel  Muon Subsystem

3.2.1	Equipment Description

3.2.2  	System Critical Elements

3.2.3	System Operations

3.2.4.	System Hazards

3.3	LINK Transfer Subsystem

3.3.1	Equipment Description

3.3.2	System Critical Elements

3.3.3	System Operations

3.3.4.	System Hazards

4.  Safety Analysis

4.1	Introduction

4.2	General Hazards (Magnetic, Mechanical, Conventional, Fire, ODH)

4.3	Endcap Muon Subsystem

    - Electrical Systems

      - Low Voltage Supplies, Distribution, Fusing

      - Signal Cable configuration

      - Grounding and Earth Shields

      - Materials

4.3.2	Laser Hazards

    - Radiation Hazards

#### TABLE  OF  CONTENTS (Con’t)

  - Barrel Muon Subsystem

4.4.1	Electrical Systems

      - Low Voltage Distribution, Fusing

      - Signal Cable configuration

      - Grounding and Earth Shields

      - Camera Systems

      - Materials

4.4.2	Radiation Hazards

  - LINK Transfer Subsystem

4.5.1	Electrical Systems

      - Low Voltage Distribution, Fusing

      - Signal Cable configuration

      - Grounding and Earth Shields

      - Materials

4.5.2	Laser Hazards

4.5.3    Radiation Hazards

5.  Quality Assurance

  - Program Description

  - Program Objectives

  - Design Reviews

  - Safety Reviews

  - Testing Procedures

6.  Conduct of Operations

6.1   Organization of  Operations

6.2   Safety Training

  - Subsystems Tests and Checks

6.4   Configuration Change Control

6.5   Emergency Plans

7.  Handling, Storage, Disposal of Replaced-Repaired Elements

8.  Figure List

9.  Tables

10.  References

##### APPLICABLE CMS MUON ALIGNMENT SYSTEM SAFETY CODES

CERN Safety Codes, Instructions, Notes to Address

A5       Safety of Apparatus for Experiments

C1       Electrical Safety Code  + IS28 on electrical dangers

F         Radiation Protection Manual (maintenance, replacement)

IS3     Transport of Radioactive Material (maintenance, replacement, repair)

IS5     Emergency Stops (key laser run control, fault trips)

IS7     Individual Protective Equipment (eyewear etc)

IS22   Safe Laser use

IS23   Electrical Cable Criteria

IS24   Electrical Installations

IS33   Voltage classification of Electrical Installation

IS36    Magnetic Field Exposure

IS41    Plastics

NS9    Protection against RF-Microwave Radiation

NS16  Transport of Radioactive Materials

1)	INTRODUCTION

1.1  PURPOSE

This Safety Assessment Document (SAD) for the CMS Muon Alignment System has been prepared to ensure that environmental, health and safety issues are identified and properly addressed in the design and proposed operation of the CMS Muon Alignment System.  By documenting the hazards in this system and showing how these hazards are to be addressed in compliance with CERN Safety Codes, Instructions, Notes, etc., the System will be a low risk facility.

1.2  The CMS Muon Alignment Project

The original conceptual description for the system is given in the CMS Muon Project

Technical Design Report, Section 7 (CERN/LHCC 97-32). The system is composed

of three interconnected subsystems; Barrel, Link, Endcap. The configuration details

and elements of the Link connection to the Tracker and the Endcap Link Lines (now called Transfer Lines) have changed. Also the optical position measurement sensors

in both of these subsystems has changed. The Link uses a Beam-splitter-CCD array

sensor while the Endcap is based on DCOPs (an open picture frame array of four

linear CCDs intercepting a crossed line laser beam).

2) Summary and Conclusions

The unusual  hazard posed by the operation of this system in the Link and Endcap

subsystems is a multiplicity of  laser beams of  varying hazard levels and residual

radiation  exposure in servicing Link elements in the inner detector regions. These

hazards are controlled by CERN codes and standards. Consequently, by demonstrating conformance to these standards even under accident conditions; potential impacts

to on-site personnel is minimized.

3) Description of the Muon Alignment Subsystems

  - The Endcap Muon Subsystem

    - Equipment Description

The primary sensor of the  subsystem is an optical laser beam relative position measuring

CCD device (DCOPS). A window frame of four linear CCDs measure the intercept of crossed laser lines at two points for each line (1). Each sensor has a local readout board with a digital signal processor that can define the intercept positions of the laser lines

in each CCD. The low power laser diode modules for the reference beams are  in the visible spectrum. There are six optical Transfer lines that cross CMS at the outer boundaries of the Endcap muon systems under the service and electronics structures

and under the MB4 barrel drift tube chambers. Measurement and transfer of  CMS coordinates are made at six locations (Transfer plates) at each Cathode Strip chamber (CSC) layer. The other optical Straight line monitors (sensors and lasers) are located within the CSC layer gaps between YE iron discs. In addition to the optical beams and sensors, a number of  electromechanical and optical analog sensors are utilized in the system. They provide Z transfer along the optical Transfer lines  and R transfer to the rings of CSCs. These include:

### Z Transfer System (Mechanical tubes, linear potentiometers, laser distance (gap) sensor)

Z4MW40 Laser gap sensor; measuring end of a tube gap to MAB surface (gap ~45 mm)

LCP8-10 potentiometer on ME1Transfer Plate; measuring the gap - opposite end of tube

LCP8-10 potentiometers on ME2 Transfer Plate; measuring the gaps to ends of a tube

AD592 Temperature sensors on the tube, Transfer Plate, ME2/2 CSC panel strip, etc.

AD592 Temperature sensors on the tubes

## Transfer Plates (Connect coordinates of Transfer line and CSC layer SLM systems)

Applied Geomechanics 900 dual axis inclinometer

AD592 Temperature sensors on the Transfer Plates

### R Transfer to CSCs (Connect Transfer plates to outer CSCs to inner CSCs)

LPX50 wire extension potentiometer; Transfer Plate to ME2/2, ME3/2 outer  CSCs

LPX50 wire extension potentiometer;  ME2/2, ME3/2 to ME2/1, ME3/1 inner CSCs

AD592 Temperature sensors on the ME2/2, 3/2  and ME2/1, 3/1 CSC panels

The sensors and distributed electronics are connected through a halogen free cable network to a few data acquisition nodes in the service and electronics structures around the YE iron discs (2). The DCOPS system (13 VDC) operates on a single low voltage in fused, low current bussing. The ANALOG  system  (13 VDC) operates on a single low voltage in fused, low current bussing.  The LASER system (5 VDC) operates on a single low voltage in fused, low current individual laser wiring network. Operation of the system is from the underground control room.

    - System Critical Elements

The safety critical elements of the subsystem are the network of lasers. There are six

distributed Transfer lasers (ME4, -4 layer)  in each CMS Endcap plus six more at SLM

endpoints (outer iron boundaries) in eachCSC layer (ME1, 2, 3, -1,-2,-3). These fourty eight lasers are individually turned on/off through a switching network under computer (operator/program) control. The laser power and switch network can be interlocked

off when appropriate. Laser operation is pulsed infrequently for a short duration to make specific optical measurements. Typically the lasers will operate only when the detector is closed and inaccessible. Some exposed operation is planned only during initial survey of CSC layer assembly or Endcap maintenance. Such activity will be coordinated with other

groups who would be impacted.

    - System Operations

Typically, after initial test, study and tracking periods, the system will be dormant except

during brief measurement cycles (~ 1-2 minutes) every ½-1 hour. This cycling will run

by program initiation and control using a UNIX computer clock. The program includes

checks on the sensor data stream and data values to insure that the system is operating

properly.  Voltages are monitored throughout the system and lasers are confirmed as

turned off using background spectrum measurements in the optical sensors.

    - System Hazards

The system nominally operates only when the detector is inaccessible. The lasers and low voltage can be interlocked off during personnel access. The potential risk to laser light exposure would only occur during test and maintenance periods. The laser light would be generally blocked and limited by apertures except potentially when the detector system is open and separated. The laser diode modules for the reference beams are visible (red)

(670 nm), low power Class 3A; which by IS 22 require warnings when in operation. The laser beams  expand into large cross lines (low power density) and operate in pulsed mode (only on during short measurement cycle). They are effectively masked by apertures and blocking plates. The system also uses aperture tubes where possible.

Laser beams at endpoints are stopped by covers/shields so they do not project past

the active measurement regions or into personnel access space.

All sensors and electronics operate at low voltage on grounded structures and electronics enclosures. The DCOPS  system has fused Interface and readout boards with temperature sensors. Analog sensors are fused through their local Interface boards in grounded electronics enclosures. The system utilizes non magnetic and approved plastic materials

and halogen free cabling; except for a small volume of standard G10 circuit boards.

- Safety Analysis

4.1 Introduction

This section discusses hazards of the CMS muon alignment system and safety assessment

evaluations for these hazards. The following hazards are excluded from this system: flammable gas, high voltage, oxygen deficiency, and toxic materials. An overview of

hazards and applicable standards is given in Table 4.1.1 below.

|  | Table 4.1.1 |  |  |
| --- | --- | --- | --- |
| Section | System Hazard | Hazard Class | Codes and Standards |
| 4.2.1 | AC Power Distribution | Routine | C1, IEC, NFC15-100 |
| 4.2.2 | Low Voltage - Medium Current | Low Hazard | C1, IEC, IS33 |
| 4.2.3 | Signal Cable System | Routine | C1, IEC, IS23 |
| 4.2.4 | Fire Hazards | Low Hazard |  |
| 4.2.5 | Radiation | Low Hazard | F |
| 4.2.6 | Magnetic | Low Hazard | IS36 |
| 4.2.7 | Laser | Medium Hazard | IS22 |
| 4.2.8 | Conventional | Routine |  |
| 4.2.9 | Plastics | Routine | IS41 |

4.2 General Hazards

Certain hazards exist because of the CMS enclosures, solenoid magnet, access and service structure configurations, radiation, activation, equipment of other systems of

the detector, etc. These hazards, warnings,  and required safety training for them is the domain of the GLIMOS and are not addressed here.

4.3 Endcap Muon Alignment Subsystem

An outline of potential hazards in elements of the subsystem is given in Table 4.3.1 below:

| Table 4.3. 1 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Element | Mechanical | Conventional Electrical | Fire | Laser |
| Geometry, Simulation, Survey, Reconstruction | n/a | present | n/a | n/a |
| DCOPS sensors | n/a | present | n/a | n/a |
| Analog sensors | n/a | n/a | n/a | n/a |
| Laser diode modules | n/a | n/a | n/a | present |
| Electronics | n/a | present | n/a | n/a |
| Power Supplies | n/a | present | present | n/a |
| Signal cables | n/a | present | present | n/a |
| Power cables | n/a | present | present | n/a |
| DAQ system | n/a | present | n/a | n/a |
| Installation | present | present | n/a | n/a |
| Commissioning | n/a | present | n/a | present |

4.3.1 Electrical Systems

4.3.1.1 Low Voltage Supplies, Distribution, Fusing

Low voltage modules (13 VDC, 5VDC) will be located in LV power module boxes

for the EMU VME crates near these crate clusters on each YE iron disc. Typically,

alignment power modules are at the 9 o’clock position near the DAQ system as

shown in Figure 4.3.1.1.1. To limit current in the perimeter LV distribution cabling

to each Transfer Plate/SLM Interface (six stations), each CSC layer splits into two branches as shown in Figure 4.3.1.1.2. This limits the current to  2.2 A in a DCOPS periphery first link of a daisy chain branch to three stations. The same  scheme for

Analog sensors limits the current to 0.5 A in an Analog periphery first link of a daisy

chain branch to three stations. Each station is fused (resetting) through the local

DCOPS Interface board at 1.25 A max for the SLM string of six DCOPS sensors

and 0.375 A for the SLM analog sensors. There will be less than a 10 degree Centigrade temperature rise in any cable at the maximum fused current. Each DCOPS sensor readout board itself is also fused (0.35 A). With a double insulated cable (DI), it qualifies as a protected extra low voltage (P.E.L.V) per IS33.  The insulation is halogen free per IS23 and samples will be given to TIS for testing and approval. At each station, the LV is daisy chained through a half SLM of DCOPS sensors (current in first jump 0.72 A)* in a double insulated multi conductor cable integrated onto the CSC surface electronics system. At each Interface card, the local distribution is again fused. The distribution is shown in Figures 4.3.1.1.3, 4.3.1.1.4. The  5  Vdc to each laser (0.06 A) is again distributed (independent pairs) around the perimeter in a double insulated cable (DI).  Conductor, cable, fusing, and current specifications are given in Table 4.3.1.1.1 below.

| Table 4.3.1.1.1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Cable | Conductor | Fused | Length | Current (max) | Voltage |
| DCOPs periphery LV | 14 ga DI | 4 A | 8 m | 2.19 A | 13 Vdc |
| Analog periphery LV | 14 ga DI | 1 A | 8 m | 0.48 A | 13 Vdc |
| DCOPS SLM Intf (I max)* | 24 ga DI | 1.25 A | 0.15 m | 0.72 A | 13 Vdc |
| DCOPS SLM 3-4 (Lmax) | 24 ga DI | 1.25 A | 3.5 m | 0.36 A | 13 Vdc |
| Analog SLM Intf (Imax) | 24 ga DI | 0.38 A | 4 m | 0.011A R1 | 11 Vdc |
| Analog SLM sensor (Lmax) | 24 ga DI | 0.38 A | 5m | 0.003 A T1 | 11 Vdc |
| Laser periphery LV | 24 ga DI | 0.1 A | 24 m | 0.06 A | 5 Vdc |

* First two short jumps between readout boards are inside a ground shield box

4.3.1.2 Signal Cable Configuration

For DAQ, the DCOPS sensor system for each CSC layer is a chain of six half SLMs as shown in Figure 4.3.1.2.1. For DAQ, the Analog sensor system for each CSC layer is a six parallel half SLM sensor groups multiplexed into ADC acquisition units.  Each half SLM is connected through DCOPS and Analog Interface cards at the periphery of each YE iron. The cables divide into two groups, periphery (between Interface stations to DAQ system located in a rack on the YE service gallery) and SLM branch. In all cases they are double insulated halogen free multi conductor cables. The SLM cables are integrated into the CSC surface electronics system. A single round shielded cable daisy chains between DCOPS sensors. Analog sensors have independent CAT4, RJ12 cable connections to the interface card. DCOPS sensors connect to the local readout electronics via a short flat plenum shielded halogen free multi conductor cable. The periphery cable configuration for a CSC layer is shown in Figure 4.3.1.2.2 and at a station is shown in Figure 4.3.1.2.3. A simpler configuration of cabling for Endcap sensor readout on external MABs on the Barrel rings MB2, -2 is shown in Figure 4.3.1.2.4. This corresponds to a two branch/station configuration. Periphery cables are given in

Table 4.3.1.2.1 below. SLM branch cables are given in Table 4.3.1.2.2 below.

MAB cables are given in Table 4.3.1.2.3 below.

| Periphery |  | Table 4.3.1.2.1 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Cable |  | Configuration | Ins | Num | Length | Shape |
| Serial DAQ-JTAG | A | 4pr 28 ga STR, F/B shield | DI | 1 | 4 m | Round, 7 mm |
| DCOPS power | B | 14 ga, 2 cond STR, jacket | DI | 6 | 8 m | Round, 6 mm |
| Analog power | B | 14 ga, 2 cond STR, jacket | DI | 6 | 8 m | Round, 6 mm |
| DCOPS Intf 422 | H | 8pr 28 ga STR,F/B shield | DI | 5 | 8 m | Round, 10mm |
| Analog signal | C | 10 pr, 24 ga STR, jacket | DI | 2 | 8 m | Round, 10mm |
| Analog signal | C | 10 pr, 24 ga STR, jacket | DI | 2 | 16 m | Round, 10mm |
| Analog signal | C | 10 pr, 24 ga STR, jacket | DI | 2 | 24 m | Round, 10mm |
| Laser power | E | 3 pr 24 ga, jacket, RJ12 | DI | 6 | 8 m | Round, 5 mm |
| T signal | E | 3 pr 24 ga, jacket, RJ12 | DI | 2 | 8 m | Round, 5 mm |
| T signal | E | 3 pr 24 ga, jacket, RJ12 | DI | 2 | 16 m | Round, 5 mm |
| T signal | E | 3 pr 24 ga, jacket, RJ12 | DI | 2 | 24 m | Round, 5 mm |

| SLM Branch |  | Table 4.3.1.2.2 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Cable |  | Configuration | Ins | Num | Length | Shape |
| DCOPS analog | J | 20c 28ga STR, F shield | DI | 6 | 0.5 m | Flat, 26 mm |
| DCOPS LVDS Int | I | 20c 28ga STR* |  | 2 | 0.15 m | Flat, 24 mm |
| DCOPS LVDS Ex | F | 10pr 28ga STR, F/B shield | DI | 1 | 0.5 m | Round,10mm |
| DCOPS LVDS Ex | F | 10pr 28ga STR, F/B shield | DI | 1 | 3.5 m | Round,10mm |
| DCOPS LVDS Ex | F | 10pr 28ga STR, F/B shield | DI | 1 | 1.0 m | Round,10mm |
| DCOPS LVDS Ex | F | 10pr 28ga STR, F/B shield | DI | 1 | 2.0 m | Round,10mm |
| Analog signal | D | 3pr 24 ga S, jacket TrP,R1 | DI | 5 | 1.0 m | Round, 5mm |
| Analog signal | D | 3pr 24 ga S, jacket R2,T3 | DI | 2 | 1.5 m | Round, 5mm |
| Analog signal | D | 3pr 24 ga S, jacket R1,T2 | DI | 2 | 4.0 m | Round, 5mm |

*First two short jumps between readout boards are inside a ground shield box

| MAB System |  | Table 4.3.1.2.3 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Cable |  | Configuration | Ins | Num | Length | Shape |
| DCOPS analog | J | 20c 28ga STR, F shield | DI | 6 | 0.5 m | Flat, 26 mm |
| DCOPS LVDS Int | I | 20c 28ga STR* |  | 2 | 0.15 m | Flat, 24 mm |
| DCOPS LVDS Ex | F | 10pr 28ga STR, F/B shield | DI | 4 | 8 m | Round,10mm |
| DCOPS power | B | 14 ga, 2 cond STR, jacket | DI | 1 | 8 m | Round, 6 mm |
| DCOPS power | B | 14 ga, 2 cond STR, jacket | DI | 1 | 2 m | Round, 6 mm |
| DCOPS Intf 422 | H | 8pr 28 ga STR,F/B shield | DI | 1 | 8 m | Round,10mm |
| Serial DAQ-JTAG | A | 4pr 28 ga STR, F/B shield | DI | 1 | 4 m | Round, 7 mm |

The halogen free cables satisfy IS23,  IS33  as Protected Extra Low Voltage (P.E.L.V).

      - Grounding and Earth Shields

All of the Interface and Readout boards are housed in aluminum grounded enclosures

with slide on removable covers. The whole assembly (cover + electronics) or boards

can be easily removed from their mounts without removing/installing any fasteners.

The grounding configuration is illustrated in Figure 4.3.1.3.1. All electronics/cabling provides power and signal ground returns. Ground connection of the cables, boards,

and covers is made locally at the mounting points with a slide on terminal connection.

      - Materials Configuration

Most of the material in the system is aluminum and other metals (copper, brass, stainless  steel). A small amount of LEXAN is used in CCD mounts in DCOPS sensors and MAB

Sensor mounting plates (for weight). Cable insulations are halogen free thermoplastics. Electronics boards (~0.005 m3 G10  ) have a small volume of standard connector plastics. Table 4.3.1.4.1 outlines the major materials used in the Endcap subsystem.

| Table 4.3.1.4.1 |  |  |  |
| --- | --- | --- | --- |
| Item | Material | Number | Volume |
| Transfer Plates and supports | Aluminum | 36 |  |
| CSC sensor mounts | Aluminum | 216 |  |
| DCOPS sensor boxes | Aluminum | 216 |  |
| DCOPS readout earth shield | Aluminum | 144 |  |
| Transfer laser mounts | Aluminum | 12 |  |
| Interface board earth shield | Aluminum | 38 |  |
| Cable insulation | Thermoplastic |  |  |
| PC boards | G10 | 300 | 0.005 m3 |

    - Laser Hazards

As stated earlier, the system nominally operates only when the detector is inaccessible. The lasers and low voltage can be interlocked off during personnel access. All lasers

will be marked with warning/identification labels and logged/listed with CERN TIS.

The potential risk to laser light exposure would only occur during test and maintenance periods when the detector system (CSC layers) is open and separated. The laser light would be generally blocked and limited by apertures, masks, and aperture tubes. The laser diode modules for the reference beams are visible (red) (670 nm), low power Class 3A; which by IS 22 require warnings when in operation. The laser beams  expand into large cross lines (low power density).  Even in testing, they operate in pulsed mode (only on during short measurement cycle). For SLM laser beams, the active measurement regions are just above the surface electronics of the CSCs (27 mm), stop at the outer boundaries of the YE iron, and do not project into personnel access space. Nominally, no one can put their line of sight into the active regions. For Transfer line laser beams, the active measurement regions are just above the outer YE surface interleaved in cable ways and services and under the MB4 Drift Chambers. These active regions are accessible only

if the YE discs or Endcap-Barrel systems are separated. Under these conditions, any of these beams would only be operated (rarely) in repairs of elements in a given line. This operation would occur only when other groups are not working in the affected areas. All

operations would include warning lights and postings in accord with IS22.

4.3.2  Radiation Hazards

The Endcap subsystem does not require the use of any radioactive sources for test or calibration. The only radiation hazard results from activation of elements. This will be dominated by cables and circuit boards. Any elements (sensors, circuit boards, mounts) removed from the activated detector will be scanned, tagged, and handled in accordance with the CERN  Safety Code F Chapter 4, Section 3.

  - Barrel Muon Subystem

  - LINK Transfer Subsystem

- Quality Assurance

5.1  Program Description

Each Alignment subsystem has developed a Quality Control and Assurance plan

appropriate to the construction, testing, calibration, tracking, and implementation

of  elements specific to that system. These plans are described in the corresponding

Quality Control Manuals.

5.2 Program Objectives

Subsystem QAQC plans are directed at providing systems with the required reliability and performance. In addition to Quality Control in the production of the subsystems,

these manuals include Quality Control during Installation and Commissioning plus

Quality Assurance during subsystem Operation and Maintenance.

5.3   Design Reviews

The CMS Muon Alignment group regularly conducts a number of workshops to review

and evaluate the developments (designs, testing, hardware) in each subsystem. This process insures a co-ordinated system. In addition, the designs are integrated with

the detector systems. Barrel and Endcap subsystems are fully integrated with the

chambers and electronics. At appropriate times, the group requests a CMS

EDR for specific parts of the system.

5.4   Safety Reviews

As the CMS Muon Alignment system goes into production  and Installation in 2002, 2003 it will initiate/participate in CMS/ TIS safety reviews, especially regarding laser, electrical, and Radiation Safety issues.

5.5  Testing Procedures

Preliminary documentation on specific system testing and certification procedures

during production, installation, commissioning, operation, and maintenance is given

in the QAQC documents of the subsystems. In preparation for installation (2003,2004),

commissioning (2003,2004), operation (2004+), and maintenance; detailed Muon alignment system plans will be developed as part of the total WEB posted CMS Installation, Commissioning, and Operations plans. This will be part of a co-ordinated Muon subsystem plan.

6)  Conduct of  Operations

6.1   Organization of Operations

The CMS Muon alignment system will be included in the CMS Operations Pages.

6.2   Safety Training

Personnel working in the CMS Muon alignment system will have special training on

laser hazards. In addition, safety trainings required for system installation, operation,

and maintenance will be established by the CMS Operations safety training pages.

6.3   Subsystem Tests and Checks

The operating DAQ  of the Muon alignment system monitors the data stream, voltage

conditions, etc. and will generate specific alarms in failure modes. The alignment

group will provide a document (WEB) that defines specific actions to be taken in

response to the alarms by CMS operations.

6.4  Configuration Change Control

Production configuration drawings and documents will be submitted to EDMS.

Any modifications to the system configuration are subject to the change control

process of CMS.

6.5  Emergency Plans

As part of the alarms action document, emergency stop and shutdown procedures

for the system will be included.

7) Handling, Storage, Disposal of Replaced-Repaired Elements

Any elements (sensors, circuit boards, mounts) removed from the activated detector will be scanned, tagged, and handled in accordance with the CERN  Safety Code F Chapter 4, Section 3  (control of radioactive materials and sources). Any repair or re-calibration of items will be done at CERN in a controlled area (3.2, 3.3, 3.4, 3.5, 3.7). If an item is to be discarded, procedures of subsection 3.6 will be followed.

8) Figure List

9) Tables

10) References