| GODDARD SPACE FLIGHT CENTER REQUEST FOR RADIATION SAFETY COMMITTEE ACTION IONIZING RADIATION PRODUCING DEVICE QUESTIONNAIRE |
| --- |
| 1. ____Krizmanic, John F _______ ____661 _ ___6-6817 ___ 01/09/2007_ Originator's Name: (Last, First, MI) Code Phone Date | Docket # |
| 2. Device Data and Description: a. Device type: ___Electron Impact X-ray Source_______________ b. Manufacturer and Address: X-Ray Technology, Inc 360 El Pueblo Road, Suite 104 Scotts Valley, CA 95066 Tel: (831) 439-9729 Fax: (831) 438-5629 e-mail: sales@oxfordxtg.com Web Page: www.oxfordxtg.com c. Model Number: __XTG5011 Apogee__ d. Operating Parameters Maximum- _____50____kVp; ______1.0 _mA Normal- ______50_kVp; _______0.8__mA Average Weekly “on” time- ___ <= 1000 minutes e. Target Materials: _Molybdenum___ f. Description of Use: ___Calibration of and testing of X-ray Optics |
| 3. Detailed Description: (attach installation plans, diagrams, applicable manufacturer info, etc) Number of pages attached: ___4_ (Include State Certifications if applicable) |
| HEALTH PHYSICS USE ONLY |
| 4. Date Received: _____________ 5. GSFC #__________________ 6. Source Type: _______________ 7. Serial Number: _____________ 8. Survey Frequency: _____Annual ______Biennial ______Triennial 9. Notes: 10. Health Physics Certification: ________________________________________ _______ Printed Name Date ________________________________________ Signature FINAL DISPOSITION: ___________________________________________________________________________ ____________________________________________________________ ________________________________ Person Authorizing Disposition |

GSFC Form 23-28I [DEVICE] (01/00)                                                                                                Return to Code 205.9

Instructions for filling out GSFC Form 23-28I [DEVICE], Request for Radiation Safety Committee Action - Ionizing Radiation Producing Device Questionnaire.  For additional guidance refer to GPG 1860.1 "Ionizing Radiation Protection," particularly paragraph 2.1.4.4.

### ALL ITEMS MUST BE LEGIBLE

1.	If you do not have a GSFC Code designation you must include a valid mailing address on an attached paper.

2.	Source Data:

a.	Device type;

b.	Manufacturer and address of manufacturer (include telephone number if known);

c.	Model number;

d.	List the operating parameters;

Maximum kVp and mA that the device is capable of

Normal kVp and mA that the device will be routinely operated

Average weekly use in minutes

e.	Types of target materials built into the system or which will be targeted by the device.

f.	Give a brief description of the proposed use.

3.	Provide a detailed description of the installation, include a diagram and specific dimensions.  Provide any manufacturer information that will be useful in reviewing the device and its installation.  Include a copy of State safety certifications if applicable.

RETURN COMPLETED FORM TO CODE 205.9 (Health Physics).

If you have any questions concerning this form, please contact

the Safety and Environmental Branch, Health Physics Department at 301-286-8482.

HEALTH PHYSICS USE ONLY

4.	Indicates the date source was originally surveyed by Health Physics.

5.	Indicates the GSFC identification number assigned by Health Physics.

6.	Indicates the source Type classification determined by Health Physics (GPG 1860.1, paragraph 2.1.4.3.a).

7.	Indicates the manufacturers serial number (if any).

8.	Survey requirements determined by Health Physics.

9.	Special Health Physics notes.

10.	Certifying Health Physics personnel.

Final disposition will be certified by persons authorized by the Radiation Safety Committee only.

Detailed Usage of X-ray Technology XTG5011 Apogee X-ray source at Area 200 X-ray Interferometry Testbed.

The MANSON MODEL 2 MINI FOCUS X-RAY SOURCE:

Figure 1: X-Ray Technology XTG5011 Apogee X-ray Tube

The source mounts on a tapped rotatable 2-3/4" UHV flange on the customer's vacuum system.  We use the high voltage/beam current power supply and cables made by the same manufacturer.  The manufacturer has provided a radiation map for this device detailed in an Application Note and it is attached to this document.

Location of this X-Ray Technology Model XTG5011 Apogee X-ray Tube:

The source is located in a locked shed at the end of a 600 meter long vacuum pipe- part of the GSFC X-ray Interferometry Testbed located at Area 200.  The beamline location is indicated in figure 2.  The shed that maintains the source is shown in figure 3.   Area 200 is a controlled access location that has very few visitors.  A properly coded keycard is required for access to the location.

The source is separated from the rest of the vacuum system by a gate valve- unless when in use.  Once a vacuum level of 50 mTorr is reached in the beamline, the source is put on manually per the manufacturer’s instructions:

- The accelerating high voltage is slowly (over 10’s of seconds) raised to 20 keV, and the electron emission current is adjusted to 400 milliamps.

- The accelerating high voltage is slowly (over 10’s of seconds) raised to 30 keV, and the electron emission current is adjusted to 800 milliamps.

- The accelerating high voltage is slowly (over 10’s of seconds) raised to 40 keV

The accelerating high voltage is slowly (over 10’s of seconds) raised to 50 keV

The chamber and the beamline is made of stainless steel and there are no ports where x-rays could exit the vacuum system.  The attenuation length of stainless steel at the Mo K line (17.3 keV) is approximately 35 microns.

The X-rays will  travel 147 meters in the vacuum system to an optics chamber where we place test optics.  The typical x-ray flux at this point is expected to be less than 1000 photons/s/cm^2 .

X-rays that pass through the optics will travel  another 450 meters to an x-ray CCD detector.  At this position we expect a typical X-ray flux of < 100 photons/sec striking the CCD.

Operations will run for typically 1 week at a time with nearly continuous x-ray usage.  We expect to have ~1 week “runs” about once a month or so.

# Text boxes

Figure 2:  The X-ray Interferometry Testbed Located at Area 200.

Figure 3:  The Source Station of the X-ray Interferometry Testbed.  The Source is located in this environment controlled building.  The Building is locked and is located in a controlled access area.