# Test of a Light Emitting Diode 
Flasher with the Fermilab Water Tank.

François Montanet2,  Hal Spinka1,  Derek Strom2

- Argonne National Lab. , IL , USA

- Fermilab, Batavia, IL, USA

October 12, 2001

# Abstract

We report on measurements made on the Fermilab surface detector prototype (FATII) of a Light Emitting Diode (LED) system. The setup was similar to future LED system, intended for the monitoring and calibration of the pre-production surface detectors. The LED was installed on one of the available water fill ports i.e. in a very much off centered position. Despite this, it was possible to show that the light scattered back by the liner was yielding well-balanced signals on the three PMT’s. We conclude from this study that it is not necessary to aim at installing the LED ports at the very center of the future liners but that an off centered position is acceptable.

## Introduction

It is planned to install Light Emitting Diodes (LED) flashers in each tank of the pre-production Surface Detectors (SD). Although they may not be required for a successful calibration, it was considered that they would ease the control and monitoring of the array (mainly fast functional tests and linearity measurements.  They would also provide a simple first order check of the entire SD station (detector+electronics) before deployment (in the assembly building, before transport and water filling). The LED system must be simple, reliable and add a minimal load to the rest of the project. Recent discussions have led to two a priori decisions:

- The LED system should consist of 2 independent flashers in order to test linearity.

- The LED’s should be installed in a dedicated port equipped with a window (to preserve the water sealed). Further more, it was advocated that in order to keep the light yield naturally balanced on the 3 PMTs, this port should be in a central position, on the top of the liner.

Installing the LEDs in a central position is not without drawbacks: this point is not accessible after the liner is installed within the tank. Therefore, it is necessary to:

- Find a procedure to install the flashers as an integral part of the liner, before the liner is placed in the tank: the flasher, the cables connecting it and the supporting port have to be compact, flexible and robust enough to survive the folding, the installation  and the inflation of the bag. It should also be made such as not to damage the bag. Although this is not considered impossible, it is not a very easy task.

- When the bag is in place and inflated or filled with water, it cannot be maintained. This raises the question of the long term reliability of the LED flashers. We plan to use blue LEDs because their emission wavelength (470 nm) is a good match to the mean wavelength of detected Cherenkov photons. These ultra-bright blue LEDs, which were still cutting edge technology not very long ago, are nowadays used in general-purpose lighting applications and progress in their manufacturing has led to highly reliable and long lifetime products. Part of the flasher electronics is the LED driver itself, which is a high voltage, high current, and very short time pulser. To avoid time smearing and EM noise emission, this tiny circuit has to be placed as close as possible to the LED itself i.e inside the dedicated port. If installed in a central inaccessible position, it would also become a non-maintainable electronic device of probably poorly know MTBF. One could live with this, assuming the LED flashers as “consumable” devices only useful during detector installation and commissioning but we will probably wish to ask more for our money!

An alternative solution to the two above problems would be to install the LEDs still in a dedicated port as central as possible but positioned in a way as to remain accessible via the largest hatch of the tank. The radius of the tank is 1.8 m, the PMT’s and the hatches are 1.2 m away from the centre. The large hatch is 56 cm in diameter and is 30 cm above the normal water level. My estimate is that a small LED port that would be situated about half way (i.e. ~60 cm) between the centre of the bag and the PMT, would still be easily accessible by a normally constituted human being (see Figure 1).

Obviously, a question is raised by this proposal: is the sharing of the light yield between the three PMTs affected by the off centred LED position? Is the unbalancing acceptable? Using data taken in September 2001 in the Fermilab water tank using a LED pulser, we have tried to answer this question.

Figure 1: On this schematic, one sees that a LED port located half way between the PMT and the center of the bag is still accessible, within a 45( angle and a distance of less than 50 cm of the hatch.

## The LED setup in the Fermilab Water Tank FATII

### The Fermilab Water Tank FATII

The Fermilab water tank called hereafter FATII, is a full size surface detector prototype identical in size to the southern observatory tanks. It is composed of a cylindrical polyethylene tank of full dimensions, 10 m2 surface area and about 1.2 m height. The tank interior is accessible through 3 hatches, above the location of the PMTs. The water is contained in a liner made of a laminate of polyethylene and Tyvek® identical to those used in the Engineering Array. This liner is equipped with dome kits with floppy windows intended for 8” PMTs. This liner was installed in the tank in March and filled with high resistivity (~18 M() purified and filtered water. The total volume of water enclosed is about 12,000 liters. The tank is linked to the control room via 30m long cables for providing the High Voltage supply and for transporting the signals back to the DAQ system. The readout and trigger electronic is base on standard NIM and CAMAC modules as well as on a 4 channel digital oscilloscope Tektronics TDS3054. The CAMAC and the scope are controlled and readout using a GPIB bus via a PCI GPIB interface card from National Instruments. The data acquisition program runs on Windows NT, it is written in C++, and uses Root. The data are saved on local disks as Root histograms and trees.

### Installing an LED in a water fill port

Obviously, the liner we are using in the Fermilab prototype was produced well before a special LED port was envisaged. We will have to use existing ports i.e. water fill ports. The liner presently used in FATII has the following particularity: the kits were assembled in China such that the 3 water fill ports are all positioned on the radius of their associated PMT and in between the PMT and the center (see Figure 2). The PMTs are 120 cm away from the center and the fill ports are 92 cm from the center. The distance from the port to the axis of the nearby tube is ~ 28 cm. The distance to the two other tubes is ~ 184 cm. PMTs are numbered 1, 2, and 3. An LED is installed in the fill port near PMT 2.

Figure 2: A drawing of the liner used in FATII. Note that the water fill ports are positioned on the radius in between the PMT and the center. The distance from the center to the fill port is 92 cm. The distance from the fill port and the nearby PMT axis and is ~28 cm. It is ~184 cm to each one of the two other PMTs.

We did not received in time a specially designed “plug” to host the LED and its driver inside the water fill port. We had instead to do a quick and dirty installation by perforating one of the spare port caps available at Fermilab with a hole of 5 mm in diameter. The LED head was forced into that hole, making the assembly reasonably watertight (see Figure 3). The LED is thus directly in contact with the water, although we cannot exclude that a small air bubble is trapped in the port. It is also important to note that as the fill port is made of white polyethylene and as there is no collimation of the LED light, the fill port itself acts as a light diffuser therefore scattered light that can go directly to the nearby PMT.

Figure 3: The water fill port and perforated cap. The LED is inserted lightly into the 5 mm hole in the cap. It is in direct contact with the water. The fill ports are made of white polyethylene.

## The Sheffield LED pulser

A complete driver and its interface to the SD Front End electronics is currently being developed by our colleagues from the Moscow State Engineering Physics Institute. In order to get an answer quickly, we used an already existing pulser, which was developed for the ANTARES collaboration by people from the Univ. of Sheffield UK [2] , from an earlier design by J.S.Kapustinsky et al [1].

This is a very compact, robust and simple driver, with a very small time jitter. Associated with ultra bright blue LEDs (for example Agilent HLMP CB30 or Nichia NSPB300A), it provides light pulses which are fast (FWHM <5 ns) and powerful (up to 40pJ or 9.5(107 photons/pulse at 465nm). The driver is triggered on the leading edge of a logic (TTL or NIM) pulse sent on a simple twisted pair cable. The pulse energy can be adjusted from 0 to 40pJ via a DC bias superimposed on the same pair (Figure 5). A firm in the UK is currently commercially producing these drivers as a small (13(27 mm) SMC board (see Figure 4).

Figure 4: The Sheffield LED driver.

Figure 5: The energy per pulse delivered by 4 different Agilent HLMP CB30 diodes as a function of the bias voltage applied to the pulser (taken from ref [2]).

## LED versus Muons

It is worth comparing the light yield delivered by a LED pulser of the kind we plan to use with the Cherenkov photon yield from a muon or from a shower.

As we already said, the energy per pulse from the LED can be adjusted from 0 to 40pJ, which is close to 108 photons/pulse.  Note that these are 465 nm wavelength photons (the spectral width of the LED is <30 nm).

A vertical ultra-relativistic muon traverses L = 120 cm of water during which it produces

,

with   and . This corresponds roughly to , assuming an effective spectral width of 1.5 eV. The number of photoelectrons detected by a PMT is given by integrating over the spectrum:

where is the collection efficiency and is the detection efficiency (quantum efficiency). For bialkali cathode PMTs like the one we are using,

.

The average number of photoelectrons measured in good running conditions with tanks prototypes using well-defined vertical muons is about 40 to 50 p.e. per PMT. One can therefore infer that the collection efficiency is of the order of:

and the total detection efficiency is:

.

Assuming that the collection efficiency is comparable for Cherenkov photons produced across the tank volume and for mono-chromatic photons produced by a LED at the top of the volume, one can convert the LED energy/pulse to photoelectrons per PMT:

p.e.

and to VEM units:

This is more than enough to test the PMTs plus electronics linearity and almost comparable to the light yield induced by a shower at some distance from its core.

## Studying the light yield balance

The Fermilab prototype is currently equipped with three Hamamatsu R5912, Ø200 mm, 10 stages tubes. These tubes are the same used in part of the EA tanks, they are high gain tubes and are therefore not suitable for high illumination. They are highly non-linear and saturate for anode current ~ 40 mA at a gain of 106. We were thus not able to explore the LED at high values of the energy/pulse. Instead we set the LED driver bias voltage such as to produce muon like pulses. Observing the PMT signals on a oscilloscope, we lowered the bias voltage until the LED induced signal was very much similar in amplitude and charge as the signal from a few VEMs. This corresponds to bias voltage values just above the threshold in Figure 5.

The gain versus HV curves of all three PMTs were measured prior to their installation in the tank, using fits to single photoelectron spectra. The HV of the PMTs was then set to correspond to a gain of 6.4(106. This gain was chosen because it is close to the one of the high gain channel (amplified dynode) when low gain tubes will be run at 2(105. It is also convenient to note that for a gain = 6.4(106, the charge corresponding to 1 p.e. is 1pC. The VEM calibration is established using a dataset of muon events that pass a minimum bias trigger (2-fold coincidence of the tank’s PMTs) and a further selection requiring some balance of the charges seen by the three PMTs. The ADC spectra corresponding to this internal trigger are fitted to obtain the number of p.e. per VEM and per tube (

Figure 6). The fit function is the sum of a gaussian for the muon peak and of an ad hoc function describing the underlying background. The fits to these spectra give mean values of 43.5, 61.1 and 40.6 p.e./VEM for PMT1, PMT2 and PMT3 respectively.

To avoid possible bias in the gain measurements and to take into account possible differences in the quantum efficiency and in the optical coupling of the PMT to the water, we will express the measured charges in units of VEM using the above individual PMT calibration.

The LED ADC spectra are shown in Figure 7. The fits to these spectra give mean values 123.2, 259.9 and 98.69 pC for PMT1, PMT2 and PMT3 respectively. When normalized to VEM, this reads 2.63 VEM, 4.58 VEM, and 2.49 VEM respectively for PMT1, PMT2 and PMT3. We recall that the LED is very close to PMT2. The solid angle ratio for direct light (i.e. the ratio of the distance square to PMT1&3 to the distance square to PMT2) is ~ 43. The observed light yield ratio between PMT2 and PMT1 or PMT3 is ~ 1.8. One can thus already conclude that the amount of direct light seen by the PMTs is small compared to the isotropized light scattered by the liner.

Figure 6: ADC spectra obtained from muons calibration (internal trigger) for the three PMTs. The abscissa is in pC (~ equivalent to photoelectrons units for the gain of 6.4(106 used here).  The gaussian fits gives mean values of 43.5, 61.1 and 40.6 p.e./VEM for PMT1, PMT2 and PMT3 respectively.

Figure 7: ADC charge spectra obtained flashing the LED. The abscissa is in pC. The gaussian fits gives mean values of 123.2, 259.9 and 98.7 pC for PMT1, PMT2 and PMT3 respectively.

### Pulse shape analysis.

One way to disentangle between the direct light (i.e. the light traveling directly from the LED port to the PMT) and the light scattered back by the liner, is to analyze the PMTs pulse shapes using the digital scope traces. As in the case of Cherenkov light, the LED light pulse is very short (few ns) but the light is scattered by the liner and bounces back and forth, producing a signal with an exponential decay with a characteristic time of the order of 100 ns. On the other hand, if part of the light goes directly from the LED to one of the PMTs, this should show up as a few ns spike in front of the signal.

The vertical muon average pulse shape is shown in Figure 8. The exponential fits give a lifetime of the light in the tank of roughly 100 ns. After-pulsing in the PMTs dominates the waveform for time larger than 500 ns after the pulse start. This tends to increase artificially the fitted lifetime. The fit is therefore restricted to the unbiased region. One can have a feeling of the remaining systematic error on the lifetime measurement by comparing the fit results on the individual PMT signals: they should a priori give the same value. From the measured differences, we estimated a systematic error of ± 5 %. One should note that two kind of data set were used and compared for this lifetime analysis: vertical muons as defined coincidence of scintillation counters (external trigger), and minimum bias muons asking for a 2-fold coincidence of the tanks PMTs (internal trigger). Both data samples gave very similar results in terms of lifetime (consistent within ± 2 ns) as can be seen in table Table 1.

Figure 8: Waveforms from PMT1, PMT2 and PMT3 averaged over 600 vertical muon events (external trigger). The time abscissa is in ns. The exponential fits gives 108 ns, 103 ns, and 96 ns. The effect of after-pulses is clearly seen for times larger than 500 ns after the pulse start.

Table 1: The light exponential decay time as fitted from averaged waveforms from different datasets: 
external trigger events are central vertical through going muons defined by scintillator paddles (fits from Figure 8), internal trigger are events from a minimum bias trigger obtained by asking for a 2-fold coincidence on the tank PMTs and a further selection requiring some balance between the 3 measured ADC charges, and LED pulses are flasher events.

|  | PMT1 | PMT2 | PMT3 |
| --- | --- | --- | --- |
| External trigger | 108.37 ± 0.41 ns | 102.96 ± 0.30 ns | 96.53 ± 0.40 ns |
| Internal trigger | 108.09 ± 0.40 ns | 100.95 ± 0.281 ns | 94.93 ± 0.39 ns |
| LED pulses | 106.51 ± 0.05 ns | 94.31 ± 0.03 ns | 99.26 ± 0.06 ns |

The same analysis was then repeated but using LED pulse events. The waveforms averaged over 1000 flashes are shown in Figure 9.  Again, the basis shape is an exponential decay with a time constant very close to the one obtained with muons. It is also clear from Figure 9 that a different structure shows up at the beginning of the waveform from PMT2. This spike is interpreted as being produced by “direct” light. If one looks closely at the pulse rise time, one finds out that the pulse from PMT2 starts also ~ 8 ns earlier than those of PMT1 and PMT3, as indeed we expect from the distance difference to the LED (1.56 m ( 5 ns/m). This time difference is more clearly seen in Figure 10 in which the three traces are super-imposed (after normalization discussed below).

### Light yield balance.

The waveform of PMT2 is thus the sum of a “diffused” contribution analogous in shape to the waveforms of PMT1 and PMT3 (i.e. exponential), and of a peak of width ~ 8 ns of  “direct” light coming in front of the diffused component. To try visualizing this, to quantify the fraction of the “direct” and to study the balance between the “diffused” components measured by the three PMTs, we have done the following:

We first normalized each one of the LED waveforms to the number of p.e.’s per VEM measured with the corresponding PMT. This accounts for the differences in the optical coupling and in the quantum efficiencies. The result is shown in Figure 10. One sees that the “diffused” component from PMT2 (in black) matches very well those of PMT1 and PMT3. We can than subtract the diffuse component from PMT2 using the waveform from PMT1 or PMT3. The result (PMT2 – (PMT1+PMT3)/2) is show by the blue histogram in Figure 11. This visualizes the “direct” light component of PMT2.

Figure 9:  Waveforms from PMT1, PMT2 and PMT3 averaged over 1000 LED flasher events. The time abscissa is in ns. The exponential fits gives 106 ns, 94 ns, and 99 ns, very close to the muon events fits with the noticeable exception of PMT2 for which the lifetime is found to be significantly shorter. This is understood to be due to an underlying contribution from short distance light trajectories. The effect of after-pulses is even clearer than with muons.

Finally, to quantify the fraction of direct light and the relative balance of the diffuse light, we simply integrate the normalized histograms. The results are given in Table 2 .

Figure 10:  The waveforms for PMT2 (in black) PMT1 (in red) and PMT3 (in green) after normalization using the calibration of p.e./VEM from section 5. The leading edge time difference is clearly seen as well as the spike of direct light. The scattered light exponential decay components are almost identical in shape and in intensity.

Figure 11: The “direct” light component on PMT2, shown in blue, has been isolated by subtracting the normalized waveform of PMT2 by the average of the normalized waveforms of PMT1 and PMT3.

Table 2: Estimating the fraction of direct light and comparing the scattered light components. Here, wi(t) is the waveform normalized by the calibration factor, Qi is the integrated signal expressed in units of VEM, D21 and D23 are the estimate for the “direct” light component estimated from the difference ot PMT2-PMT1 and from PMT2-PMT3, <D2> is the average of the previous estimates. Finally, S2 is the estimate of the “scattered” light component for PMT2. On can see that the scattered light component is indeed balanced on the three tubes to better than 10%.

| Q1 = 3.09 VEM | Q2 = 4.60 VEM | Q2 = 2.71 VEM |
| --- | --- | --- |
| Q2/Q1 = 1.48 | Q2/Q3 = 1.69 |
| D21 = 1.51 VEM | D23 = 1.89 VEM |
| D21 / Q2 = 32 % | D23 / Q2 = 41 % |
| <D2> = (D21 + D23) = 1.70 VEM |
| <D2> / Q2 = 37% |

## Conclusions

Although the LED was installed in FATII in a very off-centred position, we found that the balance between the PMTs is surprisingly good. The nearby (28 cm) PMT is affected by an excess of “direct” light which amounts to ~ 37 % of the total light. After subtracting this component, the remaining scattered light is balanced between the three PMTs to about 10%.  Further more, assuming that the direct light component can be eliminated in the future by improving the LED support and window, and that a the dedicated LED port can be installed somewhat half way between the centre and the PMT’s, we are confident that the light of the LED pulser will be naturally balanced between the PMTs to a very good precision.

## Bibliography

[1] J.S.Kapustinsky et al [1]. Nucl. Instr. Meth, A241 612 1985.

[2] J.E.McMillan, L.F.Thompson, R.A. Brook, private communication and 
ANTARES internal note CALI/2000-008 and technical note 3-INS-03-01-A
ANTARES Technical Design Report   http://antares.in2p3.fr/

# Text boxes

120 cm

92 cm

~ 60 cm

30 cm

~ 60 cm

LED port

Large hatch

PM domes: 
radius = 120 cm

Fill port with LED: 
radius = 92 cm

driver

LED
light cone in air ~ 15°

#### Fill port

GAP-2001-55