WINDHOEK-NAMIBIAN-OPERATIONAL TRAINING ON THE RADIO THEODOLITE AND MAINTENANCE TRAINING ON THE HYDROGEN GENERATOR

26TH MARCH – 01ST APRIL 2004

# INTRODUCTION

The Intermet Upper Air System and Proton Hydrogen Generator for the Namibian Meteorological Services in Windhoek were simultaneously installed June 2005. Training of participants for both equipment was also done in the very same week after the installations. However participants felt that scheduling installation and training of the two systems at the same was not suitable as the same officers had to participate in the two activities, hence this did not give them enough time to attend both training activities as they were running concurrently. It was on these bases that a second training had to be provided for participants and to eliminate other operational and malfunction of the two systems, which are indicted to be continuously recurring.

PROBLEMS REAISED

- Weak signal realized in the first few minutes of ascent.

- TEMP format not matching WMO standard code format

- Code E-13 error in the hydrogen generator.

Intermet Equipment

Workstation                            Decoder & Data Converter     Intermet Dish

Figure 1

Proton Hydrogen Generator

Hydrogen Generator Opened     Water Deionizer

Figure 2

CONFIRMATION

Intermet Equipment

Two successful ascents were launched with both Intermet sonde and Sippican Mark II sonde. The third launched with the Sippican sonde confirmed the problem. The signal was lost some few minutes after launching. Troubleshooting was done and detected that the receiver in the antenna was faulty.  It was found that moisture had gathered inside the housing of the receiver and that the gasket had blocked the breathers. A new receiver was flown in from Cape Town South Africa and was immediately installed. The system resumed normal operation. With the TEMP format it was a matter of making changes in the software.

Faulty Receiver Board                                  Mike holding the faulty receiver

Figure 3

Proton Hydrogen Generator

The E-13 error code indicated a leakage in the system. Gas leakage tester was used and no leakage was found except on the balloon-filling valve. On observation, it was found that the system error comes when the plant is in the idle mode. When the system is restarted the error disappears and resurfaces when all the tanks are filled and is back to idle mode.

TRAINING

As the training, was mainly for operators on the Intermet Upper Air System it covered the operation of the system, the antenna and IMETOS software, proper preparation of the Sippican Mark II radiosonde and manual insertion of headers on the TEMPS. On the Proton Hydrogen Generator it was resolved that the operators needed no training but technicians. The technicians who could attend were on field trip and when they came, seemed reluctant to participate until the very last day when one technician surfaced after persuasion. The training started on the 27th March 2006 to 31st March 2006.

PARTICIPANTS

Participants were mostly operators from Upper Air staff of Namibia Met office and myself (under the GCOS project for SADC countries). Although the Namibian Met Office authorities had requested the training, attendance was very poor especially on the maintenance of the hydrogen generator.

Mike showing use of CDU    Ben with Mike Releasing balloon   Operators practicing

Figure 4

#### IMETSO Software

This software forms the foundation of the operation of the whole upper air system. The software generates a number of products including TEMP, PILOT, T-PHIGRAMS and TRACKING PLOT of the radiosonde. The antenna controls are also included and hence the antenna can be controlled in the working station but cannot be started automatically from the working station. The IMETSO is capable of storing flight records that can be played back as simulated flights and any other reports of previous soundings can be retrieved.

IMS700 Antenna

The antenna can be controlled in two ways, one of which is with the receiving computer and the other by a Crew Display Unit (CDU). The CDU is weather prove handset attached to the arm of the antenna with control buttons on it. The CDU can also show the strength of the signal, which is very important. The CDU proved handy in daytime or whenever the balloon can be seen or the general direction the balloon is known. At night the use of nightstick attached to the balloon can be useful in manual tracking of the balloon.

SIPPICAN MARK II Radiosonde

The radio sonde plays a vital role leading to a successful ascent. The Sippican mark II radio sonde comes with a set of instructions on how to prepare it for the flight. The battery terminal cable if not twisted creates electro magnetic force (e.m.f) that would hamper the performance of the transmitter. The operators were strongly encouraged to follow the instructions on this radio sonde. Some failures of a successful ascent were as a result of improper sonde preparation. Before launching communication has to established first between the sonde and the workstation.

PROBLEM REALISED

- Side lobe tracking

- Unsuitable launch site

- Code E-13 error on the generator

- Indicators not reading the same values

- Leakage on the balloon filling hose.

Product pressure on Gen    Pressure on tank 1                  Pressure on tank 2

The figure below show one of the valves suspected not to be closing hence a leak on the balloon filling. The valve was tried to disassembled but appeared to be a fixed valve had to open. No spares available.

Figure 1

Figure 2 shows a new filter immediately after installation. This filter filters water from the tap before it gets to the water deionizer

Figure 2

On observation, the Hydrogen Generator plant appeared dusty and power cables hang lose from the wall. It appears no one has been assigned to take care of the hydrogen equipment.

## CONCULSION

At the end of the training operators felt that the problems they had are likely to come to an end. Michael Talbot promised to go and look further into WMO TEMP reports and try to compare. He would also to urge the software people to finish the correction in the TEMP as soon as possible. While still waiting for the correction in the TEMP software, the operators would manually enter the TEMP headers. As regards the maintenance technicians for the hydrogen generator they feel they still need to be trained at factory level.