Photogrammetrically Measured Distortions of Composite Microwave Reflector System in Vacuum at ~90K

Peter Mulé, Michael D. Hill, Henry P. Sampler

Goddard Space Flight Center, Greenbelt Maryland

# ABSTRACT

The Microwave Anisotropy Probe (MAP) Observatory, scheduled for a 2001 launch, is designed to measure temperature fluctuations (anisotropy) and produce a high sensitivity and high spatial resolution (better than 0.3° at 90 GHz.)  map of the Cosmic Microwave Background (CMB) radiation over the entire sky between 22 and 90 GHz.  MAP utilizes back-to-back composite Gregorian telescopes supported on a composite truss structure to focus the microwave signals into 10 differential microwave receivers.  Proper position and shape of the telescope reflectors at the operating temperature of ~90 K is a critical element to ensure mission success.

We describe the methods and analysis used to validate the in-flight position and shape predictions for the reflectors based on photogrammetric metrology data taken under vacuum with the reflectors at ~90K.  Contour maps showing reflector distortions were generated.  The resulting reflector distortion data are shown to be crucial to the analytical assessment of the MAP instrument's microwave system in-flight performance.

Keywords: reflectors, distortions, and photogrammetry

1. INTRODUCTION

The MAP instrument is designed to measure the anisotropy of the Cosmic Microwave Background (CMB) radiation over the full sky while operating at a temperature of approximately 90K. The instrument's Thermal Reflector System (TRS) consist of back-to-back Gregorian telescopes with reflectors manufactured by Programmed Composite Incorporated (PCI) of Anaheim California, and are constructed of SF-70A-75/RS-12D facesheets with a shaped Korex core sandwich bonded with a FM-73 film adhesive.  Each reflector has an aluminum honeycomb backing structure with XN-70A/RS-3 facesheets.  The truss structure on which the reflectors are mounted is constructed of XN-70A/RS-3 and M46J/RS-3 flat laminates bonded with EA 9394 adhesive.

Figure 1 shows the positions of the primary and secondary reflectors on the composite truss structure comprising the Thermal Reflector System   (TRS).  A flight unit as well as a Reflector Engineering Unit (REU) was manufactured identical except that the REU included only one reflector pair and part of the truss on the opposing side was removed.  Both the flight TRS and the REU were instrumental in characterizing the reflector system.

Figure 1: Thermal Reflective System (TRS)

Accurate predictions of in-flight reflector distortion and reflector position values were necessary to predict the instrument antenna patterns. Distortion and position values are input into Diffraction Analysis of a Dual Reflector Antenna (DADRA) code (a product of YRS Inc) to provide the antenna pattern at the various frequencies.  Distortions were initially derived with finite element analysis by applying flight temperatures to the structural finite element model.  In parallel, we were investigating photogrammetry as a technique to verify the reflector shapes during thermal vacuum testing.  A close range photogrammetry system1 used to acquire precise coordinates from photographs was acquired and initially used during thermal vacuum testing of a Reflector Engineering Unit (REU.  Dramatic differences in Finite Element Model (FEM) and photogrammetrically (PG) measured shapes were realized (Figure 2).  Attempts to make sense of the FEM were unsuccessful.  Photogrammetric measurements were adopted as the method of deriving reflector shape and point position values to be used for the flight antenna pattern and beamwidth analysis of the reflectors.

Figure 2: REU Primary Reflector FEM vs Photogrammetry Deformations from Ideal at ≈90K

(Plots are of z-axis deformation in the reflector coordinate system)

The desired surface figures (ideal) were defined for a cold reflector by Princeton U., and actual surface figures were verified at room temperature by the REU and TRS manufacturer. They performed a continuous surface scan with an SMX TM laser tracker prior to coating the surfaces and subsequent integration of the reflectors as a telescope unit. Tooling balls located on ears extending from each reflector were measured and valued at this time.  The coordinates of the reflector tooling balls with warm offsets were used to align the reflectors in a local instrument reference frame during integration at room temperature.  Tooling hole boss (THB) targets on the TRS truss structure were also added and valued in the local reference frame of the instrument to provide a means of reestablishing the reference frame after removing the tooling ball targets on ears extending from the reflectors. A combination of the warm (atmospheric pressure) and cold (vacuum) PG measurements, the defined ideal surface figure, and a knowledge of the warm surface positions with respect to ideal (from the TRS manufacturer), allowed us to describe the reflector positions under flight-like conditions. DADRA code was then used to analyze the effects of the relative positions and distortions of the reflectors on the optical performance (beamwidth and antenna pattern).

2.  Photogrammetry Measurements

The target and point position measurements reported here were acquired from retro-reflective photogrammetry targets designed to be placed in the THBs and from stick-on photogrammetry retro-reflective targets placed on the front surfaces of the reflectors (two primaries -35 targets each, 2 secondaries-13 targets each).  Measurements were performed at ambient temperature at atmospheric pressure and in vacuum at ambient temperature and when cold (Table 1).  The photogrammetry camera used to make the measurements was mounted in a thermally controlled canister on a carousel mounted atop a helium shroud surrounding the TRS.  The carousel itself was remotely operated to position the camera in 42 locations and 2 orientations to allow for triangulation and resection of target coordinates and camera positions in the photographs taken. PG target positions were measured with a precision of 0.002 inches.

Table 1: TRS Photogrammetry Test Conditions

| PG Run | Pressure | Temp (K) | Comments |
| --- | --- | --- | --- |
| 1 | Atmospheric | 290 | Established reflector target baseline values |
| 2 | Vacuum | 290 | Identified effects of vacuum (none) |
| 3 | Vacuum | 91 | First cold deformation case |
| 4 | Vacuum | 95 | Second cold deformation case |
| 5 | Atmospheric (N2) | 294 | Identify residual effects of thermal vacuum cycling (none) |

To verify the precision of the PG system measurements, ambient temperature and pressure PG measurements of the TRS THB targets were compared with measurements made with a Leica Laser Tracker (LT) system and measurements made using a theodolite network system (the Analytic Industrial Metrology System (AIMS)). The three-metrology systems agreed to within 0.007 inches (0.178 mm) worst case. The PG system measurements when compared to each other were repeatable to 0.002 inches (0.051 mm) and agreed with the values established in the final Laser Tracker measurements performed by the telescope's manufacturer.

Analysis of the deformations of the TRS in thermal vacuum required only minor changes in the methods of establishing the local instrument coordinate reference frame.  The low coefficient of thermal expansion (CTE) and symmetry of the truss structure allowed the use of room temperature THB coordinates measured on the truss when performing a transformation of coordinates into the local coordinate reference frame.  The truss THB targets were located symmetrically such that no biasing would occur in their alignment when cold. Vertical translations of the truss at its Titanium flexure interface due to flexure contraction (0.0046 inches (0.117 mm)) were noted. The final coordinates when cold were established by least squares fitting the truss THB target measurement to the room temperature measurement values. Then scaling back using scale-bars of known CTE and temperature (the scale factor was very near 1 confirming a low CTE for the truss), together with applying predicted flexure contraction translation.

3.  Reflector Distortions

Figure 3 shows the placement scheme of the PG targets on the surfaces of the TRS primary and secondary reflectors. Contamination concerns as well as line of sight concerns restricted the number and location of surface targets.  Target locations were chosen to best describe the overall distortion based on the distortion pattern identified during Reflector Engineering Unit (REU) testing. Targets were not placed on the upper portion of the secondary reflector due to the camera could not view it in the test configuration.

Figure 3: PG Target Positions on Primary Reflector 2R and Secondary Reflector 2

To develop a description of the distortions of the primary reflectors when cold and due to the DADRA code requiring a specific input format, the following process was used:

- The measured coordinates of the reflector targets when cold were least squares fit to their warm positions. This placed the cold surface of the reflector in a position known with respect to the ideal surface figure.

- Knowledge of the reflector warm surface position with respect to an ideal surface figure, (from telescope manufacturer data) allowed us to transform positions from step 1 to their locations for ideal surface figure.  Steps 1 and 2 provided knowledge of the measured reflector cold positions with respect to an ideal surface figure.

- The target positions were then compared with the ideal surface figure using Imageware surface fitting software where a shift in the standing axis direction was allowed to partially correct for the 0.010-inch (0.254 mm.) target thickness.

- Residuals from the least squares fit) from step 3 were used to develop distortion values in the reflector local coordinate reference frame (the format required for DADRA code).

- Fine grid interpolation was performed using NASTRAN modeling. DADRA requires a x, y, and dz input so that an interpolation of dz could be done using a flat plate model with enforced displacement.

We compared the distortions measured with the photogrammetric camera (PG) with the warm distortions measured with the PCI laser. Table 2 is a summary of LT and PG based surface distortions for the TRS primary reflectors. This comparison highlighted differences in the outer annulus that has the largest unsupported areas and the least stiff backing structure. These differences are likely due to relaxation of the reflector material between the initial LT measurement and the PG measurement. LT measurements were made prior to the surface coating and integration of the reflectors onto the truss structure.

Table 2: TRS Primary Reflector RMS Distortion Summaries (i)

|  | Primary 2R | Primary 3 |
| --- | --- | --- |
|  | r < 9.8 | 9.8< r <19.7 | r > 19.7 | r < 9.8 | 9.8< r <19.7 | r > 19.7 |
| Requirement | 0.0015 | 0.0020 | 0.0030 | 0.0015 | 0.0020 | 0.0030 |
| LT Warm + (ii) | 0.0024 | 0.0021 | 0.0034 | 0.0018 | 0.0017 | 0.0031 |
| PG Warm (iii) | 0.0024 | 0.0024 | 0.0052 | 0.0022 | 0.0037 | 0.0060 |
| PG Cold (iii) | 0.0069 | 0.0053 | 0.0117 | 0.0047 | 0.0064 | 0.0111 |

- Values represent surface normal errors.  Distortion contour plots represent z component error of the optical coordinate system.

- LT measured points and additional derived points included for more uniform and complete distribution.

- PG measured points and additional derived points included.

Table 3 below is a summary of surface target residuals for a secondary reflector in the local coordinate reference frame of the TRS after least squares fitting to the warm data. The small residuals agree with measurements made on the REU.

Table 3: TRS Secondary Cold Fit to Warm PG Residuals (*)

|  | Secondary SN 2 | Secondary SN 3 |
| --- | --- | --- |
| No. of Targets | 9 | 8 |
|  | dX | dY | dZ | dX | dY | dZ |
| Standard Deviation | 0.0012 | 0.0015 | 0.0009 | 0.0021 | 0.0008 | 0.0010 |
| Minimum | -0.0025 | -0.0027 | -0.0010 | -0.0040 | -0.0010 | -0.0018 |
| Maximum | 0.0015 | 0.0022 | 0.0015 | 0.0030 | 0.0013 | 0.0011 |

(*)	Three upper targets on the secondary reflectors were at extreme angles and are included in a limited number of photos.  These targets were not included in summary.

Contour maps were created to visually represent the distortions of the reflectors.  The contours represent actual surface as compared to ideal surface deviations in the z axes direction in the reflector coordinate system.  The local coordinate axes for the reflectors can be seen in Figure 4.  DADRA code requires the x, y, and delta z format. Figure 5 shows the relations between the surface normal and the reflector z translation direction for a single point.  An Ideal reflector would appear as a flat disc of uniform color.  Images are orientated such that we are looking at the reflective surface along an axis nearly parallel with the individual reflector z-axis (a slight x-axis rotation is applied to provide view of distortion).  The offset angles are used in the images to give a better idea of the distortion patterns and do not show reflector shape.

Figure 4: Reflector Coordinate Systems

Figure 5: Relations between Surface Normal and Reflector dZ for a single point in the local area of the point

The contour maps in figure 6 show warm (as-built) distortions of primary reflector 2R as derived from laser tracker and photogrammetry measurements. The 35 PG targets produce a relatively good representation of the distortion over the entire surface.  LT data consisted of approximately 1000 points distributed over the entire surface.  Differences in these images, and those for reflector SN 3, are mainly near the outer perimeter and are likely due to extrapolation in areas of high distortion gradient and minor changes in the reflectors shape during the surface coating and integration process following the initial LT measurements.

Figure 6: Primary Reflector 2R Laser Tracker and PG Camera Comparisons (Warm)

The contour maps in figure 7 show warm (as-built) distortions of primary reflector 3 as derived from laser tracker and photogrammetry measurements. There is an apparent difference in the lower section of the contour maps.  Later THB measurements show the balls in this area to be out of position by approximately 0.010 inches when compared to the baseline data measured prior to delivery of the reflector system.

Figure 7: Primary Reflector 3 Laser Tracker and PG Camera Measurements (Warm)

The contour maps in Figure 8 shows the primary reflector distortions based on the cold PG data.  The distortions when cold are nearly identical for the two reflectors.  These images reflect the DADRA distortion input values used to predict flight performance.

Figure 8: TRS Primary Reflectors PG Measured Distortions at 90K

The contour maps in figure 9 were created from the PG measurements performed on the REU in a cold vacuum environment.  These are included to show the similarity to TRS distortions and how a reduced number of targets were able to create a good shape.  When compared to the TRS cold images it revealed that a hat ring stiffener added to the outer perimeter of the TRS reflectors slightly reduced the distortion along the outer rim, especially in the middle and upper rim areas where the ring properties were the stiffest.  The primary reflector back rib pattern is apparent in this cold distortion image.

Figure 9: PG Camera Measurements at Room Temperature and at 90K

The secondary reflector contour maps of figure 10 are based on the REU measurements. The secondary reflector surface was completely targeted (81 PG target positions marked) for the REU testing. The images show that the overall distortion of the secondary after the cold transition is negligible supporting the limited TRS secondary reflector measurements.

Figure 10: REU Secondary Reflector PG Measurement Distortion Contour Maps

4.  Photogrammetry Based Performance Predictions

To derive the flight beamwidth predictions and antenna patterns for the MAP optical system (feedhorns and reflectors), the TRS reflector shape and position data were incorporated in the DADRA code input.  Secondary reflector warm distortions were used for flight predicts since negligible distortions were measured for the Secondary and their full surface shapes could not be obtained. Flight performance predictions were within the criteria set by MAP systems engineering analysis. A comparison with ground test results was made to give confidence in combining the PG measured shape, position predictions, and DADRA code.  For the integrated telescopes and feed horns, the comparison showed excellent agreement between test measurements and analytic predictions at the various feed-horn positions.   Results are shown in Table 4 where the X-Beam FWHM is the lateral beam width and Y-Beam FWHM is the vertical beam width.

Table 4: Comparison of Test and Analytical Performance at Ambient

|  | Pointing (1) | Beam Size |
| --- | --- | --- |
| Frequency | Feed Horn Position | Elevation (°) | Azimuth (°) | X-Beam FWHM (°) | Y-Beam FWHM (°) |
|  |  | Test | Analysis | Test | Analysis | Test | Analysis | Test | Analysis |
| 30 GHz | Ka1A | -22.26 | -22.28 | -2.34 | -2.34 | 0.551 | 0.561 | 0.684 | 0.672 |
| 40 GHz. | Q2A | -17.77 | -17.77 | 1.97 | 1.97 | 0.569 | 0.543 | 0.491 | 0.469 |
| 60 Ghz | V2A | -19.66 | -19.69 | 2.04 | 2.03 | 0.304 | 0.317 | 0.307 | 0.313 |
| 90 GHz. | W2A | -19.04 | -19.04 | -0.56 | -0.56 | 0.190 | 0.205 | 0.187 | 0.191 |
| 90 GHz. | W4A | -20.14 | -20.16 | 0.60 | 0.61 | 0.202 | 0.210 | 0.201 | 0.208 |
| 30 GHz. | Ka1A | -22.26 | -22.27 | -2.42 | -2.43 | 0.549 | 0.560 | 0.690 | 0.672 |
| 40 GHz. | Q2A | -17.72 | -17.72 | 1.95 | 1.92 | 0.573 | 0.541 | 0.491 | 0.465 |
| 60 GHz. | V2A | -19.65 | -19.64 | 2.00 | 2.03 | 0.293 | 0.313 | 0.304 | 0.306 |
| 90 GHz. | W2A | -19.05 | -19.01 | -0.61 | -0.62 | 0.189 | 0.201 | 0.189 | 0.193 |
| 90 GHz. | W4A | -20.13 | -20.11 | 0.55 | 0.55 | 0.193 | 0.205 | 0.199 | 0.207 |

- Test pointing values correct for approximated offset in chamber boresight

5.  Conclusions

Photogrammetric measurements have allowed us to obtain the shape and position values of our reflectors in vacuum at cryogenic temperatures.  Empirical data of the shape when cold was a significant improvement over analytical predictions and also highlighted shape errors due to fabrication where analytical predictions assume a perfect initial shape.  When compared to Laser Tracker metrology, photogrammetric measurements captured the as built distortions of the reflectors and identified a possible relaxation of the reflector material.  Photogrammetry together with beam mapping and DADRA code enabled us to quantify and qualify the ground and on-orbit characteristics of the microwave reflectors.

ACKNOWLEDGMENTS

The authors would like to thank Acey Hererra (Swales and Associates), John Brown (Geodetic Services, Inc.), Programmed Composite, Inc., and John Beggs (Goddard Space Flight Center).

## REFERENCES

- Automation in Digital Close-Range Photogrammetry, Clive S. Fraser, First Trans Tasman Surveyors Conference, 12-18 April, 1997, New Castle