Pierre Auger Project Electronics

Front End (FE) Board

Kwame Bowie

September 1, 2000

##### Revised: 4/17/2001

- Introduction

The Front End (FE) board is the data acquisition board for the Surface Detector Electronics system.

The Surface Detector Electronics system consists of a CPU board (CPU), a Slow Controls board (SC), a Time Tagging board (TG), an Ethernet board (ETH), a GPS receiver (GPS), and the Front End (FE) board. The FE board is directly linked via the Local Bus to the all the subsections except the Slow Controls board, which is accessible via the CPU. The block diagram Fig. 1 shows the basic layout.

Figure 1:  Surface Detector Electronics Interconnections

- Operation

The FE board is the data-acquisition board, the data storage board, and the trigger generation board for the Surface Detector. The main functions of the board are listed below:

- FE Board must take 6 channels of analog data from photo multiplier tubes and convert it to digital data.

- FE Board must store this digital data in memory deep enough to hold 25.6 (s of data (6.4 (s pre-trigger and 19.2 (s post-trigger).

- FE Board must transmit this data to the CPU based upon trigger conditions.

The following are the specifics of the board operation based upon the Surface Detector and Front End board specification.

- The analog to digital data conversion (ADC) is to be done at a sampling rate of 40MHz and with a precision of 10 bits per channel.

- There will be six (6) channels of analog inputs from the photo multiplier tubes (PMTs). The three (3) high gain channels will contribute to the analog trigger, while all six channels will contribute to the digital trigger. All six channels will be stored as data within the FIFO memories.

- There will be a total of three (3) types of triggers: analog, digital, and external

- The memory will be double buffered so that an event may be accepted while data is being read out.

- The FE board must be air-cooled  with no fan.

- The board must operate at temperatures between –20C and +60C.

To ensure that a prototype will be available for initial insertion, two (2) versions of the FE board will be produced: Version1.0 and Version 2.0.

### 2.1	Version 1.0

The first version, Version 1.0, of the FE board will include only the motherboard and the analog daughter card. However, the Version 1.0 will support the use of a digital daughter card by providing a connector and signal paths for the digital daughter card. The digital daughter card is to be developed in parallel at a different location. The digital daughter card will produce a digital trigger signal, which will be OR’ed with the other two triggers (Analog and External) to control data flow into the FIFOs.

To improve development time of Version 1.0, the FE board has been split into three (3) separate boards: the FE motherboard, the Analog Daughter Card, and the Digital Daughter Card. By separating the FE Board design into distinct parts, several engineers may work on various portions of the board in parallel. By making these distinct parts separate boards, a delay on one board will have minimal effect on the others

.

### 2.2	Version 2.0

Version 2.0 of the Front End board will have the same physical dimensions and stacking- connector positioning, but will eliminate the FIFOs and the PLD. In place of the FIFOs and PLD will be an ASIC socket for an ASIC designed to integrate all of the digital logic into a single chip. As a result, the digital daughter card connector along with all digital logic will be gone from Version 2.0. Also, the ASIC will handle a more advanced triggering scheme which will be all digital so no analog trigger will be present.

This document is to be used as a user’s manual for the FE board and will discuss only sections common to the FEv1 and FEv2 boards.

- Input signals

Input signals to the FE board originate from two (2) PMTs in groups of 3 analog signals. The two PMTs are set up so that their output gain is different by a factor of 32. To simplify nomenclature we simply call the two PMTs high gain and low gain. The PMT signals are appear at the FE board as six single-ended, negative edged signals on 50 Ohm cables from the PMT bases. The cables are terminated with high frequency SMA sockets. The high gain channels have an input voltage range of 0 to –0.3V and the low gain signals have a voltage span from 0 to –1.5V.

- 6 single-ended signals

- 50 Ohm cable impedance from bases

- High gain range (0 to –0.3V)

- Low gain range (0 to –1.5V)

The FE board has the ability to handle the 6 input signals in either purely single ended form or in a pseudo-differential form via jumper selection. The pseudo-differential method provides the ideal solution in low-noise environments as it eliminates possible ground offsets between the FE board and the PMT bases. To connect with the coax cable, we have selected standard, right-angle, PCB-mount, SMA connectors from Amphenol (part # : ).

- Shaping/Amplification/Filtering

5.1 Overvoltage protection

Before any filtering, shaping, or amplification is done, the amplifiers must be protected from over-voltage conditions. This actually serves 2 purposes: first, it prevents large voltage spikes from exceeding the maximum op-amp voltage ratings. Reliability will certainly be a major concern in this project as the stations are predicted to be serviceable only at very long intervals. Replacing op-amps after any large event or noise spike is not an option. The second reason for including surge protection on the analog input signals is to lower the power consumption within the analog portion of the FE board. To achieve the over-voltage protection goal, we use series-connected dual pin diodes. For the high gain channels we only use 2 per channel since we can clamp after one diode drop since the signal is limited to 300mV. However, for the low gain channels, the signal range is much larger (1.5 V) so we need at least 3 diode drops in the signal direction before we clamp the signal. So for the low gain channels, each channel receives 3 pin dual diodes. The pin diodes that were selected for this application are part #: .

5.2 Amplification

The very next step after surge protecting the inputs is to amplify the inputs. The amplification enables the analog inputs to match the ADC input range. We have selected the AD9203 ADC, and are using a 0 to 1V input span, so the inputs must be amplified accordingly. I use the term amplification as a generalization of the operations performed on the inputs. The operations are actually polarity inversions, amplifications, and/or attenuations. All three of these terms may be generalized and considered amplifications by a gain factor, A. The sign and absolute value of this gain factor determines whether the amplification is a polarity inversion, an amplification, or an attenuation. In our case, all high gain channels have a gain factor -10/3 and all low gain channels have a gain factor of -2/3. Thus the high gain channel is actually inverted the amplified, whereas the low gain channel is inverted then attenuated. The part selected for the amplifier section was the AD8011 current feedback op-amp because it provides excellent gain vs. bandwidth characteristic and is also a very low power device.

5.3 Filtering

After the amplification stage, the inputs are filtered. This filter is a simple, 2-pole, anti-aliasing filter. The signals that we intend to recover span a wide range of frequencies and have spectral components that also span many frequency ranges. The desired signals range from DC to the maximum resolvable by the ADC. As a result, under-sampling techniques must be rejected, and an anti-aliasing filter provides good rejection of unwanted noise signals above the Nyquist frequency. The sampling frequency of the ADC is 40 MHz so a simple 2-pole active filter with cutoff frequency at 20 MHz works very well.

- Analog to Digital Conversion

The signal lines are single-ended analog waveforms which represent the filtered, shaped, and amplified PMT outputs. These six signals are the inputs for the analog-to-digital converters (ADCs). The ADC’s ground connections are isolated from each other and only return to their own ground plane in order to prevent digital ground currents returning to the analog input.

Figure 2:   Analog-to-Digital Converter Supply Isolation

The sampling device that was selected was the Analog Devices AD9203 flash ADC chip. This chip was selected due to its very low power consumption with the high conversion rate and decent precision. The AD9203 utilizes both edges of the SYSCLK signal. The ADC inputs are clocked into the flash converter’s sample and hold amplifiers on every trailing edge of the SYSCLK signal. The digital outputs are latched onto the data bus at each rising SYSCLK edge. To enable such a high sampling rate at decent precision and low power, the AD9203 incorporates a pipelined flash architecture. The intrinsic delay of each input sample is 7 rising SYSCLK edges.

- Digital Interface

Due to the large difference between the FEv1.0 and FEv2.x boards, only a very brief synopsis will be given of the digital interface will be covered here. To fully understand the digital interface for each board, the Appendices (FEv1.0 and FEv2.x) must be referenced.

All FE boards must interface with the PowerPC (PPC) 403GCX local station CPU in order to transfer the accumulated data. The transfers occur once a data buffer is filled and ready to transmit the ADC data. The PPC is signaled that the FE is ready when the FE asserts EVTCLK. The PPC processor then determines when it would like to conduct a data transfer. To ease the load on the processor, the DMA fly-by-burst memory transfer mode was selected. By deciding to use the DMA transfer mode, the ALU and registers of the CPU can still execute instructions uninterrupted.  The fly-by-burst mode, is the fastest and simplest of all of the DMA transfer modes, which is very important as the CPU data bus is locked out at this time. The actual transfer uses 1 wait state for the memory to enable adequate time for the FE boards to access the necessary data. As a result, the data is put onto the bus at a 20MHz rate. Each buffer consists of 2 FIFOs therefore reading both of these 1024 word buffers will take roughly 102 us. However, after several CPU write cycles, the CPU intervenes to allow the DRAM to refresh and row switch so the full data transfer time ends up being a little bit longer than that.  Specifics of the DMA fly-by-burst protocol can be obtained from the PPC403 User’s Manual.

# Text boxes

SURFACE DETECTOR ELECTRONICS BOARD

Pwr.

Supl.

GPS

C3

C3

## CPU

## TG

## ETH

## SC

C1

C2

C1

## FE

C2

AVDD

VSS

AVSS

VDD

Ferrite

ADC