DEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR DEVICES AND RADIOLOGICAL HEALTH

CIRCULATORY SYSTEMS DEVICES PANEL

OPEN SESSION

Monday, March 4, 2002

10:10 a.m.

Gaithersburg Marriott Washington Center

9751 Washington Boulevard

Gaithersburg, Maryland

PARTICIPANTS

Warren K. Laskey, M.D., Acting Chairperson

Lesley Ewing, M.D., Executive Secretary

VOTING MEMBERS

Salim Aziz, M.D.

Janet T. Wittes, Ph.D.

CONSULTANTS

Anthony Comerota, M.D.

James A. DeWeese, M.D.

Michael J. Domanski, M.D.

Francis Klocke, M.D.

Marvin Konstam, M.D.

Steven Nissen, M.D.

Pilar Ossorio, Ph.D., J.D.

Ileana Pina, M.D.

CONSUMER REPRESENTATIVE

Robert A. Dacey

INDUSTRY REPRESENTATIVE

Michael Morton

PATIENT REPRESENTATIVE

Joseph Knapka, Ph.D.

FDA

Bram Zuckerman, M.D.

C O N T E N T S

PAGE

Call to Order:  Warren Laskey, M.D.	                     4

Conflict of Interest Statement        	                  4

Open Public Hearing                  	                   9

Sponsor Presentation: Thoratec Corporation

P910014/S016 HeartMate VE LVAS

Introduction

Donald A. Middlebrook      	                 9

Device Overview

Victor Poirier           	                  14

REMATCH Trial Clinical Results

Eric Rose, M.D.             	               22

REMATCH Patient Population

Lynne W. Stevenson, M.D.       	            29

REMATCH Trial Clinical Results

Eric Rose, M.D.                   	         39

Summary and Closing Remarks

Donald A. Middlebrook           	           55

FDA Presentation

Michael Berman, Ph.D., Lead Reviewer 	        57

Julie Swain, M.D., Clinical Reviewer    	     68

Gerry Gray, Ph.D., Statistical Reviewer   	  83

Questions to the Committee            	       96

Open Discussion                               	         100

Panel Recommendations                   	              222

Open Public Hearing  	                                 260

Sponsor Comments        	                              261

Panel Voting               	                           262

P R O C E E D I N G S

Call to Order

DR. LASKEY:  I would like to call this panel meeting session to order.  This morning's topic will be a discussion of the supplement to a Premarket Application for Thoratec Corporation's HeartMate VE Left Ventricular System.

I would like to now have the Executive Secretary read the Conflict of Interest Statement.

Conflict of Interest Statement

DR. EWING:  Good morning.  I would like to welcome everyone to the meeting this morning and especially I would like to thank the panel members for their time in reviewing this application.

The following announcement addresses conflict of interest issues associated with this meeting and is made part of the record to preclude even the appearance of an impropriety.

To determine if any conflict existed, the Agency reviewed the submitted agenda for this meeting and all financial interests reported by the committee participants. The conflict of interest statutes prohibit special government employees from participating in matters that could affect their or their employer's financial interests.

The Agency has determined, however, that the participation of certain members and consultants, the need for whose services outweighs the potential conflict of interest involved is in the best interests of the Government.

We would like to note for the record that the Agency took into consideration matters regarding Drs. Salim Aziz, Francis Klocke, and Marvin Konstam.  Each of these panelists reported interests in firms at issue, but in matters that are not related to today's agenda.  Drs. James DeWeese and Anthony Comerota reported that their institution has an involvement with the firm at issue.

The Agency has determined that all of these individuals may participate fully in the discussions.

In the event that the discussions involve any other products or firms not already on the agenda, for which an FDA participant has a financial interest, the participant should excuse him or herself from such involvement, and the exclusion will be noted for the record.

With respect to all other participants, we ask in the interest of fairness that all persons making statements or presentations disclose any current or previous financial involvement with any firm whose products they may wish to comment on.

DR. LASKEY:  Thank you, Lesley.

I would like to now have the panel members introduce themselves starting on my right.

DR. ZUCKERMAN:  Bram Zuckerman, Acting Director, Division of Cardiovascular and Respiratory Devices, FDA.

DR. WITTES:  Janet Wittes, panel member.  I am a statistician at Statistics Collaborative in D.C.

DR. KONSTAM:  Marv Konstam from Tufts University, New England Medical Center.

DR. COMEROTA:  Anthony Comerota, vascular surgeon from Temple University in Philadelphia.

DR. NISSEN:  I am Steve Nissen.  I am a cardiologist from The Cleveland Clinic.

DR. AZIZ:  Salim Aziz, Associate Professor at the University of Colorado Clinic.

DR. EWING:  Lesley Ewing, Executive Secretary.

DR. LASKEY:  Warren Laskey, interventional cardiologist from the University of Maryland.

DR. PINA:  Ileana Pina, Cardiology, Case Western Reserve University, Cleveland.

DR. OSSORIO:  Pilar Ossorio, University of Wisconsin at Madison, bioethicist and lawyer.

DR. DeWEESE:  Jim DeWeese, cardiac and vascular surgeon, University of Rochester, New York.

DR. KLOCKE:  Francis Klocke.  I am a cardiologist from Northwestern University Medical School.

MR. DACEY:  Robert Dacey, Consumer Representative, Boulder, Colorado.

MR. MORTON:  Michael Morton.  I am employed by Alcon Laboratories, and I am the Industry Representative.

DR. KNAPKA:  Joe Knapka, Patient Representative, from Olney, Maryland.

DR. LASKEY:  Thank you, members.

Dr. Ewing, would you now read the voting status statement, please.

DR. EWING:  This is the Appointment to Temporary Voting Status.  Pursuant to the authority granted under the Medical Devices Advisory Committee Charter, dated October 27, 1990, and as amended August 18, 1999, I appoint the following individuals as voting members of the Circulatory System Devices Panel for this meeting on March 4, 2002:  Pilar Ossorio, Michael Domanski, who will be coming soon, James DeWeese, Francis Klocke, Anthony Comerota.

For the record, these people are special government employees and are consultants to this panel under the Medical Devices Advisory Committee.  They have undergone the customary conflict of interest review and have reviewed the material to be considered for this meeting.

In addition, I appoint Dr. Warren Laskey to serve as Panel Chair for the duration of this meeting.

I also have another Appointment to Temporary Voting Status to read into the record.

Pursuant to the authority granted under the Medical Devices Advisory Committee Charter of the Center for Devices and Radiological Health, dated October 27, 1990, and as amended August 18, 1999, I appoint the following individuals as voting members of the Circulatory System Devices Panel for the meeting on March 4 and 5, 2002:  Steven E. Nissen, Ileana Pina, and Marvin A. Konstam.

For the record, Dr. Nissen is a voting member, and Drs. Pina and Konstam are consultants to the Cardiovascular and Renal Drugs Advisory Committee of the Center for Drug Evaluation and Research.  They are special government employees who have undergone the customary conflict of interest review and have reviewed the material to be considered at this meeting.

Open Public Hearing

DR. LASKEY:  I would like to now open this morning's session for the open public hearing.

Are there any individuals requesting floor time?

[No response.]

DR. LASKEY:  If not, then, I would like to close the public hearing and move on to the sponsor's presentation.

DR. EWING:  While the computers are being set up, I would like to remind all the speakers, especially the speakers for the sponsor, to identify themselves and to state their conflict of interest.

Each time you speak, because we are labeled up here on the panel, but the people that are not so labeled, if you could state your name, so the transcriptionist can write who is speaking.

Sponsor Presentation

Thoratec Corporation

P910014/S016, HeartMate VE LVAS

Introduction

Donald A. Middlebrook

MR. MIDDLEBROOK:  Thank you, Dr. Laskey, and good morning to all.

My name is Don Middlebrook.  I am Vice President of Regulatory Affairs and Quality Assurance for Thoratec Corporation.  I am also currently a shareholder.

I would like to begin this morning with a word of thanks to the FDA, especially to our FDA reviewers, all of the FDA invited panel members, and our clinical investigators and experts that are here on our behalf.

All of you, and certainly to all my Thoratec associates, all of you have worked hard to prepare for today's discussions, and those efforts are very much appreciated.

[Slide.]

This is a brief outline of our presentation that we will be making this morning.  We will be discussing the very important results from the REMATCH trial, a first of its kind clinical trial to compare a medical device to a drug therapy for the treatment of end stage congestive heart failure in patients ineligible for cardiac transplantation.

[Slide.]

Our speakers this morning are:  Victor Poirier, who will provide a brief overview of the HeartMate VE LVAS; Dr. Eric Rose from Columbia Medical Center, who will discuss the safety and effectiveness results documents in the REMATCH trial; and Dr. Lynne Warner Stevenson from Brigham &  Women's Hospital in Boston, who will provide information on the patient population study, and also discuss medical management in the control treatment group.

[Slide.]

We have also invited a number of independent REMATCH experts who may be called upon to speak on our behalf, who were involved in some way in the study, either in the study design, the study management, or in the analysis of the results.

[Slide.]

Just a brief word about our company Thoratec Corporation.  The company was founded in 1976 in Berkeley, California.  We merged with Thermo Cardiosystems in February 2001, to be known thereafter as Thoratec Corporation.

Our product focus is circulatory support, vascular grafts, and diagnostic blood testing.

Our corporate offices are located in Pleasanton, California.  We have 800 employees worldwide, and we are a world leader in cardiac assist device technology.

[Slide.]

We will be presenting a summary of the contents of the PMA that contain the results from the REMATCH trial. This is a landmark multicenter, prospective, randomized, controlled trial that looked at the HeartMate VE LVAS versus optimal medical management.

This study was initiated under a cooperative agreement between Thoratec, the NIH, and Columbia University.

The PMA we have submitted is seeking approval to expand the current HeartMate VE LVAS indications for use to include patients with end-stage left ventricular failure who are ineligible for cardiac transplantation.

Before I conclude my opening remarks, I would like to state that the REMATCH study has been the subject of a number of important scientific publications.  I would like to just mention a few of those.

In January 2001, the ACC published a consensus report entitled, "Mechanical Circulatory Support 2000, Current Applications and Future Trial Designs."  This was the proceedings from a conference sponsored by leading cardiology and heart failure professional societies and included participants from the FDA and the NIH.

In this report, the REMATCH study was cited several times as a benchmark study for evaluating mechanical circulatory support device treatment outcomes in patients with end-stage congestive heart failure.

Also, in November 2001, the REMATCH study results were published in The New England Journal of Medicine and were presented at the American Heart Association annual meeting that same month in Anaheim, California.

In December of 2001, the American Heart Association named the REMATCH study the number 2 in the top ten research advances for 2001.

With that said, I would like to ask Victor Poirier to come to the podium, who will provide some information on the device overview.

Device Overview

Victor Poirier

MR. POIRIER:  Thank you.  I am Victor Poirier.  I am an employee of Thoratec and a shareholder of Thoratec.

[Slide.]

To begin with, the device that was used in this clinical study is the HeartMate Left Ventricular Assist Device.  The device is implanted either in the peritoneal cavity or in a surgically-created pocket in the abdominal wall.

The inlet of the device is inserted into the left ventricular cavity while the outlet is anastomosed to the ascending aorta.  With this arrangement, blood simply drains from the natural heart into the device, and we provide the energy to propel blood through the body.

The device has a volume of 83 ml and can produce in excess of 10 liters per minute above flow.  Power is delivered to the device through a percutaneous lead.  This is a coaxial lead.  The center of the lead is an electrical conductor, and surrounding that is a cavity to propel air in and out of the system.

This lead is attached to an external controller that controls the function of the device, allows the patient to operate in an automatic mode or a fixed rate mode, has all of the alarms that are required, and that controller is then attached to the power.

The power is obtained either from batteries worn by the patient, there is one on each side of the patient, or it can be attached to a bedside console.

[Slide.]

When we started the trial, we started with the HeartMate VE system, and in that configuration, there are screw rings that are used to attach the components.  To lock those screw rings in, we utilized locking sutures, which were either applied by the physician during implant or at the factory.

In reality, we learned that these were not 100 percent safe.  In addition to that, from pump migration or patient movement, we also uncovered that in some patients, a kink would develop at the outflow.  This would be a complication which we would have to increase the pump chamber pressure to overcome this resistance to propel the blood out through the outflow.

In addition to that, rubbing of the adjacent areas of the graft could cause abrasion, which could cause a hole to develop.  So, we made modifications to correct that, and we configured the SNAP device.

What we did is we changed the screw rings to use locking screw rings, which essentially replaced the suture with a metal lever to lock the rings in place.  In addition to that, we added a bend relief.

So, the differences between the VE system and a SNAP system are only those two things.  There is nothing else that has changed.  The internal components are identical. So, any information that we gain from the endurance testing or the evaluations that we did in the past are still valid.  The only changes are the rings and the bend relief, and these were not evaluated in the in-vitro life test.

[Slide.]

Four and a half years ago, we started a reliability test.  We took 15 systems and we put them on test.  What we learned from that test was that there is an 88 percent chance that the LVAD will be free of critical failures at one year, and a 76 percent chance that LVAD will be free of critical failures at two years, and this is at a 90 percent confidence interval.  The mean time to failure is 3.1 years.

If you look at these results and compare them to what we experience clinically, you will see that there is a difference.  These results are better.  So, what happened clinically?  There are three factors clinically that affect  reliability.

One, it's the device itself, is there a design flaw in a device, or, the second factor that you need to consider is patient comorbidity, do the patients have infection which can affect the long-term durability of the tissue valves, or patient management, which is the third factor that you have to consider, the patient's arterial pressure being maintained at low levels, so that the pump does not have to produce high pump chamber pressure.

So, those three factors affect in-vivo reliability.  In the in-vitro test, we only evaluated the device, and if you only consider the device, this is what I think you can expect in terms of reliability.

[Slide.]

We are obviously continuously improving the device.  When you develop a new technology, improvements are necessary, and if I just give, for an example, the pacemakers, the original pacemakers were external and the patients had to carry batteries on a cart.

We all know how that technology has developed through continuous improvements, and the same with the implantable defibrillators.  Well, the same is true here. We started off with a system that incorporated a double lead.  There was one lead coming out of the system for the wire and one lead for the pneumatic operation.

After we gained some experience, we concluded that one penetration would be better than two, so we developed a coaxial exit line, and we exited through the lower left quadrant of the patient.

With more experience we gained knowledge that this was not the most optimum site in terms of exit sites, so we changed the exit site and we reverted to an exiting at the right upper quadrant of the patient.  This helped in reducing exit site infections.

As I spoke before about the SNAP VE, we changed the outflow bend relief, we changed the locking screw ring, and we improved the system controller battery module.  These are the result of experience that was gained in the field, and we haven't stopped there.

We have looked at all of the malfunctions that we have experienced in the clinical program, the REMATCH with the program in the bridge to transplant in the commercial application of these.

We have addressed all of the malfunctions that we experienced, and we made modifications to either eliminate those malfunctions or to greatly reduce the possibility of those occurring.

This is in the XVE configuration, which is now being used in the bridge to transplant field.  We have extended the leads, we have made them smaller, more supple, so that the exit site would heal better.

We modified the way that we hold the diaphragm in place to eliminate the stresses that had caused us fractures in the clinical use.  We have changed the leads on how we attach them to the controller.

We changed the materials at a wire to prevent the wire breakage, we use stronger, more durable materials throughout the system, so we are continuously improving the technology to try to make it better and to have these systems go longer and longer.

I think if you look at the evolution of this technology and where we have come from and where we are going, I think it is quite impressive.

[Slide.]

Our first trial that we were involved with was in 1975, even before FDA had any involvement with devices.  Our average duration in that trial was four days.  Our longest patient support was 41 days.

The second trial we did, the HeartMate pneumatic device, our average duration was 69 days.  Our longest duration was 344.  The first trial on our electric system, 113 days.

What we tried to do is develop technology that we could improve and get longer and longer durations.  If you look at what our experience is in terms of longevity, we went from four days to 69 days, to 113 days, to 276 days, and on the REMATCH program, our average duration, 344 days, and a maximum duration of support now past 3.1 years, and that patient is still ongoing.

We have made significant progress.  This represents 3,100 patients, 3,100 implants worldwide, a vast amount of experience and a consistent improvement in performance and in reliability.

[Slide.]

So, in conclusion, clearly, the VE LVAS is clinically proven technology for bridge to transplant.  Worldwide, the VE LVAS experience provides strong platform for expanded indication, and Thoratec is dedicated to circulatory support and heart failure patients, and they are committed to continuous improvements of this device.

If you look at the technology today, you might say yes, we have had malfunctions, yes, there have been adverse events, yes, there have been complications, but even considering all of those things, if you look at how has the patient done with the device compared to a control group, the patient with the device has done significantly better even with all those complications.

Can we make it better?  Of course, we can make it better.  We know how to design these, we know how to improve them, and we will do that.  We can make these devices better.  We can learn better on how to manage patients, we can learn how to pick patients more appropriately for this patient category.

So, for going forward, I think there is a lot of up-side potential with this technology.

With that, I would like to turn the podium over to Dr. Eric Rose.

REMATCH Trial Clinical Results

Eric Rose, M.D.

DR. ROSE:  I am Eric Rose.  I am the Milstein Professor and Chairman of the Department of Surgery at Columbia University, and the principal investigator for the REMATCH trial.  I have no financial interest in Thoratec, and my presence here today is supported financially by the company for being here.

[Slide.]

To get to the punch line first, the REMATCH trial makes the following critical clinical findings.  There is a clear survival benefit, which we strongly believe, as well, is clinically meaningful and quality of life in device patients is at least equal to, if not better, than medically managed patients in the control group.

With regard to safety, the incidence of adverse events in context of the mortality reduction and the quality of life improvement trends provides very reasonable assurance of the safety of the device.

In conclusion, we believe that the VE LVAS is a scientifically validated alternative, in fact, now the only scientifically validated alternative therapy for end-stage heart failure patients who are not candidates for cardiac transplantation.

[Slide.]

What I would like to walk through in my talk today is the history of the trial itself, the trial's design and how it was administered.

Lynne Warner Stevenson, a trial investigator and leader of the medical management group of the trial will summarize the description of the patient population and the nature of medical management, which is a critical limb in this trial.

I will then come back and review the effectiveness results, the safety results, and summarize our views.

[Slide.]

Anyone who works in the field of management to patients with end-stage heart disease, and certainly the patients themselves, recognize the enormous gap in our treatments available for those with end-stage disease.

Heart transplantation has shown enormous promise, but it is only benefitting approximately 2,000 patients per year compared to the many more who could potentially benefit.

We have had enormous experiences Vic just described with more than 3,000 patients worldwide using the device that we tested in REMATCH for bridging to transplantation.

Based on that experience, we did a pilot trial we called PREMATCH in 1996 to 1998, in which we randomized 10 patients to a control group and 11 patients to LVAD, and found that randomization was feasible.

We found that doubling in one-year survival from 20 to 40 percent in the device patients, which was not statistically significant, yet, these observations generated the necessary clinical equipoise to support a full-blown randomized REMATCH trial, and REMATCH trial enrollment began in May of 1998.

[Slide.]

The trial itself is the product of a cooperative agreement between Thoratec, Columbia University, and the National Heart Lung, and Blood Institute.  The trial was coordinated by an independent coordinating center, the International Center for Health Outcomes and Innovation Research at Columbia led by Annetine Gelijns and Alan Moskowitz.

It was a multicenter, randomized, controlled trial.  Patients and physicians were not blinded to treatment in this obvious major surgical intervention.  There was a prospective plan for interim analyses to be done by an independent Data Safety Monitoring Board.

The analysis was done by an intent-to-treat, and the primary statistical analysis employed the Kaplan-Meier actuarial survival method and using a log-rank test to assess differences between the entire area under the Kaplan- Meier curve.  Paul Meier himself was one of our statistical advisers.

[Slide.]

In order to control bias, first and foremost, this was a randomized trial.  Again, we used an independent coordinating center independent of the sponsor.  The sponsor Thoratec was blinded to control data.

Investigators and the InCHOIR staff except for the statisticians were blinded to the overall data until the trial ended.  We used credentialed investigators, both cardiologists and surgeons, at the trial sites.  An independent gatekeeper reviewed each patient's eligibility prior to entry into the trial to ensure that patients did indeed meet the stringent entry criteria.

We used an independent Data Safety Monitoring Board appointed by the NHLBI, as well as an independent Morbidity and Mortality Committee to adjudicate the etiology of adverse events and causes of death, and there were Medical and Surgical Management Committees that actively met during the course of the trial to ensure that patients got optimal care in both limbs.

[Slide.]

The key study objectives were to evaluate the efficacy of the HeartMate on survival of patients with end-stage heart disease who were ineligible for cardiac transplantation, and with regard to safety, to document and analyze adverse events and the incidence of device malfunction and failure, of course, in the context of the disease population we studied.

[Slide.]

The secondary endpoints of the trial included quality of life, the patient's functional status, the patient days in and out of the hospital, cardiovascular mortality, and cost.  We will not address cost in this discussion today.

[Slide.]

The key assumptions guiding this trial was that patients and clinicians would not adopt use of a device like this unless all-cause mortality over a two-year observation period was reduced by a third or more.

Conversely, I think there was generally strong feeling that demonstration of this kind of benefit would indeed lead to adoption when indicated for this type of technology.

We also assumed that quality of life with the VAD should equal or exceed the OMM group in order to essentially allay the concern that a device like this could literally prolong death instead of prolonging and enhancing life.

[Slide.]

The trial was powered to survival and not to secondary endpoints.  We also hypothesized that survival over time would decay roughly exponentially and that the hazard ratio between the two groups of LVAS to OMM would be approximately 0.56.

Using these assumptions, the endpoint is the number of deaths in this trial, not the number of patients enrolled, and we calculated that we needed to see 92 deaths in order to have an 80 percent power in a log-rank test to find the hypothesized benefit.

We estimated that we would need up to 140 patients in order to prove that benefit, but again the 92 deaths was the decided prespecified endpoint to this trial.  I want to emphasize this trial was not stopped early.  This trial was stopped when we reached the predefined endpoint for the study.

[Slide.]

The patients were randomized between the two limbs in a 1 to 1 ratio.  The randomization was stratified by center, and this was put in place by blocking on a center-specific basis, and the block sizes were also randomly selected so that investigators could not gain the assignment to patients.

[Slide.]

REMATCH study sites included 20 highly experienced centers in cardiac transplantation and medical management to patients with end-stage heart disease.

I am going to turn the podium at this point over to Lynne Stevenson, who ran our Medical Management Committee in order to describe the patient population and medical management.

REMATCH Patient Population

Lynne Warner Stevenson, M.D.

DR. STEVENSON:  Thank you very much.  I have no financial interests in Thoratec.  They are reimbursing me for the cost of travel to this meeting.

[Slide.]

To see where the REMATCH population fits in, let's look at the big picture of heart failure.  It is estimated that there are 4 to 5 million patients in the United States with heart failure.  Of those, about two-thirds have heart failure with low ejection fraction or systolic dysfunction.  The other third have heart failure with a preserved ejection fraction which dominates in the elderly.

Some patients over 80 also have heart failure with low ejection fraction.  In general, they have such comorbidities that they would be considered appropriate candidates for cardiac replacement therapy, so we will focus on this group here in the turquoise of patients under 80 with ejection fractions that are reduced.

It is estimated that approximately one-quarter of those patients have Class III to IV heart failure with symptoms that limit their daily lives.  Multiple models have demonstrated that approximately 50 to 100,000 of these patients might be considered as candidates for cardiac replacement therapy, but probably a smaller number of these are truly appropriate recipients.

[Slide.]

The REMATCH trial aimed to identify a population of those patients for whom primary cardiac replacement could be considered.  The initial eligibility criteria were New York Heart Association Class IV symptoms for 90 days on standard therapy, left ventricular ejection fraction less than or equal to 25 percent, peak VO2 less than 12 ml or inotrope dependent.  These criteria were relaxed in late 1999, but, in fact, 123 patients enrolled met the original criteria rather than the revised one shown in blue.

The mortality estimates at the time this trial was designed were that there would be a two-year mortality of approximately 75 percent.  This was based primarily on data that we had for potential cardiac transplant candidates, but since we additionally required that these patients be ineligible for cardiac transplantation, that, in fact, was an additional contributor to their mortality.

[Slide.]

The reasons that patients were not transplant candidates are shown here.  It was predominantly for age and for diabetes.

[Slide.]

Let's now look at where these patients fall on the spectrum of escalating therapies for heart failure.  Standard therapies for heart failure are angiotensin converting enzyme inhibitors, beta blockers is tolerated, digoxin, diuretics for fluid retention, and spironolactone when renal function is acceptable.

We had anticipated that patients enrolled in REMATCH would fall somewhere in here, such that multiple additional interventions could be made in expert hands.  In fact, the patients referred and ultimately randomized in REMATCH more often fell approximately here.

In many of these patients, beta blockers would not even be considered due to hemodynamic decompensation.  Fluctuating renal function frequently made spironolactone not possible, and many of these patients had even become intolerant to ACE inhibitor therapy.

Inotropic therapy was frequently employed in these patients with the hope that a brief infusion might turn things around.  Often that does not happen, and one is left with the difficult problem of trying to wean therapy to allow hospital discharge, which in some cases was not possible.

[Slide.]

If we look now at the REMATCH therapy at baseline in the OMM arm, off of medical management, 75 percent of patients were on IV inotropes at the time of randomization, 53 percent were on ACE inhibitors, and another 18 percent were on A2 receptor antagonists.

[Slide.]

If we look at ACE intolerance in advancing heart failure, this again defines where we are with this population.  This is now looking at a number of trials of patients unable to take ACE inhibitor therapy.  You can see in the two trial here, this is the ESCAPE trial and the OPTIME trial in which patients are hospitalized with heart failure at the time of treatment, that rates approximately 20 to 25 percent of being unable to tolerate ACE inhibitors. In this trial, in fact, it's 29 percent when you consider the patients who were able to tolerate A2 receptor blockers.

[Slide.]

Let's look now at other REMATCH patient baseline characteristics.  You can see the age is 67, which is a higher average age than most heart failure populations studied.  The cardiac index is 2 liters per minute per meter squared, and you need to remember that 75 percent of patients were on inotropes at the time that this very low cardiac index was measured.

The Minnesota Living with Heart Failure score of quality of life shows a higher degree of disability than in any trial previously reported.

[Slide.]

If we look now at renal dysfunction, we are just coming now to realize how important a predictor this is for outcome in heart failure.  Every tenth of a milligram predicts a higher mortality in heart failure.  REMATCH, with an average baseline creatinine of 1.8 is higher than any trial previously reported in heart failure.

[Slide.]

Let's now line up the profile of REMATCH patients and compare it to other trials of heart failure that is considered to be severe.  The closest population we can come up with is in the first trial.  This is a trial in which IV prostaglandin therapy was administered chronically through an in-dwelling catheter.

If we look at other trials, CONSENSUS from 1987 is shown here.  This is the initial trial of ACE inhibitors in Class IV heart failure.  If we look at the ejection fraction, it wasn't routinely measured at that time, but if we move across in trials of what are considered to be severe heart failure, you can see that REMATCH has the lowest ejection fraction of any of these trials.

Serum sodium remains a very robust predictor of outcome in heart failure, and you can see that the serum sodium is lower in REMATCH than in any other patient population studied.

Perhaps most telling is the systolic blood pressure, which was 103 in the REMATCH population, much lower than any other trials.  I would point out in COPERNICUS, a trial of beta blockers in severe heart failure, the average baseline blood pressure was instead 123, which is 22 mm higher.

[Slide.]

In terms of the medical management of the OMM population, this was the result of very intensive review and approval by the Cardiology Committee of all the investigators with additional input from Dr. Gary Francis outside the trial.

The primary goal was survival without suffering.  Specific goals were to maintain perfusion for organ function, to relieve congestion, to treat exacerbating factors, and then to maintain stability with a regimen of neurohormonal antagonists which would be considered standard therapy.

In fact, however, most of these patients were beyond what we considered to be standard medical therapy.  Inotrope infusion, the net use decreased by the first month, but had increased again by the last follow up.

In terms of the ACE inhibitor use, the use increased, but, in fact, by the last follow up was lower than at the time of initiation because the patients continued to deteriorate.

[Slide.]

We specifically thought long and hard about what we should say about the use of IV inotropic agents.  It was recognized that we should make every attempt to try to discontinue these agents and use any other regimen if possible.

It was agree this therapy would not be used unless everything else had been tried.  In fact, there was considerable effort to try to modulate the inotrope use.  During the initial hospitalization, 20 percent of patients who started out on inotropic infusions came off of the infusions, another 10 percent had the number of drugs infused decreased, so they were on fewer inotropic agents, however, 40 percent of the patients who were not on inotropic agents at the randomization were found to need inotropic agents.

It was well recognized by all the investigators that inotropic therapy is only palliative in this population, that it does not improve survival, and, in fact, might even worsen mortality.  The problem was that current medical therapy had reached and end, and there just were not other options.

[Slide.]

Understanding the profile of the REMATCH population now allows us to look at where the mortality of this group fits in.  With every factor being worse than in the previous trials, it is not at all surprising that there was a very high one-year mortality of 76 percent in the REMATCH population.  They had moved beyond what could be offered by the medical therapy.

[Slide.]

So, in summary, REMATCH patients define a new profile of severe heart failure, different than what had been considered severe in previous trials.  By the time of randomization, REMATCH patients had already received optimal management in the most cases and had moved beyond current medical therapy.

[Slide.]

How are we to determine what would be a meaningful benefit in this population?  In trying to decide, this committee accepts a major challenge.  We don't really have previous precedence that should guide us in how to decide this.

Let's review what we do know of therapies that are accepted as beneficial in heart failure.  If we take ACE inhibitors, the cornerstone of our therapy of heart failure at the SOLVD trial is considered to be the landmark trial for this, if we look at the one-year mortality, it was decreased from 14 to 11 percent.

It has become very common in drug trials to look at percent, of a percent improvement, which in fact magnifies a benefit that is relatively modest.  This impacts 3 patients per 100 treated in the first year.

A similar relative benefit was observed in the CONSENSUS trial, an ACE inhibitor in Class IV failure, but because the one-year control mortality was higher, there is an absolute benefit that is greater in terms of the number of patients treated, which here is 17.

Dosing of ACE inhibitors has been felt to be important in optimal management of heart failure.  That is based primarily on the ATLAS trial, shown here, with the one-year mortality, a relative small difference between a high and very dose ACE inhibitor, which in fact led to a survival benefit of less than 1 patient per 100 in the year treated.

The COPERNICUS trial, demonstrating unequivocally the benefit of beta blockers in severe heart failure, had a mortality in the control arm that was less than 20 percent, and the RALES trial, showing a similar benefit with a slightly higher baseline mortality.

Accepting these benefits for therapy, how should we now evaluate a population in which the one-year mortality on controlled therapy is 76 percent, do we look at relative benefit, do we look at absolute benefit, what will our patients look at?  We are stepping into a new horizon.

I would like to turn it back to Eric Rose, who will further discuss the results of the trial.

REMATCH Trial Clinical Results

Eric Rose, M.D.

[Slide.]

DR. ROSE:  Almost 1,000 patients were screened for REMATCH, of whom 128 were randomized, 67 to the device limb and 61 to medical management.

[Slide.]

The baseline characteristics here in sample for critical ones, age, ejection fraction, cardiac index, serum creatinine, IV inotropes, Minnesota Living with Heart Failure score, across virtually all baseline characteristics.  There were no differences between the two groups, illustrating the success of the randomization.

[Slide.]

This is a Kaplan-Meier plot illustrating survival of the VAD group compared to the OMM group at the time that the study was stopped in June of 2001.  The difference, the area under the curve, under the LVAD curve, between the two curves, shows a 46 percent reduction in mortality hazard over the two-year observation period, again, the primary endpoint of the trial, the p-value being 0.003.

In addition, at this time of analysis, there was a 52 percent, one-year survival in the device group compared to 26 in the control group, and 23 percent at two years compared to 8 percent in the control group.  The one-year difference was statistically significant, the p-value at two years was 0.09.

[Slide.]

In terms of causes of death, the overwhelmingly most common cause of death in the OMM group was heart failure.  In contrast, in the device group, the most common cause of death was sepsis related typically to the device, following by a range of mortality related to either failure or cessation of function of portions of the device.

Two devices failed completely to support the circulation, while three had partial failures of the device not meeting the formal definition, but did require device replacement ultimately leading to death.

In addition, fortunately, there was only one death due to bleeding.  The next most common cause of death was cerebrovascular disease.

[Slide.]

To summarize these effectiveness results, the one-year survival doubled.  There was an absolute reduction of mortality of 27 percent at one year meaning that for each 100 patients treated, we would estimate that 27 deaths would be averted during that first year.

The two-year survival was tripled, and the median survival time increased to 408 days for LVAS patients compared to 150 days for OMM patients.

[Slide.]

As we said, the all-cause mortality was reduced by 46 percent exceeding the primary objective of a 33 percent reduction.

[Slide.]

Now, with regard to safety, looking at serious adverse events, these were more common in the device patients, in fact, they were 2.3 times more common, a high significant difference.

Particular types of adverse events, namely neurologic dysfunction and bleeding, not surprisingly were also more common in the device patients, but interestingly, complications like infection, sepsis, and thromboembolic events, these were not statistically significantly different between the two groups.

[Slide.]

Also, of interest is the fact that adverse events are generally concentrated early in the postoperative course of these patients, and decay off to a lower incidence over time.  It does not reach zero, but it does decrease over time.

[Slide.]

Looking in particular at neurologic events, there were 28 neurologic events in 67 VAD patients, a total of 40 events in those 28 patients, but very importantly, only a quarter of those events were permanent or disabling.  The overwhelming majority of these events were toxic encephalopathy typically early in the intensive care unit after the operation, or transient ischemic attacks.

[Slide.]

Looking at the sum total of serious neurologic events, there were 5 ischemic strokes, one intracranial hemorrhage, and 2 patients who had serious morbidity from air embolism.

[Slide.]

Also looking at this particular adverse event, we see that the majority of events or the highest incidence occurs early in the postoperative course and falls off over time.

[Slide.]

In terms of bleeding, the majority of events were associated with VAD implant or reimplantation, which is quite similar to the bridge experience.

Infection was a specific complication of VAD use.  There was initially in this trial an unappreciated association of infection with malnutrition.  Patients who were nutritionally depleted generally could not heal the drive line sites well, and investigators only began to pay specific attention to this late in the course of the trial.

Infection guidelines were developed in the midpoint or in the earlier course of the trial with the adoption of specific guidelines that investigators adopted on the recommendation of the Surgical Management Committee.

[Slide.]

There were 156 malfunctions of the device reported, 70 were confirmed by actual analysis of the device component, 50 of these were external components, 20 internal components.  The most common malfunction was the controller, malfunction with broken lead wires in 27 percent, 14 percent had the serious complication of inflow valve incompetence seen in 14 percent of patients in this trial, and 9 percent of patients had a broken Y-connector.

The top 1 and 3 of these were external device malfunctions, which allowed their replacement.

[Slide.]

Median time spent in and out of the hospital was dramatically different between the two groups.  As I said before, median survival of 408 days compared to 150 at the time of the 92nd death.

VAD patients enjoyed almost a full year of days out of the hospital during a two-year period of observation, where, control patients enjoyed approximately 100 days out of hospital.

VAD patients also spent more time in the hospital with readmissions for device complications or other complications, higher than controls, and index hospitalization days, as would be expected for the major procedure of VAD replacement, the length of hospitalization as 29 days for the VAD patients, which was virtually identical to the length of stay that we typically see for bride to transplant patients.

[Slide.]

With regard to quality of life, again, our hypothesis was that quality of life with the VAD should equal or exceed the OMM group.  Toward that end, we carefully chose the instruments that we used.

In particular, we used the SF-36 Health Survey, a general health measure, and prespecified two very specific domains - physical functioning and role/emotional scores.  These were prespecified before we even began the trial that these were the particular SF-36 portions we were going to look at.

The Minnesota Living with Heart Failure scale is a disease-specific quality of life measure for patients with heart failure.

The New York Heart Association functional class is probably the most broadly used functional status estimator for heart failure in the world.

The Beck Depression Inventory is a widely employed scale for measuring clinical depression, and the EuroQOL scale for patient preferences.

We imputed no values for quality of life for dead patients.  Had we done so, it would have created an even greater benefit documented by the device patients, so we specifically did not input these values in order not to further bias the data.

[Slide.]

With regard to role/emotional scores, in this slide and the succeeding slide, the VAD patients are in yellow, OMM in purple.  A higher score here is better, reflects improved quality of life.  We have very little data after one year, but we see that by 12 months, there is a highly statistically significant difference between VAD and OMM patients with regard to role/emotional functioning.

[Slide.]

Similarly, with regard to physical functioning, at 3, 6, and 12 months, again, a higher number is better quality of life, and what this means in terms of the questionnaires is can you walk a block, can you walk up a flight of stairs.  There were significantly more positive answers to those questions in the VAD patients compared to the OMM patients, and at 3, 6, and 12 months, this P was less than 0.05.

[Slide.]

With the Minnesota Living with Heart Failure scale, we did see at one time point significant improvement, and at 3, 6, and 12 months, a trend to improvement for the VAD patients.  On this scale, improved quality of life is a lower score.  The 75 at baseline, to our knowledge, is the highest MLHF score in a heart failure population subjected to a randomized trial.

[Slide.]

In terms of New York Heart Association functional class, the percentage of patients in Class IV at the time of observation is plotted here.  There were only three patients alive at 24 months in the OMM group, one of whom had improved out of Class IV.  Otherwise, at virtually all other times, the Class IV representation in the control group was almost 100 percent compared to a marked decrement, statistically significant decrement in the VAD group.

[Slide.]

In the Beck Depression Inventory, both groups at the time of enrollment met the definition of clinical depression with scores of 17 or higher.  There was a highly significant difference between the VAD and OMM groups at 1 and six months, and the p-value was 0.055 at 12 months between these two groups, suggesting that these patients were indeed less depressed.

[Slide.]

To summarize these Quality of Life findings, VAD scores were never worse than the OMM group.  I think the argument that these devices prolong death is simply not tenable based on this data set.  They prolong meaningful life.

The only negative Quality of Life impact was the issue of short-term, postoperative pain, which is an obvious result of the operation that the patients had.

The LVAD general Quality of Life was better than OMM at 12 months in the key, prespecified SF-36 domains, a physical functioning and in role/emotional scores.

The LVAD disease-specific Quality of Life measurement with the Minnesota scale was improved, but not statistically significant at a year.

LVAD functional status measured by the New York Association classification was significantly better, and VAD patients had reduced depressive symptoms to normal range, below that of clinical depression, a phenomenon which was not seen in the OMM groups.

[Slide.]

To put this in context, the LVAD physical function scores, while improved, were not normal, but these scores are analogous to patients who were receiving long-term hemodialysis, of which there are 300,000 now in the United States, and moderate heart failure patients.

Similarly, the role/emotional scores are better than those reported for clinical depression and similar to those with moderate heart failure.

[Slide.]

The most recent Kaplan-Meier plot is depicted in this slide with very little difference essentially in the shapes of these curves.  The area between them is now significant at the 0.001 level, and the difference at two years, the p-value here is now 0.054.

[Slide.]

In terms of efficacy, REMATCH has shown that the VE LVAS exceeded the primary objective of the trial by demonstrating a reduction in all-cause mortality of 46 percent in patients who were not candidates for cardiac transplantation.

With regard to safety, the incidence of adverse events associated with implantation is higher than OMM patients, but the incidence of overall adverse events we believe is clearly acceptable when compared to the natural history of this terminal illness.

In addition, even in this small early experience, multiple opportunities for improvement have been identified.

[Slide.]

To return to Lynne's question as what is clinically meaningful benefit, we did see an absolute reduction in one year mortality of 27 percent in this patient population, which in terms of simple magnitude is obviously more than comparable, more than favorably comparable with those that we have seen for other heart failure therapies that are now considered standard.

[Slide.]

Now, we also would like to address the seven questions which were posed to the panel by the FDA to offer our views of these very important issues that were raised by the FDA examiners.

[Slide.]

First, there was the question of whether or not there now are good end of pump life indicators.  Before I speak to this about the VADs, I think it is important to recognize that it is not just the mechanical circulatory support that has the issue of end of pump life indicators. Heart failure itself, we have difficult issues with regard to end of life indicators, whether or not serum sodium or creatinine or inotrope dependency themselves are absolute indicators of end of life is something that is under intense investigation and deserves further, but even in the context of this trial, there are patient indicators that can suggest that a person is reaching end of life of the device.

For example, increased alarm frequency and changing pump sounds are just simple clinical indicators, and there are device indicators.  This device's motor current waveforms can be measured on an ongoing basis and the waveforms analyzed to indicate earlier forms of wear in the device, and in addition, particulate matter can be analyzed in the filters from the air vent side of the device.

[Slide.]

In terms of device reliability, are we seeing enough reliability with this device, so that is appropriate for destination therapy?  In the context of the available alternative therapies and the terminal illness, the observed failure rates in the REMATCH trial essentially define what is acceptable at this time.

The reliability, even though flawed, is sufficient to produce a very measurable survival benefit.  Furthermore, we see multiple opportunities that address device improvements, patient selection, and patient management.

[Slide.]

The other important question, is the survival benefit that we documented clinically meaningful?  Now, this was the guiding question for the REMATCH trial.  Before we embarked on this, we tried to reach a consensus among many groups as to what would be a clinically meaningful indicator and what could be done practically.

There was agreement between the REMATCH investigators at 20 sites, the NHLBI, the FDA, and Thoratec that the endpoint that we defined for mortality was indeed a clinically meaningful benefit if we could document it.

We documented that as a 33 percent reduction in mortality during two years with Quality of Life that was greater than or equal to that of OMM.

Indeed, we found that median survival more than doubled in this two-year observational trial.

[Slide.]

The other question raised, is the effectiveness of the system on functional status clinically meaningful, in particular because the data with regard to a test like six- minute walk and PPO2 were confusing.

The trends with regard to all of the Quality of Life indicators that were used, that are subjective, admittedly subjective, are consistent favorable.  With regard to peak VO2 and 6-minute walk, there are not benchmarks available that have been validated in order to look at those and compare them to patient's assessment of Quality of Life or survival.

The New York Heart Association class, while flawed, is still the most widely used classification of functional status of heart failure patients in the world, and the prespecified physical function domains and role/emotional scores, again admittedly subjective, we are asking the patients how did you do or how do you feel were highly significant at 12 months.

To quote the President's Commission for the Study of Ethical Problems in Medicine and Biomedical and Behavioral Search, "Quality of life is an ethically essential concept that focuses on the good of the individual.  What kind of life is possible given the person's condition and whether that condition will allow the individual to have a life that he or she views as worth living, not whether some dispassionate third party views the life as worth living," and to make the judgment that on the basis of a 6-minute walk or a peak VO2, a life is not worth living, I think is more than a stretch.

[Slide.]

In terms of destination therapy - do the benefits outweigh the risks?

The adverse events in this trial were not calculated independent of survival and Quality of Life. Their impact is calculated into the survival and Quality of Life analyses, and even in the presence of those adverse events, it is clear that the risks outweigh the benefits.

[Slide.]

With regard to the labeling, is it adequate?  We feel that reasonable labeling has been proposed, but certainly I think anyone reasonable will view that recommendations for improvement would be well accepted.

[Slide.]

In terms of postmarket evaluation, I do believe there is a widely held view in this field that this is essential going forth and that Thoratec shares in that belief.

I thank you for the opportunity and am proud, on the part of the REMATCH investigators, to present this work, and turn it back to Don Middlebrook.

Summary and Closing Remarks

Donald A. Middlebrook

MR. MIDDLEBROOK:  Thank you, Lynne, Eric, and Victor for an excellent presentation.

[Slide.]

I would like to summarize and conclude with the following remarks.  It is our belief that this study has scientifically validated the safety and effectiveness of the HeartMate VE LVAS for the proposed indication.

Every aspect of this study has been thoroughly planned and carefully executed.  The REMATCH study has demonstrated strong scientific evidence of clinically meaningful survival benefit.

If you take a look at the experience we have in the previous clinical trials in our commercial use and in this REMATCH study, the VE LVAS is a well characterized and a proven technology.

It is our belief also that the REMATCH trial has demonstrated reasonable evidence for safety particularly in the context of terminal illness.  As Eric said, even in the fact that the REMATCH patients faced a higher opportunity for an adverse event, all of the Quality of Life instruments showed sustained improvement trends over the optimal medical management group.

The device, as Dr. Lynne Warner Stevenson alluded to, provided unprecedented reduction in mortality in end-stage congestive heart failure patients when compared to the landmark drug studies.

The VE LVAS is now the only proven alternative therapy for non-transplantable end-stage congestive heart failure patients.

[Slide.]

The threshold for PMA approval is reasonable assurance of safety and effectiveness.  If you consider all of the data we have presented here this morning, and contained in the PMA in relation to the patient population, the conditions of use, the probable benefit versus the probable injury, and the demonstrated reliability, we strongly believe that there is clear and compelling scientific evidence that the HeartMate VE LVAS should be approved as a new treatment option for end-stage heart failure patients ineligible for cardiac transplantation.

Thank you very much for your time and attention.  This concludes our presentation.

DR. LASKEY:  Thank you all for a most informed and articulate presentation.

I would like to move now to the FDA presentation starting with Dr. Berman.

FDA Presentation

Michael Berman, Ph.D.

DR. BERMAN:  Good morning.  My name is Michael Berman.  I am the FDA lead reviewer for this PMA supplement. For the record, this is Supplement 16 to PMA P920014.  It is being brought by the Thoratec Corporation for the their Thoratec HeartMate VE LVAS System.

I will be speaking to the engineering review for this PMA supplement.  I will be followed by Dr. Swain, who will address the clinical review, and then by Dr. Gray, who will address the FDA statistical review.

After that, I will read the questions to the panel into the record.

[Slide.]

These are the folks on the FDA review team for this supplement.

[Slide.]

In this supplement, the sponsor is proposing an expanded indication for use for the currently approved VE LVAS system.  The current approval is as a bridge to transplantation in patients who are transplant-eligible.  This is the language the sponsor is proposing.

"The HeartMate VE LVAS is indicated for use as a bridge to transplantation in cardiac transplant candidates at risk of imminent death from non-reversible left ventricular failure."  That is the approved language.

This is what the sponsor wishes to add.  "The HeartMate VE LVAS is also indicated for use in patients with end-stage left ventricular failure who are ineligible for cardiac transplant."  And for both indications, "The HeartMate VE LVAS is intended for use both inside and outside the hospital."

[Slide.]

This is what the FDA has to do during this review. We have to determine if the sponsor has provided a reasonable assurance of safety and effectiveness.  That is the law.  And we are obligated to consider the following factors:  the patient population, in whom will this device be used, the conditions of use, how will it be used, what is the probable benefit versus the probable injury to the patients using the device, and finally, what is the reliability of the device when used as indicated.

[Slide.]

I would like to remind you of the device description.  The device basically has implanted and external components, the implanted components being the blood pump, the valved conduits, and part of the percutaneous tube.

The external components are the controller, the battery packs, and what we will call accessories.  The accessories are described in your panel pack in Tab 3.2, pages 5 and 6.  It is just external components.

[Slide.]

This is a figure which the sponsor allowed me to use.  This is their figure.  Just to remind you, this is the blood pump.  It sits in the chest.  It is a rigid titanium shell.  It is divided into two halves internally.  One half contains the electric motor and is connected by means of a vent tube to the outside.

The other half is the blood side.  Blood comes from the apex of the left ventricle through the valve conduit into the pump.  The electric motor rotates.  There is a cam that converts that rotation to linear motion, pushes a pusher plate, pushes the blood out of the blood side into the aorta.  This also is a valve conduit.  So, this is all implanted.

The percutaneous tube carries an air line and an electric line out.  This is the skin breakthrough site.  There is a connector here.  There is a vent line which allows air to vent in and out of the motor side of the pump. There is an electric line, which connects to the controller, which has alarms, which conditions power, and so on.

There are two battery packs, only one of which is shown.  The system can also be powered from a bedside console.  The system also comes with a battery charger with a system monitor, and so on.

[Slide.]

This is a very complex system, as you might well imagine.  During the review of such a complex system, these are the things that the FDA will consider.  In this case, the sponsor is asking for an expanded indication for use for an existing system which is currently approved for a bridge to transplant, so all of the things shown here have been reviewed in detail, and we have no concerns.  Nothing will change for these things for the proposed expanded indication for use, the longer term use.  We have already reviewed this, and we have no concerns.

[Slide.]

These are the things that we do have concerns with:  the difference between the proposed long-term use, the destination therapy versus bridge.  The thing that we looked at very hard was device system reliability, in particular the reliability of the internal components since they would require that the patient have surgery should anything go wrong with them.  We are less concerned about the externals.  It has been demonstrated quite readily that they can be changed out rather easily and aren't a problem.

In particular, we are concerned with the motor, this is what pumps the blood, and the valved conduits as Mr. Poirier spoke to and I will speak to again, and we are concerned about a Device End of Life Indicator.  The sponsor noted that there are several things which could indicate end of pump life, but there are no objective indicators or at least nothing the sponsor has proposed by which you can tell that this particular pump is at end of life and should be replaced and that particular pump has six months left.

[Slide.]

I would like to address the reliability testing, the bench testing that was performed by the sponsor, which was described briefly by Mr. Poirier.

Fifteen units were put on test.  All of them were VE LVAS, none of them were VE SNAP.  They were put on test on a mock circulatory loop.  The implanted components of the system were in water at 37 degrees Centigrade to simulate the environment they would see in the patient, the temperature.  They were pumping water.

The external components were in air at room temperature as they would be in clinical use, and the pumps were run at what is described as worst, average, and minimum operating conditions, and they were cycled from condition to condition to condition every week, and so cycled around and around throughout the course of the test.

The things that were changed were beat rate, outlet pressure, and flow that the pumps were generating. This is the way it is done.  This is the way this reliability testing is done.  The number of units is good, the conditions chosen were reasonable or apparently reasonable.  This testing was conducted in a most appropriate way.

[Slide.]

This is the results of the testing, which, by the way, is still ongoing.  As of this past summer--this testing began in 1997--as of this past summer, 10 pumps had failed, 8 of those failures were main bearing failures, 1 was a diaphragm failure, 1 was a commutator failure.  The commutator is part of the motor.  Five units, as of last summer remained on test.  That is, they still continue to function.

[Slide.]

Looking at the main bearing failures, this is the motor.  If the bearing fails, if the motor fails, you don't have electric pumping.  The mean run time before the observed failure was about 136 million cycles at 75 beats a minute, that comes to about three and a half years.  That is roughly 40 million cycles a year at 75 beats a minute.  The standard deviation for that was about half a year.

The soonest a bearing failed was at about 2.25 years, the longest one ran before failing other than the ones still running was about 4 years.  The median is about 3.5 years.  What I would like you to look at, look at the mean and look at the median.  It is out at around 3.5 years.

The sponsor has initiated a corrective and preventive action to address the problem of the bearing failures.  This is an investigation, there is no mitigation yet.  This is an open investigation, no steps have been taken to address the failures, no specific engineering changes have been made that we are aware of.

[Slide.]

Using these numbers and using a mathematical model, a Ybel [ph] model, one can predict what the reliability of the system should be.  This is a prediction.  86 percent reliability at 2 years with a 60 percent confidence, 76 percent reliability at 2 years with a 90 percent confidence.

The mean time to failure for the pump with 90 confidence should be 3 years.  Keep those numbers in mind, 3 years.

[Slide.]

These are observed end of pump life events.  They occurred at times ranging from 460 to 779 days.  These are events observed in patients in the clinical trial.

The cycles are in million of cycles.  For instance, looking here, this is about 50 million cycles at 75 beats a minute.  At 100 beats a minute, it is about 67 million, and what was observed, there was dust, and when the pump was looked at, it was a bearing failure.

Similarly, for the other ones, the number in parenthesis is either the higher estimated rate or, here, for these two, these are numbers of cycles estimated by the sponsor based on records of what the beat rate was when those patients were seen at follow-up.

The point here is that these occurred early in the process compared to what the bench testing predicted.  The sponsor is saying the bench testing predicts end of life at between 80 and 120 cycles for the pump, and these are occurring early.  The clinical observation is they occur early.

[Slide.]

Another malfunction, not failure, a malfunction, is inflow valve incompetence that Mr. Poirier spoke of.  This has been confirmed, 12 events in 11 patients.  One patient experienced this twice.  Six of those devices were VE, five were VE SNAP, which calls into concern the effectiveness of the 90-degree elbow, which is one of the things that differentiates the VE SNAP from the VE.

The idea was to not allow the outflow valve graft to kink, so that the pump would not develop high pump pressure and thereby high back pressure across the inflow valve.  Roughly half of the events observed were in the modified pump.

We wonder whether this might be related to end of pump life.  One of the symptoms of inflow valve incompetence is an increase in pump rate up to on the order of 100 to 120 beats a minute.  That means for a given absolute length of time, for a day, the pump will undergo more cycles at the higher rate than at 75.

Consequently, the pump will reach its end of life sooner in time.  The clinical consequences of inflow valve incompetence can be surgery for replacement of the inflow conduit, and so that, by definition, puts the patient at some risk.

[Slide.]

The engineering summary is that the bench testing, in our eyes, did not account for all of the observed clinical conditions.  In particular, it did not account for the elevated pump chamber pressure which was seen clinically and/or for the high beat rate which was seen clinically, the elevated pump chamber pressure could well be related to inflow valve incompetence, which resulted in a higher beat rate.

The observed pump end of life events were at the low end of the bench testing reliability prediction.  These things may not last as long as the bench testing predicted, and there is no objective device end of life indicator.

It was pointed out by Mr. Poirier that, yes, you can look at the dust in the vent, and, yes, you can look at the pump current waveforms, which by the way, cannot be done by the patient, it does require some training to understand and to interpret, and you can listen for the pump sounds.

There is no objective device end of life indicator, and replacement of the pump requires major, major surgery.

At this time, I would like to turn this over to Dr. Swain, who will discuss the FDA clinical review.

Julie Swain, M.D.

DR. SWAIN:  My name is Julie Swain.  I am on detail to the FDA for the last several months.  I am a cardiothoracic surgeon and a faculty member at Mass. General Hospital.

[Slide.]

I would like to present our review today.  The clinical review was performed by Dr. Ewing, cardiology, and Dr. Sapirstein and I in cardiothoracic surgery.

[Slide.]

As Dr. Berman has said, the indication is an expanded use for patients with end-stage left ventricular failure who are ineligible for cardiac transplantation.

[Slide.]

The primary effectiveness endpoint was survival benefit, as Dr. Rose has said.

[Slide.]

This was estimated based on mortality at 2 years, and the estimates were based on the literature, as Dr. Stevenson has said, based on a mortality of 75 percent in the OMM group and 50 percent in the LVAS group, calculated for 92 study deaths.

The worst case power calculation for a power at 80 percent was calculated as 60 and 40, so a 30 percent improvement in survival.  When you look at the results of the study, the observed Kaplan-Meier curves, the mortality was 91 and 76 percent at the two-year mark.

[Slide.]

The data that I will be showing you are based on three separate submissions from the sponsor.  The first was in June, and in each slide, I will attempt to have that marked, a June submission.  There was an update on mortality curves in November, and another update in February.

So, the survival will be based on the February update from the sponsor.  We received the serious adverse event update about three working days ago.  It was too late to include in the presentation, but on first pass, it appears to be not substantially different.

[Slide.]

The inclusion criteria used originally for when the study was designed, which was used for 124 of the 129 patients, is as shown, that the patients were ineligible for transplantation, NYHA Class IV for greater than 90 days, 70 percent of these patients turned out to be on inotropes, intensive medical therapy, LVEF lower than 25 percent, VO2 max less than 12.  VO2 max was performed at baseline on about 50 percent of the patients in the study.

[Slide.]

Later, in order to increase enrollment in the study, the criteria were changed slightly, and this was for five of the patients in the study.  It was NYHA Class IV for greater than 60 days as opposed to 90, and added NYHA Class III or IV for greater than 28 days, and inotropic therapy or balloon pump.  Also, increased VO2 criteria to less than 14 rather than 12.

[Slide.]

The exclusion criteria are as shown, in particular, patients smaller than 1.5m2, history of stroke, neurological complications, things of that sort.

[Slide.]

Baseline characteristics were equal between the two groups or nearly equal in the number of patients able to do the 6-minute walk, the distance of the walk, NYHA IV, 65 and 2 in Class III versus 60 and 1, so, in general, this was a study of NYHA Class IV patients.

[Slide.]

This is based on the February 02 data, and the study design, as Dr. Rose has said, 968 patients were screened as possible enrollees in this trial, and patients enrolled, 129.  That gives you a ratio of 7.5 to 1.  You will see slight differences in our slides, 128 versus 129 patients, 67 versus 68.  This reflects one patient who was consented for the trial before the end of the trial, but received his device after the end of the trial, and these are included in the survival data, this one particular patient.

What we see is that 7 patients reached the two-year end of study mark in the left ventricular assist group, 45 patients died at less than two years, and there are, as of February 02 data, 16 patients alive less than two years.

For the medical management group, 3 patients survived to the two-year mark, 5 are alive less than two years, and 53 patients have died.

[Slide.]

All of my slides that have percents are zero to 100 scale, those with number of patients are zero to 70 scales just to keep things consistent.

Looking at the Kaplan-Meier curves for the data that you just saw, one-year survival, n equals 16 for medical management, and medical management in my slides are in gold as the gold standard, and blue for the device, and 33 patients were the basis of the KM curve, survival curve at one year at 57 percent.

At two years, the curves are based on  3 patients and 7 patients respectively in the two groups.  By 12 weeks later, at the end of the trial, this is the February 02 data again, 4 of the 7 device patients had died, and all 3 of the medical management patients had died by 12 weeks past the end of the two-year trial.

[Slide.]

Serious adverse events occurred in 64 of 68 of the device patients, 36 of 61 of the medical patients, P less than 0.001.

[Slide.]

Then, comparing for safety, the use of this device in destination patients versus bridge to therapy patients, there are several problems as the sponsor has pointed out in making the comparisons.  They are certainly different patient populations, different definition for many of the SAE's, the studies were done in different times periods with a different patient care team, which makes comparison of SAE's between the bridge data set and the current data set to somewhat problematic.

[Slide.]

I am going to discuss just a few of the SAE's, and the slides all contain two graphs.  One is the percent of patients experiencing the event, the other is per 100 patient days, as Dr. Rose has shown.  I think it is important to do it both ways.

There are certain complications like the need for transfusion that would probably be more appropriate looking at patient days, events per 100 patient days.  However, other complications, such as a stroke, it might be more important or clinically more important I think to look at the percentage of patients experiencing the event it would tend to overshadow other events occurring, and not be able to detect lesser strokes and lesser neurologic dysfunctions in this case.

Also, for situations like sepsis, a patient who has sepsis and then has sepsis for the remainder of the time in the study would not be reflected in events per 100 days, but probably would be reflected more clinically relevant in percent of patients.

So, when we look at neurologic dysfunction, we find a statistically significant difference between device versus medical therapy, 27 percent of the patients having a serious adverse event in neurologic dysfunction.

It is important to note that this study was designed in the early-to-mid-nineties when we had less of a knowledge of neurologic events than we do now in cardiac surgery, so the definitions of neurologic events are really based on the NIH stroke score and a clinical examination, there are virtually no cognitive function being done other than what is present in the NIH stroke score, which is truly a stroke score, and that is really two questions relating to alertness and ability to function.

These are the results of neurologic dysfunction.

[Slide.]

When you look at local infections, we also look at a P equals 0.07 difference, less of a difference between the two groups, and less than 0.1 when look at it per 100 patient days.

[Slide.]

When you look at sepsis, and you look at the individual patient summaries, you will see very often that local infections may lead to implant infection, sepsis, and that sort, so a lot of these SAE's are not independent at all.  They are very much interrelated.

When you look at sepsis, the difference is 31 percent versus 13 in the two groups, and percent of patients, which is a 0.02 difference significantly.  Likewise, when you look at per 100 days, a P less than 0.2 difference in that particular SAE.

[Slide.]

Now, for complications that only the device patients could have, looking at percutaneous or pocket infections, 24 percent of the patients and a 0.11 event per 100 day level, and this is again based on the June data.

[Slide.]

Looking at pump housing inflow or outflow tract infections, which could be extremely serious, 13 percent of the patients had that complication.

[Slide.]

Look at bleeding.  This is not perioperative bleeding, this is bleeding after the perioperative period, a statistically significant difference between the two groups here with the device patients having 25 percent versus 3 percent incidence of bleeding, and again, a statistical difference when you look at event rate.

[Slide.]

Perioperative bleeding, which obviously can only be present in the device patients, 34 percent of the patients and this event rate.

[Slide.]

When we look at SAE's of all types, and look at operations, and this excludes the original implant operation obviously, it includes all operations in the medical group and every operation after the original implant in the left ventricular assist group, we can see the number of reoperations, the number of patients having 1, 2, and up to 10 reoperations.  This includes all reoperations, everything from a Hickman catheter to replacement of the device itself.

[Slide.]

What about the device malfunction analysis that Dr. Berman has mentioned?  When you look at the various types, implant, as he has explained, which is the most serious, external element replacement, and then device malfunctions, and the definition of malfunction, 20 elements in 19 patients had implant element replacement.  Of the 12 pumps removed, 8 pumps were replaced, 7 of those 8 patients died, 6 of those 7 in the postoperative period.

Of 4 devices removed, but not replaced, all of those patients died, 3 of them were removed for unrelenting sepsis, and 1 patient chose not to have the device replaced.

[Slide.]

When we look at withdrawal from treatment in the two groups, 4 of the medical management patients chose to have the treatment withdrawn within one month of randomization, of choosing to be in the study and being randomized to medical therapy, they chose to have treatment withdrawn.  Eight others chose to have treatment withdrawn at some time later in the study.  So, a total of 12 of the 61 patients in optimum medical management chose to have their treatment withdrawn.

In the left ventricular assist group, 7 patients or their family chose to have the device turned off or did not agree to replacement when replacement was recommended, and 6 more chose to have treatment withdrawn, so a total of 13 of the 68 patients chose to have their treatment on the device withdrawn.

[Slide.]

Now we look at secondary endpoints, which were originally part of the study design, and that includes NYHA class, QoL's, functional status by way of 6-minute walk, VO2 max, number of hospitalizations, and what we have just spoken about, adverse events and device malfunction.

[Slide.]

When you look at NYHA class, as Dr. Rose has shown, there is an excellent result in comparing the device, that is statistically significant at 6 and 12 months, and there was very little data after 12 months.  It is important, as Dr. Rose said, that no values were imputed.  We standardly looked for values that are imputed, and that can either be done by taking the last result in the chart that was done and imputing results, which would give you different results from here, very good, or to look at data that was not collected because the patient died or could not do the test and impute the worst possible result, which would also give you very different results from shown here.

So, no data was imputed, these are all the real data that were collected.

[Slide.]

When you looked at Quality of Life results, the problem in an unblinded study, and necessarily this was unblinded to both physicians and to patients, and what is the effect on both the physicians and the patients of not being selected for the test device or the test therapy.

We are often concerned about the placebo effect in any unblinded study, and this may be very important, up to a 30 percent difference in results.

There may be sample bias for missing data.  A patient who has had a large CVA would not necessarily be expected to be able to complete any of the QoL questionnaires or the walk or VO2 data.

We also expect consistency, statistical consistency between the Quality of Life and functional measures, and clinical concordance, you expect things to move in the same direction with the same relative magnitude.

[Slide.]

What about the data, what amount of data do we have for Quality of Life?  When you look at the questionnaires, we can see these number of patients had completed QoL questionnaires.  At one year, the LVAS patients, 22 of them, and 4 at two years, 6 and 3 in the OMM patients completing the questionnaires, of which we are looking at the QoL data.

[Slide.]

When we look at 6-minute hall walk, the number of tests performed, 37 of the 68 patients in LVAS either never or only once had a 6-minute hall walk, and 29 had multiple performances, two or more performances of this test.

OMM patients, 44, and 17 had multiple performances of the test.  So, there really are inadequate data for a comparison of these groups.

[Slide.]

When you look at the data, this is clear.  Of the patients who did have the test performed, these are the median distances in meters performed.

[Slide.]

Look at peak VO2 data, which is one of the inclusion criteria for the study.  About half of the patients in both groups had baseline performance of data. When we look out to one year and two years, 15 of the LVAS groups and 3 of the OMM group had data, 2 and 1 at two years, and you look at the data obtained, and again no statistical difference in this small sample size.

[Slide.]

Another way of look at hospitalization is the percent of the remainder of life out of the hospital.  According to the June 01 data, when we look at the average and median for the device group, 50 percent of the remainder of their life after randomization was found out of the hospital, median of 64 percent.

On the OMM group, 71 percent of the remainder of their life was spent out of the hospital versus 83 in the median numbers.

[Slide.]

Death during hospitalization, 20 of the LVAS patients and 6 of the OMM patients died during their initial hospitalization.  Of the OMM patients who died, of the 6 who died, several had treatment withdrawn, 1 patient had ventricular fibrillation on the day of randomization and died on that day.  Another patient within a week of randomization was found to have critical aortic stenosis, underwent urgent cardiac valve replacement, and died in the postoperative period.

[Slide.]

Our clinical summary is that in a very advanced heart failure population, the device used produced a survival benefit.  The mortality and morbidity associated with use of the LVAS was considerable.

Interpretation of functional testing data is limited by the amount of data available.

Thank you.

Gerry Gray, Ph.D.

DR. GRAY:  Good morning.  My name is Gerry Gray.  I was the statistical reviewer for this device.

[Slide.]

Just give a synopsis of a couple of things you have already seen in this study so far, patients were randomized in this study 1 to 1 to Optimal Medical Management versus Left Ventricular Assist System.

The primary endpoint of the trial was two-year mortality.  By that, I mean the trial was designed to detect a difference in survival over the course of the two-year period between the two arms of the trial.

There were three interim analyses designed into the trial at 23 deaths, the trial was designed to stop after 92 deaths.

For the survival analyses we have complete follow-up.  There is no missing data for survival.

As of the end of last June, there were 128 patients enrolled, 61 OMM, 67 LVAS.  At that time, we had 40 deaths in the LVAS arm, and the remaining 52 in the OMM arm.

Some of my slides will use the updated survival data, that was updated as of the beginning of February of this year.

[Slide.]

The primary endpoint for the trial was all cause mortality.  This is a graph of the survival curves showing the all cause mortality for the two arms of the trial.  The X axis shows months since randomization, the Y axis, the estimated survival probability.

The thick red line is the LVAS arm, the thick black line is the OMM arm.  The thin line show 95 percent confidence intervals at each point over the course of the trial.  The blue circles indicate the estimated median survival for the two groups.  That was 150 days for the OMM arm versus 408 days for the LVAS arm.

Below the graph are the numbers of patients who survived up to 0, 6, 12, 18, and 24 months.  So, all the results past that are primarily on those patients.

Some features to note in this graph, you might note the survival up until the 30-day point was pretty much identical for the two groups at which point they diverged. The LVAS arm stayed above OMM up until maybe around 22 months, where there was somewhat of a decline in LVAS survival, but again, that is only based on a very few patients.  I think there were only 11 patients who survived up to 22 months.

As has already been stated, the log-rank test for the overall difference is significant with a p-value of 0.003.  The two median survival times of 150 and 408 days are significantly different.  At the one-year point in the June 2001 data, there is a significant difference in survival.

[Slide.]

The updated survival curves shown on this graph are pretty much the same.  I am going to go back and go forward, and you see that the only real difference is that the confidence bounds got slightly smaller because we had more follow-up.

This data of February 2002 had seven months more follow-up than the previous.  The basic features are still unchanged.  We still have a significant difference between the two curves using the log-rank test.  We still have potential drop-off in LVAS survival around the 22-month point.

Using these data at 24 months, pointwise, there is a marginal difference in survival.  The p-value is around 0.05.

[Slide.]

One of the secondary endpoints of the trial was cardiac mortality.  This graph shows the deaths due to cardiac mortality in the two arms of the trial.  Again, red line is LVAS, black is OMM.

In this analysis, deaths that are not due to cardiac causes are considered to be censoring.  Those patients are censored.  When you look at the data in this way, the LVAS arm is very much superior to OMM.  There still might be a drop-off in LVAS survival at 22, but again, this is very, very few patients.

There is a problem in looking at this kind of analysis.  In this case, the cardiac and non-cardiac mortality are not independent.  In statistical terms, you would say that the censoring is independent of the mortality.

[Slide.]

This graph shows the converse of the previous mortality due to non-cardiac causes.  When you look at it this way, the OMM arm looks to be superior to the LVAS arm. You will recall that only 1 out of 52 deaths at the stopping point were due to non-cardiac causes in the OMM arm, whereas, 29 of the 40, about three-quarters of the LVAS deaths, were non-cardiac causes.

So, you have to look at these two graphs sort of together to interpret meaningfully.  So, again, it would appear that the causes or mortality, at least cardiac versus non-cardiac, are not independent, and it would appear that the LVAS device reduces the cardiac mortality and, at the same time, increases the non-cardiac mortality to some degree.

[Slide.]

So, to summarize the mortality results, there is a significant increase in the median survival time, 150 days versus 408 days.  There is a significant difference between the two survival curves using the log-rank test with a p-value of 0.003.

There is a significant difference in the mortality at one year, 50.8.  That says "mortality," it should say "survival."  The survival at one year was 50.8 percent versus 24.4.

Using the updated data, updated as of early February, there is marginal statistically significant difference in mortality at the two-year point.

The cardiac and non-cardiac mortality don't look to be independent, and there might be a drop-off in LVAS survival starting at around 22 months, but again, that is based on only 11 patients who survived after that point.

[Slide.]

Also, as has been previously noted, survival benefits didn't come without a cost.  There were significant increases in the number of adverse events and serious adverse events per person and in the rates per 100 days in the LVAS arm.

[Slide.]

This graphic shows the death and serious adverse event rates per 100 days in the two arms of the trial, the red bars on the left LVAS, the gray bars the OMM arm.

The top five bars in here are device-related events that could only have happened in the LVAS patients.  Overall, as has also been stated, the LVAS rates were significantly higher than the OMM, especially the bleeding and neurologic dysfunction down at the bottom were very different.

[Slide.]

Cutting the time at 30 days or less versus events that happened at more than 30 days is shown in this graph here, the lefthand side being what you might call peri-operative events that happened at 30 days or less, the righthand side being events after 30 days.

Again, these show rates per 100 patient days, so they are adjusted for the different survival times for the two arms of the trial.

You will note that the X axes are not the same in these two, and that is because the rate per 100 days did drop off significantly for both groups after 30 days.  Bleeding and neurologic dysfunction stayed high in the LVAS arm, and for both of these time periods, the event rates were significantly higher for the LVAS patients.

The mortality at the bottom, the death rate for the two arms was nearly identical for the first 30 days. After that, the LVAS rates were much better.

[Slide.]

Associated with this increase in serious adverse event rates, the hospitalization time was significantly higher for the LVAS arm, a median of 61 days in the hospital versus 16, 24 percent of the total days in the LVAS arm were spent in the hospital versus 15 percent in the OMM arm.

I am not sure about this 141.  I got it out of the panel pack, and the sponsor seemed to indicate that was more like 300 days, so I am not totally clear why that says 141.

[Slide.]

One of the panel questions asks the panel how to combine or how to make the tradeoff between mortality benefit and adverse events.

I am going to take a little sidetrack here and talk about two possible ways that you might formally try to combine those into one analysis, the first one being an analysis based on a hierarchical ranking of patients, the second one being survival to not just death, but death or some other bad thing.  These are by no means the only two possibilities.

[Slide.]

First, using a hierarchical ranking, what you might think of doing is trying to rank patients by whatever the most important outcome is, and then if there are any ties, to break those ties using secondary outcomes.

In this trial, the obvious example is to first rank patients by how long they survived, and if they survived to the end of the trial, to try to rank them by some other secondary objective like how many days they spent in the hospital or perhaps how many serious adverse events they had.

The bottom line here is if you believe in this kind of analysis and you can meaningfully rank patients using this hierarchical method, if you use the death time as the first ranking factor, then, regardless of how ever else you rank patients after that, there is always a significant improvement in favor of the device.

So, if you like this kind of analysis, the device would always look better.

[Slide.]

The other method you might do is to try to combine death and other serious adverse events into one analysis. This is again  Kaplan-Meier survival curve that shows the survival up until the time of the first serious adverse event or death.  In this analysis, death and serious adverse events all count the time, there is no differential.

The median LVAS survival time in this analysis is much less, 13 days, versus 87 days in the OMM arm.  After about five or six months, the rates level out to be about the time.  The medians again are shown by the blue dots.

[Slide.]

If you are wondering how we reconcile this with the mortality results, remember that although the time to the first bad event, the first serious adverse event in the LVAS arm is much less, the classification of that bad event is much different.

So, out of the 58 people who had death or some serious adverse event, the first thing that happened to them in the LVAS arm for two of those people was death, whereas, in the OMM arm, the first thing that happened to more than half of the patients was they died, so although the rate was higher in the LVAS arm, the seriousness of the event I guess you would say was different.

[Slide.]

Now, there was a clinical impression from the clinical reviewers that when they read the case reports, that the first bad event initiates this cascade of things that leads ultimately to the death of the patient, and the question to me was, is that impression borne out in some formal analysis, and if it is, is there is some difference between the two groups in the timing of those subsequent events.

[Slide.]

This graph shows the results of such an analysis. Again, it is Kaplan-Meier survival curves of conditional survival to death or serious adverse event.  The two upper lines here are the time up until the person had their first serious adverse event, and those are the two lines that we have already seen in a previous graph.

The two lower lines show if you survive that first adverse event, how long was it until you had some other bad thing happen to you.  The point here to note is that after the patient has their first bad event, the second and third and fourth events come very rapidly, and actually, they come at about the same rate for the two arms.

So, the main difference between the two groups is the time up until the first event.  After that, events occur rapidly at median times of about three to five days, at about the same for the two groups.

[Slide.]

This table summarizes the functional status and Quality of Life results that we saw in the trial.  The entries in the table are the p values for tests between OMM and LVAS improvement over baseline.

These results generally favor the LVAS arm especially for NYHA, where LVAS consistently improved significantly more than the OMM arm.  The hall walk and the peak VO2 are pretty much inconclusive partly because there were very few patients who could complete those.

The Quality of Life results somewhat favor the LVAS especially for the EuroQOL, less so for the Beck and less for the Minnesota Living With Heart Failure and the SF-36.

I would note that the samples sizes are very small at the 6- and 12-month point.  At 6 months, there were 30 LVAS and 21 OMM patients, and at 12 months, there were 22 LVAS and 6 OMM patients who are represented here.

[Slide.]

To summarize the statistical results, there was a significant decrease in mortality for the LVAS arm whether you measure by median survival time or you use a log-rank test, or compare pointwise at one and two years.

The significant adverse event rates were much higher in the LVAS arm.  The LVAS treatment resulted in decreased cardiac mortality rates and increased non-cardiac mortality rates.

Survival much past two years was poor in both groups.

There may be some indication of a relative LVAS drop-off in survival at about 22 months, but there were very few patients, only 11, that survived that far.

The difference between the two groups is almost entirely in the time up until their first event, not to the time between subsequent events after that, and any stage, the odds of death versus some other serious adverse event is always higher for the OMM arm.

The functional status results favor the LVAS arm, but not consistently.

Thank you.

DR. LASKEY:  Mike, did you want to read the questions for the panel?

Questions to the Committee

DR. BERMAN:  For the record, I am Michael Berman. I am going to read into the record the questions.

Panel, as you deliberate this afternoon, we ask that you please keep these questions in mind.

1.  The bench testing performed to assess device reliability did not account for all observed clinical conditions, in particular, higher than expected pressure in the pump chamber and higher than expected beat rates. Accordingly, the observed times to device failure and/or device malfunction seen in the clinical study are less than those predicted by the reliability model.  As well, there is no reliable end-of-pump-life indicator.  Please discuss the clinical implications of the observed reliability for the indication proposed, which is long term.

2.  Are the device failure and malfunction rates and their times to occurrence appropriate for a device intended for use for destination therapy?

3.  Given the Kaplan-Meier survival curves and the fact that 7 device patients and 3 control patients, as of February 02, had survived to 24 months, have enough patient data been reported to demonstrate a clinically meaningful survival benefit?

4.  The New York Heart Association, the Quality of Life, and the functional testing results are not consistent. From these data, can we determine that there is a clinically meaningful improvement in functional status?

5.  This device demonstrated an increase in median survival time and showed an overall difference in survival. However, this benefit diminished at two years and was associated with serious adverse events and hospitalizations throughout the course of the study.  Do the benefits of this device outweigh its risks?

6.  One aspect of the premarket evaluation of a new product is the review of its labeling.  The labeling must indicate which patients are appropriate for treatment, identify potential adverse events with the use of the device, and explain how the product should be used to maximize benefits and minimize adverse events.

6(a).  Please discuss the appropriateness of the proposed indications for use for this device, which reads:

"The HeartMate VE LVAS is indicated for use as a bridge to transplantation in cardiac transplant candidates at risk of imminent death from nonreversible left ventricular failure.  The HeartMate VE LVAS is also indicated for use in patients with end-stage left ventricular failure who are ineligible for cardiac transplantation.  The HeartMate VE LVAS is intended for use both inside and outside the hospital.

6(b).  Does the labeling accurately inform patients of the risks of the device?

6(c).  Does the labeling adequately inform patients of the expected duration of use for this device?

6(d).  Are there any other issues of safety or effectiveness not adequately covered in the labeling?

7.  Based on the clinical data provided in the panel pack, do you believe that additional clinical follow-up or postmarket studies are necessary to evaluate the long-term effects of this device?  If so, how long should patients be followed, and what endpoints and adverse events should be measured?

Thank you.

DR. LASKEY:  Again, thank you all to the presenters and the FDA reviewers.  A well-deserved lunch break, and we would like to reconvene promptly at 1 o'clock. Thank you.

[Whereupon, at 12:00 noon, the proceedings were recessed, to be resumed at 1:00 p.m.]

AFTERNOON PROCEEDINGS

[1:10 p.m.]

Open Discussion

DR. LASKEY:  I would like to start off the afternoon by having Drs. Aziz and Konstam deliver their reviews and ask questions of the sponsor.

Dr. Konstam, do you want to begin?

DR. KONSTAM:  Sure.  First, I just want to say a couple of things.  One, I want to congratulate Dr. Rose and his colleagues and the sponsor for conducting this study.  I think it is going to stand as a landmark study and showing that a study like this can be done, and I just want to compliment them for embarking on it and carrying it out as successfully as they did.

I also think that the FDA reviewers were fabulous. I think that they have hit all the critical issues, at least that I am aware of, and maybe some that I am not aware of, but I think they did a fabulous job of summarizing everything.

I find sort of six categories of discussion, and just to go through what I think they are.  The first is the primary endpoint met.  I think there is a little bit of discussion that is required here, maybe I can ask Eric.

As I read the protocol, it seemed as though it was your intent to show a significant improvement at two years.

Do you want to comment on that because that isn't quite met and maybe you could comment on that.

DR. ROSE:  In earlier incarnations of the protocol, that was indeed the intent.  When we formed a DSMB, and actually once we had NHLBI involvement, as well, and a DSMB had been constituted, there was a lot of consideration of what the primary endpoint ought to be and concern that a discrete, single endpoint would pose potential difficulties in monitoring data or the safety of the trial in terms of outcomes, and instead, the area under the curve would be something that would be more easily monitored and monitorable, and for which you could set stopping rules as time went on.

So, the statistical plan and the primary analysis then focused on a two-year duration observation, and that indeed became the primary analysis, and the final statistical plan was agreed to by the NHLBI, the Date Safety Monitoring Board, and the FDA.

So, I think it is fair to say that that evolved over time.

DR. KONSTAM:  Personally, I can accept that.  That might be some cause of discussion on the part of the panel. I think we are going to circle back on that when we deal with the clinical relevance of the finding more than the statistical significance, but I guess it is worth just mentioning the fact that the timing of the primary endpoint has changed a little bit.

I guess the statistics of the log-rank test seem pretty strong.  One must ask whether the two groups are similar.  They seem similar in most regards.  There are a couple of things a little bit different between the two groups.  One I noticed is difference in warfarin use between the two groups, and you might want to comment on that.

The other thing which is a little bit more troublesome is that there is a discrepancy between the two groups in the number of--as I read it, unless I am misreading it--in the number of patients enrolled within the last year at least by my count.

Looking at the FDA document, 34 percent of the LVAD patients were enrolled with less than one year of follow-up, was only 23 percent in the medical management arm, which leads to a couple of issues, but one is, does this suggest some difference between the two groups, and why is that, what would there have been that difference?

DR. ROSE:  On the warfarin issue, I have no idea.  You look at 20 variables on a randomization, and you use a p-value of 0.05, one of them is going to turn out to be that way.  How that one ended up being the case, it is not the issue that clinicians have not viewed them as candidates for some kind of anticoagulation because the converse aspirin use was higher in the device patients.  If anything, that might have disadvantaged the device patients because going into the surgery, which had to happen within a brief time after randomization on aspirin is something that I think we would find less than desirable, knowing the frequency and severity of the bleeding issues.

With regard to that difference over the later part of the trial, as I said, we did block these by center, and the blocks were random.  I don't think that this reflects any degree of bias, and just is the way the cards fell out.

DR. KONSTAM:  Well, I suspect you are right.

DR. ROSE:  We also still had, in terms of the blocks, you know, the enrollment after the trial was stopped before the 140 that we expected to enroll.

DR. KONSTAM:  Well, those are the caveats I find regarding the decision that the primary endpoint is met.  I think that the statistics are strong that there is an effect on survival.  Those are the only holes I can find to pick at, but I am not concerned about those greatly.

I think the second issue to discuss, and to me, in fact, I think the whole decision about approvability centers around this, which is the clinical relevance, and then where the clinical relevance comes up for discussion is what is going on at two years.  To some extent, this is going to be a very subjective decision about, you know, what does it mean that there is a statistically significant effect at one year, and you lose it at two years, but more than lose it at two years, most of the patients essentially are dead at two years.  They are either dead at two years or they are dead within a month or two of two years.

Now, there are a number of aspects of this.  One is the reliability of the device and to what extent is a less than anticipated reliability of the device contributing to the mortality of two years.

The second problem is we don't have that many patients reaching two years, in part because again 34 percent of the patients in the LVAD group were enrolled within less than one year from the time the study ended.

Maybe I just can ask you again, Eric, about this two-year thing.  When you set up the protocol originally, the primary endpoint was I think you said envisioned at the two-year endpoint, but there is a couple of aspects of that.

One is that when you looked at your pair analysis, the pair analysis was constructed saying that you wanted to see, you anticipated 75 percent more mortality at two years in the control group.  You hoped to achieve a third reduction.  That gives you 50 percent survival at two years.

Now, I guess the question I have for you, is that just a numbers game, but you had to put some numbers in to have a pair analysis, or are you implicitly making some kind of value judgment about the survival of 50 percent at two years, are you sort of saying--I guess what I am leading to is does this help us in terms of the overall judgment about the clinical relevance?

Were you saying at the beginning that we wanted to have a device that is clinically relevant by virtue of the fact that half of these patients are going to be alive at two years?  I don't know whether you can comment on that.

DR. ROSE:  I think the issue of clinical relevance has been part of the discussion throughout the generation of this trial.  We initially did a preliminary randomization where our initial modeling of mortality, we anticipated that it would be worse in the VAD patients and that the two curves would cross late, that that would be the way this would shake out.

After pre-match, we actually revised our thinking in that regard, and that is what allowed the more appropriateness of the log-rank as opposed to other types of analysis.

With regard to whether or not we wanted to see 50 percent of the patients alive at two years, of course, we would have preferred that, but our thinking also early was that the control group mortality was not going to be as high as indeed we saw.

That was a surprise to us even knowing how sick these patients were.  We thought that that longer term survival of the control group at a rate higher than 8 percent was likely to result.  Now, that clearly was not the case, and I think, you know, would we have preferred to see 50 percent survival at two years?

Certainly, but to come up with tripling of survival at two years with a quarter of the patients still alive, and a large number still ongoing, and this field being in its infancy, I would have to say that for patients now, I think this is clinically relevant and important survival benefit that we have documented, and it is time to let this field begin what I believe will be a ascent to better outcomes.

With regard to the issue is device reliability accountable for the lower survival rate than I think a lot would have like to have seen at two years, that is only part of it.  Don't forget that half of these patients, as well as they did do in the first, half died in that first year, as well, and the overwhelming numbers of death occurred during the initial index hospitalization. About a third of patients with this type of therapy die in hospital with the index hospitalization.

I believe that we can cut that down substantially as we learn more over time, but that way you end up with more patient eligible--

DR. KONSTAM:  Let me ask this then.  I mean it seems what happened is the device was exceedingly good at preventing cardiovascular deaths, and the fact that the mortality was so high in the medical management group means that this is an incredibly sick population, and so they were appropriate for consideration of this type of treatment, so that it good.

But since the device was so good at preventing cardiovascular deaths, the fact that you have so few people alive at two years, I guess implies there are more other kinds of things happening than you would have anticipated.

DR. ROSE:  Another way to view it would be, though, to say that half the people that are alive at one year, are alive at two years, and in that regard, the device has a degree of reliability that is still potent.  That is not to say that reliability is not a problem, but there are other issues that limit survival in this patient population if we are going to holistically improve it.

DR. KONSTAM:  Well, part of the problem is we have so few patients at two years.  I mean we don't really know what the right number is at two years because there aren't that many patients there.  So, I just would comment on that.

Just to ask rhetorically, I mean if you are putting in a device that is clearly life-sustaining, what would you like out of that device in terms of its mechanical reliability, and what would you like out of that device in terms of warning signs that it is going to fail?

These are I think again questions that are going to come back.

The next category I think is the whole issue of Quality of Life assessment.  I see some very reassuring things in the data set.  I see some confusing things in the data set.

The confusion I think stems from the fact that there is such a high dropout rate in large part due to mortality and to some extent due to perhaps patient incapacity that it is extremely difficult to know exactly what the numbers for each of these tests are saying to us.

I can, however, accept the fact that things look better in the device group than in the medical therapy group, certainly in terms in your Heart Association class and in some of the Quality of Life parameters.

I guess one question that I want to ask you about, though, and the problem then, of course, in any of these, so part of the problem is the dropout rates that we are seeing and what do you do with that, but the other problem is that we don't have a blinded study, and we are going to be stuck with that.

The specific thing I wanted to ask about that is if you look at the numbers, let me say a positive thing.  I am really reassured by the fact that a lot of the Quality of Life scores, forgetting about the comparisons, but if you just look at the LVAD group, a lot of the Quality of Life scores and a lot of the New York Heart Association class scores are pretty good.

What we can glean out of those numbers, it really seems to be, despite all the complications that we do see, there still are a number of patients, a fair number of patients at three months, six months, and 12 months, who seem to be doing pretty well.  I get a lot of reassurance from that.

The one specific question I was to ask, though, is if you look at the number of patients who are said to be New York Heart Association Class I and II in the LVAD group, it seems that that number is disparate from the number of patients participating, for example, in the 6-minute walk.

The number of patients having repeated measures--I don't have it in front of me exactly--but the number of patients having repeated measures in the 6-minute walk is very small, I think, and smaller than the number of patients reported to be in New York Heart Association Class I and II, whom I would have expected could participate in the 6-minute walk as an objective question.

Can you comment on that, and why shouldn't we be concerned from that, that there is just some kind of a bias being injected in the New York Heart Association class grade?

DR. ROSE:  Specifically, with regard to the 6-minute walk and the peak VO2 measurements, I think it is not unfair to say that a lot of investigators viewed the tests themselves as tests that could impair patients' Quality of Life.  They didn't like doing these things.

Secondly, I think that we did come to the conclusion that the more subjective measures of Quality of Life were a lot more reasonable and reflected more what is relevant to the patients than anything we would have measured in a 6-minute walk or a peak VO2 measurement.

So, in all those regards, we did not push investigators in the way I wish we had, and the way that the data has turned out to generate those data.  That, notwithstanding, also, the correlation between 6-minute walks and peak VO2's and Quality of Life for any degree of heart failure are not well established, and we certainly did not have a large enough N to do that in this trial, and their correlation with survival is not well established, if at all, as well.

So, for all those reasons, I would say we would have to be criticized as remiss in terms of not pushing to get them, but I don't think it reflects any bias on our part towards one group or another in assessing them.

Actually, if you look at the data, you don't see much of a difference between the two groups.

DR. KONSTAM:  Bias is not necessarily a bad thing.  It is just a reality.  That is why we blind drug trials.  You can't escape it.  If you know what group of patients it is in, it is a potential bias in how you grade it.  That just has to be said.

I think, however, I just give you my conclusion again, which is that it certainly does not appear as though everybody alive with the LVAD is alive, you know, miserable in bed, it does not appear that, and I do get that out of the Quality of Life data and the New York Heart Association classes.

The next broad category is the adverse events.  I think that is really there juxtaposed against the survival benefit, and it was nicely sort of delineated by the statistical reviewer about different ways of looking at this.

If you look at all serious adverse events including death as equal, you have got a problem on your hands from the perspective of this trial.  I certainly don't look at it that way.  I would create a hierarchy where survival is better, and so we are really going to have to face that.

I have some specific questions about the adverse events.  There seem to be a high number of neurologic adverse events, and there seem to be also a higher number--there are some septic events--but there are a higher number of bleeds than I would have anticipated even remote from the time of the operation.

I just wanted to go through your thoughts about a few of those things.

DR. ROSE:  You mean non-neurologic?  There was only one intracranial.

DR. KONSTAM:  Non-neurologic bleeds, yes, two different things.  So, there is strokes, which I guess are predominantly embolic.  There are other neurologic events, and then there are some bleeds, non-neurologic bleeds.  I guess those are the things that caught my eye.

Just to take them in order.  Do you want to comment on the strokes?  I guess I would have anticipated with the non-thrombogenic surface that we are dealing with, that we wouldn't have seen this many cerebrovascular accidents.

DR. ROSE:  I think that the disparity in that regard is that many of the strokes occurred in a setting of device infection, so they are septic emboli.

DR. KONSTAM:  Really?

DR. ROSE:  Yes.

DR. KONSTAM:  Is that just your impression or do we have some evidence of that?

DR. ROSE:  The particular question is with the textured surface and everything else, should we have seen an overall smaller incidence of stroke.

DR. KONSTAM:  You are saying it in a different light than I was thinking.  I thought that these were for embolic events.

DR. ROSE:  We have a summary slide on neurologic adverse events.

[Slide.]

The frequency actually of disabling stroke was, at least for the device literature, in this trial remarkably low especially considering the length of time that the patients were on devices.  All neurologic events here were not strokes.

DR. KONSTAM:  I understand that, and I was going to come to that, too, but there are strokes.

DR. ROSE:  Yes.

DR. KONSTAM:  And there are more strokes, if you sort through it, you are still left with a certain number of strokes that I am inferring at thromboembolic strokes.

DR. ROSE:  Yes.

DR. KONSTAM:  And I guess it would be important to note this, and I would presume, unless we have information otherwise, that the device is somehow contributing to this, because I think if that is the case, then, going forward, we ought to think about ways or preventing that including one question I would have is should anticoagulation be reconsidered with the device.

DR. LONG:  I am Jim Long.  I am the co-director of the surgical arm of this REMATCH trial.  I guess for the sake of completeness I ought to also declare that I do not have a conflict of interest financially with Thoratec although they are paying my expenses at this meeting.

There are a number of factors that went into the stroke, a couple of which were not specifically device related, but management related.  For example, two of those were air emboli that occurred at the time of operation, so included in this are not only device-specific issues, but management issues.

Also included in this, and I am sorry I don't have those numbers at my fingertips right now, but also included in this are patient factors, and as Eric alluded to, a fair number of these strokes occurred in the setting of sepsis. That is a fairly well documented occurrence in this field with this technology.

DR. ROSE:  With regard to the question of anticoagulation, obviously, one of the unique aspects of this device is that we did not routinely use systemic anticoagulation here, and whether or not it would have made a difference in stroke rate, we don't know the answer to that.

Certainly, the N that we studied here with 68 device patients, even had we randomized them, it was not even close to something that would have allowed us to power this to judge whether or not anticoagulation ought to be part of the management of these patients.

DR. KONSTAM:  I understand.

DR. LONG:  Here is a summary of the neurologic events just for frame of reference.

[Slide.]

Eleven strokes, 10 in the LVAD area.  Two of those were, as I indicated, air emboli, two to intracranial hemorrhage, and two presumably due to thromboembolic events, six of those in the LVAD arm, and that is by way of frame of reference, relative comparison to the VE bridge to transplant experience, which you see in the bottom, 0.45 events per year in the REMATCH group versus 0.51 events per year in the VE bridge, so actually a lower incidence.

DR. KONSTAM:  I accept that.  I am just making note of the fact that there are strokes, that there are events that seem to be thromboembolic strokes, that is far from a zero rate.  It's a significant difference between the two treatment groups in this regard.

There are other events that are TIAs that I think many of which are likely to have been thromboembolic.  So, I think it is worth noting this because I think if that is the case, and if some of the patients are really--if this is a cause of morbidity and perhaps mortality that is contributing to the overall problem, well, then you have to think about that, and perhaps even rethink the whole anticoagulation question in some regard, you know, whether it be antiplatelet or whether it be warfarin anticoagulation.  It is worth rethinking in future trials or future experiences.

I guess that sort of brings me to the other part, is that there seem to be some bleeds that occurred, that were not limited to the perioperative period, non-intracranial bleeds.

Do you want to comment on those?  What were they and why were patient bleeding?

DR. LONG:  The vast amount of the non-neurologic bleeds were associated with operations or reoperations for the device specifically, but there were bleeding episodes in addition to that, that were counted in that, for example, femoral puncture leading to a retroperitoneal hemorrhage got included in that as a bleeding episode.

There are a fair number of those, but the vast majority of them are perioperative bleeding episodes either at the time of device implantation or at the time of device replacement.

DR. KONSTAM:  I couldn't tell that from the data because many of them were temporally removed from the time of the initial implant.

DR. ROSE:  But some of those were device related, though, too, if we can get the device slide comparing the VE SNAP to the old VE.

DR. KONSTAM:  The overall bleeding rate looks like about 0.15 per 100 patient days compared to something like 0.01 in the other group.  If you look at rate after the first 30 days, it is still about 0.12 or something per 100 patient days.

DR. ROSE:  One of the mechanisms you see is illustrated in this slide.

[Slide.]

Without strain relief, you essentially have a kink in the graft, and you then have friction rubbing against it, so that what can and did happen with both this patient population and the bridge patient population is we would see delayed bleeding from erosion of the outflow graft due to this kind of kinking.

DR. KONSTAM:  I see.

DR. ROSE:  That is what generates this kind of strain.

DR. KONSTAM:  If that is the case, you might be able to see a difference in the VE SNAP patients when you analyze them separately.  I mean that is something that seems like it would be worth looking at.

As I read it, both of those things concern me, the strokes and the bleeds, and I think it is worth some detailed analysis, and I also think that, again, it is worth rethinking the anticoagulation question in my mind in light of those things.

If you believe that the bleeding problem is solved technically from this, and if you know that most of the other bleeds are occurring at the time of an operation, then, maybe the patients could benefit from some form of anticoagulation to prevent some of the thromboembolic events that do seem to be occurring.

DR. ROSE:  That is absolutely fair.

DR. LASKEY:  Marv, excuse me.  Just in fairness to process, we are going to try and finish.

DR. KONSTAM:  That's fine.  I am sorry for going on.  I will just skip some of the other things.  The only thing else I would say is that if we go forward, if this device is approved for this indication, I guess I would like to see a lot more clearly defined for whom is it indicated.

Certainly, the proposed wording by the sponsor--and we can talk about this at more length--doesn't come close to satisfying what I would like to see, and certainly doesn't recapitulate in any way the entry criteria of the study.

So, particularly in light of the large number of adverse events and the fact that at this point in time, very few patients are alive at two years, there is going to need to be an awful lot of consideration given to exactly who is a candidate for this and make sure that the genie doesn't get out of the bottle in terms of overusing this device.

Sorry for going on.

DR. LASKEY:  No, no, please don't apologize.  This is a terribly important discussion particularly for the lead reviewers, but I will ask subsequent reviewers to just try and limit their time to about 15 minutes if possible to get through the table.  We will have a second round.

Did you have more questions, Mr. Konstam?

DR. KONSTAM:  Nothing worth going on about.

DR. LASKEY:  Well, you never know.

Dr. Aziz.

DR. AZIZ:  Let me also echo the early statements. I think the presentation was extremely well done by Dr. Rose, the company, and I think it is a good model to have industry, academia, and the NIH involved in helping to answer some important questions.

I have a lot of questions, but I will try to limit them in the interests of time.

Going through the sites, first of all, the Cleveland Clinic was involved and then it withdraw.  I wonder if I can ask what the reason behind that was.

DR. ROSE:  At the time we began the trial, they began their interest in the Batista operation, and accumulated an experience over a couple of years of in excess of 50 patients for the Batista.

I think they viewed us potentially competitive with this protocol and withdrew for that reason.

DR. AZIZ:  In Table 4, it says, "Reasons Patients were not Enrolled, 187 patients refused the device."  This is in your selection criteria.  Was there any particular sort of events, or were they afraid of an operation?  It seems a high number of patients refusing such a therapy.

DR. ROSE:  You mean from the screening criteria?

DR. AZIZ:  From the screening criteria, yes.

DR. ROSE:  I think based on everything we have shown, it is not surprising that there is a considerable degree or range of patient responses to the prospect of this kind of therapy.  At the other end, we saw patients who absolutely wanted the device and others who absolutely didn't.  Certainly, at least willingness to have a device insertion was essential for the informed consent process in a patient who categorically wouldn't want a device, we excluded from the trial.

DR. AZIZ:  Going to certain other issues, in the neurological sort of events, in one of the tables they mentioned there were two patients that had seizures.  When I was going through the patients who had expired, actually, there were at least sort of four patients who had grand mal seizures.

Were there any indications as to the etiology of that?

DR. ROSE:  They were toxic metabolic-related seizures.

DR. AZIZ:  Because I mean those were four patients that had actually died.

Again, going to the patients who had actually died, I came across at least seven patients who, at autopsy, had evidence of thrombosis either in the graft or on the valves.  Now, that again seems like a fairly high incidence.  I mean I could understand the value obviously of no anticoagulation and having a textured surface, but do you think that this is, on the other side of the equation, where not being anticoagulated does result in sort of some of these other dramatic events?

DR. ROSE:  I don't think we can dispute that we don't know the potential value of anticoagulation in these patients, but counterpoised against it also is there is a lot of bleeding morbidity in these patients, as well, so what we would rather see is a better anticoagulant.  I think if we had that, that we wouldn't hesitate to use it.

DR. AZIZ:  Also going to the sort of patient profile, there were two patients that had aortic dissections in the surgical group.

DR. ROSE:  Yes.

DR. AZIZ:  I think one of them ended up dying.  I think both of them actually died.

DR. ROSE:  Yes.  That is another late bleeding complication from these devices, and there is tension, gravitational force that is delivered to the ascending aorta by the outflow graft that I think was an under-appreciated potential complication, and the way that outflow grafts are sewn to the ascending aorta I think evolved considerably during the course of the trial to a much more meticulous technique with attempts at strain relief at the aortic anastomosis.

DR. LONG:  Another factor that contributed to that is small patient size seemed to be a predisposing factor, small aorta with a fairly good-sized graft.

DR. AZIZ:  What was the incidence of hypertension, bad hypertension postoperatively in the majority of these patients?

DR. ROSE:  Uncontrollable hypertension was not common, but the need for antihypertensive medications in these patients was very common.

DR. AZIZ:  Was uniform.

DR. ROSE:  I can't say that it is uniform, but I would say it is fairly typical of a patient on a VAD that they need a lot of the agents, for example, ACE inhibitors and beta blockers they were on preoperatively, they go back on them for other reasons.

DR. AZIZ:  Let me move on to certain issues related to device dysfunction, either the motor or the valves.  The valves are stented valves, I presume, or unstented valves, so they are exposed to a very high pressure particularly because of the pressure generated in the LVAD itself.

Now, compared to what the bench testing showed, the longevity was I think about three and a half, four years, but what we actually see in reality clinically, and I am sure there are a number of factors, is that it really isn't that long.

What troubles me a little bit is I think this device is clearly good as a bridge to transplantation.  I think what we are asking of this device now is to really function for a much longer period of time with the repeated stresses on the valve and other components, and I am not completely convinced that something like this could be, let's say, advocated for a four-year period or a five-year period, the reason being I think again the bench testing values defer from what we find clinically, and particularly there were some patients where the inflow valves I think become dysfunctional fairly quickly.

How would you address that, because I mean what we are trying to do is advocate this as a sort of destination therapy.

MR. POIRIER:  If you look at the valve dysfunction, the primary reason for the valve dysfunction was excessive pump chamber pressure, and that came about from several factors, one, outflow kinks, patient's high arterial pressure, any occlusions in the anastomosis, the anastomosis to the aorta was done with just a simple slit as opposed to removing material, there was a pressure drop across that area.

So, unlike the human body, that has only one pressure drop in the ventricle to the aorta, which is the aortic valve, these devices have many, many pressure drops, the anastomosis, the size of the graft, the outflow valve, any kinking, so there were many factors.

All of those are additive.  The net result is the pump chamber pressure has to increase, which puts stress on the valve.  We have taken great pains at improving that.  We put the strain relief bond to reduce that pressure drop, we have instructed and written papers on patient management in terms of the anastomosis, how the device should be installed.

We have looked at changing, modifying the way the device is controlled to reduce the pump chamber pressure, so we have done many, many things to try to reduce pump chamber pressure, to increase the longevity of the tissue valves.

I would like to clarify one point.  The in-vitro test does not evaluate the tissue valves.  It is impossible to do a long-term test with tissue valves in mock circulatory loop because the valves become contaminated and the tissue deteriorates.

So, the valves are never tested.  They are tested separately in a separate circulatory loop that can be kept in a sterile environment.  So, the results of the in-vitro tests do not include the tissue valve durability, so there is a difference there.

I think that should answer some of your questions.

DR. LONG:  A couple of other thoughts from a clinical point of view.  The inflow valve failure problem with the HeartMate device is relatively new within the last couple of years.  During its IP era, when it was a pneumatic-driven console with lower pressures, there were virtually no inflow valve failures.

It appears that a change in configuration of the inflow conduit to reduce bleeding has been associated with an increased risk that pressures inside the pump will lead to inflow valve failure.  It's a new phenomenon.  It appears to be correctable by the measures that Vic is introducing.

I guess the other issue of relevance is what are the consequences of that inflow valve failure.  We don't have a lot of that, the bulk of that data within the REMATCH trial itself per se.  From other experiences, we know the natural history of inflow valve failure is slow failure with patients who progressively increase in their heart failure, their pump goes up.

It may lead to a decrease in longevity of that pump, as it was alluded to this morning, but in the vast majority of experience, we do not experience catastrophic outcomes with these patients, and they can be managed medically over time, observed, and then finally committed to pump replacement.

DR. AZIZ:  Obviously, the other issue, which is responsible both from morbidity and mortality, is sepsis.  I accept that, I think as most of us, that that is clearly multifactorial.  So, there are two questions.

One, was there a correlation between the albumin or pre-albumin levels and the incidence of sepsis?

DR. ROSE:  To show you how ignorant we were of this problem, we didn't even collect prospectively albumins on the patients when we started the trial, obviously, highly relevant data and something we are trying to recreate retrospectively, but I think it is clear serum albumin is the best single predictor in the VA study of operative mortality for all patients having any kind of surgery regardless of what the reason was.

I do think that we paid much too little attention to perioperative nutrition in this trial.  We had a symposium of investigators to look at just this subject shortly before the trial reached its endpoint, and I think going forward that nutritional management is going to be key in diminishing this frequency of infections.

DR. AZIZ:  The other, which is probably more than a theoretical suggestion or it may be an etiological factor is the nature of your textured surface.

I think people from your own institution have very nicely shown that the new intima, although it actually may prevent thromboembolic events, does cause an elevation of PRAs in a T-cell separation and B-cell hyperactivity.

What role do you think that might have played or do you have any evidence that that could have also played a role in the genesis of sepsis?

DR. ROSE:  I do believe there is firm evidence that these devices do modulate immune function, and, if anything, it seems that they up-regulate humoral responses and perhaps down-regulate cellular responses.

I don't think that that is unique to the textured surface.  When we have looked at patients with smooth surface devices, we have seen similar perturbations, but I think there is certainly more than the possibility that these devices are, in themselves, immunosuppressant, and can lead to increased susceptibility to infection aside from the issues of access through the transcutaneous ports, and it is a very important set of issues for us.

DR. LONG:  This is an extremely important part of the future for the field, and if you think about this as an issue-compromising outcome, thinking about it in three categories, one is what might the device do; two, what does the patient bring to bear, and three, how does management affect it, you have already alluded to the possibility that the device may play some role in terms of altering an immune system.  Our general consensus is that that is probably a very low role relative to the other two at this point in time.

Patient factors, we think are substantial, and there is an analysis ongoing now trying to look at a subset that might predict risk of infection including nutritional factors.  These patients in this group were absolutely dramatically different than any we had ever faced in the bridge to transplant group.

We thought bridge to transplant patients were sick because they were acutely decompensated with balloon pumps, drips and all sorts of other things.  These patients would walk in, but be terribly cachectic, have lost 20, 30, 40 pounds over the last few months, terribly malnourished, as you know, from that, and immune compromised.

These are the kinds of patients we were dealing with here, and we believe that to be a very significant factor, patient-host issues that contributed to our infection rate.

It goes beyond that, into the third category, though, of management, and let me put a slide up.

[Slide.]

When we looked at a couple of centers that were responsible or that took on themselves extremely aggressive infection control efforts, which I will try to describe in a minute, we found that infection can be controlled, and we think this infers that management played a significant role in this process.

If you look at two centers here that did 16 LVADs relative to 52 for the rest, or 24 percent of the total LVADs done, with a median duration of use, even greater than the rest of them, so a longer time of exposure at which they could be more susceptible to infection, we find that in these two centers, we have zero percent of serious infection, 0.06 serious drive line or pocket infection, and zero serious pump housing inflow or outflow infection relative to the others that you see there.

This is certainly a small and not absolutely demonstrative that management does make a difference, but a strong inference that that is the case.  So, out of the three categories of this very, very important topic, device probably plays a small role, patient factors in this patient population played an immense role, and management, we think plays  very strong role, as well, and very significant room for the future in both of those latter two categories for improvement.

DR. EWING:  Excuse me for a moment.  I would like to remind the sponsor and the panel members that this is new information to the FDA, we have not reviewed this, and we would like to keep in consideration we are to be discussing material that is presented in the application.

DR. AZIZ:  One other quick question.  How many of the patients at two years had had the original device?  In other words, how many patients who went on for a long period of time had had to have the device changed or something done technically?

DR. ROSE:  A total of eight pump replacements as far as I know.  Also, a number of patients who could have had their pumps replaced, who opted not to.

DR. LASKEY:  Before we move on, Dr. Rose, your comments remind me of a small book I read as a surgical intern, written by Frannie Moore, about the metabolic control of the surgical patient, and I think it still remains true, perhaps more so.

It is my understanding that in a randomized trial by the toss of a coin, these predisposing factors should certainly sort themselves out in either arms, so there is something to be said here.

Moving on with the other panelists, Dr. Wittes, please.

DR. WITTES:  First of all, I want to echo everybody's statement that the presentations, both in the panel pack and today, were very clear, and it helps us a lot when there is clarity.

I also want to just point out what a good thing it is that you randomized this study, because had you observed 75 percent mortality in your treated group, which is what you expected in the control, I don't think we would be here today.

I do see three main themes.  One, of course, is the mortality, which has been already brought up in various ways.  The other is the issue of device failure.  The third is sepsis, which has been brought up, but I only have one extra thing to add.

Obviously the question of this very high mortality rate in the control group raises the question of who were these patients, and the device appears to have high risk, raising the concern of this device in a population of healthier patients.

I am just stating that, it is not really a question.  It's a statement of concern.  My own feeling is that death trumps nearly everything, so I am not all that concerned about Quality of Life and things like that.  This may not go over so well around here, but it seems to me that when we have a big effect on mortality, that we have to sit up and take notice.

But I am troubled by the dearth of data at two years.  I understand what happened, that you powered the study in a way that if you have 92 deaths, you stop the study at 92 deaths, and because the mortality rate was so high, the deaths occurred before two years, so the follow-up didn't go as long as you planned.

I am not looking actually for statistical significance at two years, I am fine with a log-rank test over the period, but I would actually like to see more data at two, and there is 25 percent--not 25 percent, more than that--I looked at the data through February, and there is still I think 20 patients in the treated group, in the device group, who are still censored, so there is data in the pipeline that will speak to two years.

What are your plans for follow-up of these patients?

DR. ROSE:  Well, the trial is ongoing.  We stopped enrollment and we reached the primary endpoint, and we did an analysis at the time of the enrollment.  It is remarkable to me how--it is not well appreciated this is now the largest series of long-term device patients ever as far as I know, and I think that there is a lot of additional useful information that will come to us.

We are funded by the NIH to continue to follow these patients, and we intend to follow them all for the duration of their lives.  That is our intent, and we intend to keep reporting that as time passes.

DR. WITTES:  The other thing, I just want to make a note about log-rank.  I am very uncomfortable with the language about the log-rank being the area under the curve measure, because it is really not, and I think that it gives a very--I mean the log-rank tests looks at data at the time of deaths, and it is not integrating area at all, and I just want to make sure that that is not the statement that you made, because it gives a very different picture.

DR. ROSE:  What was the statement I made?

DR. WITTES:  Well, I don't know who made it.

DR. ROSE:  It probably was me.

DR. WITTES:  I heard it several times.  The first time I didn't pay any attention to it, but it came up a second time.

A question about sepsis.  I may have misinterpreted the data in the panel pack, but it looked to me as if the number of cases in the two groups was the same, but the case fatality rate was higher in the device group.

So, my first question is, is that true, and the second question, if so, why?

DR. ROSE:  Do you mean the mortality rate from sepsis?

DR. WITTES:  Yes.  If my first statement is correct, and I am not sure it is, but it was hard for me to find that, that the rate of sepsis is the same in the two groups, and we know that the mortality from sepsis, the split was 17-1 in the device versus the control.  That would imply a higher case fatality rate.

So, the question is, what is it, why would that have been true?

DR. LONG:  The overall sepsis rate, total rate was not different between the two groups, but when you look at device-associated sepsis, that led to a significant impact on adverse event that exceeded that in the medical management arm.

DR. WITTES:  I don't think you are answering my question quite.  What it seems to me it is saying is that somehow the sepsis, if you had sepsis in a device, your probability of living through that episode was much lower than if you were in medical management and had sepsis.

DR. LONG:  That is correct.

DR. ROSE:  The ratio of the incidence of sepsis in the device patients compared to OMM was 2.  It was double in the device patients compared to the controls.  I don't think that reached statistically significance in isolation, but it still was double.

DR. WITTES:  But the death was 17 to 1.

DR. ROSE:  Excuse me?

DR. WITTES:  Wasn't it 17 to 1?

DR. ROSE:  As a cause of death, it was significant.

DR. WITTES:  Well, let me ask in a different way.

DR. ROSE:  This is hot off the presses March 1st.

DR. WITTES:  What?

DR. ROSE:  This is March 1, 2002, data, so that difference was significant.

DR. WITTES:  I am not interested in significance.

DR. ROSE:  But another reason for that, though, the medically managed patients died from heart failure.

DR. WITTES:  But the ones in medical management who got sepsis didn't seem to die of sepsis, is that right?

DR. ROSE:  It may not have been the exact proximate cause of their death.

DR. LONG:  One patient died of sepsis.

DR. ROSE:  That was ascribed to sepsis.

DR. LASKEY:  Are these not all adjudicated by the DSMB in terms of the cause of death?

DR. ROSE:  By the M&M Committee, yes.

You are correct, there is one death attributed to sepsis.  Whether that means that sepsis is worse in the device patients or heart failure is worse in the OMM patients and supersedes it, I don't know.

DR. LONG:  Your observation is correct, that there are more deaths due to sepsis in the device group than there are in the medical management group.

DR. WITTES:  Let me say it once more, though, because that is clear, but the other issue is, if the data as reported are correct, and I realize that you could have differential reporting in the classification, and so forth, then, it seems to say that if you are in a septic episode and you have the device, you are at much higher risk for death than if you have a septic episode and you are on medical management.

The relevance to me, and I obviously defer to all of you guys, is that that seems like it's a very high risk and maybe some of the discussion that came up before in terms of management of these episodes needs to be addressed.

DR. ROSE:  I don't think it is surprising that sepsis in the presence of a large intravascular foreign body results in a higher mortality rate than in the other patients.  I think that is part of the risk profile.

DR. LONG:  Plus some of the patients who did get infected, LVADs underwent replacement in the face of sepsis and replacement, ended up with mortality.  So, yes, we concur entirely with you.

This is an area of extreme concern for the field, as I discussed earlier, out of the three things that you look at in terms of etiology, device patient issues and management issues, we consider the latter two to be very substantial areas for targeting improvement.

DR. WITTES:  The other issue has to do with patient reported outcomes.  I guess I have to confess I find them questionable in an open trial especially with huge differential mortality in the two groups.

I am comforted by the people who go from Class IV to Class I and II, but other than that, I think it is really hard to tease out what these mean.

DR. LASKEY:  Dr. Domanski.

DR. DOMANSKI:  I would kind of like to back off a little bit to a big picture and make sure that I really understand fully.  I think, first of all, on the Quality of Life data, I guess I am taken with the fact--I am sort of coming backwards to things that I think are more important--but I think the small numbers, the unblinded nature of the trial, and so the potential for investigator bias, the potential for placebo effect, the small numbers, from my standpoint, make the Quality of Life data far less than compelling.

I share for different reasons Dr. Wittes's lack of interest in it actually, because I don't think it is terribly useful, but I may be wrong, and I am willing to accept an argument relative to that.

But I would like to come back to the mortality and also to the serious adverse events.  It seems to me that if you look at death, there is a difference, certainly at one year.  If you look at death plus serious adverse events as they have been defined, there is no difference, but I guess some of the serious adverse events strike me as not all that tremendously serious.

I mean extending hospital stay a bit is unfortunate, but I think that is one of them, and I guess that is not something that you live with long term even being rehospitalized, while it is a significant event, is not I think on a par with, for instance, a cerebrovascular accident.

While I don't think, in fact, I know these data weren't tabulated in this way.  I guess things that do strike me as being irreparably serious adverse events are things like cerebrovascular accident.

So, if one were to look at death plus cerebrovascular accident, I wonder if you have a back-up slide that actually looks at that.  If you don't with Kaplan-Meier, can we at least add them up and see what the difference is between the two groups, because that is in the submission, it is just not parsed that way.

I think if there weren't a difference, then, I start to wonder.  So, that is Question 1 and I have a follow-up.

DR. ROSE:  We don't have such a slide.

DR. DOMANSKI:  Is it easy to add up?  I know you can't get the Kaplan-Meier stuff this easily, but a reasonable approximation to it would be to look at one year at your deaths plus cerebrovascular accident rate.  Do you have that for the two groups?

DR. ROSE:  We have the raw data.  Whether or not we can create freedom from death plus CVA at one time--

DR. DOMANSKI:  Because that is something I would really like to see.

DR. LASKEY:  You mean death or.

DR. DOMANSKI:  That's correct.  That is what I mean by death plus CVA, that is what I mean, death or cerebrovascular accident.  I am just curious because it seems to me that those are the two--

DR. ROSE:  I assume you mean cerebrovascular accident also with--

DR. DOMANSKI:  With a bleed, with an ischemic stroke.

DR. ROSE:  With a permanent deficit.

DR. DOMANSKI:  That's correct, or with a deficit that lasts beyond 24 hours.  I will spot you even the TIAs.  I will even spot you the seizures.  The reason I am so curious about it is that one is talking about, you know, if one asks whether something should be approved for use, one can reasonably ask whether it is demonstrated to be clinically very useful, and what you are talking about is taking people who are more or less at the end of their lives and saying maybe if I give you this device, I will give you a big operation, I will give you a big, expensive device, and I buy you a year of something, and it is not entirely clear to me what you have bought, all right, so there are a fewer of them dead at one year, but I wonder.

You know, if there were no difference between death or CVA, even at one year, then, I guess I would wonder what you have done.  I still wonder a little what we have done doing it for one year, but that is the pith of the question.

DR. ROSE:  I think it's an excellent question.  It is something we are going to try to tabulate during the time that we are here.  I want to make the point, though, with regard to the analysis of AE's of the two groups, that part of the design and reporting spotted the OMM group, worsening heart failure, as an AE.  That was not considered to be an AE in the control group, had hospitalizations for it, you know, getting short of breath, need to be intubated, having to be hospitalized again, I think is a considerable adverse event when you have severe heart failure that was not totalled in the OMM group as adverse events.

If we were to do a fair comparison, we need to count everything from a patient perspective as an adverse event, not--

DR. DOMANSKI:  Why didn't you then?

DR. ROSE:  Excuse me?

DR. DOMANSKI:  So, why didn't you count it?

DR. ROSE:  The sense was that this was tautologic, we knew that these patients had heart failure and they were going to die from it, so that the fact that they had exacerbations even of their severe heart failure was something that there was not much--in retrospect, I wish we had.

DR. DOMANSKI:  In order for me not to perseverate and hold the committee up, I would just like to, you know, again, just back off.  I would like to see that tabulation, and I see the folks doing it over there, and I think the challenge that I would offer to the group that has to sit here and decide whether to approve it is whether or not one has done a clinically significant deed with this.

I have some concern about it even if the tabulation is in favor of the device.  Certainly, if it's not, my concern would be raised.  I think with that, I would pass it to the next person just while they are going ahead with that.

DR. LASKEY:  Dr. Comerota.

DR. COMEROTA:  Thank you.  Also, I would like to extend my compliments to all the presenters for an excellent job.  Many of my questions have already been addressed, so I will funnel them down to a couple comments and perhaps a couple questions.

Eric, in the manuscript, I think it was stated that there was no significant learning curve to this or that was a non-issue, but then also it was stated or the observation presented that in the patients enrolled in the last year, there was a 25 percent decline in mortality relative risk per year when survival was adjusted for date of entry.

Are those contradictory statements, so is there truly a learning curve or perhaps is there a change in patient care that resulted in this observation?

DR. ROSE:  I do believe that there is a learning curve.  I don't think REMATCH is a large enough trial powered to assess it.  It was that difference of survival, adding a year to the date of enrollment we found did occur was not statistically significant.

When we tried to remeasure it again, it still does not reach statistical significance, but I think for anyone who gets involved in this field and starts doing it, the collective wisdom of the investigators I think is increased considerably during the course of this trial.

Also, I want to point out that before this, there was no experience of any large number with destination therapy.  PREMATCH was a total of 21 patients.  That was essentially the American experience, and it was only 11 patients that got devices, so the entire American surgical experience with destination therapy is now 79 patients.

I do think that there is a lot more to be learned at this point.

DR. LONG:  Also, this is the first trial ever to look at a patient population of this age, which gives us a new learning insight, as well.

DR. COMEROTA:  Correct, and that age issue obviously is unique to this trial.

DR. ROSE:  Encouragingly, in patients that were younger than 60 years old in this trial, the one-year survival was 74 percent, which approaches outcomes at least at one year for the process of transplantation.

That is not to say that this is up to that, but that observation in that subgroup is a very encouraging observation.

DR. COMEROTA:  That is a good segue to my next question.  It was stated this morning that the mean time in your bridge to transplant group was about 29 days.

DR. ROSE:  Right.

DR. COMEROTA:  In those patients in the bridge to transplant group that exceeded the 29 days, that may have gone out to two months, have you looked at the survival curve in those patients and compared them to the survival curve in the REMATCH patients?

DR. ROSE:  Meaning bridge to transplant patients with a longer hospital stay than--

DR. COMEROTA:  Well, the bridge to transplant patients that had their heart transplants delayed, so had the device implanted for a much longer period of time, and how did those early curves fit with the early curves for a similar time frame in the patients in the REMATCH study who were randomized to the device?

DR. EWING:  I was just going to say that he could answer, I would recommend that he could answer the question, but you need to remember that we have not read the transplant the data.  So, we can't tell you our opinion on that data.

DR. COMEROTA:  Perhaps it is not an appropriate question for the purpose of the deliberation of this panel.

DR. ROSE:  We have published on the subject of the long-term results of bridge to transplant patients, and there is an incidence of adverse events that I think is analogous to what we saw in the REMATCH data.

With regard to N's, though, that are out more than a year as a bridge, that N is pretty small as I remember.

[Slide.]

DR. ROSE:  There it is.  It is 104 patients.

DR. LASKEY:  It is of interest, but not germane.

DR. COMEROTA:  I would just echo my concern as to who this is indicated for.  That question was raised, so obviously, it need not be addressed.

The concern in terms of the potential patient bias and the Quality of Life was also raised, and I think there is an enormous potential bias because those patients who were randomized to the device are going to get a great deal more attention than those that are randomized to ongoing medical care, feel, well, perhaps I am abandoned and I don't want to live like this, so I will just die.

Well, obviously, that bias is there, and my final comment was since most--this is infection related--since most of the infections were in the drive line tract, are there any specific recommendations coming from the manufacturers regarding how to manage this drive line infection?

DR. LONG:  There are several issues that are very critical.  The drive line does play an important part of that.  One is the location of the drive line as it exits allowing for immobilization at the exit site.

One of the keys to success when we have looked at subsets within the REMATCH group, and again this is outside the data you have access to, but it will answer your question, is those who paid a great deal of attention to immobilization and proper exteriorization of the drive line had a very much lower rate of infection.

We also believe that the stiffness of the percutaneous lead has been a factor, and that is being addressed with a lead that is now being introduced that is more flexible and less prone to trauma at the exit site itself.

There are also measures being looked at in terms of antimicrobial impregnation.  So, this is an area clearly that is very, very important for the future, that has to be considered strongly when you look at attempts to correct infection in this patient population.

DR. STEVENSON:  If I can just interject, with regard to the issue about attention, we agree completely that we cannot eliminate the patient bias in terms of how they feel about having a device or not.

On the other hand, we controlled what we could, the frequency of clinic visits, the frequency of telephone contact was virtually exactly the same per patient month of follow-up in the two groups.  We were very careful to try to achieve that.

DR. HEITJAN:  I am Daniel Heitjan, the REMATCH statistician from Columbia University.  I have no financial interest in Thoratec although they are paying my expenses to attend the meeting.

Dr. Domanski had a question about the stroke and death.  Everyone who had a stroke died, so if you consider stroke or death an event, it doesn't change the event rates compared to the death rate.  So, the death fractions are exactly the same.

DR. DOMANSKI:  So, you had no non-fatal strokes.

DR. HEITJAN:  I believe that is not the case, but they all eventually died.  Some of them died from other causes, they may have died from sepsis or heart failure

DR. DOMANSKI:  That is interesting.  Okay.  Got it.

DR. LASKEY:  I mean we have to live by what they are adjudicated as.  Isn't that the key question here, you had an independent committee that assessed etiology and causation, and so forth, so if I understand you correctly, every stroke was fatal--

DR. HEITJAN:  No.  I believe the question was, Dr. Domanski's question was instead of considering an event death, let's say death or stroke as an event, does that change the picture at all, and the fact is that everyone who had a stroke in either arm died, although they didn't necessary die of stroke, they may have died of sepsis or other causes.

DR. DOMANSKI:  Yes, but it might well change the time to event analyses.

DR. HEITJAN:  That is correct.  I have never computer the Kaplan-Meier curve death or stroke, and I don't think TCI or Thoratec has done that either.  In principle, it could change it.

DR. DOMANSKI:  What I was trying to get at with that question is I don't buy into the Quality of Life data in this study.  Frankly, I don't think it is useful.  But I also don't buy into the death plus serious adverse event rate as it is defined.  I think that attacks the machine inappropriately because I don't think some of the serious adverse events are so serious.

So, what I was trying to do is pick out something that I knew was serious and look at it at that point, and I guess you have answered the question.  The bottom line is that if we approve this device, and you can go to somebody and say at one year, if you are willing to pay for this, I can buy you less death at one year.

I don't think you can say that for any other time. I don't think you can say that for two years really, I don't think, but certainly at one year you can.

DR. KONSTAM:  He hasn't really said that, because he hasn't told you when they died.

DR. LASKEY:  We are getting out of sequence here, so we will have a chance to reapproach this again after everyone has had their first round.  I think this is very, very fruitful, but we need to at least complete round one.

Have you at least delivered the totality--

DR. HEITJAN:  Yes.

DR. LASKEY:  Thank you.

DR. COMEROTA:  I have finished my comments, thank you.

DR. LASKEY:  Dr. Nissen.

DR. NISSEN:  I will try to keep us on time.

DR. LASKEY:  Well, after you, we will take a break.

DR. NISSEN:  Lynne, you characterized the medical treatment of these patients very well, but there is one thing I didn't see, and I may just have missed it in here, but AICD use.  I would like to know about AICD use in this trial.  What can you tell me about that?

DR. STEVENSON:  If you want to wait a couple minutes, I can give you slightly more thorough, but as I recall, there were 15 ICDs in the medical management arm at the time of randomization.  There were several put in after that.  I would have to look up the data.

Of interest, though, I think we have got to get back to how sick this population is.  In fact, five of the people who had ICDs and requested that they be turned off, so they would be allowed to die without any further ICD firings.  So, I think this is not a population in which sudden death, in fact, is our major concern.  We are more concerned with kind of the slow, miserable dying from heart failure.

DR. NISSEN:  Let me just take issue with you on that.  During the course of this trial, there was a lot of evolution of data on the benefits of implantable defibrillators, and remember now we are talking about very small numbers.  We had seven patients alive at the end of two years in the LVAD group, and we had three patients alive in the medical group.

All it would have taken was survival of four additional patients in the medical group to have equalized the actual two-year mortality.  I would argue that 15 patients, only 15 patients in implantable defibrillators represents a shockingly under-utilization of a therapy that has been shown to prolong life in these patients.

DR. STEVENSON:  I actually have the correct number here.  Fifteen had it at baseline, an additional 10 had it inserted, so that would be 25 patients in that arm.

DR. NISSEN:  It's a little closer.

DR. STEVENSON:  Again, I would emphasize that none of the ICD trials have addressed Class IV heart failure at all.  They have been specifically excluded from every ICD trial done.  So, we, in fact, have no data that ICD would prolong survival in Class IV heart failure, and certainly even less concept that it would prolong meaningful survival in Class IV heart failure.

DR. NISSEN:  Let me move on for the moment, but I do think that currently, patients with low ejection fractions are getting a lot of ICDs put in.  Wouldn't you agree that today, more of these patients would arrive into the trial with an ICD in place?

DR. STEVENSON:  They would arrive into it, but wouldn't necessarily get it if they didn't arrive with one.

DR. NISSEN:  But I am just saying it is a low utilization for this type of population by contemporary standards.

DR. STEVENSON:  I don't know.  In fact, the patients referred to us for transplantation, approximately 15 to 20 percent have ICDs in, which would be very comparable with what you see here.

DR. NISSEN:  I am also interested in understanding how New York Heart Association classification was determined and who assessed it.  Let me tell you why I want to try to understand this.  There is a divergence in the data with the best data being the NYHA classification for the device.

So, in the medical patients and in the surgical patients, who determine what their New York Heart Association classification was, how was that done?

DR. STEVENSON:  At baseline, that was done by the cardiologists at the time of enrollment, and certainly I think all the baseline qualifications indicate that Class IV was appropriate at baseline.

DR. NISSEN:  No, no, no, I am talking about during the course of the trial.  There was an improvement in classification in the group that got the LVAD, and what I want to know is how was NYHA classification determined during the course of the trial.

MS. TIERNEY:  I am Anita Tierney from the Data Coordinating Center.  I have no conflict of interest with Thoratec.

We credentialed both the surgeons and the cardiologists at the beginning of the trial.  The patients were seen on follow-up generally by both the surgeon and the cardiologist.  It was one of the two who determined the Class IV heart failure or the class of heart failure during the duration of the trial.

DR. NISSEN:  You understand why I am asking the question, and that is, that if I am a surgeon and I put an LVAD in a patient, and I am asked to assess whether their NYHA classification has improved, my answer may be very different from--I mean I think I would have been much more comforted if you told me you had an independent person not involved in the trial, and I think it is how you should have done the trial, is someone not involved in the direct care of the patient should have interviewed them about their symptoms and determined their classification.

So, to me, this process really doesn't validate that aspect of the results.

DR. ROSE:  But the Quality of Life aside from NYHA classification, which was not made by the patient, but is done by the clinician seeing them or the nurse clinician seeing the patient, the Quality of Life questionnaires themselves were not administered by the investigators themselves.

DR. NISSEN:  No, but you see the data there is very limited and much weaker.  The strongest data you presented on Quality of Life was New York Heart Association classification, and when I saw that, and I saw this big divergence between the Quality of Life questionnaires and the 6-minute walk, and all the things that might be quasi-objective, I said to myself I wonder if the surgeons weren't determining the Quality of Life, you know, the NYHA classification for surgical patients, and the medical people determining of medical patients, and didn't that, in fact, bias the assessment.  So far, nothing is addressed.

DR. ROSE:  Your point is well taken, and it very well may have.

DR. NISSEN:  I think we really have to not consider that data to be valid for purposes of this study.

I guess the other question I had was, were any patients deemed Functional Class I if they were in the hospital, could you be Functional Class I and be hospitalized?

DR. ROSE:  I would think not, but I don't feel like I can assert for sure.

DR. NISSEN:  I guess at some point, if there is an answer to that, I would like to know that.

DR. PINA:  You can have an infection and still be a Class I, you can still walk around and have an infection.

DR. NISSEN:  Again, I think it would be just helpful to me to understand that.

The final question that I had was what was the mean time to internal device failure and what was the mean time to any device failure for this device in this trial.  I saw the 3.4-year figure, but I would like to know what the mean time to internal device failure was for the use of the device in this trial.

DR. ROSE:  Well, something is still changing because there are devices in place.

DR. NISSEN:  Yes, but I mean all the other data you have given us, have given us as of a specific time.

MR. POIRIER:  We have not calculated that.

DR. NISSEN:  If it would be possible to determine that, it seems to me it would be useful to know, because if we are going to label the device, I mean if you are going to use a device, say, for an indication for a two-year period of time, it would be nice to know what the actual mean time to device failure is in actual clinical use as opposed to the data we heard from bench lab testing.

I mean you have that data now for a significant number of patients.

MR. POIRIER:  We are going to try to do that calculation at break.

DR. LASKEY:  Your timing is propitious.

I would suggest we break.  I have 2:30.  May we reconvene at 2:45.  Thank you.

[Break.]

DR. LASKEY:  Congratulations to all for keeping us on schedule.

We would like to continue on with Dr. Pina.

DR. PINA:  I add my congratulations since I know how very difficult this is to do, and I think that for once we have data on probably the sickest group of heart failure patients that have ever been looked at or studied.  I am amazed, as many of you are, at the incredible mortality rate of this population.

A lot of my questions have already sort of been answered, but I want to come back to the Quality of Life for a few minutes.  Another recent trial of a lot less sick group of heart failure patients has equated their lack of worsening of Quality of Life, and there is a difference obviously between improvement in Quality of Life and lack of worsening to being out of the hospital.

So, my question is, in the patients whose Quality of Life was able to be assessed at multiple times, were those most likely to be the patients who were out of the hospital when those Quality of Life assessments were done?

DR. ROSE:  The completeness of the patients surveyed Quality of Life data actually was much greater in the VAD patients compared to the OMM, and they also were the ones that spent more time in the hospital, so I don't think it is the issue of whether or not you are hospitalized that determines whether or not we had completeness of data.

We do have some data as to the reasons that we didn't have responses to the Quality of Life data.  Those we have in a back-up slide, and we can show.  Before, as they bring up the slide, most of that data that is missing is in the OMM group, and a typical reason was that they were too sick to want to respond to the questionnaire or to come in for a visit.

DR. PINA:  You also wonder how much the depression of not having been chosen to get the device counted in it, because I also noticed that there were 12 patients who withdrew care in the medical treatment group or asked for withdrawal of care, and at least five of those were within 30 days of randomization, so you wonder how much the Beck would pick up the depression of not being chosen since I think there is the perception from patients that you are doing something to them rather than just sticking in an IV and sending them home.

DR. ROSE:  I don't know if we have Beck data at 30 days, do we?  I think the earliest that we had the follow-up was at three months, but your point is very well taken.  I think there was in some patients a sense of letdown, but I would say also in some, there was probably a sense of relief.

DR. PINA:  It depends upon how you look at it.

DR. ROSE:  That's right, and patients have a range of subjective views,  you know, going forward.  Some are going to want to do this, and some are not, and the way it is perceived is either as a victory or a loss.

There is some element we found of cognitive dissonance in patients, you know, who would I got randomized to the better limb.  That was their own perception.  That was keeping them going with the sense that the process worked for me, and I got the better limb, and that could be either.

DR. PINA:  Just in concept.

DR. ROSE:  Yes.

DR. PINA:  My other question relates to the implant element replacement of which there was really one patient alive, so it seems that that complication of the device failure seems to be the greatest.

How many of those were in the SNAP version versus the previous version?

[Slide.]

DR. ROSE:  Apparently, there is 30-day data on Beck depression.

DR. STEVENSON:  [Off mike.]

DR. EWING:  I am sorry.  Could you repeat that into the mike, so we can get it into the transcription.

DR. STEVENSON:  Looking at the one-month score, in fact, the OMM arm was less depressed at one month than baseline, which suggests that the randomization outcome did not significantly increase their depression.

DR. ROSE:  Going from 16 to 13 on the Beck compared to 16 to 10.

DR. PINA:  It's a small change on the Beck, but nonetheless.

Do you have that data about the SNAP, how many of the device failures, the implant element failures were?

DR. ROSE:  I don't believe we have that data immediately at hand, if it's obtainable.  I think your point, though, about reoperation regardless of whether it was the VE SNAP or the other device is a highly risky procedure at this point, but again, the world experience now is 8 in REMATCH plus a small N, as well, in the bridge to transplant experience, it's a major reoperation.

DR. LONG:  And that is an area that is subject to evolution in the management of that.  While there were 7 out of 8 in the REMATCH trial that died with device replacement, there has been experience outside that now that suggests by doing it entirely extra-thoracic approach to device replacement, not doing a median sternotomy, you can see much better results.  In fact, in one series, 6 out of 6 patients survived to do well with hospitalizations half of the length of their initial implant experience.

DR. PINA:  I think two items that I know that we don't have the data for the submission, but the issue of malnutrition in this population, being unable to handle another invasive procedure added to whatever the immune response is.

I wonder if you looked at the women separate from the men, did the women have more septic complications?  You have a very small number of women.  My sense is that the women's immune response is very different particularly if they have had children.

DR. ROSE:  We are in the midst of a multivariate analysis of all of the potential demographic factors, procedural factors, process issues for death and adverse events.  My suspicion is that few of them will shake out as significant determinants just simply because the N is too small to make that judgment.

DR. PINA:  That would be an interesting point to look at.

DR. ROSE:  We are in the midst of the analysis.  We do not have that data here.

DR. OSSORIO:  First of all, I would also like to add my congratulations to those who sponsored and conducted this trial.  It is really clear that you are certainly adding to the knowledge base that we have in this very important and under-studied area.

My first question has to do with DNR orders, particularly in the Medical Management population.  It seems to me in looking through the case reports that a number of those people had DNRs, and that could certainly influence the time to death in the sense that those people, on their first event in which they had a DNR, and since most of them died of heart failure, this is going to influence the results.

So, I am wondering how many of those people had DNRs and in how many of the instances that was really relevant and how many of the people with the device had DNRs.

DR. STEVENSON:  In terms of the medical arm, there were five patients during the index hospitalization who had DNR at some time during that hospitalization.  I don't remember exactly, I believe two or three of them died, the other two went on to hospice, but this, in fact, is I think indicative of the overall feeling of patients with heart failure, which is that frequently, at some point, during a hospitalization, they decide there are certain things they don't want done again, whether that is intubation or defibrillation.

In many instances, however, DNR doesn't necessarily mean do not treat, it just means they don't want to be resuscitated, and the efficacy of resuscitation in this group when something happens in hospital, is in general the outcome is pretty poor.

DR. OSSORIO:  I think that is right, but I think that could just add to a particular bias, which is that for those patients who are on medical management, they are not perhaps so interesting to medical staff, and there already is I think good evidence to suggest that resuscitation is not very effective in these patients, and physicians often just don't want to do it, right?  So, the effort might be pretty slender on those who are just on medical management, whereas, those who have the device are much more interesting to the physicians around them, and the sort of effort put in to keeping them alive may be much greater, and also because the patients on the medical management arm were dying particularly of heart failure, where the DNR really is relevant as opposed to the ones who are dying of other things, like sepsis, where that is a little different issue because it is not a do not resuscitate, so that is why I was thinking that could have actually contributed to--I guess I am wondering with respect to efficacy, whether the difference in time to death really can be attributed to the device as opposed to these other things that might be related to have the device in you or not.

DR. STEVENSON:  First of all, it is only 5 of the 61 who requested DNR, so it is a small number during that hospitalization.  With regard to the intensity of their therapy, certainly when a patient finally comes to a DNR decision, we begin focusing more on comfort issues than a prolongation of life.

In fact, however, that is what we are treating all the way along are comfort issues.  With regard to the intensity of therapy in this population, I think all of the cardiologists involved felt extremely committed to provide the patients who got randomized to medical therapy with the optimal medical management that, in fact, we had promised them in that arm.

All I can say is that one has to at some level have some faith in the integrity of the investigators.

DR. OSSORIO:  Oh, I am not actually questioning anybody's integrity.  I think it is just a cognitive bias that ends up coming into play.

I would also like to sort of second, I guess, Dr. Domanski's interest in knowing the time to stroke, as well as the time to death, and in part because I do think that to assess efficacy, we have to decide, okay, whether the device, if used as intended, really produces clinically significant results.

That must mean something more than just keeping a warm body around, and so from my point of view, the Quality of Life measures really have a lot to do with understanding whether or not we have achieved something clinically significant and the sense of whether those adverse events were really--how many of them were so adverse that they could really have quite a negative impact on somebody's Quality of Life.

I won't belabor the issue of the weakness of the Quality of Life data that we have, and I commend you for doing what you did because I know it is very hard to do, and there aren't baselines aren't out there, and so forth.  I guess for me, I have a hard time really assessing whether we have achieved something clinically significant without those data.

That was just a comment.

Also, this is perhaps also a sort of seconding of something that has already been said, is that it seems to me that the indications, if this were approved, there would have to be something much more clearly stated about the indications for which it would be used and in which patients it would be indicated, because it seems that the reasonableness of the risks in a situation like this, where you have got patients who are very sick, the reasonableness of the risk depends on the situation the patient is facing.

Nobody is going to say that having this implanted in you is not very risky because it is.  It is just that it may be worth it if your other alternative is even a more likely death.  But if it were approved with very sort of slender guidance as to which patients were appropriate for this, it may be used in a much wider variety of patients than it was in this trial, people who are much less ill and perhaps have other alternatives, or in people for whom the reasonableness of the risks might not be as compelling as it might be in a case where somebody has no other alternatives.

Certainly, I would want to see more about that.

Also, in terms of sort of labeling and guidance, both the physicians and patients, I think particularly for patients, the manual was very technical for one thing, and it wasn't clear to me whether you did anything to assess patients' understanding or whether they even read this manual, whether you did anything to assess their understanding of that, whether you did anything to assess their understanding of the training that they received, and, in particular, with respect I think to infection control, that might be very important.

Did you do anything to assess that?

DR. ROSE:  I think one of the lessons that we learned in PREMATCH is that these patients are so sick, that the original Quality of Life battery that we had for them, they couldn't endure even those levels of measurement.

So, to say that we were in a position to test their cognitive function and say how much of the manual did you read and how did you score on a test, I hope the investigators would have done better on such a test than the patients would have, but it is not something that we added as a test to this.

I don't think there is any question that the labeling on this and beyond the labeling, the interaction between clinicians and patients, it is going to have to happen, and it does happen over this kind of therapy, is intense, very individual, and has to happen before a decision like this is made.  If it is not, I don't think that a standard of care in general for a terminal patient is being met if those types of options are not being discussed.

Whether or not that is an issue that can be addressed in labeling, I personally doubt it.  I don't think that is going to have the impact.  I think the major issue is really whether or not the panel has faith enough in the community of physicians doing this that this type of discussion will indeed happen.

I am certain that this happens in heart failure programs and in transplant programs, which is the kind of group that we had doing this kind of study, I think it is reasonable to say, will that continue if it becomes more generalized to other types of institutions, for bridging, for example, centers had to be trained in the use of the device.  I think it is perfectly reasonable to say that part of the training ought to be a patient education piece around what patients, you know, ought to be told in this context.

I think those are issues for going forward.  We don't have data on the issue of how much they grasp.  Unfortunately, I think they are not in a position to grasp as much, and there are profound ethical issues that that raises in people that are so helpless that often they are going by instinct rather than by analysis.

DR. STEVENSON:  I would just like to address briefly your previous point, which I think is the most pertinent one in front of us, which is who are the best patients to benefit from this therapy, and there is no easy answer to this.

For this trial, in addition to the criteria we had here, which were really very general, there were gatekeepers, there were three of us, one of whom reviewed every single case and often made recommendations for further therapy before randomization, particularly early in the trial.

Our best precedent may be looking at what happened with cardiac transplantation.  Over the experience, we gradually assembled a set of contraindications and indications for cardiac transplantation.  Regardless of how rigorous these were, however, they really could not employed except in the context of people with extensive experience.

I don't think one can come up with a labeling set of indications that could be delivered divorced from the concept of expertise in the clinical hands of the physicians advising the patient.  So, this is clearly a major issue, there are no easy answers, but we certainly can rely on some of the lessons that we learned from transplantation.

DR. OSSORIO:   I also had just one other question, which has to do with recruitment of people into this study. I noticed that 90 percent of the patients in the study were white.  It seems to me that if this goes out onto the market, well, at the very least we know that there is a higher rate of diabetes and serious complications from diabetes in people of color, and there may be other reasons why people of color are less likely to get transplants in general, at least with kidneys.

So, it may be that there would be a disproportionate number of people of color who might actually be at least potential patients for this kind of a device supposing they could get it paid for.

Do you think that there is a problem?  Do you think that if is widely used among people of color that there would be any differences?

DR. STEVENSON:  I think at this point, based on at least in the medical management, you know, 61 patients, and 68 in the VAD arm, we clearly have tremendous amounts to learn about all the different subgroups.

Minority enrollment in this trial, in fact, is very comparable to what happens in most NIH trials in which minorities are under-represented, a specific focus of the NHLBI at this point, but we all recognize that there is a lot of information that will come in the future as well get more experience with this.

DR. OSSORIO:  I am finished.

DR. DeWEESE:  Thank you for the excellent presentation.  One quick question.  Of the patients who did not receive the implant, how many of them had DNRs before they had their procedure?

DR. STEVENSON:  That was five during the initial hospitalization.

DR. DeWEESE:  No, the five were for those who did not get the implant.

DR. ROSE:  I think I misunderstood the question.  The people who did have the device?

DR. DeWEESE:  Yes.  Did any of them have DNRs prior to the procedure?

DR. ROSE:  Prior to procedure, I don't believe any of them had--the procedure itself in many respects, you could consider a resuscitation.

DR. DeWEESE:  That is what I wondered.

DR. ROSE:  But late in the course of the trial, there were patients on device with complications that made that conclusion.  We certainly were committed to them enough to want to resuscitate them, though, perioperatively.

DR. DeWEESE:  Infection is a real problem obviously, 17 out of 65, about 25 percent.  How about in the studies to bridge, how many of those developed infections?

DR. ROSE:  Those rates are comparable in the bridging experience with this type of implantable VAD.

DR. DeWEESE:  That would argue against what you said, that the cause of death might be due to malnutrition.

DR. ROSE:  Well, the rescue from infection in the bridge patients, though, is transplantation, where if you can suppress the infection and get a donor heart, you can take the infected device out and put a biologic replacement in instead, which is a definitive solution in general to the problem of infection.  That is not an option for patients here.

DR. DeWEESE:  Device failure was listed as a cause of death was only two patients, I think, or none in one of your slide, two patients in Table 14, a summary of the cause of death.  There were only two who were attributed to device failure?

DR. ROSE:  Two devices which could not generate a cardiac output.  That was the strict definition of device failure.  Can you bring up that slide, my causes of death slide?

[Slide.]

Three deaths in patients perioperatively at the time of device replacement for a component failure of the device.  For example, if these there are a bearing failure and the patient were being pumped pneumatically, which this device allows, as opposed to using the electric pump, you can see those.

There were three patients who died at the time of operation or a VAD replacement, where using a strict definition of the pump is not pumping was not the case because the device in which a component had failed was still working well enough to sustain the patient to get to the point of being able to replace it in the operating room.

But these are devices in which we knew it did not pump and that the pump was explanted and it was found to have failed.  That was the strictest definition of VAD failure.

We showed you before there were device malfunctions.

DR. DeWEESE:  I know there were a high percentage of that.

DR. ROSE:  One definition, you could say that inflow valve insufficiency is device malfunction, that generally allowed time enough to replace it.  By some constructs, you might say that is a device failure because you need to replace the device, but these are the malfunctions, 70 were confirmed, as we said, 50 external components, 20 internal, the controller most commonly followed by the inflow valve.

Actually, the mechanism of failure in the in-vitro studies was something that we saw rarely in the trial.

DR. DeWEESE:  But the cause of death was 6 device failures, is that correct?

DR. ROSE:  The nomenclature that the M&M committee has used is root cause of death, and, yes, it is more than you two patients, I think you could say it is probably in the range of 6 or 7, the root cause of death was pump failure and the need to replace it.

DR. DeWEESE:  How about the patient who refused another procedure where it had failed?  Where was his death put?

DR. ROSE:  How did we adjudicate that?  This is a patient who is a candidate for device replacement because of an inflow valve, for example, failure, and then elects not to have the reoperation.  How was that death classified?  The cause of death would be, what?  That is what Dr. DeWeese is asking.

DR. DeWEESE:  That's all right.  There could be an additional.  We are getting more.  How about the patient who went ahead and had the implant and died postoperatively within three or four days?

DR. ROSE:  Those are these three here.  We are looking to see what that death was classified as ultimately. Actually, those may be the LVAD dysfunctions.  One of these is a pump that is turned off.  There are two deaths from heart failure in the VAD group, and that is where those are.

DR. DeWEESE:  I just wanted to clarify.  I had the impression that there were none, there are 6 or 7.

DR. ROSE:  Yes, in that range.

DR. DeWEESE:  The other is a question about the air venting.  Did that cause any infection problems?  I know it is not touching the blood, but did the passageway for the air venting get infected or not?

DR. ROSE:  Well, to know whether or not the channel for infection was either through the air vent or around the periphery of the drive line, I don't know that we could discern.  One of the air embolism events, though, was in a lining rupture that resulted in air embolization with communication with the outside world.

The hypothetical you raised, that specific patient was classified as LAD dysfunction.

DR. DeWEESE:  Thank you.

DR. KLOCKE:  Eric, like everyone else, I recognize this is a landmark study, and I do want to express my admiration for the people who conducted the study and the way in which it was conducted.

In the discussion, I still find myself most focused on the infection issue.  I sense and realize that you are also.  Frankly, you have demonstrated a temporary survival benefit that clearly, by 30 months or something, has essentially closed.

My concern is, I am trying to assess and would appreciate your comment on whether the continuing infection is perhaps the dominant role in that.  That is, you have made so many advances with the pump and everything else, but it seems to me that that is perhaps the biggest unsolved issue, and that at least gives me the most pause in trying to decide how far are we ready to go forward.

I think the slide you had up there, in the first 40 deaths, there were 17.  That is 40 percent, and actually, as I read through the individual patient summaries, and obviously, the care is remarkable, but I have the impression and realize how adjudication of deaths goes on, but it seems to me that infection, continued serious infection played a prominent role in many of the cases where it was not a primary cause of death, for instance, if a VAD had to be removed because of infection and the died of perioperative bleeding, I have no quarrel with the classification.

So, I am troubled by that, and I presume from the course that the infection business wasn't just in the first 30 days.  In terms of the Kaplan-Meier curves, that had to be a continuing problem over time, and I don't know if you have any data.  Those are the first 40 deaths.  I realize you have more.  I am not sure if things have been adjudicated, but there are now 65 deaths, if I understand correctly, in the LVAD group.

DR. ROSE:  No, there are still 18 ongoing.

DR. KLOCKE:  Good.  So, there are 50 deaths.

DR. ROSE:  Right.

DR. KLOCKE:  So, in the first 40, it was 40 percent, and I guess the infection rate is continuing on.

DR. ROSE:  Right.

DR. KLOCKE:  Jim asked you about the bridge to transplant.  I guess I would have liked to hope that with the benefit in LV function, which I am sure the LVAD patients got, that apparently, that benefit in overall, even those these were terribly sick patients, but that benefit apparently is not sufficient to reduce the infection problems.

So, that is the one that I am sort of left struggling with the most, and I sense that you are.

DR. ROSE:  I think we are struggling, but I don't think we are in the dark at this point as to what it is that we need to deal with.

DR. KLOCKE:  Okay.

DR. ROSE:  As I said, malnutrition was not something that was even on the investigator's radar screen when we began this, because our assumption was that if you fix the cardiac output, the nutritional depletion of these patients would reverse itself.

I think it is clear to us at this point that that is not the case, that simply restoring an adequate cardiac output in a nutritionally depleted patient may not be enough, and there was no particular attention paid to preoperative nutrition, as well.

DR. KLOCKE:  I hear you, and I guess in the bridge to transplant patients, the same nutrition issue occurs?

DR. ROSE:  Except for the sense that the entry criteria here was initially that you had to be sick with heart failure for three months, which actually I hope in the labeling will not be there because that allows an awful long interval for people to really waste away.

DR. KLOCKE:  I hear you.  I guess the question I am sort of left with is how strong is the information we have at the moment to think that the current infection rate, which I see at this stage as the dominant limiting factor, and we have ideas how to address it, but we have limited experience in addressing it successfully.

DR. ROSE:  I want to sing Jim Long's praises.  There are two centers in the trial that have taken that on as a very specific area of interest, and they have offered a lot of guidance to us, and now with a small, but very encouraging experience with us, I think will become a model for the rest of those of us involved in the field.

As well, I think the drive line in the VE was too rigid, and is a source of contamination because the inflexibility of the drive line would crack the junction between the device and the skin, which a softer drive line, which the newer iterations have, we believe will also help deal with this.

DR. KLOCKE:  I am confident that you will and we will go out and we will solve it.  I guess the problem that we have is at this point in time, where do we stand in terms of having satisfactory evidence to approve the request that has been made to us.  That is really what I am struggling with.

DR. ROSE:  The issue there is that even in spite of all the infections and all the other adverse events, there is still a highly significant statistically valid survival benefit to the device, so in spite of all these troubles, these people are still living longer, and for those of us who do believe the Quality of Life, they are feeling better, as well.

That, I think is the overwhelming evidence, and I think the very reasonable expectation that these results are the floor of what we are going to be able to achieve going forward.

DR. LONG:  Some evidence of that is a shift in infection rates between the precursor to this trial, which is called the PREMATCH trial, and the REMATCH trial, and there was a very significant reduction in infection between those two with efforts to try to improve management.

Now, that is limited in its data, but is one more inference that management plays a significant role in helping control this.  In essence, in considering the etiologies of this, it is device related, patient related, or management related, and our sense is that it would be unfair for us to place the burden of this entirely on the device.

There is a very substantial portion of this that is related, in fact, to management and to the patient especially in that early upfront period while we have so many comorbid factors ongoing.  A very important topic, though.

DR. LASKEY:  Mr. Dacey.

MR. DACEY:  I want to salute everybody involved with this because I don't think, at least when you get out in the general population, that people will fully capture in their daily lives how much the landscape is changing and how much work that you are doing is contributing to that change.

When I first read over this material, the first thing that occurred to me was is this a heart prosthesis, because if it is, this also is a change in landscape and understanding.

If my prosthetic leg malfunctions, I can still hop around, I can use crutches.  If a heart prosthesis malfunctions, then, it is a much more serious problem.  I know you understand that, but speaking as a consumer, as a representative of the consumer, trying to capture these ideas in ways that are very understandable to them is part of my mission.

Also, I see this remarkable change, both quantitatively and the velocity of change in the body of knowledge especially around something like CHF.  There is a few of us in this room who remember this, which was the ACPR guideline for heart failure, which is only eight years old.  This is just eight years.  That is not a long time really, yet, in some people's eyes, it is an eternity.

But at the end of the algorithm for end stage, refer for evaluation for heart transplant, and now you have got to add a couple more steps.  There is no reference there to bridge to transplant.

Then, also, on the patient handbook, and I know this is a work in progress, there was a time in patient education and information where the pass of distribution of information, words on paper, was enough to satisfy legal requirements, informed consent requirements, and I have heard very reassuring words today from the professional perspective there is a great deal of intensity at trying to work with patients, because these are very, very sick patients.  I have known some of them.

But as we know, if you are going to get people to the point where they understand, and it is not just the patient, it's the family, it's the caregivers, it's the people around them who have to take care of them, they have to have skill training.

When this was published, 24 pages of fifth grade text, now we are seeing 50- and 100-page technical documents that are being given out to satisfy this requirement.  There is no substitute for that one on one, and I know you know that.

I hope if this should be approved and moves out across the country, that all the colleagues involved understand it, as well.

Finally, as I end my four years on this panel, there is just one editorial comment I would like to make. Having been an advocate for the subject of Quality of Life for such a long time, going back a very long time, and maybe my age is showing, but I sure wish we could find another term.

I know you have to test and measure it, but it has become a marketing term also, and I sure would hope to see it convert to a marketing term within the context of this subject.  I understand all the methodology involved, SF-36, and the well life expectancy model and the general health policy model, and I have worked with them, and it's a daunting task, and it's horribly subjective.

You are at the cutting edge of some new work here also, and for this I salute you.  So, in summary, my perspective says you have got a heart prosthesis, and it is part of a process that is going somewhere, and I don't know fully yet where that "where" is, but I salute your effort.

DR. ROSE:  Thank you very much, appreciate it.

DR. LASKEY:  Mr. Morton.

MR. MORTON:  Thank you.  Thank you, Bob.  Thank you, Dr. Laskey.

Just to follow up, Mr. Dacey, on your comments about this being a heart prosthesis, I would like to bring up something that was discussed during the break, and something that is of interest to me, and that is, we have heard several comments that the device, as presented to us, does not have an end of pump life indicator on it.

I would respectfully like to remind the panel that many other heart prostheses and critical devices also do not have an end of device life indicator, and I would ask us not to view that lack of such an indicator for this device as a negative, but review it as tested, designed, and presented to us today.

That ends my comments.

DR. KNAPKA:  I am supposed to be representing the patients, and as a heart patient myself, I do have a defibrillator, a pacemaker like Vice President Cheney, you know, both waiting for them to go off.

I just have to make several comments.  I think from a patient's standpoint, strictly not so much as a scientist, just as a patient, I would probably be scared.

We talked a lot about first effects and then Quality of Life, and to me this was kind of contradictory because it was saying with this implant or with this device, there is all these adverse effects, and then we say, well, this is improved Quality of Life, and to me, that is a little contradictory.

I mean maybe when you first start, and as you said, you are very sick, and you are glad to have something, but you go through all these adverse effects, and then maybe that last year, maybe the Quality of Life is good, maybe there is a balance there.

Now, that being said, I think it is good that many, many of us patients would be very glad to have another two years, praying that there is a new development that is going to help us.  So, from that standpoint, I think again this becomes a very, very individual decision.

I know I was very sick, and quite frankly, I didn't care if I lived or not, so maybe in that case, those people, they don't want to live, and that is their choice, but I think it is good that people have a choice to maybe have two more years, and, gee, maybe NIH is going to come to something that is really going to be top-notch, and I will live the next hundred years, which may or may not, probably not, because even some of the people that Jesus saved died, too, you know.

[Laughter.]

DR. KNAPKA:  I would like to just make one other comment about the patient handbook, and I would agree it is very technical, but I think if I were a patient reading it the first time, every other page there is warning, caution, that would just make me more scared, and I think if you redo it, and I think it is very honest, don't get me wrong, I think it is very honest, but I think some of these maybe could be tuned down, and I would agree that probably there should be a training session with the physician and the patient rather than just hand it, you know, here is your handbook.  It is really highly technical, and I don't think most people would know.

I got a lot out of it.  The other one thing that surprised me, and again I am going to talk a little bit as a scientist, and not as a patient, but this issue came up that there wasn't much until now, much concern about the management, the 20 different centers that were involved, they all seemed to have different management this and this, and to me, that is Statistics 101.  You identify all your variables and measure.  I am a little disappointed in that because to me, management of any disease, any experiment is very important, and I would think that would be, as I say, that is Stat 101.  You identify your variables and start considering it.

Thank you very much.

DR. LASKEY:  Thank you, sir.

I think before we get on with the nitty-gritty, I think there has been additional food for thought and some people have some follow-up questions, so I would like to limit that to perhaps five minutes per person, also, with the hope that maybe some of the information that was requested before the break is available.

Janet.

DR. WITTES:  I actually have one other thing that I think we haven't really addressed very much, and that is the distribution of time to critical failure.  What we have is information about estimated mean and estimated median time to failure, but not a distribution of time.

I wonder if you have that both in vivo, especially in vivo, but also in vitro, so not just the point estimates of failure.

MR. POIRIER:  We don't have all our data here.  We will be looking at it, and we can certainly get that information back to you, but we don't have the data that we can come up with any mean time to failures in the clinical environment.

Also, I think FDA frowns upon look at clinical data for reliability because it is uncontrolled.  Every patient is different.  If you look at reliability, you really want to look at a controlled environment where you control the parameters, and you can compare one device to another.  You cannot do that clinically.

So, that is why we have not made those calculations because FDA frowns on that.

DR. LASKEY:  Mike.

DR. DOMANSKI:  No.

DR. LASKEY:  Dr. Konstam.

DR. KONSTAM:  I have just a few things that I want to touch on that I don't think we have gone into at any length.  One is some help about indications for use, what patients are we talking about is one topic, and the other topic is, Eric, you mentioned training, and I am going to want to pick your brain about what ideas you had for that.

With regard to indications, I guess I have got some deep concerns on both sides of utilizing this device in patients who are not as sick as you have in this study given the mortality that you see with the device at two years, we had better be reserving the use for patients who are really sick, really have a very, very limited life expectancy.

I am worried about the other side, too, but, first, let me ask you about that one.  How would you propose this indication be written, so that we are sure that patients who get it have a very limited life expectancy?

DR. STEVENSON:  Well, as I said before, this is probably the most pertinent question that I think some of us are focused on at the moment.  I would suggest that we perhaps add language that says "refractory to all current standard therapy," and in addition, including the proviso that patients be ineligible for cardiac transplantation implies, in fact, that they will be evaluated for cardiac transplantation by people who are accustomed to do that.

I am hopeful that perhaps including both of those, refractive to all currently accepted therapy and ineligible for cardiac transplantation may, in fact, help us to get to the severity of illness here, but I agree it is a major challenge.

DR. KONSTAM:  Do you think that will do it or do you think you would want to be a little bit more rigorous at defining what that means?

DR. STEVENSON:  It is exceedingly difficult to be more rigorous.  I think judging from our experience with cardiac transplantation, we can certainly put in some of the constraints that we put in for transplantation, we can think about that.

If we look at the indications for entry into this trial, Class IV, EF under 25 percent, peak VO2 less than 12, we all know plenty of patients like that who, in fact, would not have a 75 percent, two-year mortality.  So, I think those may be helpful, but they are not adequate yet to get the severity.

DR. KONSTAM:  I don't think we are going to wordsmith it now, but certainly there are guidelines for indication for transplant, and I just leave it with the comment now that I think that if we are going to approve this, I think it will, in my mind, require some work to define, to make sure that patients who don't have, based on what we see here, an extremely limited life expectancy, don't get this device.  Let me just say that, and say we are going to need to work on that.

The other side of it is a more societal question of putting it in people with comorbidities or advanced stage, and that can be more difficult to deal with.

Let me just ask, because it is sort of interesting, one of the entry criteria here was comorbidities.  I mean that was a permissive entry criteria. You had to not be a candidate for transplant.  That is one of the ways you could not be a candidate for transplant.

So, it certainly didn't exclude patients with, quote "comorbidities," but do you want to comment about what type of patients and what type of comorbidities you would not want to see receiving this device clinically?

DR. STEVENSON:  At this point, I don't think I want to list them for the record.  I think we all agree that there are a number of organ system degrees of failure that we would not want to see.  For instance, part of the success of the bridge to transplant program has been, in fact, that people have to be eligible for transplant to get into the bridge, and I think that some of those criteria, for instance, in this one, you couldn't have a creatinine over 3.5 to get into this trial.

I think that we will end up with some organ system function parameters, and I would, in fact, propose that hepatic function be a very important factor in here, but I don't want to actually suggest specifics at this point.

DR. KONSTAM:  Who is going to do that, though, do you think that is something the panel is going to be able to do?

DR. STEVENSON:  I think this is going to be a work in progress along with the patient manual, will be how indications should be specified, and I think it is impossible to divorce it from the people who, in fact, will be evaluating these patients, and I think that those should be people with expertise in evaluating end-stage heart failure.  I don't think there is any way we can write something out here that we could give to someone in the community and have them make this decision.

DR. KONSTAM:  Okay.  The other big concern I have is who is going to be putting these in.  Let me just preface it by stating the obvious, that the results that you have here with--I guess I can't get away from it--a less than anticipated survival rate in the treatment group, nevertheless were achieved with extremely expert investigators, all of whom I guess had experience with this device as a bridge, I assume, so maybe I could just open that to your comments.

How do we assure ourselves that we are going to have this device used by clinicians who are capable of achieving at least this same degree of success with it?

DR. ROSE:  I think the best way to assure that that is happening is with postmarketing surveillance, to know, we can all want that to happen, but I don't see how you can know that without doing it.

DR. KONSTAM:  You wouldn't propose some kind of a certification process, a training program?

DR. ROSE:  Absolutely.  Aside from that, even for bridging, the use of these devices requires mandatory training for centers, and even retraining for centers in the use of this device, and all of them now I would expect are going to participate in the ISHLT registry of devices, as well, which is the community's attempt to--

DR. KONSTAM:  Not to take too much time on this, because I think this is a concern to me, and may be to other people on the panel, could you just expand a little bit, what might be the nature of a training program that people might have?

DR. ROSE:  It is generally two to three days of going to an expert center where a device will be implanted in an animal typically, or two animals, as well as didactics.  We teach several of these courses at Columbia, and there is a curriculum, a syllabus, indications, process of management to the patients, and the ability to participate in an animal implant with experts in device insertion is all part of the course.

Frankly, after the course, there is often a lot of dialogue, as well.

DR. KONSTAM:  What about patient experience, what about actually assisting or participating in actual patient operations?

DR. ROSE:  Why don't you speak to that for the bridging.

MR. POIRIER:  In the bridging program, we made the determination right upfront that no one would be allowed to use this device without going through a formal training session.  The training session consists of two parts.

One is that we have a training center, and Columbia is one, where the physician and his team go to Columbia for two to three days, they implant the device in two different animals, they go through all of the issues in terms of patient management, patient care, patient selection, all of those things.

In addition to that, the company sends to the specific hospital that wants to use this, a team to train the nurses and the people who will take care of these patients, go through the operation of all of the systems with the engineering people on site, so that everyone is thoroughly trained.

There are manuals, operating manuals, patient manuals, a whole variety of different manuals that are used in this training.  So, nobody will touch this device until we are convinced that they are adequately trained.

DR. KONSTAM:  And in terms of other elements at the site, you know, in terms of the heart failure care or the ability to select patients, any thought about that?  Should it be limited to certain sites or only certain surgeons?

MR. POIRIER:  One of the advantages that we have is this is a close-knit community, everyone talks to each other.  If there are any questions, people call each other and discuss it, and we have a whole network of people willing to do that.

If there are any issues on patient selection, there are many people who will discuss that.  I think the physicians here will back that up.

DR. KONSTAM:  Thanks.

DR. PINA:  But I think you have to extend that a little bit differently, because up to now you are bridging classes, which I am familiar with, have been for transplant centers because that has been the approval, and now if a community hospital wants to put this device in, if it gets approval, they are going to be able to do it.

My worry goes even before the surgical expertise of the people that are putting it in, it is the people who are treating the heart failure, and what we don't want to see in the heart failure community--and I am sure Lynne would echo this--is people being inappropriately treated for heart failure, and not being offered the therapy that sometimes is tough to do, but if you are persistent about it, you can get people on therapy, for example, beta blocker use.

It may be simpler to say, well, look, we now have this device approved.  So, I think that Lynne's point about somebody not being a transplant candidate means that they must have gone through some process of being looked at as a transplant candidate, for whatever the reason, comorbidity, age or whatever.

That is my bigger concern even before, because I think you can train a good surgeon to do this and gather experience, but I want to go one step before this.

MR. POIRIER:  We agree with that, we agree with Lynne.  Don't think that there is going to be an avalanche of implants tomorrow.  That won't happen.

DR. PINA:  But that is something that has got into your training thinking beyond what you have done now.

MR. POIRIER:  Yes, of course.

DR. PINA:  I know your programs right now, and they are terrific.

MR. POIRIER:  I mean as a company, we are concerned with that more than you are, because the results will be detrimental, and that will hurt us.  So, obviously, we don't want that to happen.  We will be very careful on how we let this out, and we will be very careful who gets it, and we will make sure that the people who are being evaluated are being evaluated properly.

We have a long track record of that.  We have not been careless.  I have been involved with this for 35 years.

DR. LASKEY:  We are not impugning your integrity either, but the nature of the marketplace is also a wild animal at times.  That is our concern, it is always our concern with these devices.  It is not all up to you always.

DR. LONG:  Clearly, responsible dissemination is utterly essential.  We would agree with controlling that and making sure that there is excellence involved in this, especially until such time as it is appropriate to expand the volumes with adequate experience.

I would like to add one other comment about the patient populations that this is appropriate for.  While we agree that this is a very high-risk patient population that should be receiving these devices, it would be unfortunate to constrain this field to serve up only patients that are very high-risk patients and patients who bring a burden, not because of the device, but because of their comorbidities to the process, so that we don't have the opportunity to improve the outcomes with these patients based on that particular feature.

DR. LASKEY:  Nevertheless, we need to evaluate what we have in front of us.  I understand you, and it would be wonderful if we had a distillation of what the gatekeepers went through, but we don't have that, and that is what we need, and I think that is what many of us are concerned about, is that that thought process has not been codified, it has not been translated into scalable covariates.

We have no idea who these patients are except for the fact that there was a 7 to 1 ratio between looking t them and putting them into this protocol, and that is not necessarily a generalizable study result.

DR. COMEROTA:  I would just echo Marvin's concern about indications, and I think inclusion criteria are one element that probably will be easier to identify than exclusion criteria, and I think it is fair to say that this panel would be very uncomfortable with defining that, and that is something that definitions need to be made and then brought to this panel.

I think many of those issues have been already enumerated, and some of them being societal, age, and are there going to be cutoffs, as well as other comorbidities.

I will just leave that as a comment.

DR. NISSEN:  I must tell you that I am terribly disappointed in your inability to provide mean time to failure data.  Let me tell you why.  We have a device here that if I read Dr. Swain's review, failed in 20 of the 68 patients with an internal failure, not an external component, but an internal component, that is a 30 percent failure rate.

Now, for us to counsel patients about whether they ought to undertake such an operation without being able to say to them, look, if we put this device in you, it has a mean time to failure of 12 months, and you didn't know that within an average of 12 months, you are likely to require replacement of the device, we have to know that.

The in-vitro testing data doesn't tell us that. Only the in-vivo data tells us that.  So, I think that we must know how long we can expect this device to function for in order for patients to make an educated decision about whether they want to undergo an implant.

I am told the FDA is not interested in such data, I don't know if that is true or not true, but I am certainly interested in knowing how reliable is the device in a clinical in-vivo setting.

Can anybody give me any insight into that?  I would certainly appreciate it.

DR. ROSE:  I think to argue that there is no insight from this based on the survival data is just not correct.  I think that patients, while they may be interested in failure rates and detailed failure rates, I think that most patients want to know even more how long can I expect to live.

That, unquestionably, I do believe we have very firm data about, firm enough that I think it is reasonable for a physician to make a recommendation or to advise a patient as to whether or not it should be considered.

Without the approval of this body, those choices can't be made out in the public, and I think it is time based on this data set, that those choices be available to patients.  The additional data, I think is desirable from the point of view of helping elucidate these issues, but from the point of there being critical to making a decision or not, as to whether or not this belongs out there for patients to benefit from is a separate question.

DR. ZUCKERMAN:  Dr. Laskey, can I just provide a point of agency clarification on this reliability issue?  I think the clinical question posed by Dr. Nissen is an extremely important one, and while the sponsor may have gotten the impression in the past that certain reliability calculation wasn't called for, et cetera, I don't think that that would be our current position.

In fact, I am going to ask Dr. Berman to better explain what we were trying to convey to the sponsor.  It is also the reason why we do have panel discussions like this one, just to better clarify what are the pertinent issues to, one, look at the device, and two, better capture information in our labeling.

DR. BERMAN:  There may be a misapprehension of a misunderstanding.  We would not accept reliability data from an animal study to demonstrate long-term reliability for a device.  That is, typically, in the past, people have done eight cows for three months, and that does not demonstrate long-term device reliability in patients, it just doesn't, and we don't like that, and we tell people we won't accept that.  We do want to see bench testing, there is no question.

We will certainly look at reliability data, data of rates of occurrence of different kinds of device malfunctions rate, time to occurrence of different types of malfunctions, and so on, as observed in a clinical trial.  We will not accept that as the only data.  We do want to see formalized bench testing.

Mr. Poirier is quite correct.  Patient data is somewhat uncontrolled.  You don't know under what conditions the device had a problem.  There were at least one or two instances in the REMATCH trial in which there was what appeared to be a device problem, but which we, in discussion with the sponsor, agreed was not, it was a patient problem. The device is not responsible for what the patient misuses.

So, to correct the misapprehension, yes, we want to know what happens, what is observed during the clinical trial, no, it is not by itself entirely sufficient, but it is very important, as in this case, to look at bench testing and how the clinical trial proves out that the bench testing was adequate or perhaps not completely adequate.

DR. NISSEN:  One other question.  I will ask the sponsor to comment on that.  The other relates to the fact that the most common reason for not being eligible for transplantation is being over the age of 65.  You suggested that there might have been some differences between how people did according to age.

So, I would be very interested in understanding whether there was a difference in the over-65 category.  I know it was most of the patients.  You suggested that the under-65 did particularly well.  Is the converse also true, if you were over 65, did you tend to do quite poorly with the device?  Is that is the case, doesn't that speak somewhat to labeling?

DR. ROSE:  We had three prespecified age strata in the survival analysis.  One was less than 60, another was 60 to 69, and the other was above 69.  In all of the age strata, there was a survival benefit of the LVAD arm.  Only in the 60 to 69 group was that benefit statistically significant, but that was also the largest group, so I think it was reasonably powered to answer that question.

The younger age stratum that I described, that difference was not different compared to the others.  I think as we accumulate more experience, that those kinds of data are going to be critical to deploying this kind of device.

DR. AZIZ:  Just a suggestion and a question.  Would it be reasonable to suggest that initially, at least for the next year or so, that the implantations only be done at centers that do transplantation?

DR. ROSE:  Excuse me?

DR. AZIZ:  That do heart transplantation rather than letting every community hospital be using this device.

DR. ROSE:  I personally think that dissemination to community hospitals at this point is not the way to go.  I am open to argument.  In particular instances, there may be a strong reason as to why a particular institution that doesn't do transplantation ought to have this available as an option, but in the early get-go, I think it is probably the wrong way to go.

DR. AZIZ:  The other thing, I know that obviously, our discussion is related to this particular device, but I do believe that I think other devices have been used for long periods of time particularly in Europe.  I don't have the numbers.  I think the Quality of Life at least for those devices, the patients have done quite well.

DR. PINA:  One small point.  We learn from clinical trials whether the trials are positive, negative, or neutral, and I think as you look at the demographic data, which you apparently have not done right now, you may be able to come up with a risk profile for the patient who would be more likely (a) to have sepsis, the patient would be more likely to develop a CVA, looking at, say, vascular disease.

I would be particularly interested in you looking at the body mass index, which you can probably calculate if you have the height and the weight--

DR. ROSE:  We have that data.

DR. PINA:  --and get some sense of muscle mass.  I mean it's a very gross sense of muscle mass to look at the rate of--if you don't have albumin or pre-albumins--to look at the rate of complications based on the muscle mass or nutritional status.

DR. OSSORIO:  I have two questions.  One goes to the informed consent issue, and you had mentioned that these patients are so ill that it is not as though they are sitting around reading manuals or whatever.

Did you do anything particular in this trial to try to ensure that the informed consent was adequate?  I am asking that question because I am trying to think about generalizing that, and thinking about perhaps unusual or special things that could be done in a non-research context, but that could help.

DR. ROSE:  Oddly enough, I think one of the confirmatory issues around informed consent is the fact that the ratio of screened patients compared to enrolled patients was so high.  I don't think that we pulled any punches with regard to describing to patients what it is that was entailed here.  Clearly, a large number of patients said with regard to a device, "I am not interested."

So, if anything, I think we bent over backwards and here the issue was clinical equipoise.  We had the appropriate degree of equipoise in our posing these issues to patients.

Early on, I think there were questions around--I remember the first investigators' committee, there was still question as to whether or not this was an ethical randomization.  I think we came to that conclusion particularly reassured when the DSMB looked at the first cut of data and just said to us, "Keep working."

That was enormously encouraging to us, so I think that we did have a reasonable degree of informed consent for the trial.  I think the nature of informed consent though now, if the device is approved, is a different issue with a lot of other considerations, particularly the issue of concerns around overselling the device, and also I think it reasonable to have concerns around underselling it, too, that patients who could benefit from it, as you mentioned before, patients of color that could benefit from it, that don't necessarily get it.

I think on both sides, we need to be particularly vigilant, and that is a challenge to us.

DR. OSSORIO:  Another question, which I don't know if this is exactly a fair question, this is more for general information.  Obviously, there are a lot of real societal concerns about investing tremendous amounts of resources extending to very, very end of life.

Did you and your company have any kind of an ethics discussion or particularly an ethics discussion about this that helped you to decide that it was a good thing to move forward with this kind of a trial as opposed to some other kind?

DR. ROSE:  I don't work for Thoratec.  The company I work for is Columbia University.

DR. OSSORIO:  Right.

DR. ROSE:  At my company, yes, we have considerable discussions around the ethics of doing this kind of dissemination.  I think at the other end of the spectrum, though, can a society as successful and productive as ours, afford not to do this ethically, I think is as good a question as whether or not we shouldn't.

DR. LASKEY:  Dr. DeWeese.

DR. DeWEESE:  I have no additional comments.  I would hope that your group would be able to provide a definition as has been asked for of just who would be accepted, but with your experience, and then it could be evaluated by the panel, if necessary, at a future date, and carried out.

DR. KLOCKE:  I am sure you will do it.  I guess I would encourage you to, if you could, one was saying codified, but you pointed I think correctly that the LDS in the Minnesota experience I understand with infection, which I have been focused on, appears to be different, and certainly anything you could do to codify that, to extend it to a larger group of patients if that really is a reasonable answer to the infection problem, it would be useful to, if the technology spreads, to be sure that other people don't go through the same learning curve that you have been forced to go through.

DR. LASKEY:  Thank you.  If there are no further questions, thank you, gentlemen, very much for a very persuasive and articulate presentation.

I am going to ask that the sponsors step back from the table at this point, so that we can go through the questions again.

Panel Recommendations

DR. BERMAN:  I am going to read into the record the questions we would like the panel to consider as they deliberate their decision for this PMA supplement.

1.  The bench testing performed to assess device reliability did not account for all observed clinical conditions, in particular, higher than expected pressure in the pump chamber and higher than expected beat rates. Accordingly, the observed times to device failure and/or device malfunction seen in the clinical study are less than those predicted by the reliability model.  As well, there is no reliable end-of-pump-life indicator.  Please discuss the clinical implications of the observed reliability.

2.  Are the device failure and malfunction rates and their time to occurrence appropriate for a device intended for use for destination therapy?

3.  Given the Kaplan-Meier survival curves and the fact that 7 device patients and 3 control patients, as of February 02, had survived to 24 months, have enough patient data been reported to demonstrate a clinically meaningful survival benefit?

4.  The New York Heart Association, the Quality of Life, and the functional testing results are not consistent. From these data, can we determine that there is a clinically meaningful improvement in functional status?

5.  This device demonstrated an increase in median survival time and showed an overall difference in survival. However, this benefit diminished at two years and was associated with serious adverse events and hospitalizations throughout the course of the study.  Do the benefits of this device outweigh its risks?

6.  One aspect of the premarket evaluation of a new product is the review of its labeling.  The labeling must indicate which patients are appropriate for treatment, identify potential adverse events with the use of the device, and explain how the product should be used to maximize benefits and minimize adverse events.

6(a).  Please discuss the appropriateness of the proposed indications for use for this device, which reads:

"The HeartMate VE LVAS is indicated for use as a bridge to transplantation in cardiac transplant candidates at risk of imminent death from nonreversible left ventricular failure.  The HeartMate VE LVAS is also indicated for use in patients with end-stage left ventricular failure who are ineligible for cardiac transplantation.  The HeartMate VE LVAS is intended for use both inside and outside the hospital.

6(b).  Does the labeling accurately inform patients of the risks of the device?

6(c).  Does the labeling adequately inform patients of the expected duration of use for this device?

6(d).  Are there any other issues of safety or effectiveness not adequately covered in the labeling?

7.  Based on the clinical data provided in the panel pack, do you believe that additional clinical follow-up or postmarket studies are necessary to evaluate the long-term effects of this device?  If so, how long should patients be followed, and what endpoints and adverse events should be measured?

DR. LASKEY:  At this, Dr. Zuckerman, would you like some consensus opinion on each to these?  I can tick these off with the help of my colleagues, so, please, feel free to correct me if I am misquoting or misparaphrasing, any of you.

For Question No. 1, on device reliability, I think we have established the fact that we would like to see more data on device reliability, that what we have seen to date indicates that in the clinical arena, the reliability falls short of the predictions made from theoretical and in-vitro testing, and that we would like to see, as requested by two of the panelists, the distribution of the times to failure, not just the medians and the means, but all the data points.

Your colleagues, please feel free to contribute.

Is that helpful, Bram, am I touching on the high points here?

DR. ZUCKERMAN:  Yes.

DR. LASKEY:  With respect to Question No. 3--

DR. DOMANSKI:  Could I just ask a question about that?  That could also be after, that is the postmarket period also or could be if we choose to approve this?

DR. LASKEY:  Yes, I am simply rehashing.

Mike, can you go to Question 3, data analysis?

DR. DOMANSKI:  Shouldn't we talk about 2?  I am not sure we really addressed 2.

DR. LASKEY:  I am sorry.

Are the device failure and malfunction rates and their time to occurrence appropriate--but we needed to see more data up to this point, is that not correct?  Dr. Nissen wanted to see means and medians to failure.  A number of us would like to see the actual distribution of all the points, not just those two.

DR. KONSTAM:  I agree, but I think we could discuss No. 2 based on best case of what we think we are seeing, that is, a device that lasts on the average about three years in vitro and appears to be somewhat shorter than that in the trial.  This is asking the judgment question of whether--I mean that is how I interpret it--whether that is an appropriate level of reliability for destination advice.

DR. DOMANSKI:  I guess the question I have is if you are going to do that, again, do you feel like there is not enough data in to consider this application, or can that be done by the FDA staff after approval?

DR. COMEROTA:  Wouldn't it be simpler to define destination?  If the definition was one year of additional life versus four years of additional life, would the answer be clearer, and then the gray area gets in the 2 1/2 to 3?

The real crux of the matter is what is the destination.

DR. NISSEN:  I guess what I was trying to get at here is that as I understand it, the patients lived an average of around 400 days, and during those 400 days of life, approximately 30 percent of the devices had an internal failure that could not be fixed without another operation.

That gives me some flavor for what the durability of the device is in a clinical setting, so I would be prepared to answer the question.  I think I would answer it as no, that it is not reliable enough for destination therapy, it is reliable enough for a bridge to transplant, but in this application, my answer would be no.

DR. DOMANSKI:  That is a fundamental question about whether or not this thing is going to be approved.

DR. LASKEY:  We are overlapping with voting now, so I think many of these issues will be more black and white as each member gives the reasons for yea or nay.  So, that ultimately may be the answer to many of these questions.

DR. DOMANSKI:  Yes, but I don't know that we have a consensus that you can give the FDA on No. 2 right now is the point.

DR. LASKEY:  Then, let the record reflect that there is no consensus perhaps due to the wording or perhaps due to the issue itself.

I think Question No. 3, we are all uncomfortable looking at 7 versus 3, nevertheless, the P-values are statistically significant.  My question to this question, is a clinically meaningful survival benefit, can it be viewed in isolation?  It needs to be viewed in relationship to the associated complication rate.  So, yes, we have demonstrated a survival benefit, but is it clinically meaningful if it  confers a hazard of adverse events, as well?

Does the rest of the committee share that sentiment?

DR. WITTES:  I don't think that the first clause and the second clause match, and I am having trouble with this because the problem is not only that there is only 7 and 3 at 242 months.  It's that there is a lot of censored data.  So, it's not as if we had all 68 patients and we knew that the rate is 7 and 3.  Then, we would have a good estimate of two-year survival rate.

The problem I think--I mean there are several problems--but one is that there is data in the pipeline that we don't know yet, and so I would like for you guys to reword that question to effect that.

DR. LASKEY:  That would indicate that not enough patient data has been reported then.

DR. KONSTAM:  Can we take a step back?  I think there is a core question that probably it might be worthwhile to sort of have the panel reflect on, that seems to me to come through as you work through these.

It relates to reliability and it relates to this Question No. 3.  That is the very essence of what REMATCH shows, and is the REMATCH result clinically significant, yes or not, and to me it all circles around whether if you have a statistically significant effect and apparently clinically very relevant effect at one year, but for the sake of argument, let's say that we are all lost at two years, because we don't have very reliable data at two years, and there are concerns about the reliability of the device at two years.

If that is what we have, is that a clinically meaningful result, and before you get into the adverse effects, yes or no.  I think, to me, from my point of view, I would love to hear the panel sort of try to reach a consensus about that.

DR. KLOCKE:  I think Marvin stated it.  I would have to vote no.

DR. DOMANSKI:  Well, I don't agree with that.  I think a year means different things to different people, and governments rise and fall in a year, our grandchildren are born, you know, it means different things to different people, and I think putting it on the market and letting people make their own decision is more to the point.

The fact is that the job of this panel isn't to make major societal decisions about resources are allocated. Our job here is limited to saying is it safe and effective, and, you know, I think the thing is safe and effective to extend life by one year.

DR. KONSTAM:  Since I posed the question, I guess I will weigh in.  I happen to agree with Mike.  I think that there are major societal questions that are hit upon by this application and I have to keep reminding myself that those questions are not before us, that the questions before us, and I think Mike sort of stated it, is this device safe and effective, and if you accept the core finding of the study--I mean one can challenge it, I mean I heard some challenges about whether there could be some bias introduced because some patients had the device and maybe were not DNR, and this sort of thing, that might be worth asking--but if one accepts the basic core finding that there is a highly significant prolongation of life even though it may well disappear at two years, I guess I cannot make a value judgment that that is not something we should offer to the patient.

DR. KLOCKE:  I would understand and would agree with that, and could imagine circumstances in individual cases where someone has a daughter who is getting married in four months and wants--I mean I would certainly do that.  On the other hand, I think, Marvin, at least for me, it depends on the meaning of the term "clinically significant," and I think that actually, I find it difficult, although I understand the survival data, it seems to me that the data we have indicate that the window may well be closing, although you don't know that until the data are in, but I also have to judge that in terms of the full complement of the device, the prolongation of life at the expense of increased adverse events, which I think really is correct, and so it's a judgment business, which I personally have no problem that reasonable people would differ.

But I don't think in this circumstance, if I were dealing with one patient, I am the advocate for that patient, and I certainly would do everything I could, but I think "clinically significant," and I don't mean to consider it in a societal text or anything else, in the overall best medical judgment case, separate from society, separate from cost, there may be a clinically significant benefit, but I am not convinced at this point.

DR. NISSEN:  I would like to weigh in on this one, too.  Let me say that, first of all, I really do think this was a valiant effort on the part of everybody involved to try to make this work, but I don't think it worked very well, and I don't think it was a clinically meaningful enhancement to survival.

I want to point out to the committee several things that we have heard today.  Thirty percent of the patients that got the LVAD never made it out of the hospital, 27 percent had a serious neurological event, 31 percent had sepsis.  Overall, 64 of the 68 patients had a serious adverse event.

So, if you said, well, we can extend your life by a year and we can avoid really large numbers of major morbidity and mortality, and we can improve your quality of life, then, I think it would be meaningful, but the Quality of Life data is very inconsistent as we have all talked about.

There is not really any solid evidence that that was the case, and I think the fact that so many people didn't get out of the hospital, so many devices failed during the course of the study, means that it was a good idea, but the device is not good enough to yet turn this thing loose on a population of people who undoubtedly are likely to be not as good at using it as the investigators in this trial.

So, I think that if we are going to let the genie out of the bottle, let the genie out of the bottle for a device that really works well, and I don't think this device worked well.

DR. COMEROTA:  I guess I need to make a comment.  I don't necessarily agree with you, Steve, because if you take it on face value, these are exceedingly ill patients, there will be an operative mortality from a large operation, that we need to accept, and I think most of us probably do.

The bottom line is if we are focused at two years and beyond, I think there is discomfort, but the discomfort should be lessening with the updated data that we are presented, obviously, not quite statistically significant, but more convincing.

The bottom line is at one year, there is a significant increase in survival, and there is no device failure at one year, which we can accept.  There must be improvement at the device level, and there will be.  I thin the quality of life does parallel the findings in the improvement in the New York Heart Association functional class although we have to accept the possibility and the probability of bias, but they are parallel.

With that said, and a significance at one year an device failure at one year, can we justify not approving it with the definition of the indication thrown in.

DR. LASKEY:  I am not sure we can approve this device only for one year, though.

DR. DOMANSKI:  Actually, it is not a matter of approving it for only one year.  I mean one doesn't usually put into an approval the survival data of any of the devices we put out.

DR. DeWEESE:  I think that we have good evidence that it does increase survival rate albeit it maybe only at one year at this time, but I think that there is going to be improvement in this device.  I think this is something that we are going to have eventually, and I think this group that has presented this and the group that worked with them, should continue to do it, and I would hope we would support them to do this and make the advances that are necessary to make it a little better maybe.

DR. LASKEY:  Maybe we could get just a little bit of help since we have really answered Question 4, as well, here, but Dr. Zuckerman, you might try and frame for us where the FDA is going with respect to changing definitions of survival benefit.

This is clearly a very different relative risk reduction or reduced hazard ratio than we are used to thinking about, and the general rule of thumb has always has been the sicker the patient, the more dramatic you want to see the relative risk, or it is always the sickest who "benefit the most."

What is that the Agency has in mind with respect to a meaningful survival benefit if 25 percent is not enough or 33 percent?

DR. ZUCKERMAN:  Can we go back to question 3, Dr. Berman.

When we look at our definition of "reasonable assurance of safety and effectiveness," we need to be able to demonstrate clinical utility.  As indicated in the opening presentations, clinical utility for this type of device in this type of population hasn't been previously perhaps well defined.

Consequently, we are looking for some panel consensus, if possible, as to what a clinically meaningful survival benefit might be - is it what we see at one year, is it a difference in the median survivals for the two-patient populations, even though the survival curves might come very close together at two years?

We don't have a priori a predefined definition, and we are looking for some help here.  The fact that this is a gray area isn't surprising, but maybe you can poll panel members again to see if there is some more of a consensus.

DR. LASKEY:  I will poll them once again.  I would like to also suggest one more index of survival benefit, which is how much longer the patient is going to live, how many more days or months of life can one expect with this treatment, and I think that needs to sometimes be added to this clinically meaningful survival benefit in addition to a P-value.  I think if we knew how many more days a patient had, that would answer the question that I think Dr. Rose posed to us, which was how much longer do I have.

DR. KONSTAM:  Let me just give my reflection on this.  I come at it the other way, which is that I see a dramatic statistical effect, a large number for risk reduction, a substantial augmentation in median survival, and a very low p-value.

So, the effect is pretty dramatic in my mind.  Again, to me, I compartmentalize the issue of adverse events, so I think Steve's points are very cogent and need to be addressed by the panel, but I think maybe separately.

To me, it is I think logically useful to first just go through the exercise of whether you think that there is a clinically meaningful effect on survival, yes or no, and to me I think there is a very dramatic effect on survival, and the question is do we feel that that is negated by the fact that we don't see it at two years.

This is to me the way I frame the question for myself, and I come away saying who am I to say that those findings are not important just because I don't see it in two years.  You know, I am not the patient really being potentially presented with that choice.

I guess in terms of the survival effect, I am impressed with it, and I can't talk myself out of that.

DR. LASKEY:  I am not sure we are any closer to a consensus.  We keep going around and around.  I must say as a clinician, I find it hard to divorce survival from the quality that goes along with that survival, and when there are so many infections and strokes and re-ops.

It just can't be viewed in isolation.  Statistically, it can, but clinically, we take care of the whole organism.

Janet.

DR. WITTES:  I would look at it as at one year, I am coming in, trying to make a decision about whether to have this implant or not, then, the relevant data it seems to me, if I don't have it, then, my chance of being alive in a year is a quarter, and if I do have it, my chance of being alive in a year is a half, and that to me seems like a big difference.

I don't care about two years, I am talking about  disease where my imminent death is--so, it seems to me that then what I personally would weigh, given those data, and I am comfortable with those data, the p-value for me tells me that those data are pretty robust, I am comfortable with those numbers.

So, then I would play into am I willing to take all these other risks to give me this benefit of mortality. My personal feeling, and I think as a panel member, that we shouldn't make that decision for other people, that I would say yes, it demonstrates a clinically meaningful survival benefit.

People may not choose it, but I would say yes.

DR. DeWEESE:  When I first read this, I thought that the persons who had had a number of adverse events, it would discourage them and make them feel they should not have done what they did, and then I find that an equal number of people who were controls and had the procedure said that they wanted to withdraw.

I would have thought there would have been a much higher withdrawal rate from those who had the procedure, had it, when they are looking back, had they been that person.

DR. NISSEN:  One comment.  I personally cannot separate survival from quality of life, and I will tell you why.  Let's just take, for instance, for a moment, that you had a therapy that could prolong survival by one year, but all the patients were in a vegetative state during that period of time.

Would you call that a clinically meaningful survival advantage?

DR. WITTES:  No.

DR. NISSEN:  Just so we all are on the same page here, I think it is one thing to say there is a statistical effect on survival, and the other is to say there is a clinically meaningful effect, and I think the word "clinically meaningful" to me implies that there is some quality of life.

DR. DOMANSKI:  What little data we have on Quality of Life, and, you know, I don't think much of the Quality of Life data in this trial, not because of any fault of the investigators, but because of just the nature of the study, but what little we have suggests, in fact, that although not absolutely consistently across things, it has improved.

Much of the chemotherapy we give is little more than chemical last rites in a setting where the patients know they are not going to get much benefit from it, so people do choose it.

Here, they have an opportunity to choose some benefit.

DR. KNAPKA:  Again, talking from a patient that was in this condition, given like 30 days to live, and this sort of thing, I don't think a panel can make a decision whether one year or two years is significant.  There is just no way anybody can make this decision unless you have been there.

I think we probably need, if we feel very confident that this device will give the majority of people, and we realize that the tests we are looking at, these are real sick people, kind of a last resort.

I think this is one of the problems.  There is a lot of these new devices and new chemicals that it is usually used on patients as a last resort, but I think we are arguing whether a year or two years is a significant amount of life, we will never come to that decision.  It ha got to be a patient's decision, and I can't make that for anyone.

DR. OSSORIO:  I just want to weigh in on the side that says clinically meaningful survival is something more than just statistically lengthened life.  I want to weigh in on that side because I think that once you do, things begin to unravel here a bit, which is unfortunate.

DR. KONSTAM:  I just want to clarify what I said earlier.  I guess I was segmenting the issues, and to me I think it is worthwhile segmenting the issues.  I took Question No 3 really just to address the one-year versus two-year issue, and really that is what I spoke to.

I certainly concur with everybody else, but I do think it is worth getting past that and just looking at that issue in isolation if for no other reason for its logical value.

I certainly, however, would not stop there.  I certainly concur that if we were extending life, but people were in a vegetative state, I would say now, okay, the numbers are okay, but forget it.  So, I certainly agree with that.

So, the next question, I think the next logical question is okay, life is extended if you accept that, and is it meaningful.  Well, I have trouble saying that we have not extended meaningful life here.  If you look at six months, for example, in the LVAD group, you have 8 patients who are classified as New York Heart Association Class I, and 19 patients classified as Class II.

The corresponding numbers in the medical management group is zero and 2 patients.  Now, I agree that this statistical analysis of the Quality of Life comparison is extremely problematic, but to me, comparing the Quality of Life in the two groups is much more important if you have most of the patients alive.

If you have one limb that is 75 percent dead after one year, to me the question changes.  The question becomes is the clear extension of life very problematic because very, very few of the patients are having meaningful life, and a best I can read into this, not having the patients in front of us, I see significant, clinically significant, in my mind, numbers of patients who are doing pretty well.

This really gets back to the comments that were made earlier, who are we to say it is not appropriate to offer that to patients.  So, I agree it is a several-step process, and I wasn't up to that step yet, but I don't think you can look at this and say well, we are keeping everybody alive, but they are vegetative.  There seem to be people who are doing pretty well here.

DR. LASKEY:  I am sure Steve used that more for hyperbole than for reality, but it certainly makes the point, and I think certainly what we are all grappling with is we can easily deal with the easiest thing to do here is to interpret the p-value, as Janet has said, and say that there is a significant effect demonstrated here.

However, we are clinicians, and I think that there is this nagging feeling that many of us can't get past when a survival benefit is associated with risks of bleeding or risks of sepsis or risks of stroke, and so on, and so forth, which is identical to the end of life chemotherapy issue, as well, and I am not sure there is a meaningful answer to that one either.

I am not sure that we can give you an answer to your question yes or no.  I think you have heard the deliberations and the really assiduous thought process that we have put into this.  This is a different category of patient, this is a different proposal here.

DR. WITTES:  Several of you have said that I responded to the p-value, and I want to make it clear what I was responding to.  I was responding to the difference in the magnitude of survival at one year, the 50 percent versus 25 percent, which I was interpreting.

You know, you may not agree with the interpretation as an important difference.  What I said about the p-value was that the p-value made me believe that that difference was likely to be real, but I was not reacting to, and I think it is important for you guys to know statisticians tend not to react to the p-value by itself, it is the magnitude of the difference.

DR. ZUCKERMAN:  The fact that there isn't one clear-cut answer is okay.  This is a difficult question that FDA has struggled with.  Would it be fair to summarize, though, that panel members come to this question with different prior beliefs, and there is a range of answers as to whether or not a clinically meaningful difference has occurred at one year and throughout the course of the study?

DR. DOMANSKI:  And you are going to get a quantitative estimate of that balance when they vote yes or no in terms of approval.

DR. ZUCKERMAN:  Right.  It just underlines the need therefore for a control in this study to better assess that risk-benefit profile.

DR. LASKEY:  I think we really have beat up Question 4.

DR. ZUCKERMAN:  Before going on to Question 5, though, can you just give us a quick summary for the record?

DR. LASKEY:  I think I can summarize the panel's feeling as saying that, number one, it is recognized and accepted among the heart failure community that measures in NYHA, 6-minute walks, et cetera, et cetera, don't always correlate.  Number two, they are soft endpoints.  Number three, unless they are blinded in their ascertainment, it is even more difficult, and the potential for bias is always there when you have unblinded observers assessing soft endpoints.

I think we would all like to read into this an improvement in functional status, but it is hard with those caveats.

DR. OSSORIO:  I guess for me this is the crux right here.  If people can't do better than this, I think we are in real trouble.  I people should be expected to do better than this in terms of measuring something about whether or not there is a functional improvement or whether there is some kind of decent functioning going on in people's lives after they have had this intervention because what patients do care about is, you know, if what somebody cares about is I want to see my child's wedding or my grandchild's birth or whatever, if you can't recognize your child or your grandchild, it doesn't really matter, you know, that you are alive at that point.

There are ways of assessing these things.  They may not be in the cardiology community, they may not be as widespread, but there are ways of assessing these things, and I think people ought to be expected to do it.

DR. DOMANSKI:  I don't think that is an entirely fair analysis, though, because, you know, these people did go through the whole battery of things.  There are certain things that are peculiar to an unblinded study where you put a device in.

DR. OSSORIO:  I know.

DR. DOMANSKI:  I think if you say that, then, you ought to offer some indication of how you think they could have done it better, because I think they did what they could, but it is the nature of the beast that makes it difficult to assess, so I don't buy into that statement.

DR. NISSEN:  Mike, how about just having somebody who is not the operating surgeon ask the question about what your functional class is.  I mean to me, you know, I can't think of a lower standard to apply than to have the person who actually did the operation asking the patient whether they feel better or not.  That is about as bad a data as you can possibly generate.

DR. LASKEY:  That perhaps will be recommendations for further study design.

DR. DeWEESE:  Where did we have that evidence that the surgeon got that information?  Was that in anything we received?

DR. NISSEN:  That is what we are told, to ask the question, and that was the general gist of the answer.

DR. DOMANSKI:  Since that is raised factually, is that really true, because actually, I didn't know that.

DR. OSSORIO:  That question was asked.

DR. DOMANSKI:  I took it at face value.  Is that what they said?  Okay.

Could I ask them to clarify that for us, please?  Could I ask them to come back to the table and answer that question?

DR. LASKEY:  The question being, so we are clear for the record, how was NYHA classification assessed and by whom.

DR. ROSE:  We didn't specifically state that surgical patients should not be evaluated for NYHA class by the surgeon or the operating surgeon, and I can't say that there are data that don't reflect that.  I think knowing how these clinics and follow-up function, the overwhelming majority of patients were assessed by the cardiologists who were following them or the nurse clinicians who were following them, as well.

So, to think that this reflects, this improvement in NYHA class reflects some enormous blinding on the part of the surgeons doing this at some level is an insult I think to the participants in the study.

DR. LASKEY:  The panel abjectly apologizes for any sense of insult.  It wasn't meant to be an insult, but I think that the answer is that the investigators and/or their associates obtained this information.  We certainly didn't mean to insult them.

DR. NISSEN:  As opposed to an independent third party.  I mean, look, I mean you asked if there could have been better methods, and the answer is that it would have been extremely easy to have a non-participating person do the assessment of functional classification, and that would have been the proper approach from a trial design point of view, and that was not what was done.

DR. DOMANSKI:  I think there are some other things, though, that came in.  It wasn't just NYHA.  NYHA class is tough, because don't forget, there is a big placebo effect.  I mean if you bury one of these things in somebody's chest, they want to feel better.

So, I would be worried about that, as well, in terms of unblinding, but I think there were also, you know, they did the SF-36, you know, they did a bunch of things and stuff.

DR. KONSTAM:  I completely agree with the points that were made, but I still keep coming back to the fact that the control group had so many dead people in it, so the issue really changes.  I mean if you had 90 percent survival in both groups, you would really want to know which group is doing better.

When three-quarters of the control group is dead by a year, I think the question changes.  The question becomes one, which is much more difficult to answer, which is how is the treatment group doing, and do we have a sense that they are all vegetative, do we have a sense that they are all hospitalized, or do we have a sense that at least a sizable number of them are actually doing okay, and looked at it that way, you know, I think the data are adequate for me to say, you know, there are sizable numbers of patients that at least at six months and a year are doing okay.

DR. LASKEY:  And that you can't be doing well if you are dead.

DR. KONSTAM:  You can't be doing well if you are dead.

DR. LASKEY:  No. 5.  I think we have been grappling with the denominator here, that is, the benefit as it relates to Question No. 3 and the data analysis and the magnitude of the clinically meaningful survival benefit, that being the denominator of the risk-benefit ratio.

Do we need more discussion about the presence of risk or do we need to focus in on the magnitude of this risk and how it relates to that ratio?

We all agree that there is risk with this device.  Where is this ratio, is it closer to zero or closer to 1, I guess is what you need to know.

DR. BERMAN:  I think the panel has discussed this question along with the others.

DR. LASKEY:  Okay.

DR. PINA:  Let me bring up one point that hasn't been brought up, and it is part of the protocol, and that is the issue of cost, which has not even entered into any of the data that we have, but the expense even on a surviving family member may be some of the risks that have to be assumed other than just the physical risk.

So, I think that is something we need to keep in the back of our mind even though that is not part of the data that we have.

DR. ZUCKERMAN:  Again, Dr. Laskey, for Question No. 5, the question is do the benefits of this device outweigh its risks.  Is there any consensus at this point from panel members?

DR. LASKEY:  What I am hearing at this point is I don't think there is consensus, you may see consensus or lack thereof with voting in several minutes, but I think right now we have aired the concerns about both the numerator and the denominator, and their relative weight will be reflected in how people vote.

DR. KLOCKE:  I would answer that as possibly, but not clear to me at this point.

DR. LASKEY:  Labeling.  I think it is fair to say we have had an extensive discussion about the nature of this patient population.  What is missing is how the investigators got there and how that translates into more detailed information in the IFU, I am not convinced that we have heard that.

DR. BERMAN:  Is it fair to say that the panel thinks there needs to be more details or more specific indications?

DR. OSSORIO:  Yes.

DR. PINA:  Yes.  I think in the patient booklet, where there is a whole series of warnings, and I understand what Eric said, that a lot of these patients are not exactly going to be reading this great detail, but the relatives will be, and the spouses very often do, that all the complications that have been found in the study need to be enumerated, so that the patient is well informed and the spouse is well informed about all the risks, and that is not in there.

DR. COMEROTA:  Which includes those who are indicated for the procedure, as well as those who are not indicated for the procedure.

DR. KONSTAM:  Can I just make sure that the point is made and see if other people agree, that based on what we see in the REMATCH study, this device should only be indicated for people who are severely ill and with an extremely limited life expectancy.

Now, there were some comments on that made both sides, but I think based on what we see in the survival data at two years, and those of us who think it is acceptable or not acceptable, nevertheless, most of the patients in the device group are dead at two years, and therefore, I cannot imagine doing anything other than approving this device for patients with an extremely limited life expectancy without this device.

Now, how you get there is for further discussion, but I think that would be the sense that I would want to inject into it.

DR. DeWEESE:  I would agree.

DR. LASKEY:  Does the labeling accurately inform patients of the risks, well, there were certainly conspicuous warnings and cautions, and so forth, which relate to I think mechanical malfunction.  I am not sure what the data are there on the risks of infection and bleeding, and so forth.

DR. PINA:  It's not in there.

DR. OSSORIO:  Along with accurately informing about risk, if there are things that patients can do to minimize those risks or their family members and people who are helping them if they happen to go home, those things ought to be in there, too.

DR. KONSTAM:  I would say that this cuts to the heart.  We talked earlier about patient choice.  If we wind up approving this, I think some of us are going to say we are doing it because we think there should be an option for the patient, but if it's an option for the patient, here, really more than any other application for anything I have ever seen, we really have to take pains to inform the patient of what he or she is getting himself into and what kind of adverse events have been observed over the two years of this study.

You know, here is one where I don't mind scaring the patient, to tell you the truth, because of the degree of uncertainty that we have and what we have seen in this trial.  I think a good deal needs to be done with this document to make sure patients adequately get a sense of that.

DR. PINA:  I think also, in all fairness to the sponsor, they probably haven't seen these kinds of complications because this is a sicker group than I have ever seen LVADs in, because most of these patients, we would not transplant, and therefore, we would have not have VADed.

So, I think in all fairness, this is probably their first experience with this very, very sick cohort where these complications were arising.  This isn't the usual.

MR. MORTON:  One point of clarification I would make is I have heard the panel refer to letting the genie out of the bottle, and that is a connotation of something that is out of control, and labeling is very much a way that release of this device could be controlled.

I know that on devices of this type, the FDA and the sponsor work together to very clearly define what the training program is going to be and what the release program is going to be, so I would like us to move away from the genie out of the bottle image.

DR. LASKEY:  Fair enough.

Does the labeling inform patients of the expected duration of use?

DR. WITTES:  It needs more.

DR. LASKEY:  Very good, Janet.  Needs more.

DR. KONSTAM:  I would be fairly draconian about this.  I think that the panel has some serious, at this point, lack of information and uncertainty about the duration of this device, and I think that somehow again, I think the patient needs to be informed of that.

I think the comments that's in the proposed wording is simply that sometime in your life, you might expect to have to have this replaced.  Well, I mean that sort of begs a lot of the issues that we have raised like there is no indicator of end of life, number one, so what is the implication of that.

Secondly, how many patients do we have now experience with, who have had successful reimplants and alive?  I think there is one such patient who has had reimplant and are alive.  Maybe there is more than one.  Two?  Okay, two.  That is a very, very, very, very small experience, and to have that reflected by a comment that says at some point you might have to have this changed really strikes me as pushing it.

So, that wording really has to be rethought, and I think based on what we see right now, it should be made very blunt that we don't fully know the expected life expectancy of this, and somehow these points needs to be brought out.

DR. LASKEY:  No. 7, PMA, postmarketing.  I think we all would agree yes.  There is a need for additional clinical follow-up specifically, and there is certainly a need for more appropriate risk stratification for who will benefit or who won't.

DR. ZUCKERMAN:  Can you comment on what the major questions would be to answer in any sort of postmarket experience?

DR. PINA:  Let me start out by again expressing my concerns about the internal device malfunction and the lack of knowing when it is going to malfunction, and I think the  company has experience in the bridging group with perhaps the same problems, and all these other variations have been made, like the SNAP, the different wiring system, and I think we need a body of data on internal pump malfunction with all the advances and all the improvements that the company has made.  I don't think we have that, and I certainly haven't seen it.

DR. DOMANSKI:  One thing you will gain with your postmarket surveillance is you will gain a sense of the time course of the different types of complications - patient complications, device complications, and so forth, with this device, and then other ones that come along.

So, actually, I think the postmarket surveillance in this case could be quite useful, particularly as the device evolves or devices like it evolve.

DR. KONSTAM:  Can I add one other thing, I guess to come back to, I had raised earlier, is the question of anticoagulation.  I just would like to see going forward some consideration of potential anticoagulation regimens.

I don't know what that would be right now, but we are seeing, to me, a higher than expected number of what I consider thromboembolic events, and I don't understand all the bleeds, but I think that that should be addressed with some postmarketing research.

DR. LASKEY:  Certainly, first of all, with collection of information to get a better idea of the rate because it is entirely possible the rate may be higher in real life than in the trial settings.  So, additional ascertainment of endpoints.  Okay.

Open Public Hearing

DR. LASKEY:  I would like for the last time today to open the forum for public hearing.

Is there anyone in the audience who wishes to address the panel on this topic?

[No response.]

DR. LASKEY:  If not, then, I will close the public hearing and request that the sponsor come forward and give us your final sentiments.

Sponsor Comments

MR. MIDDLEBROOK:  Certainly, all aspects of the study have been discussed thoroughly here today, and we just want to thank all the panelists for their insight and their analysis.

As a company, we are committed to, as Vic said, continuously improve this product, and if we look at our experiences from the bridge to transplant, generally speaking, our results have improved from the time when we did the clinical trial.  We would hope that that improvement would be seen here as we look at this nascent therapy.

Again, I would like to thank the panelists, the FDA, and certainly all of our presenters.

Thank you.

DR. LASKEY:  Thank you.

I would like to ask Dr. Ewing to read the voting options.

Panel Voting

DR. EWING:  The panel recommendation options for premarket approval applications are the Medical Device Amendments to the Federal Food, Drug, and Cosmetic Act as amended by the Safe Medical Devices Act of 1990, allows the Food and Drug Administration to obtain a recommendation from an expert advisory panel on designated medical device premarket approval applications that are filed with the Agency.

The PMA must stand on its own merits and your recommendation must be supported by safety and effectiveness data in the application or by applicable publicly available information.

Safety is defined in the Act as, "Reasonable assurance based on valid scientific evidence that the probable benefits to health under conditions of intended use outweigh any probable risk."

Effectiveness is defined as, "Reasonable assurance that in a significant portion of the population, the use of the device for its intended uses and conditions of use, when labeled, will provide clinically significant results."

Your recommendation options for the vote are as follows:

Approval if there are no conditions attached. 	Approvable with conditions.  The panel may recommend that the PMA be found approvable subject to specified conditions, such as physician or patient education, labeling changes, or a further analysis of existing data.  Prior to voting, all of the conditions should be discussed by the panel.

The third option is not approvable.  The panel may recommend that the PMA is not approvable if:  the data do not provide a reasonable assurance that the device is safe, or if a reasonable assurance has not been given that the device is effective under the conditions of use prescribed, recommended, or suggested in the proposed labeling.

Following the voting, the Chair will ask each panel member to present a brief statement outlining the reasons for their vote.

DR. OSSORIO:  May I ask for a point of clarification?

DR. LASKEY:  Yes.

DR. OSSORIO:  That very last thing that you just said, under the proposed labeling.  Now, we have all made quite a number of suggestions about what we think ought to go on the labeling, so when we are voting, are we voting assuming that those suggestions would be incorporated, or are we voting based on what we have right now in this packet?

DR. EWING:  If the panel feels that to be approvable, then, the changes in labeling would be necessary, then, that could be approvable with conditions.

DR. LASKEY:  It is your prerogative, Dr. Konstam, to make a motion.

DR. KONSTAM:  I move approvable with conditions.

DR. LASKEY:  And they would be?

DR. KONSTAM:  I have a bunch.

DR. LASKEY:  First, we need a second.

[Second.]

DR. LASKEY:  You have a second, so you might want to delineate the conditions.

DR. KONSTAM:  I have a bunch of conditions, and I think they have all been touched upon.  One is additional analysis of existing data.  The two that occur to me are revisiting the whole question of reliability, bringing data up to date, and getting a clear indication of reliability  at two years.

Secondly, that analysis that Mike Domanski had suggested of time to death or stroke be looked at to be certain that it is at least consistent with the observation with regard to survival.  So, existing data would be one.

Secondly, the indications for the device be much more extensively delineated particularly to denote a population with a very limited life expectancy without the device implanted.

Third, that there be fairly rigorous criteria for implantation from the perspective of both the surgeon and the facility at which it would be done to be certain that both patient selection and expertise in the procedure and in patient follow-up meets a high level of acceptability.

Fourth, that there be a significant amount of postmarketing work be done surveying patients in whom these are implanted as least out to two years with analysis of reliability and analysis of all of the other adverse effects that we saw here and their rates.

I had mentioned earlier I would include in that some kind of consideration of the need for anticoagulation.

Finally, that there a lot of work go into a detailed set of information for the patients that really, as best as can be accomplished, delineates for the patient what the tradeoff is that he or she is getting himself into with this.

DR. LASKEY:  There are five conditions then, and we need to vote on each one separately.  Can we have some discussion amongst the panel members about Dr. Konstam's first condition, which is the requirement for more data analysis from the current data set?

DR. NISSEN:  Could I understand this, does that mean that it would not be approvable until the data is analyzed, is that what you are saying?  I am not sure I know what you mean by that.

DR. KONSTAM:  Well, I guess I would want to ask the Agency what the options are in that regard.

DR. ZUCKERMAN:  You are making a recommendation whereby potentially the device is approvable if these conditions are met.

DR. DOMANSKI:  But I thought that what you meant was that these analyses would take place over time.  You know, if you don't think it's approvable until they meet it, that is different than saying we are going to do postmarket surveillance after.

DR. KONSTAM:  Let's take the one that you suggested, the analysis of time to death or stroke.  You made a compelling argument for that.

DR. DOMANSKI:  I think that is fine.  I think that is an easy one for them to come in with, and I think the numbers of events are such that that is not going to be a land mine in the field for them, but that is something they can go over with FDA staff, so that is fine.  They are not going to have a problem with that.

I will just finish discussing the question.  I would suggest that they not be forced to go out and gather new data.

DR. KONSTAM:  What I meant with regard to the reliability is more detailed analysis of existing data.

DR. EWING:  If you do not believe that there is sufficient information currently here, then, it would make more sense to vote not approvable, if you are talking about you need the results of current analysis.

DR. KONSTAM:  At least what I had intended was pretty much along the lines that Mike was suggesting.  I do think that the first part of that is the analysis of death or stroke.  I would like the Agency to hear my sense that I am assuming, Mike is assuming that that analysis will not radically change the overall survival analysis, and if it did, if it looked like all of a sudden it was substantively different--and I don't know how to advise you better than that--then, I would reconsider my approval.  I don't know how to convey that.

DR. DOMANSKI:  I have a suggestion.  I will tell you what I really think.  I think the number of events is not going to be sufficient to make much difference, and I would leave that out, delete that from your motion, and let the FDA just look over the entire application, because this is a recommendation, and not put this in as some kind of a firewall.  I can't imagine it would, and the FDA is going to look at this thing.  I just don't want to see them disapproval over something like that.

DR. KONSTAM:  I think we are saying the same thing.  I think the Agency has the prerogative of looking at the data and say you know what, I mean they can do whatever they want with it.  They don't have to accept our final vote, so I guess we are saying the same thing.

DR. DOMANSKI:  One way of doing this would be to structure the recommendations on things that need substantive change.  I mean the FDA is going to go over this whole application again, but there really are some substantive things that they have to do.

The panel wants different labeling, so they are recommending that.

DR. LASKEY:  He has put five recommendations, five conditions on the table for his recommendation.  We need to either take them apart or just whittle that down to one or two.  Can we try and do that, so we get a coherent, articulate vote?

DR. KONSTAM:  Let me then clarify.  I guess I would still stick to my first one, but I am not basing approvability--

DR. LASKEY:  The first one being--just so we are all on the same page here--is additional data analysis as pertain to device reliability and?

DR. KONSTAM:  And analysis of time to death or stroke.

DR. LASKEY:  Time to death or stroke in the current data set.

DR. COMEROTA:  And the purpose of that is not for device approval, but better judgment on the basis of physicians and patients in the future upon whether their decision should be to move ahead with their own decisionmking or not, more informed consent on the patient's part.

DR. LASKEY:  Yes, it is to be added to the labeling, is that correct?  Okay.

Enough discussion on that condition?  Shall we vote?  Voting on these two conditions to Dr. Konstam's motion for approval, the first condition being additional data on device reliability and additional data on time to death or stroke.

DR. WITTES:  Analysis.

DR. LASKEY:  Analysis, yes, of the current data set.

DR. OSSORIO:  Can I ask for a point of clarification?  Could we vote, say yes on the various amendments, and at the end, still vote no on the motion?

DR. LASKEY:  Yes.

DR. EWING:  For the first condition, we need to go around the room, start with Dr. Wittes, please.

DR. WITTES:  I am going to vote no.  I think it should be a recommendation.  I don't see it as a condition.

DR. DOMANSKI:  So, this isn't framed as a recommendation, is that right?

DR. EWING:  This is the first condition of approval.

DR. DOMANSKI:  I will vote yes to that, fine, do the analysis.

DR. KONSTAM:  I vote yes.

DR. COMEROTA:  I vote yes.

DR. NISSEN:  Either we need more data or we don't need more data, and since I would like to see more data, I am not sure whether voting yes is going to let us look at more data before we make a decision, so I am going to abstain.  I mean either we need these data in order to make a decision or we don't.  I think I need these data in order to make a decision.  Lacking those data, I don't think the amendment helps me.

DR. AZIZ:  Yes for Dr. Aziz.

DR. PINA:  Yes.

DR. OSSORIO:  Yes.

DR. DeWEESE:  Yes.

DR. KLOCKE:  Yes.

DR. EWING:  Thank you.  I will tabulate that that is 8 yes, 1 no, 1 abstain.  The condition carries.

DR. LASKEY:  Are you happy with the distillation of the conditions?

DR. ZUCKERMAN:  Yes.

DR. LASKEY:  Any discussion on Dr. Konstam's recommendation for additional clarification or new indications for use?  Do we need to go through that?

DR. PINA:  No, I think we have discussed that pretty well, and it may merit some other meeting at some point to sit down and really go through, and the Agency can do this, to go through the patient population that, in fact, would be eligible for this.

DR. EWING:  Okay.  We can take a vote although Dr. Domanski has stepped out.  We will come back to him.

Dr. Wittes.

DR. WITTES:  Yes.

DR. KONSTAM:  Yes.

DR. COMEROTA:  Yes.

DR. NISSEN:  Yes.

DR. AZIZ:  Yes.

DR. PINA:  Yes.

DR. OSSORIO:  Yes.

DR. DeWEESE:  Yes.

DR. KLOCKE:  Yes.

DR. EWING:  That is almost unanimous for No. 3, the conditions for clarification of indications of use, clarification of the patient population indicated and excluded.

The next was the postmarket study.

DR. KONSTAM:  No, the next is criteria for use both in terms of the operator and the site be clarified to be certain of high quality in patient selection and in performance.

DR. LASKEY:  And training and experience.

DR. KONSTAM:  But it is a little bit more than training.  Training and--

DR. PINA:  Setting standards?

DR. KONSTAM:  Setting standards for both the operator and the site.

DR. LASKEY:  So, what you would like to see the Agency receive is a document outlining the criteria for credentialing, if you will.

DR. KONSTAM:  Right, and the spirit is that it be fairly rigorous.

DR. PINA:  And could I add that it is not just limited to surgeons, but also to cardiologists taking care of this patient population.

DR. KLOCKE:  And could I ask that it include the group's best recommendations at that point in terms of infection control in relationship to the issues that I have been talking about?

DR. KONSTAM:  That is fine with me.

DR. AZIZ:  Do you want to restrict it to transplant centers?

DR. KONSTAM:  Well, I held back from saying that.  Maybe it is worth some additional discussion on the part of the panel about how far we want to go and what we mean.  I mean I would be happy with limiting it to transplant center if other people on the panel would.

DR. LASKEY:  I think in deference to Mr. Morton who raised the point that we really should not concern ourselves with how this device is used or abused, we just need to go forward in good faith and outline criteria for training and experience of these centers and individuals, realizing it is a system as much as people.

DR. KONSTAM:  I thought he was saying that it is these very processes of labeling and training that give us a comfort level that we are not "letting the genie out of the bottle" by doing these things.

MR. MORTON:  Thank you.  That is true.  Again, from my experience on a device of this type, it is a collaborative process with the Agency of exactly what the training program is going to be and exactly how the release will happen.

DR. NISSEN:  I am glad you all have faith that nobody ever uses devices off label.

DR. ZUCKERMAN:  I think the issues from the Agency perspective are, one, two, consider Dr. Konstam's motion that the training program be rigorous and looked at by FDA. The second part is that potentially, as a postmarket requirement, one can look at how new sites are brought up to speed, et cetera, but there are two parts, and Dr. Konstam is first just asking about the review of a training program.

DR. KONSTAM:  Did you want more specification to that or do you feel like leaving it as clear as we have is sufficient?

DR. ZUCKERMAN:  Why don't you try restating it, so that then we can have a vote.

DR. KONSTAM:  I think the spirit of what I would like to convey is that this device be implanted in centers and by individuals who are highly trained and specialized in the management of patients with end-stage heart failure both from the perspective of patient selection and surgical procedure, and perioperative management and long-term management.  I guess that is really what I meant now.

DR. ZUCKERMAN:  Are you asking FDA to review the company's training program as a condition of approval?

DR. KONSTAM:  Yes.

DR. LASKEY:  That is fairly straightforward.  Do you want to just do a hand vote?

DR. EWING:  Sure.

DR. LASKEY:  All in favor of this condition?

[Show of hands.]

DR. EWING:  I believe that is unanimous.  Is that correct?  Okay.

DR. LASKEY:  As I recall, the last condition was--

DR. KONSTAM:  There were two more.

DR. LASKEY:  Sorry.

DR. KONSTAM:  Next was the postmarketing, and then the final one was information for patients.

DR. LASKEY:  Let's take the first.  Maybe I misunderstood.  The information requested from postmarketing surveillance would be?

DR. KONSTAM:  Well, we talked about it before, I think developing a sizable body of experience at least out to two years, tracking patient survival, tracking complications, tracking device reliability and durability. There may be others people want to add to that.

DR. COMEROTA:  Marv, would you be comfortable in extending that to death since a large percentage of those who have lived two years, died shortly thereafter?  I would like to see it extended out to death.

DR. KONSTAM:  Sure, so indefinitely.

DR. COMEROTA:  Yes.

DR. KONSTAM:  To death, yes.

DR. COMEROTA:  Well, death is rather definite.

DR. KONSTAM:  Definitely.

DR. LASKEY:  Marv, given the limited capability of the Agency to do these kinds of surveillance, what would you recommend that they focus on, that they require?

DR. KONSTAM:  You mean in terms of duration of follow-up or in terms of what they are looking at?

DR. LASKEY:  More what they are looking at.

DR. KONSTAM:  I think the things that I listed shouldn't be too onerous.  I mean patient survival, device survival, and major complications, at least the major complications as were demonstrated in this study.

DR. PINA:  Implantation rate.

DR. KONSTAM:  That is device survival.

DR. LASKEY:  This is all in the form of a registry, or are you recommending the conduct of an additional trial, if you will?

DR. KONSTAM:  Well, I wasn't recommending a randomized trial.

DR. DeWEESE:  Can't we just ask for what they have done in the first two years?  That is what we want them to do, we want to get the same information we had before.

DR. WITTES:  We are asking for less, I think, than what they did in the first two years.  It seems to me we are talking about a registry.

DR. ZUCKERMAN:  At this point, Dr. Konstam, perhaps you can help us with what major questions you would like to see answered in a new cohort experience, and then the Agency and sponsor can work together on what type of trial design might be optimal, but if we can first define the questions.

DR. KONSTAM:  I would like to know the life expectancy after implantation.  I would like to know reliability and longevity of the device.  I would like to know the frequency of the major complications that were identified in this trial, and I guess I would add to it and ask for some discussion or comments, you know, some indication of patient function, some indication that patients are doing well, and I don't know if people want to discuss what that should look like, but something more than that they are just alive.

DR. AZIZ:  Also, I think it would be important when a patient dies, that the device be examined by one center specifically, so any valve problems or any motor problems could be documented, because I think just because the patient dies, there may be different reasons, maybe sepsis.  I think the devices must be examined and explanted.

DR. PINA:  I would like to echo the functional capacity assessment because there are a lot of data on functional capacity and bridge to transplant, and that is well published and well known.  So, I would like to see how this compares.

DR. KONSTAM:  I also again would like to see something explored with regard to anticoagulation.  I guess I don't want to make that more definitive right now except to say that I would like some sort of study proposed in the next six months, say, to explore--

DR. LASKEY:  You need a handle on the rates.

DR. KONSTAM:  My feeling would be that we already have from the data set in front of us, evidence that there is an excessive rate of thromboembolic events, more than anticipated.  That is my interpretation of the data.  Maybe there will be some differences of opinion on that.

DR. COMEROTA:  I thought they were less for a totally implantable prosthetic without anticoagulation.

DR. KONSTAM:  Less than what?

DR. COMEROTA:  Less than anticipated.

DR. DOMANSKI:  I have a sort of process problem with that, and the process problem is this.  We are approving this.  All of the data that have come in here, have come in without the anticoagulation.

Now, we are approving the device, and you are trying to mandate a study that basically send them off label to do some anticoagulation.  I don't think we can do that.  I mean I think it is a good research study to do, but I think somebody ought to apply to NIH.  I don't think we should be asking the sponsor to try to do something that is off label.

DR. COMEROTA:  It is part of the condition, what you are saying, it would be appropriate to document that platelet inhibitors, what anticoagulants the patients are on, and monitor the outcome as a registry.

DR. KONSTAM:  That's fine.  If there is no objection, I would amend it to say that special attention should be placed in the postmarketing survey to examine rates of thromboembolic events with an eye toward considering subsequent investigation based on what events rate is seen.

DR. LASKEY:  I think that is fine in concept, but there is not going to be a DSMB, there is not going to be adjudication committee.  This may be a tough one.  I agree with you it is important, but it could be an under- or an over-estimate unless somehow it is ascertained appropriately or accurately.

DR. KONSTAM:  Can't the FDA--

DR. LASKEY:  You need to make that recommendation clearly in this condition.

DR. ZUCKERMAN:  Yes, in the condition of approval, you can also ask that an independent CEC, similar to the original randomized study, continue to look at events to better clarify what are actual rates.

DR. KONSTAM:  I could live with that.

DR. LASKEY:  The registry's requirements are getting longer.

DR. KONSTAM:  That's fine.  I don't have any objection.  That doesn't bother me.

DR. LASKEY:  Are we all clear on what we are voting on then for Condition No. 4, which is the recommendation to the Agency in terms of establishing a registry for purposes of postmarketing surveillance to assess rates of--

DR. KONSTAM:  --survival, device failure, and major adverse events including thromboembolic events.

DR. LASKEY:  All in favor?

[Show of hands.]

DR. EWING:  So, everyone except for Dr. Wittes?

DR. WITTES:  I am going to abstain.

DR. LASKEY:  And your last condition was the instructions for the patient information pack?

DR. KONSTAM:  Yes, that the patient information package really--

DR. LASKEY:  Serious buffing up.

DR. KONSTAM:  Yes, make clear the risks and what the patient is getting into.

DR. LASKEY:  So, rates of adverse events and certainly perhaps the addition of some help rewriting it.  It is highly technical, I will agree.

Can we vote on that?

DR. AZIZ:  The device once it's explanted, should be sent to a center.

DR. LASKEY:  I like that idea.  I don't know how you can mandate-- you mean explanted--

DR. AZIZ:  When the patients die.

DR. LASKEY:  Well, when they die, doesn't that require consent or permission?  I don't know if you can mandate that.

DR. DOMANSKI:  I wonder if they are not getting those devices back anyway.  Maybe we could ask them that, because that may be an easy one actually.  Are you getting them back?

MR. MIDDLEBROOK:  Yes, we do like to get the devices back, and we do take them apart and disassemble them, and we do an analysis on them.  We collect that data and analyze it periodically.

DR. LASKEY:  Can you require that, though?

MR. MIDDLEBROOK:  I don't think it can be required because they refuse to return it, and we can't mandate that.

DR. AZIZ:  That is really the only way, if these devices are going to be in for a long time.

DR. LASKEY:  I would agree.  I can see a family just refusing permission.

DR. AZIZ:  That is a different issue, but I think all efforts should be made, because that is the only way, if these devices are in for four years or three years, we are going to learn something.

DR. LASKEY:  I would agree.  Should that be in the patient package then?

DR. DOMANSKI:  I think it is great to inform.  I was actually part of a trial where we did that once.  I would stop short of informing them about what should go on at their autopsy.  That probably is unreasonable.

DR. LASKEY:  So, voting on the condition for the modification of the patient information package as currently written.

DR. KONSTAM:  Let me just add to that, that based on what we see right now, there should be some indication of the limited present life expectancy of this device.

DR. LASKEY:  Full disclosure.

All in favor?

[Show of hands.]

DR. EWING:  That is unanimous for the last condition presented so far.

DR. LASKEY:  That's the end of Dr. Konstam's list.

I shudder to ask this question, but are there any other conditions?

[Laughter.]

DR. LASKEY:  No.  Well done.

DR. EWING:  I would like for the panel members to go around now and just state their vote for approvable with conditions.  We might as well start with Dr. Wittes again.

DR. WITTES:  Yes.

DR. DOMANSKI:  Yes.

DR. KONSTAM:  Yes.

DR. COMEROTA:  Yes.

DR. NISSEN:  Do you want yes or no, or do you want explanations?

DR. EWING:  I want a yes or no first.

DR. NISSEN:  No.

DR. AZIZ:  Yes.

DR. PINA:  Yes.

DR. OSSORIO:  No.

DR. DeWEESE:  Yes.

DR. KLOCKE:  Yes.

DR. EWING:  So, that is 8 yes and 2 no.

DR. LASKEY:  If each panel member could take 60 seconds to defend their position.

DR. WITTES:  A short essay, right?

DR. LASKEY:  Well, short.

DR. WITTES:  To me, the data showed convincingly a sizable benefit, mortality at one year, but the device is risky, and so I believe the patients and their families should be informed of the likely risks and the time course of those risks.

DR. DOMANSKI:  I think they demonstrated safety and efficacy per the statutes.

DR. KONSTAM:  I think that survival was extended.  I didn't hear any disagreement about that.  I think there were sizable numbers of patients who it appears from the data set were doing fairly well with a very, very high percentage of mortality in the control group.

I am concerned about the adverse events that were seen.  I am concerned about the durability of the device and the fact that there was such low survival at two years, but that was not enough to negate the basic underlying finding of efficacy in terms of survival.

DR. COMEROTA:  We have a treatment that doubled survival in one year with no device failure.  There is an accompanying body of data that shows substantial improvement in the functional status of the patients albeit it that has come under some criticism.

There was a parallel increase in quality of life.

DR. NISSEN:  I think this is a very promising therapy for heart failure, but this device is not ready, and I find it hard to accept a device that during an average duration of survival of 400 days had a 30 percent internal failure rate.

The reason I am so concerned about that is that we all know the impact of approving a device, is there is a trickle-down in the device to patients very quickly beyond the group of patients studied in a trial like this, and labeling does not protect patients from that.

That trickle-down would be less concerning to me if the device were reliable, but the problem is it is not, and so if you put this device in patients that have a 92 percent risk of dying in two years, it is not too bad a bargain.  If you put this device in patients that are a little bit less sick, now what you have done is replaced good medical therapy with not such good surgical therapy.

So, I think getting the device reliability up to a higher level of reliability would be essential for this to be a meaningful advance in the clinical treatment of such patients.

I think that that can be done, and I think it will be done, and I am concerned that based upon the outcome in 67 patients with a very high device failure rate, I do think we have opened Pandora's Box, and I don't want to be the purveyor of doom and gloom, but I think this will be a decision that many people in the medical community will come to regret.

DR. AZIZ:  I think that device therapy as an alternative to transplantation is going to be here to stay. I echo some of the comments of my colleague that I think the device should be carefully monitored.  I think it does have a fairly high incidence of malfunction or dysfunction.

I think at one year, clearly, it works well, but I think it will be very important in the postmarket surveillance to watch that carefully, that this really does not get out of hand.

DR. PINA:  I share Dr. Nissen's concerns about the device falling into the hands of people who are ill equipped to use them, and who are not going to apply proper medical therapy to the advanced heart failure patient.  However, that does not negate the survival number at one year, which was their endpoint, and they have met it, and therefore, I don't think that we can sit here and say no, do not approve it.

I do share some of the very concerns that Dr. Nissen has expressed, and I think we have couched this with a lot of conditions, and I am hoping and really hoping beyond hope that the company will take this to heart and apply it only to those populations that need it, and put it in centers where people do know what they are doing, and try to control the trickling-down effect which we have already seen with other devices.

DR. OSSORIO:  While I was impressed by the survival numbers, I am still not convinced that there is real clinical significance, and until I see more and better data, and I feel very torn about this and knowing that other people had already voted in a way that it would be approved with these conditions, it left me open to express my feeling that this really should have come to us a little bit later where there were more data for us to evaluate.

I think the number of conditions that were put on this actually was part of what convinced me that a lot of the other panel members feel strong discomfort about what was before us today.  I think we should be voting to approve or not based on what we have seen before us, what we could evaluate.

We don't have I think adequate Quality of Life measures to evaluate.  We don't have adequate description of what kind of training and labeling there is going to be to evaluate.  So, I voted no.

The other thing is that part of our assessment of safety has to be an assessment of whether the harms have been reduced as well as they could be, and I think that is where the adverse events data, and so forth, call into some serious questions with respect to the standards that we are supposed to be applying.

DR. DeWEESE:  I am convinced by the survival information.  I feel confident that the final group deciding these things will be sure that these are performed in transplant centers by capable people, and that it will be limited to people who cannot be transplanted.

DR. KLOCKE:  My question is that the long-term beneficial effect of this device will depend crucially on the degree to which the current incidence of adverse events can be reduced.  If it can't be reduced, the data we have will still stand, but my hunch is that after a period, that clinical acceptance will be in fact limited, and if the genie does get out of the bottle, we make go through a period, as we did with transplant, where we have a disappointing experience because the genie out of the bottle.

That will be unfortunate, but it will be transitory, but I believe that the key event for this long term is the degree to which we can improve further on the incidence of adverse events.

DR. OSSORIO:  Can I say one more thing?

DR. LASKEY:  Yes.

DR. OSSORIO:  I also just wanted to say that part of the reason I am very torn about this is because in other contexts, in the cancer context, I have dealt a lot with patients who are really at the end of life, and I think it is very important for us not to be making value judgments as to whether or not one year is long enough.

If I really believed the data, or it's not that I don't believe the data, if I really had enough data to convince me, then, I would have voted for approvability, but giving patients additional choices, especially additional choices with lots of uncertainty, is also a burden to them.

It is not merely a benefit to have a lot of choices.  It also adds then responsibilities to their lives, decisions they have to make, which we have data on the fact that these decisions are very stressful for people to make.

So, I don't want it to seem as though I am insensitive to the needs and desires of patients, but I think it is not always doing anybody a favor to give them an option where we can't tell them enough about it, and we can't tell them the kinds of things they would like to know to help them make their decision.

DR. LASKEY:  Mr. Dacey or Mr. Morton, any final thoughts?

If not, I would just like to reopen for one last time the public forum.  Are there any public comments?

[No response.]

DR. LASKEY:  If not, then, I thank all participants, in particular the presenters, for their participation and for their endurance.

I close this portion of our panel meeting.

[Whereupon, at 5:50 p.m., the proceedings were recessed, to be resumed on March 5, 2002, at 8:00 a.m.]

- - -