UNITED STATES OF AMERICA

DEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR DEVICES AND RADIOLOGICAL HEALTH

+ + + + +

CLINICAL CHEMISTRY AND CLINICAL

TOXICOLOGY DEVICES PANEL

+ + + + +

MEETING

+ + + + +

TUESDAY,

NOVEMBER 14, 2000

+ + + + +

The Panel met at 9:00 a.m. in Salons C and D of the Gaithersburg Hilton, 620 Perry Parkway, Gaithersburg, Maryland, Dr. Martin H. Kroll, Chairperson, presiding.

PANEL MEMBERS PRESENT:

MARTIN M. KROLL, M.D., Chairperson

DONNA M. BUSH, Ph.D., D-ABFT, Guest/Federal

Liaison

JAMES EVERETT, M.D., Ph.D., Member

CASSANDRA E. HENDERSON, M.D., Member

THOMAS L. KURT, M.D., M.P.H., Consultant

FRED D. LASKY, Ph.D., Industry Representative

SHERWOOD C. LEWIS, Ph.D., Consultant

BARBARA R. MANNO, Ph.D., Member

STANLEY  M. REYNOLDS, Consumer Representative

ARLAN  L. ROSENBLOOM, M.D.,  Member

DIANA G. WILKINS, Ph.D., Consultant

VERONICA J. CALVIN, M.A., Executive Secretary

FDA REPRESENTATIVES:

ALBERT PEACOCK, PH.D., Scientific

Reviewer Clinical Chemistry and Clinical 	Toxicology Branch

STEVEN I. GUTMAN, M.D., M.B.A., Division

Director

JEAN M. COOPER, M.S., D.V.M., Branch Chief

CLARA SLIVA, M.P.A., Acting CLIA Coordinator

LINDA R. MANLEY, Secretary

PUBLIC SPEAKERS:

DAVID BRILL, M.D., M.A., M.P.H

ROBERT DUPONT, M.D. (MRO),

DUPONT ASSOCIATES, P.A.

THOMAS CAIRNS, PH.D., D.SC., Senior

Scientist, Psychemedics

CARL SELAVKA, PH.D., Director,

Massachusetts State Crime Lab

RICHARD NEWEL, M.A., Professor of

Statistics, Research, Methodology and

Forensics, University of South Florida

WILLIAM THISTLE, ESQ., Vice President

of General Counsel, Psychemedics

I-N-D-E-X

AGENDA	PAGE

Call to Order	26

Opening Remarks, Introductions, Conflict of	26

Interest Statement

Open Public Hearing

David Brill, M.D., M.A., M.P.H.	31

Robert DuPont, M.D. (MRO),

DuPont Associates, P.A.

Sponsor Presentation

Thomas Cairns, Ph.D., D.Sc., Senior	68

Scientist, Psychemedics

Carl Selavka, Ph.D., Director,	103

Massachusetts State Crime Lab

Richard Newel, M.A., Professor of	53

Statistics, Research, Methodology and

Forensics, University of South Florida

William Thistle, Esq., Vice President	159

of General Counsel, Psychemedics

FDA Presentation

Albert Peacock, Ph.D., Scientific	178

Reviewer Clinical Chemistry and Clinical 	Toxicology Branch, Division of Clinical 	Laboratory Devices

Open Committee Discussion	130

Open Public Hearing	242

Closing Remarks	254

P-R-O-C-E-E-D-I-N-G-S

9:12 a.m.

DR. KROLL:  I am Martin Kroll, and I am a Chairperson for this session.  I'd like to call the meeting to order.

And first we have some opening remarks from our Executive Secretary, Veronica Calvin.

MS. CALVIN:  Good morning and welcome to this meeting.  Today the Committee will discuss and make recommendations on a pre-market notification 510K for a first of a kind prescription use screening device in human hair.

When yesterday the Committee discussed prescription use and OTC use guidance documents for drugs of abuse, and information where you can obtain transcripts or summaries from the meeting is on a salmon colored sheet outside on the table, and that also goes for today.

Also, we had indicated persons may submit comments in writing.  And where you can send that is on a green sheet outside on the table.

At this time, I would like to ask the panel members to introduce themselves starting with Dr. Lasky.

DR. LASKY:  Fred Lasky, Director of Regulatory Affairs, Ortho Clinical Diagnostics.  I'm the Industry Representative.

DR. REYNOLDS:  Stanley Reynolds, Supervisor at the Immunology and Virology Pennsylvania Public Health Laboratory.  I'm the Consumer Representative.

DR. EVERETT:  James Everett.  I'm the Director at Madison Memorial Health Care and member of the panel.

DR. WILKINS:  Diana Wilkins, Assistant Director of the Center for Human Toxicology and Research Associate Professor of Pharmacology and Toxicology at the University of Utah.

DR. KURT:  Tom Kurt, Medical Toxicologist and MRO, who is a founder at the Certified North Texas Poison Center at Parker Memorial Hospital in Dallas, and a Clinical Professor of Internal Medicine at that Institution.

DR. KROLL:  Martin Kroll.  I am the Director of Clinical Chemistry at the Dallas VA Medical Center, and I'm an Associate Professor of Pathology at the University of Texas Southwestern Medical Center in Dallas.

DR. MANNO:  I'm Barbara Manno.  I'm a Professor of Psychiatry at the Louisiana State University Health Sciences Center in Shreveport, Louisiana.

DR. LEWIS:  I'm Sherwood Louis, Director of Toxicology at the Office of the Chief Medical Examiner in the state of Connecticut, and faculty member in the Department of Laboratory Medicine at the University of Connecticut Health Center, consultant to the Committee.

DR. CLEMENT:  Steve Clement, Associate Professor in Washington, D.C. at Georgetown University and practicing Endocrinologist.

DR. ROSENBLOOM:  Arlan Rosenbloom, Distinguished Service Professor of Emeritus at the University of Florida of Pediatrics, and Director of Children's Medical Services in Gainesville.  I'm a regular voting member of the panel.

DR. GUTMAN:  I'm Steve Gutman.  I'm Director of the Division.

MS. CALVIN:  I will now read the conflict of interest statement.  The following announcement addresses conflict of interest issues associated with this meeting and is made part of the record to preclude even an appearance of an impropriety.

To determine if any conflict existed, the Agency reviewed the submitted agenda and all financial interests reported by the Committee participants.  The conflict of interest statutes prohibit special government employees from participating in matters that could affect their or their employer's financial interest.

However, the Agency may determine that participation of certain members and consultants, the need for whose services outweighs the potential conflict of interest involved, is in the best interest of the government.  We would like to note for the record that the Agency took into consideration certain matters regarding Drs. Martin Kroll, Arlan Rosenbloom, and Diana Wilkins.

Drs. Kroll and Rosenbloom reported unrelated interests with firms at issue.  Since the interests are unrelated to the issue before the panel, the Agency has determined that they may participate fully in today's deliberations.

Dr. Wilkins reported an imputed interest with a firm at issue.  Since the interest is imputed through her employer, and is not her personal direct interest, the Agency has determined that she may participate fully in today's deliberations.

In the event that the discussions involve any other products or firms not already on the agenda for which an FDA participant has a financial interest, the participant should excuse him or herself from such involvement, and exclusion will be noted for the record.

With respect to all other participants, we ask in the interest of fairness that all persons making statements or presentations disclose any current or previous financial involvement with any firm whose products they may wish to comment upon.

Thank you, and I'll turn the meeting back over to Dr. Kroll.

DR. KROLL:  All right.  Thank you.

At this part of the meeting, we have our open public hearing.  And as you can see in the agenda, public attendees who have contacted the Executive Secretary prior to the meeting, will address the panel and for that information relevant to the agenda.

I'm urging the speakers to state their name, where they're from, and whether or not they have any financial involvement with the manufacturer of the product being discussed, or with their competitors.

The first speaker we have on the agenda is David Brill.

DR. BRILL:  Good morning.  I'm Dr. David Brill.  I'm Boarded in Internal Medicine and Occupational and Environmental Medicine.  Until June, I was the Corporate Medical Director at Michelin of North America.  That's the tire company without the recall.

And I was also a representative on the hair testing work group for Donna Bush, who I see isn't here today, which was evaluating hair drug testing for the Drug Testing Advisory Board, SAMSHA.  I represent an industry interest in that group.

I have a conflict of interest.  I do have shares of diagnostic laboratory companies.  I do not have any interest whatsoever in the laboratory here today.

What I want to tell you today is that hair is critical -- absolutely critical component of the employer's tool box.  And I'll share a couple of anecdotes as to why.

So that you understand, hair testing for drugs tests long term use.  When you clip a hair from the head, an inch an a half tests about three months worth of drug use.  It cannot test acute use.  It takes time for the blood, or any involvement of drug with hair, to come into the hair.

It simply cannot test for acuteness.  It can only test for chronic use.  Obviously, it can also not test for immediate impairment.  Again, because it takes days to move from the blood to the hair.

However, for employers who want to use effective tools in the workplace, hair is critical because of its effectiveness and especially its transparence.  And what do I mean by transparence?  You'll see in a few minutes.

My past to this position right here in front of you today began shortly after I began at my former employer.  We used urine testing, and I've been taught that you don't use hair testing in MRO training because it's a new technology which is not method tests and so on.

Shortly after I came, we had a 26 year old young man who passed his screening urine test.  He was a contractor.  Again, he passed his scheduled screening urine test.  He passed his scheduled screening urine test, and then two days later, died on the floor of a cocaine overdose in the plant.

I repeated that three times.  After we put in hair testing for new hires, after investigating it thoroughly, both on scientific, legal, and other basis.  After we put in hair testing, suddenly our positive rates with no other changes in the program tripled the same number of times that I just repeated that statement.

Every time that we were doing the urine test, two other -- and finding one positive -- two other people who were using drugs were entering the company routinely.

Two years after that, we had implemented a random hair testing program where we had the CEO of the company giving the first sample with a video tape of that event distributed to all employees, and required to be viewed by all employees so that they could see that everybody was in this together.

And it starts with the top, with the CEO actually giving the sample with one of my nurses taking the collection.  Try that with urine.  In fact, at the time in Oklahoma, we had a plant where we couldn't use drug testing.

The Oklahoma Board of Health has now changed that.  We can use hair testing in Oklahoma.  But at the time, we had to put in a urine random testing program.  And there the plant manager, when coming to the same moment, had to discreetly turn around and walk away toward the lab where he would give his urine collection.

But you get the point.  Hair testing doesn't involve genitals.  It also can't be adulterated.  It also can't be substituted because it's a witness collection.

In urine testing, the person is alone with the sample 95 percent of the time.  And of course, you know what wonderful creative things are on the Internet.  Many of you may or may not, but for instance, 30 miles away from where our headquarters were, there was a creative young gentleman who was marketing his own urine on the internet.

And further, he had developed a device which would strap the urine to the leg, warm it to body temperature so that it could pass SAMHSA temperature requirements.  And then move along the genitals, so that even if it were a witness collection, you couldn't see the tube.

I was asked to comment on this.  And frankly, I said it's essentially irrelevant to us because we've used a hair testing program.

Now, let me help you understand workplace drug testing if you're not familiar with it already.

There are essentially two critters here.  Scheduled tests and unscheduled tests.

The unscheduled tests tend to be reactive.  That's if there's an accident in the workplace and you have a trigger for a drug test after that accident.  That's catching -- or perhaps catching something after the damage has been done.

And 40 percent of the time, that's involving harm to someone not using the drugs.  It's not the person who's driving the forklift who gets hurt.

Reasonable suspicion testing is another one.  If someone is tottering in the plant, and you have a drug test, that's a reactive.

Now, the scheduled tests are proactive, or can be viewed as proactive.  That would be an applicant drug test, and let me again repeat, as an applicant drug test, that's not a new hire drug test.

Because Congress expressed its will very clearly in the Americans With Disabilities Act saying that a drug test is not a medical test.  In fact, they went so far as to say that it can be done before an offer is made.  Whereas, every medical test has to be done after the offer is made to the applicant.

So you say applicant drug testing, not new hire drug testing, because you can do the drug testing before as per the will of Congress.

The second scheduled type of test is random or periodic testing where an employer either periodically says, "We'll test at such and such a time," or, "We will have random testing."

Let me describe to you what the difference is for an employer between random urine testing and random hair testing.  I ran five random urine testing programs under the DOT airline -- you know, aircraft pilots -- everybody who puts a tire or who makes a tire for an airplane is covered.  Anyone who puts a tire on a bus is covered, and so on.

If I'm an employer, or if I'm the Medical Director of the company, and I'm running a random urine testing program, I am immediately going to be a secret son of a gun.  I'm playing the keystone cops, because I have to keep that darn list secret.

Because if anybody knows that they are going to be either giving a random urine test, they can tip the person off.  Or, if they're going to be receiving the random urine test, obviously they have a thousand and one ways to adulterate and substitute and so on.

So, we have to be secret, which is a pain.  You don't want to be secret with your employees and your staff.  So, the list has to be really, really secret.

Second, everything has to be incredibly abrupt.  Once it goes down to the nurse that there is going to be a urine test, they have to go out on the floor, grab the person off the floor.

Usually there's two people off the floor.  The person whose coming off of the job, and another person to accompany them back to the laboratory and make sure they don't stop at their locker, don't stop at the drinking fountain, da da da da da.

So, you have two people off the floor and the Supervisor, who is a little PO'd that he suddenly lost somebody and he has to scramble to get somebody to get on that machine.  It's very abrupt.  It's really not fun.  A lot of workplace disruption.

So, you've got these two people off the floor.  They come back.  You have to have it in a place with a laboratory because genitals are involved.  You have to have security in the laboratory as witnessed every time, except for the moment when the person is alone with the sample, of course.

And then the sample is put in a Fed Ex box or whatever -- three samples typically, and you send the biohazards through the mail.  If it happens to be the CEO on the other hand, if you want to have a fair program, and the CEO is just about to fly off to General Motors to clinch that contract and his number or her number comes up, you have to say, "I'm sorry.  You really can't fly on that plane right now.  You really in all fairness to all the other employees, you have to give the sample now just like anybody else.  Oh, you just peed?  Well, here's eight ounces.  Drink this and we'll wait.  We'll just keep the pilots waiting."

That's urine testing in a random person.  That's the reality.  Now, look ahead.  I don't have to keep it a secret.  If anybody gets that list, three weeks in advance, I don't care.  We just took the list, put it on an Excel spreadsheet, send it out, anybody can know who's on that list a couple weeks in advance.

In fact, there's some preferences to do that because then they can schedule their vacations or whatever else is going on.  If somebody comes back, the Supervisors have no trouble with people off the floor.

They schedule it in advance.  You don't need two people off the floor at any time.  You don't need a laboratory.  You don't send biohazards through the mail.

And if somebody says, "Oh, I'm on vacation that day."  Or the CEO is heading off to Detroit, just say, "Sure.  Schedule it to come back at your convenience any time in the next couple of weeks."  No big deal.

You can't substitute the sample.  You can't adulterate the sample, and it's a long window of testing.  So, it's wonderfully transparent.  It's beautiful for an employer, because you can be really up front with employees.

That's what I wanted to tell you.  Along with short window tests, which are wonderful.  For instance, saliva, urine, breath for unscheduled tests, reactive tests.  Hair is great for scheduled tests.

In fact, if you do have a scheduled test, everybody knows how to beat it with urine.  If you're going to have a scheduled test, really employers have very little choice but to use something like hair.

In fact, hair is the only choice available.  The long range winners in the new media will probably be hair, saliva, followed by urine -- or breath, urine, and blood.  Each one with a descending level of ability to substitute the sample or difficulty in attaining the sample.

So, what I wanted to tell you today is that hair is a critical component of an employer's tool box.  Again, it has wonderful uses.  It's effective as opposed to the others.  And it's transparent.  It lets you be a human being with your employees.

Thank you very much.

DR. KROLL:  Thank you.

The next speaker is Dr. Robert Dupont.

DR. DUPONT:  Thank you, Mr. Chairman.  I apologize for the voice.  I got all prepared except for one problem for this, and that was I got laryngitis last night, so I apologize for any difficulty you have hearing.  And please put up a hand and I will say it again.  It's distressing but not painful, so I'm happy to mostly to listen.

Let me say about the conflict of interest issue that came up that I do own stock in the Psychemedics Company and many other companies, and I also serve as the Scientific Chairman of the Scientific Advisory Board for the Psychemedics Company.

I have been a strong advocate for drug testing in general, including hair testing.  And I'm also active in supporting all the other companies that promote or develop hair testing.  And I don't think of myself as an employee or an advocate particularly for Psychemedics, but for drug testing generally.  And not even just hair testing.

Now, my background -- I began in this field more than 30 years ago working for the District of Columbia government, and drug testing was central to what we did.  We did a study in D.C. jail using urine testing that identified heroin use as a major factor in the crime problem in the District at that time.

That was published in the New England Journal of Medicine, and was one of the central pieces of evidence in the identification of the heroin epidemic in the early 1970's and the late 1960's in this country.

And I have been active in the drug testing field generally since that time.  I was the first Director of the National Institute on Drug Abuse, the second White House drug czar, then called the Special Action Office for Drug Abuse Prevention.

And since 1982, I've been Vice President of Benzinger Dupont and Associates, a national consulting firm advising workplace programs dealing with drug and alcohol abuse.  And in that role, I've been a Medical Review Officer, and have been active as a Medical Review Officer not only supervising other physicians, but also doing medical reviews myself.

And I speak primarily as a Medical Review Officer to you today.  I came to the Psychemedics Company into hair testing from working in my own practice with my own patients who were addicted.  And to send off samples of their hair to be tested when I knew what their drug use was, including sending samples of people who I knew were not users.

And I was impressed by the fact that the hair testing got the information right.  As Dr. Brill said, the phenomena of a long surveillance window of 90 days as opposed to three days -- one to three days for urine testing -- was very attractive from a clinical point of view in my work.

But also the resistance to cheating.  I want to emphasize that that also is a very important factor.  And it was on the basis of that personal experience as a physician with my patients that I contacted the Psychemedics Company and asked them what I could do to help encourage wider use of hair testing in this country.

Now, I wanted to say two things by way of background with respect to the opiate testing, which is our subject today.

First of all, drug testing is extremely important to demand reduction in this country.  It is absolutely central to the national effort to reduce the demand for drugs.  This is the most important part because a central element of the drug experience is dishonesty -- is lying about it.

And the .8 million current illegal drug users in the country are spending 63 billion dollars a year on drugs.  And the central technique for being able to reduce that demand is drug testing linked to consequences in all areas of our life including workplace, which is what we're talking about primarily today.  But the criminal justice system is another very important area.

The second point of background that is going to come up in a moment that may not be obvious is that heroin is a relatively new product.  It was introduced in 1898 by the Bayer Company, by the simple manipulation of adding two acetyl groups to morphine.

And the addition of those acetyl groups was very important in terms of what we're going to talk about in terms of the identification of heroin use through hair testing because of the identification of 6-MAM in that.

That was introduced by the Bayer Company at the same time they added one acetyl group to salicylic acid to create aspirin.  And that's why the two names heroin and aspirin are so similar.  Because they were introduced at the same time by the Bayer Company.

With respect to the opiate testing, as an MRO, the practical matter is -- and practical reality is that all opiate positive are reversed on MRO review of the opiate positive.  The only exceptions are if a heroin user admits use freely.  I have never had that happen as an MRO.

And the other possibility is that you identify 6-MAM in the urine result.  I have rarely seen that happen.  Almost all opiate testing using urine is reversed by the Medical Review Officer which means urine testing is almost completely ineffective at identifying heroin use in workplace drug testing.

That's very important to understand that that is the reality that is faced today by anyone who's using urine testing to identify heroin in workplace or other testing.

And that is because of the poppy seed problem.  Poppy seeds contain morphine.  And they will trigger a positive test on a urine test.  And because of the way the testing is done, if there is any possibility that the testing could be innocent, even if the person does not claim to have eaten poppy seeds in the prior three days, the MRO is duty bound to declare that test a negative test.

Now, hair testing is very different in two regards with respect to this problem with the poppy seeds.  First of all, poppy seeds do not produce a positive hair test even when large amounts of poppy seeds are consumed.  They do not produce a positive test at the cut off levels that are traditionally used.

And sencon, 6-MAM rather than being rarely identified, is routinely identified on hair testing.  That means that hair testing can identify heroin use very efficiently.  And urine testing cannot identify.  That is very important in terms of this particular application of this technology.

Now, I want to mention three areas of -- let me call it concern about hair testing.  The first has to do with the concern about external contamination as an explanation for a positive result.  		And in thinking about this, I call your attention to the fact that heroin is not a common contaminant in the environment.  It is not easy to get heroin on your hair or anywhere else.  So, that's probably the most important point.

Second of all, should one get heroin on his or her hair, the wash techniques that are used by Psychemedics of a minimum of three hours and 45 minutes of wash are sufficient to exclude external contamination as a source of a positive test.

I think that that's very important about the external contamination -- that the wash is very important to exclude that rare possibility.  An example being a police officer working in a narcotics unit where heroin is an environmental contamination.  That is important as a protection.  So, the external contamination issue is dealt with by the washing.

The second issue has to do with bias with respect to hair color.  This is very important that you understand this.  First of all, there is no possibility of a person who does not use having a positive test as a result of hair color, or race, or any other factor.

So, non-use is zero for everyone.  The only argument exists about if a person has used the drug heroin, is the concentration in the hair the same or different depending on the hair color.  So, this only relates to people who have used heroin.  And I want to emphasize that point.

Now, with respect to the question of whether there is a "bias" based on hair color, I call your attention to other drug testing, which is very extensive including the most striking example to me, is alcohol testing.

In alcohol testing, we make no effort to level the playing field in the sense that each alcohol drinker can drink the same amount to get the same blood level.  We know that gender is a major factor in the relationship between blood level of alcohol and alcohol consumption.

But we also know that weight is a major factor.  And there are dozens of other factors.  And there is no attempt to normalize or equal the amount of alcohol consumed in terms of declaring a positive test.

The same thing is true with urine testing.  The biggest determinant of whether a person is going to reach that cut off, or at a close call which is only what we're talking about by the way for hair either -- at a close call, near the cut off level, is how much fluid the person has consumed.

We don't make any attempt at urine testing to normalize for fluid consumption.  The cut off is the cut off regardless of how much you've consumed.  And the other determinant for urine testing is how many hours it's been since you've used.

We don't make any attempt to try to normalize that, to say a person who has used six hours ago is going to be judged equal for someone who used one hour ago.  None of that is done as we use a per se standard for both urine testing for all drugs, and for breath testing and blood testing for alcohol.

No attempt is made to level that playing field based on the amount that the person has consumed.  And that's very important because I think this whole argument is without precedent in the drug field, and throws a curve ball into the equation that is very harmful to the process, and bears no relationship to some sense of equity.

Now, in addition to that, the -- and you will hear more about this -- the tests that have been done looking at hair tests around the question of racial bias that is now among the people who have used the drugs, is there a difference in positive rates between urine self report and hair have shown that there is not.

Whatever the laboratories studies show, this is the result of a real world study.  And it shows no difference in the ratio between hair, and urine, and self report.  That's very important data to take into consideration about this.

Now, my third concern has to do with cheating.  And that is -- I think those of you who are not expert in the drug testing area -- Dr. Brill did a wonderful job of explaining the fact that most drug testing in the workplace is pre-employment testing, as he called it applicant testing.

That is a scheduled test.  That is more a test of intelligence than it is a test of drug use.  The window for a urine test is one to three days.  That's very important.

That means that the major testing that is done in the workplace is ineffective in a scheduled test with a one to three day window.  All the user has to do is not use for one to three days, and the test is negative.

It is also very easy on a scheduled test to consume extra fluids.  Never mind the complicated kinds of problems as he was talking about a sample substitution.  The problem of cheating in urine testing is an enormous problem.

And it is particularly a problem for the most serious drug users.  The ones who are most important to identify are least identified by urine testing because of the ease of cheating on the test.

Hair testing cannot be cheated on as you heard.  There is no possibility of sample substitution.  No issue of additional hydration, for example, or using an adulterant.  These are not relevant matters.  This is extremely important in terms of the integrity of the test and the effectiveness of the test in being able to identify drug use.

Now, for these reasons, I believe that hair testing deserves a place along with urine testing and other testing.  Because I think saliva testing is particularly important in an application like highway testing, which as you may or may not know, there is no testing now for drugs in the highway which is a terrible national problem right now.  And saliva is the obvious means of doing that.

But hair testing does belong in there with urine testing.  And of the hair testing, the opiate test is particularly important for the identification of use.  In fact, I think that one of the things that would be a very practical thing for somebody who has a positive opiate urine test where there is any concern whatsoever, is to require a hair test.

Because as Dr. Brill said, you've got 90 days.  And the fact that you get a positive urine test makes it fairly easy because you can't substitute anything.  You've got 90 days to look at it.

So, I think hair testing does add substantially to the ability to detect drug use.  And it should be approved in the national interest to be able to reduce the demand for drugs.  And it is in that spirit that I address you.

Thank you very much.

DR. KROLL:  Thank you.

Now we proceed to the sponsor presentation from Psychemedics Corporation.  The first speaker is Dr. Thomas Cairns.

MR. THISTLE:  Actually, the first speaker will be me.  I'm Bill Thistle.  I'm Vice President and General Counsel for Psychemedics.

And before we start the presentation, I'd actually like to go through and introduce the people that we have that will not necessarily be presenting.  You should have a list of presenters there.  But we brought people who are available to answer questions, and I'll just go over them briefly so you know who is here.

And the people here, just acknowledge who you are.  We'll get that out of the way first.

We have Werner Baumgartner, who is a Scientific Director of Psychemedics Corporation.  He is a founder of Psychemedics and a pioneered hair analysis for illicit drugs over 20 years ago.  Additionally, along with working with Psychemedics, he has served until recently for the last 24 years as Director of the Clinical Radioimmunoassay lab in Los Angeles for the VA Medical Center.

We have Thomas Cairns, Senior Scientist at Psychemedics, an adjunct Professor at Pharmaceutical Sciences at USC.  He served 20 years with the FDA.  Part of that time was as Director of the FDA National Center for Toxicological Research.  He continues to be a Senior Science Advisor to the FDA, and he was at one point the recipient of the FDA Office of Regulatory Affairs Regulatory Scientist of the Year Award.

We have Michael Schaffer, Vice President of Lab Operations at Psychemedics.  He's a licensed NIDA Inspector.  He was formerly Director of Toxicology at the Smith Klein Lab, and formerly Chief Toxicologist for the Office of the Medical Examiner in Cook County, Illinois, where he made the determination several years ago that the Tylenol scare we had of tampering occurred post manufacturer.

We have John Irving, Deputy Director of Laboratory Operations at Psychemedics.  Again, a licensed NIDA Inspector.  John Irving was a member of the first drug testing Advisory Board for NIDA, now SAMHSA.  He participated in drafting the initial federal drug testing guidelines and he was formerly Director of several Navy drug testing laboratories.

We have Virginia Hill, a Research Scientist with Psychemedics Corporation who worked directly with Werner Baumgartner on hair analysis at the VA Medical Center.

We have Carl Selavka, who's not with Psychemedics, but is Director of Massachusetts State Crime Lab.  He is co-chair of a hair testing working group created by the drug testing Advisory Board to advise SAMHSA on inclusion of hair testing.  He has published one of the seminal studies on poppy seed ingestion and the issues that that creates, and to serve as a consultant in that regard to HHS.

He was formerly Director of Forensic Services for New York's 25 public forensic science labs.  Formerly Operations Officer for the U.S. Army Toxicology Drug Testing Lab in Hawaii.  And has testifies as an expert witness in over a dozen military cases involving hair analysis.

We also have Richard Newel.  He's a Professor of Statistics Research and Methodology in Forensics at the University of South Florida.  Along with Dr. Mieczkowski, Mr. Newel has amassed one of the largest databases of its kind regarding race, drug testing, and drug prevalence.

We have Jonathan Ma, Manager of Biostatistics at Engine X Pharmaceutical Sciences, formerly a Mathematical Statistician with the FDA.

And myself, I'm General Counsel with Psychemedics.

We -- is there technical difficulties with this?  Is that why it keeps flipping?

As we approached this, we found it difficult to condense two decades of research and over a decade of use into a one hour presentation.  We tend to provide a background on hair analysis as done by Psychemedics, and then get into the submission specifics.

And already we're out of sequence here.   But our Corporation -- is that legible at all?

Our Corporation was founded in 1987 to analyze specifically drugs of abuse in hair following years of research by Dr. Baumgartner.  We test for the NIDA five drugs, opiates, cocaine, methamphetamine, including MDMA, ecstacy, which should be part of the covered guidelines in the future, just for marijuana and PCP.

We are licensed as a -- you should have this incidentally in front of you, so follow along as we go through this.  The audience can pretend they're watching this.

I think it would work better if we could get somebody clicking up there.  Is there a problem with the length of the cable?  Why don't we do that?  Why don't we have someone clicking up there.  We're going to take a couple minute break and just set it up.

(Whereupon, the foregoing matter went off

the record at 10:52 a.m. and went back on

the record at 10:56 a.m.)

MR. THISTLE:  We are a licensed high complexity clinical lab under federal CLIA standards.  We are CAP certified with the limited urine testing that we do.  We are licensed and certified by numerous states, including states that have specific criteria for hair analysis labs, Florida, Maryland, New York, Oklahoma.

We have three U.S. patents on aspects of the methodology unique to Psychemedics.  We have European, Japanese, and Canadian patents as well.  And we have over 1,800 corporate accounts from a wide variety and cross section of U.S. businesses.

And by corporate accounts, I'm not talking about doing one or two tests.  For instance, for General Motors, we do all their pre-employment at a 104 North American plants, all the North American operations.

We do Anheuser Busch, testing for Anheuser Busch, FIC, Rubbermaid, Toyota, BMW, Michelin, Goodyear, Whirlpool, Gillette, Steelcase, U.S. Steel, Federal Reserve Banks.  In addition to the major corporations, we also -- I know you thought it was going to be all 1,800, but in addition to those major corporations, we're also utilized by law enforcement agencies, including some of the nation's largest, NYPD, Chicago PD, Boston PD.

We're used in prisons in several states.  We're utilized by court systems in diversionary and first offender and probationary programs.  These are programs designed to keep people out of jail through hair testing.

We're utilized by state use and family service agencies, in child custody issues, chemically abused children.  We're utilized by schools.  We have over 100 schools in 26 states utilizing hair testing.

We've been upheld in federal courts.  This is a June 2000 decision.  The court agreed there was no scientific support for the claim that the plaintiff could have tested positive due to her race.

Over ten years ago, we were upheld in federal court where the court recognized that RA hair testing -- they're recognized through liability and acceptance in the field of forensic toxicology when used to determine cocaine use.

We've been upheld in state courts.  We have a number of police officer cases where there's a tremendous impact on society.  These are people involved in drugs and protecting the public.

We've been upheld in state appellate courts, where the courts have found that the radioimmunoassay analysis of human hair to determine cocaine use is generally accepted in the scientific community.

We've been upheld by state supreme courts.  This Nevada case concluded that hair testing, especially when coupled with a confirmatory GC/MS test was an acceptable and reliable scientific methodology for detecting illicit drug use.

We've been upheld by administrative agencies, unemployment boards, unemployment appellate boards.  We've been upheld in union arbitration decisions, including Teamsters, United Steel Workers, Operating Engineers, fairly tough unions.

Every issue that can be brought regarding a hair test has been brought and is brought, and we get decisions similar to this where they find Psychemedics' wash procedures are effective in removing environmental contamination.

The cut off level for cocaine is appropriate in the field studies.  There was no bias here on the basis of race or hair color.  The chain of custody was unbroken.

Many of these companies use hair analysis because of the advantages.  The hair analysis -- and you've heard some of them earlier today.  There's a wider window of detection.

Approximately three months for morphine, and six for acetylmorphine in hair.  Urine is 24 to 72 hours for morphine, several hours for 6 acetylmorphine.  Six acetylmorphine, or other clinical evidence of use is necessary to confirm heroin use.

We have significantly increased detection rates, because you're looking at several hours.  You're comparing several hours to several months.  In real life, the purpose of testing is for a third party to determine the likelihood of future use based on approximate past behavior.

It's certainly more effective and safer to the general public to make that decision based on a long period of time as opposed to a shorter period of time.

There are many reasons to use tests with a shorter period of time, but certainly in pre-employment, there's a reason to use a test with a longer period of time.  Many of our clients have done side by side hair and urine tests, and have found tremendous differences in the detection rates.

We are also a less intrusive collection.  It's a small snip of hair, the thickness of a shoelace tip.  People regularly will get their hair cut in public, at least with a barber present.  Sometimes in front of a plate glass window at the mall.  It is not likely you'll find people urinating in front of the plate glass windows at the mall.

In fact, the U.S. Supreme Court has long held that the collection of urine was a search, and found that there are few activities in our society more personal or private than the passing of urine.

In contrast, again the U.S. Circuit Court of Appeals Federal Court, concluded that the collection of hair in 1982 was akin to obtaining fingerprints and did not rise to the level of a search.  We are also, as mentioned earlier, an observed collection.

One of the biggest threats to urine testing is adulteration and substitution.  That can't be done at the collection site.  And these are just two companies off of hundreds that we've pulled off the internet.

The one on the left is the gentleman that provides the clean urine and heater pack.  And the tubing guarantees 100 percent satisfaction, because obviously a lab cannot tell the difference between if you submit clean human urine that's of the proper temperature when it's collected.  A lab cannot identify that.

The second on the right is an adulterant, Spectrum Laboratories.  The gentleman that runs Spectrum Laboratories has been interviewed on television, and he says, "I get the same publications as the urine labs.  Obviously when they begin to test for my products, I change my products."

We have the ability to repeat the collection.  And this is a huge, huge advantage for the employee or the donor.  You can collect a wholly new sample that can encompass the original time frame after the original result is known.

This eliminates concerns over sample error or lab error, or sample mix up.  It's not possible with liquid matrices as the original window of detection has passed by the time the result is obtained.

Getting a second urine sample four days later, or five days later, or three days later -- even two days later -- is worthless in terms of seeing what happened during that original time frame.

This provides unprecedented fairness to the donor.  We have repeatability with a new sample, with a new collection, with the union representative present.  Unprecedented fairness to the donor.

We identify 6 acetylmorphine, establishing heroin use.  Hair testing eliminates the poppy seed issue.  Our studies that have been submitted with this submission have shown that even massive ingestion of poppy seed products will not cause a positive hair result.

In contrast, in '95, HHS reported that 87 percent of confirmed urine opiate positive were overturned largely due to the inability to distinguish that use from poppy seed use.  That was at the cut off level of 300, which does not, as it turns out, protect the individual because it could be a poppy seed positive.

The level has been raised to 2000, which at this point may not protect the public, because now you've raised it to the point where you may be eliminating heroin users.

We also have ease of shipping and storage.  It's not a biohazard.  They have much more safety -- increased safety in shipping and handling the specimen.

There are several unique aspects to our submission.  It may be somewhat different than what you're used to seeing.  For one, it's been used in millions of tests over 13 years.  We have a track record of safety and effectiveness.

It is an in house assay, a home brew.  It's not sold to other labs.  And this is a test where the donor knows the results before taking the test.  The donor knows clearer than everybody what the results should be.

You're not advising the donor that they have cancer.  You're letting them know that they engage in activities that they know certainly well that they engaged in.

It's a forensic test that provides a third party with information on past illegal activity.  Heroin is a Schedule one drug.  This test does not show intoxication, under the influence, or addiction.  It's a forensic test.

At this point, I'm going to turn it over to Dr. Cairns who will get into more of the laboratory operations and the science.

Thank you.

DR. CAIRNS:  Good morning, Mr. Chairman, members of the panel.  I welcome this opportunity to do a presentation before you, not only on an overview of our hair test procedures, but also on the seven important questions that have been posed by the panel.

I need to go open my own presentation.  Apologize.  We used to get along so better with these slides in the old days, and I wonder whether technology should take a back seat and go back to slides.

However, what I wanted to do for you just for a few moments is to place into a proper perspective setting where the RA assay for morphine sits in our strategy of hair testing.

Now, the next slide will show basically that there are three major platforms here.  The first platform is in fact the primary stringing via RA to identify presumptive positive based on morphine.

If at that point, the assay -- the screen is negative, the sample is reported negative.  If however, that primary screen -- and by the way, that primary screen involves only eight milligrams of hair digested or liquified for the analysis.

And that's a five drug analysis panel, similar to NIDA.  The next step, if a presumptive positive is triggered by the screening assay, then we go into a second sample weighing from the original envelope that contains the hair and the chain of custody.

Except in this case, this second weight of hair is in fact extensively washed, washed to remove external contamination.  Once the hair is washed -- and I'll be going into the exact details of the wash later -- that hair is then liquified, it's digested, and sent on for structural confirmation.

Now Psychemedics has developed mass spectrometry that goes back to in fact just the simple gas chromatography mass spectrometry, but more recently has elevated the specificity by looking at these compounds, meaning morphine, 6-MAM, and codeine, by a liquid chromatography mass spectrometry, mass spectrometry.

So, what are we looking for?  We're looking for morphine and the heroin metabolite.  And in that process, we're looking at the analysis of the last wash in our review of the data before we release a positive.

Now, let me just go over the criteria for properly conducted hair testing.  We want the sample under chain of custody.  We want the digestion procedure, that is the process of dissolving the hair to release the drugs from the matrix to allow the screening.

But it's important to stress at this point, that you must use rigorous forensic standards that is both quality control and quality assurance.  Now, by that I mean that in the RA assay, we have in fact the standard at the cut off.

We have a high control at 25 percent above.  We have a low control at 25 percent below.

We have a negative and several blind samples in each batch.  Now if in fact that is then satisfied, the next step is the re-weigh and the washing procedure.

And that is an extensive wash lasting three and three quarter hours plus an analytical extrapolation to have five additional hours.  Then we move into structural confirmation.

Depending on the result, we will review the wash criteria.  And then the final review by a certifying scientist before the issuance of a positive result.

Looking at all of the QA/QC blanks and blinds that go into the batches, whether it's the RA batch or whether it's the GCMS batch.

Now, let me give you just one last slide on our overview.  We accessioned the hair.  That's the process of weighing the hair, giving it a bar code.  At that point, the paperwork is checked for valids and invalids, and fatals.  Maybe that the seal -- the integrity seal was not initialized by the donor.

Weigh the eight milligrams, liquify the hair, go to radioimmunoassay.  Five panel drug, marijuana, cocaine group, amphetamine group, opiates, and PCP.

The negatives are released, the presumptive positive move forward by a new sample being re-weighed and washed, then digested, and moving into structural confirmation by mass spectrometry, whether it is GCMS, or the more sophisticated tandem mass spectrometry associated with either gas chromatography or liquid chromatography.

Once that data is forwarded, the wash data is also reviewed in conjunction with the mass spectrometry result, and the certifying scientist reviews the positive before release.

Now, with that little background, let me move into the next part of the presentation.  And this is to focus basically on the issue before us today.  That is regarding the safety and effectiveness of an RA screening assay for morphine using a cut off of two nanograms or ten milligrams of hair.

Now, question one revolved around clinical studies.  If you were to use the traditional perspective clinical controlled study, there would be two major impediments, and one is an ethical problem.  That is, can you ethically administer multiple and or chronic doses of heroin for a period of 30 to 90 days to mimic drug abuse?

The second one is an analytical problem involved in the time frames between urine and hair.  And we've heard that's a few days, and the hair is 90 days.  The administration of a single low dose would not challenge the cut off of the sensitivity of the assay.

Therefore, the design of the clinical trials in Volume 3, which is a response to FDA, was that the study hypothesis was to demonstrate substantial equivalence to a urine predicate device.

Number two.  The positive clinical gold standard was a positive urine and a positive self report.  A negative clinical gold standard is a negative history of hair analysis and a negative urine.

And the inclusion exclusion criteria were the same for all the studies we submitted in this area of the application.  Now, from the statistical point of view, the standard error of clinical sensitivity and specificity has been determined and submitted.

The next slide will show you the five independent studies that were submitted to support the performance of the assay.  These are five heroin user field studies, and they are involving Study A involves 22 individuals, Study B, 93, Study C, 94, Study D, 20, and Study E, 16.

Now, I'd like to go through each of these in turn to explain perhaps a few more details on each of these clinical studies.  The first one was published in Journal of Forensic Science in 1989.  The hair testing cut off was two nanograms per ten milligrams.

Again, 22 subjects.  The use by self report was somewhere between 6.4 and 288 milligrams per day.  The morphine range found in hair was from the two nanogram level to 130 nanograms per ten milligrams of hair.

This permitted a mean detectable heroin dose use resulting in the value of 173 milligrams of heroin per month.  And the clinical sensitivity derived from the data using 17 true positive, and five false negatives, with 77.2 percent.

Study B, from the National Development and Research Institute to Mount Sinai Hospital and various clinics surrounding the New York area, once again, 93 subjects self report heroin.

The morphine levels found in hair by RA assay ranged from zero all the way up to 605 nanograms per ten milligrams of hair.  Cutting through the calculations, the clinical sensitivity reported with 90.3 percent.

Study C, a NIDA Research Monograph Publication in 1997 from the Center for Substance Abuse.  Here again, 94 subjects.  The clinical sensitivity calculated from the study is 90.4.

Study D, published in the International Journal of Addictions.  The study conducted at the Van Etten Hospital Drug Treatment Program in the New York area.  Only 20 subjects, a sizable number, morphine range two nanograms to 98 nanograms per ten milligrams of hair.  Again, the clinical sensitivity calculated from the study results, 80.0 percent.

Study E, by Baer et al U.S. Probation in Santa Ana Southern California, in conjunction with the Veterans Administration Hospital in Los Angeles, California.  Sixteen subjects in the study.  The morphine range found in here, two nanograms to 130 nanograms per ten.  The clinical sensitivity calculated 75 percent.

Now, taking all five studies as submitted in the application, and doing a combined approach to the calculation of clinical sensitivity from such five independent sources resulted in a clinical sensitivity calculation of 87 percent.

However, that clinical sensitivity really does not in fact address the issue of the effectiveness of using hair versus urine.  In Attachment 22 in our submission, we list the following data.

We looked at two matrices, urine and hair, and the number of tests performed within a large population.  Seventeen hundred and eighty-three urine tests performed, eleven positive were identified yielding a positive rate for morphine and for heroin use of 0.6 percent.

In the cases of hair samples taken from the same population, only 187 hair samples were taken.  But the number of positive were 15, a dramatic shift. Looking at the hair results, hair is more effective than urine in this environment.

Next slide.

Now, looking at analytical sensitivity and specificity for the assay admitted most of the data. The detailed data is in our Volume 3 response, but essentially, analytical specificity is 96.6 and the analytical sensitivity is 100 percent.

Now, the negative portion of the clinical study, we used a group of Psychemedics' employees, 81 employees.  And the clinical specificity is 100 percent.

Now, the questions the panel had regarding the clinical sensitivity and specificity was in the inclusion criteria.  And the cross question was negative urine plus negative self report.

Psychemedics did not simply use those two criteria.  We did not simply use a negative urine and a negative self report.  We used an employee group with a history of negative hair tests and a contemporary negative urine test.

So, the features of our negative controlled study where yes, they were volunteer employees.  They were required before entering employment to have a negative hair test for Psychemedics.  And the company policy is that we conduct randomly monitored hair tests.

The urine samples were also monitored for any immunoassay response as compared to the negative control.  We did not apply a cut off there.  If the drug was present, even below cut off, it was pursued by mass spectrometry for identification.

The next question asked, "Was positive urine not always confirmed plus a positive self report?"  And here again, we also went beyond those two criteria by looking at clinical evidence of drug use.  Hence, heroin treatment programs basically for the controlled studies.

There is a limited ability in urine to confirm the heroin metabolite, 6-monoacetyl morphine, or 6-MAM for short.  And this tends to make the urine test the weakest piece of information.

Hair, on the other hand, does provide for storage of 6-monoacetyl morphine, which can define heroin use.  Therefore, in urine, 6-MAM will clear in less than eight hours.

Now, if we look at the recent history with morphine at the old cut off of urine of 300 nanograms, it took 24 to 72 hours to clear.  When we moved to the higher cut off of 2,000 nanograms, 24 hours to clear. And 6-MAM, that unique descriptor of heroin use, less than eight hours to clear.

Now, the minimum dose issues.  The estimated mean minimal detectable heroin dose from the field studies, that is A through E as outlined to you, results in a value of 173 milligrams per month.  And that's in Volume 3.

This translates to a choice for use of a cut off of two nanograms of morphine.  Heroin addicts, in a paper by Kintz, used 900 to 24,000 milligrams per month.  So, you can see already that the average heroin addict is way above the mean minimal detectable heroin dose, and i.e., well above the cut off.

Now, the poppy seed controversy.  Psychemedics has presented data on the poppy seed issue.  And by feeding 150 grams of poppy seed, the morphine level in hair was only 0.17 nanograms per ten, well below the cut off of two nanograms.  And of course, no 6-monoacetyl morphine.

Pharmacokinetic considerations.  Again, I go back to the impediment to controlled clinical trials is that it is not ethical to conduct single, multiple, or chronic dose levels studies with heroin.  The minimal detectable dose information is all that should be really necessary due to the proven accumulation and stability of both morphine and 6-monoacetyl morphine in hair.

That is, the ability for hair to act as a trapping device, not only for the metabolite directly from heroin, but also for morphine.  The RA assay using hair can detect heroin use via the mass spectrometry confirmation of any presumptive positive above the two nanogram level by detecting the metabolyte, 6-MAM.

So clearly, the combination of morphine and 6-MAM is found in the hair of heroin users.  The combination of morphine and 6-MAM is not found in the hair of non-users.

The bias issues.  I'm merely going to summarize some of the bias issues.  We have submitted a hair color and curvature study, meaning curly hair versus straight hair, in Volume 3.

The large multiple statistical studies, that is those studies involving large populations and a statistical analysis, are not significant.  In other words, no bias.

And these studies are the Pinellas County Probation Study, the Pinellas Cleveland Juvenile Study, the New Orleans Diversion Data, the National Institute of Justice Study, the Pinella County Drug Use Forecast Emulation.

Eighteen hundred participated.  And finally, a large metropolitan police department.  Again, very large study of 1,800.

Methological differences within research reports can create variances in results that appear to be color related, but are not.  By that I mean that the study protocol may have used inefficient extraction procedures, failure to remove melanin and sweat, ineffectual wash procedures.  These are all part and parcel of methological differences.

The bias studies have previously not been required for submitted applications with immunoassay applications for urine to correct for either diet, gender, muscle mass, creatinine level, body weight.

And finally, individual biochemical variability.  That issue of the bioavailability within the individual within any population can create large variances.

Now, at this point I'd like to introduce Professor Newel to discuss in detail the statistical evaluation of these large populations as well as the small.

Professor Newel has the largest database on hair to look at statistical profiles with color versus hair.

Professor Newell.

MR. NEWELL:  Thank you and good morning.

I do not have any conflict of interest or financial ties to the Psychemedics Corporation, nor any other laboratory.

May I have the next slide.

I work at the University of South Florida.  I've been doing research involved in hair testing with funds primarily from the Department of Justice, National Institute of Justice for about ten years.  So, I appreciate the opportunity of being able to come before you and share with you some of the work that we've done.

First I'm going to start with a kind of conceptual transition slide.  The question that was posed by the panel regarded racial bias.  Originally, racial bias evolved from some initial concerns that there were associations between type of hair and its ability to record exposure to drugs.

People wondered if racial bias would be part of that problem.  And I think, probably first that was because there was an erroneous perception on the basis of the general public that there are biochemical racial differences.

In fact, I think over time this has evolved further, and that now racial bias issues are handled primarily as issues in the type of hair, that is color of hair, or curvature of hair, other descriptors of the hair itself.

Also, there have been some ethical limitations.  Research has generally relied on an epidemiological approach to this issue.  That typically has involved taking hair analysis results and comparing them to self report drug use and urinalysis results.

For studies that have proceeded that way, no such race effects have been confirmed for either of these drugs, cocaine or heroin.  And more recently, this issue has been addressed I would say more scientifically than calling it racial bias, by calling it differences due to hair color or hair type.

Can I have the next slide?

I want to share with you very briefly a secondary analysis, a statistical analysis, that was conducted on five data sets that have appeared in the literature.  All five of these were characterized with small-n.

Could we have the next slide?  Thank you.

Small-n is just basically the number of people that participated in the study.  The basic question is, is there a statistical relationship between hair color and assay concentration value?

In these five small-n studies, we had anywhere from nine to 20 subjects.  Analysis is done for studies recording analyte concentration along with color.  Now, that's important.

A lot of studies may have reported only the analyte concentration, and we can't analyze for color unless the authors either initially reported it or we were able to recover that data from the authors subsequent to a paper's publication.

There are three methods of analysis that we used.  These are statistical techniques that we used through the analysis.  I'll go through these in a minute.  But they include things like analysis of variance, when we had two comparison groups and two piece procedures when we had three or more.

And a linear regression, we had independent and dependent variables are both continuous in a racial level data.

Next slide, please.

This is a one slide summary of these five small-n studies.  You see they've started in the early 90's and up through the Kronstrand study at the bottom, published recently in 1999.  The n's refer to the number of subjects, 20, ten, 15, nine subjects.

And the analytes refer to the kinds of drugs that were being examined in the hair by these respective authors.  What we did is a secondary data analysis, taking those original articles and those original data points, and subjecting them to additional tests to see if there were any associations in those early works with hair color and with drug concentrations.

Those statistical tests included analysis of variance, Tukey's HSD, and some linear regression.  Significant question mark as the column total there refers to after the test, did we find any significant effect?

And you see for four of the five, we did not.  The fifth one, the Kronstrand study actually does show a strong significant effect.  The significance ranges are on there.  We use an a priori alpha of .05.  Most of those are nowhere close.  The Kronstrand study is a departure from that.

Can we have the next slide, please.

This is just the summary of basically what I just said, four of the five reviewed small-n studies do not show a statistically significant effect associated with color for either cocaine or heroin.

Now, why is this a likely result?  Why would we expect to see this?  Well, we have small-n studies, very small-n studies.  And initially, I think the authors, when they reported their studies, looked at mean values and just reported, "Well, the mean is bigger here than it is here."

The use of mean values is typically a very poor measure of central tendencies since these scores are very sensitive to extreme scores.  High values for dispersion mitigate all the systematic mean differences, especially in these small-n studies.

And of course, not all influencing variables can be controlled or even enumerated in a laboratory design study, and hence the need for epidemiological research.

The Kronstrand's study has significant results may indeed be due to unique features of codeine.  We don't know yet.  Tommy, just a moment ago, suggested some alternatives that included some important analytical differences of what the lab does that may account entirely for this particular effect.

May I have the next slide, please.

I'm going to spend just a few minutes reviewing what are characterized as large-n studies.  We have five data sets here, each of which has 70 or more cases.

Next slide.

These studies contain approximately 2,900 data points.  Currently, we have about 100,000 cases right now in our database, and I don't have all of that analyzed.  But we have substantially more than the small-n studies.

The advantage of this is that generally larger-n studies provide a more stable database, and they reduce the likelihood of a beta error.  Large-n studies are also more likely to find significant effects.  They have higher statistical power, and they allow a measurement of their strength.

Here's basically a list of the five studies.  One from Glasgow on MDMA and analogs.  One from Pinellas County, Florida on cocaine.  One from Las Vegas, Nevada, cocaine and cocaethylene amphetamine.  Another from Los Angeles, cocaine and so on.  University of South Florida, morphine study that we recently concluded and I'll report on last.

All of these statistical analysis utilized analysis of variance into these procedures.

The first study, the Glasgow study, collected hair samples from volunteers who attended rave parties.  Hair samples were analyzed for the presence of amphetamine, methamphetamine, and so on.  There were 232 subjects who donated the hair samples.  And the specimens were analyzed at two laboratories, 139 at the University of Glasgow and 93 at Tricho-Tech in London.

However, and unfortunately, hair color was only recorded for the 139 specimens that were tested at the University of Glasgow.  Nevertheless, there were no statistically significant relationships between hair color and drug concentration.

The Pinellas County, Florida study.  We had specimens collected from adult probationers who were asked to voluntarily participate in a six month project.  For each volunteer, at least monthly urine and hair specimens were collected.  Psychemedics Corporation did do the hair analysis with RA, GC/MS, and tandem MS confirmation.

We had a total of 589 hair samples, but only about 95 were testing positive for cocaine.  Hair -- brown hair was the most prevalent color.  And when we looked at these 95 cases again, we found no significant associations relating color to cocaine concentration.

The data from 500 randomly selected cocaine positive, and 500 amphetamine positive subjects were collected and analyzed by the same techniques.  Hair assays were performed for the parent cocaine, and cocaethylene, and amphetamine by hair color.

Hair color was arrayed over seven categories, red, grey, blonde, light brown, medium brown, dark brown, black.  When you have this many people involved, basically you have enough people in each one of those categories for a sufficient statistical power.

The hair specimens at extracts were screened by radioimmunoassay, cocaine metabolites, hair specimen extracts were confirmed by GC/MS.  And again, there were no significant relationships uncovered for these analyzed drugs.

Out of these 998 people from the Psychemedics Laboratory -- again analysis was done by RA, confirmed with GC/MS, or GC/MS/MS.  Samples were taken as a consequence of employment, those people who were applicants to employment.

Because this is an applicant test, only 72 tested drug positive.  That's typical for pre-employment screening procedures.  Among these 72 people, we had no statistically significant associations between color and concentrations.

And finally, the last study -- one we just concluded -- random study of 95 morphine positive clients.  Hair color was arrayed in five dimensions, and curvature in two, curly and straight.

Statistical associations were tested by my colleague, Tom Mieczkowski.  He has an article in the present Journal of Analytical Toxicology, and myself presented here using a variety of statistical techniques.

Analysis that I ran included non-parametric Kruskal-Wallis one way analysis of variance.  It's actually a more powerful technique when we can't make assumptions about the standard normal distribution in the parent data set for a population.

And even with this, we've again found no statistically significant results either between hair and -- hair color, rather, and morphine concentrations, or curvature and morphine concentrations.

This is a one slide summary of each of these five studies.  You see in the first column, the original place where the data was gathered.  The n refers to the number of people.  This would be drug positive that were analyzed.  Analytes were the drugs that we were looking for.

The test is a statistical test run on these data to look for associations between color and  curvature, and drug concentrations and significance.  Again, question mark.  They all were non-significant.

The significance ranges, as you see up there, are way different from our priority cut off of .05.  Nowhere close.

Next slide.

So, as a summary, all of the large-n studies in most of the small n-studies reviewed, do not demonstrate a statistically significant effect between hair color or curvature and drug concentration.

Small-n studies were often initially reported showing differences between mean concentrations, but mean value comparisons can be deceptive.  Significance cannot be determined by visual inspection.  That means they're sensitive to extreme scores, and that mean differences cannot be evaluated without consideration of deviation values.

And finally, codeine concentration may be uniquely related to melanin, but effects are probably small.  In fact, in the Kronstrand study, I believe, the most extreme effect showed that the difference was less than one half of one nanogram.

Next slide.

I guess before I get into this slide, I just want to make one additional comment.  There was another study that I did not summarize on here by Ben Hoffman, who also had 1,800 people, and again found no significant difference, no significant association between hair color and drug concentration.

DR. HENDERSON:  Where was that study from?

MR. NEWEL:  I'll have to get the records.  I have the citation with me.

This is something completely different.  This is a little study that we have done outside of this bias issue that I wanted to share with you.

It has to do with external contamination, and I think Tommy made some initial comments about studies that they've done with external contamination.  We did this one with 110 undercover police officers who logged their every day exposure to cocaine and other drugs over a 60 day period.

The hair samples were analyzed for both cocaine and benzolecgoning by Psychemedics.  In our study, no officers were found to have drug contaminated hair despite their frequent contact with cocaine, drug paraphernalia, drug users, and their environments.

That's all I have for you today.  Thank you for your attention.

DR. CAIRNS:  Thank you, Richard.

If we can return to -- we just had done the biased issues, the two slides on the biased issue.

Now, getting back to another theme and that theme is the removal, the effective removal of external contamination.  Now, let me begin by just outlining the analytical procedure whereby we do washing of the hair.

The first wash is an isopropanol wash, and that wash is conducted mainly to remove any cosmetic preparations that may reside on the external surface of the hairs, such as gels, mousses, hair sprays.

The isopropanol is removed.  The hair is then placed in a phosphate buffer with agitation, and there are three such phosphate buffer washes at 30 minutes each.  After those, the hair is put into a phosphate buffer for one hour again with agitation and again for another hour.  At that point, the hair is dried and then sent for liquefaction or digestion.

However, I want to put on an additional safeguard here on the washing procedure, and that is that we also add a mathematical extrapolation as if we were mimicing an additional five hours of washing.

Now, let me explain that.  I'm going to show you a graph.  I understand where you're coming from.  I'm going to show you a typical wash profile for an opiate user.  Now, you'll see the yellow bar diagram.  The IPA wash is not analyzed in this particular case.  But the first buffer wash, which was the half hour wash, contained 4.65 morphine.  Removal from that buffer and placement in the next buffer for 30 minutes resulted in a wash value of 1.3 nanograms.  Then into the third wash buffer for another half hour, 0.54.  Then buffer 4 is a one-hour wash with agitation, and buffer 5 is another one-hour wash.

Now, I think you can see that the washing sequence is such that each wash contains a diminishing amount of morphine.  And we're rapidly approaching a plateau in buffer number 3, 4, and 5.  To add an element, a strong element of conservatism, we take the fifth wash value, which is 0.52, and multiply that by five.  What we're saying here is that if we were to wash five more hours, the drug concentration would certainly not be greater than 0.5.  It would probably be less than 0.5.  So, we multiply the 0.52 by five.  The hair is then dried, moves into digest and confirmation, and has a value of 31 nanograms of morphine.

Now, if you look under the heading profile, what I'm going to do is the digest is 31.  I'm going to subtract from that five times the value of buffer 5, which was 0.52.  And, so the level after the extensive wash and the mathematical extrapolation is 28.4.

So, what we've done is to add a highly conservative buffer zone as if we were washing five more times.  Now, that's important, because when I show you the profile for a soaked hair, a deliberately contaminated hair, you will see the fundamental differences in profile.

Now, here's the soaking experiment.  We took 33 different hair samples.  We soaked them for one hour at room temperature in an aqueous solution containing 1,000 nanograms of morphine per milliliter.  We applied the wash procedure, that is the 3.75 hours of washing, and applied the mathematical extrapolation of five more hours, and all samples were negative.  But let me show you a typical profile.

This is the decontamination profile for a deliberately contaminated soaked hair one hour in aqueous morphine.  Now, in this case, we analyzed the isopropanol wash for this experiment, and you'll see that the isopropanol wash, the wash was only 1.2.  However, the first buffer, that is the half-hour washing, 480 nanograms; the second buffer half-hour washing, 98.1 nanograms morphine; the third buffer, 30.9; the fourth buffer, 12.9; the fifth buffer, 7.2.  This sample, externally, contains a lot of morphine.

Then we digest -- dry and digest the sample, we find we have 20 nanograms in the digest.  So, if we perform our routine wash criteria, you have the digested 20 minus five times the last wash, which is 7.2, would actually result in a value of minus 16.3, showing you how conservative the analytical extrapolation procedure is.  We've overcompensated.

But if you look at the profile here, the washes contain basically a lot more than the digest.  This is a typical contamination profile.  The previous profile was a typical user profile.  So, we are able, clearly, to use the wash procedure to distinguish or differentiate a user from a contaminated hair.

Now, the sweat study.  We were asked to respond with this type of study, and it's in volume 3.  Here we took three groups of 13 different hair specimens and we soaked at room temperature for 15 minutes in a chloroform solution containing this time 10,000 nanograms of morphine and 10,000 nanograms of the metabolite 6-MAM.

The specimens were dried, they were then put in contact with synthetic sweat for time periods one, three, and six.  So, 13 samples in each of these time zones.  The Psychemedics wash procedure applied as I've outlined, 3.75 hours of washing and an analytical extrapolation of five additional hours.  The results were that all 39 samples were negative.

Now, we were also asked to comment on the cosmetic treatment; that is, can you affect the hair testing result.  Data would show that certainly perming, relaxing, and to a lesser extent dyeing of the hair, could reduce the amount.  However, the samples at risk if such procedures were performed would only be those samples that are existing between two nanograms and 2.6, because such treatment would take them below the cutoff.

So, a calculation of the morphine positives that we have in our database, it would only affect about three percent of all positive heroin tests.  And you compare that with a urine test where adulteration, substitution, and hydration, all samples could be at risk.  In here the only samples at risk are those residing approximately 30 percent above the cutoff, meaning from two nanograms to 2.6.

Now, conclusions.  To summarize, in the application, we have given you five user field studies that we consider have provided more reliable data from multiple sources than a single controlled clinical study may have provided if it were ethical possible.

Secondly, the establishment of a dose-response curve strictly not necessary for an assay where the decision is are you above the cutoff or are you below the cutoff.

On the pharmacokinetic issues, yes, the key issue is to detect morphine via the screening method application to identify heroin users.  As with urine testing dose-response was not required.  Regarding the minimal detectable dose, all the information that is necessary is contained in the hair, because it has the ability not only to store, to trap, but to stabilize over long periods of time.

And finally, urine testing for drugs of abuse did not provide such detailed information as part of any FDA submission.  The ratio and hair color bias you've heard from Professor Newel.  With valid techniques, meaning the analytical assays in conjunction with statistical populations, there is no race or color bias.  Statistical treatment and review of small and large population revealed no evidence, as you've seen from the previous presenter.  There are no conflicting large population studies in the literature.  The small studies may have conflicting conclusions that result from different methodologies, and these are influenced, as we've said, extraction, efficiency of wash, non-ingestion.  And these results are not unexpected.

And, finally, with valid techniques such as the washing procedure we've outlined, there is no external contamination issue.  In fact, the research presented in our application has demonstrated that the exhaustive level of washing effectively removes externally the positive drug and can ensure with confidence the differentiation between a heroin user and a non-user.  There are no conflicts in the literature regarding external contamination provided the correct techniques are applied.

Finally, research models without adequate washing procedures do not remove contamination.

I think that's my last one.

Oh, hair treatment.  Again, cosmetic applications, we've seen it can reduce the amount near the cutoff, and that would only affect three percent of the samples.  And this is a far less effect than that created by simple hydration of adulteration.

And for the last comments, I'm going to invite Professor Selavka to the podium.  We've saved a little treat for last.  Professor Selavka is in fact co-chair of the Small Working Group for Hair Analysis under SAMHSA, and he's going to share with you his views on hair testing and the poppy seed controversy.

Professor Selavka.

DR. KROLL:  Hi.  Before we begin, I would remind you it's about 10:45, which is getting into our scheduled break, so use your time wisely.

DR. SELAVKA:  Say again how long I have.  So, I'm between you and food.

(Laughter.)

You probably thought I was here just to run the projector.  My name is Karl Selavka.  I'm the Director of the Massachusetts State Police Crime Laboratory, which has nothing to do with hair drug testing.  So, since you wanted to see somebody outstanding in their field, there you go.

I do have to disclaim I'm not here representing the state or the executive office or any laboratories per se.  I am hopefully representing some insight into the operating process of the Hair Testing Working Group that was sponsored by the Drug Testing Advisory Board of HHS, as well as poppy seed studiers.  Some day after my wife's vet bills are paid off I hope to have a conflict of interest.  Of course this is the land of disclaimers.

Anyway, I was one of the co-chairs along with Dr. Donald Kippinberger of the Hair Testing Working Group to DTAB.  We're asked by them to review the field, and there really are a world of different samples other than those currently being applied in workplace testing that could allow for some insight into a drug use history for a person being tested.  Of those, urine drug testing, sweat, oral fluid, and hair predominate the matrices that are being studied currently for applications in the workplace.

They all are providing complimentary information.  The complimentariness may be related to time periods and sensitivity to different substances and graphically speaking some look into the past, and one can look into the future.  If you were to use all of the complimentary matrices for each of the applications in a workplace, a business setting, forensic examinations, and even post-mortem work, you'll find that all of them have unique specialties that they can offer to the particular investigation.

But the federal government is now looking outside the box.  They are not feeling that they have to use a single-matrix approach anymore, and because of that, they also recognize that the old approach is the larger approach, and it's going to take some time to demonstrate adequately to everybody's satisfaction what needs to be demonstrated.

The Hair Testing Working Group was made up of a number of different kinds of people.  Those actively working in hair drug testing laboratories as well as interested outside scientists and medical review officers.  A number of federal and non-federal observers and ghosts and some people that were providing specific insight into issues since it was federally funded, we were fueled mostly by doughnuts.  The federal government graciously paid for travel, but otherwise we were basically on our own.  But they did well for us.

We had three meetings, most of which were attended by volunteers.  We put out to the hundreds of people that have expressed interest in the past about hair drug testing, a notice that we were going to have these meetings, an invitations for them to come.  We also had a number of contrarians, people who had published in contravention to other people in the field of hair drug testing to try to get their insight into the process and to try to get their consensus with us.

But after the third of these meetings, I was invited to also give this to the international perspective.  It's important to recognize that America is not the only place where hair drug testing occurs.  In fact, in Germany there's routine use, as well as France.

The substances that are listed are those that are routinely reported in the literature, early in the literature that comes from those countries.  Canada has presented results in the literature, Japan, the United States' copious data, and the laboratory where I used to work, National Medical Services, is actively involved in esoteric drug and other analyte testing in hair.

This summarizes the drug classes that have been detected in hair.  It's virtually all of them that can be ingested.  But we formed consensus on all of the components for the most part that we were asked to by the federal government -- how much hair to collect, what type of hair to collect, the length to collect, stability of drugs, collection integrity, screening cutoffs, which screening precisions and accuracies were necessary.

One of the elements asked us to define the immunoassay cutoffs and what the analytes should be.  Those are listed here for you.  We also formed consensus on the confirmation cutoffs that should be chosen to limit any claim of external contamination.  These were chosen much higher than the analytical sensitivities of the assays, but building in factors recognizing that this particular matrix had not been used routinely in federal workplace testing, and therefore we should be overly cautious, using specific metabolites for amphetamines, opiates, cocaine, and THC to rule out the presence of external contamination as the only source of drug being determined.

We helped design what we thought would be adequate blind PT programs and sources for open quality control materials.  The confirmation cutoffs for opiates are shown here.  The difference between what you've seen the rest of the day and these is that you have to multiply this times ten, because this is nanograms per milligram as opposed to nanograms per ten milligrams.  I'm not a mathematician so I always have to do this for myself.

These are the other confirmation cutoffs in case you care.  I know today we're here to talk about opiates.

We were able to reach consensus on interpretative guidelines, and that is a positive finding in hair represents chronic and repetitive exposures and uses during the time period represented by that hair.  It's not a single dose giving rise to these findings; it's repetitive uses.

We talked about what alternative medical explanations you might have for a heroin-finding hair.  Well, there shouldn't be any.  There are no alternative medical explanations for that.  We discussed dose-response relationships, helped design what we thought would be a comprehensive PT program, and for the first time for workplace drug testing defined what we thought were adequate standards for tandem mass spectrometry for confirmation.

About dose-response relationships, I think it's important to recognize heroin is an illegal drug, there's no therapeutic drug monitoring requirement here.  It's sort of counterintuitive to deterrence to expect there to be a dose-response relationship.  In effect, we haven't held any of the other matrices to that equivalent standard, and therefore detection is appropriate, but cautious application of cutoffs is also appropriate.

More consensus.  You've heard this well from Dr. DuPont so I won't repeat it, but bias is a normal component of all biometric testing.  We recognize that.  We also believe that it's important to have discretion for MROs in order to take into account different factors of a person's history when relating drug positive results.

Animal models do not generally correlate well with human use, and low-dose studies should not be used dispositive of the issue.  We know very well that there's wide differences between the hair on animals and the hair on humans.  Therefore, animal-use studies have certain application, but they're not dispositive in any way on the overall studies in the field.

The other reasons I'm here is because I've done one of the studies on poppy seeds that's been referred to by some, and also I've recently updated this and presented at the American Academy of Forensic Scientists Meeting earlier this year.  As part of the presentation, basically, I'll go over what foods are out there, the toxicology involved, the fluctuating policy we've seen, and that there is help.

I was the Operations Officer for the Army's drug testing lab in Hawaii for four years.  I asked for Ft. Meade, Maryland, but they gave me Hawaii.

(Laughter.)

In this laboratory, we were at that moment in time, in the mid-80s, we were bringing opiate testing on board, and we had to recognize that there were some issues.  One of the issues we got inspected too often.  I did learn the official military vegetables for urine testing -- leeks and peas.  If you don't get it, I'll explain it later.  And you should never eat anything with the word "whiz" in it.

But anyway, opium poppies can give rise when ingested to morphine findings in urine.  There was some literature at that time, but we recognized that we might have a problem.  We didn't want to ignore the problem; we wanted to use reasonable cutoffs in order to define the difference between poppy seed ingesters and the contributions those might have to positive findings and those that were using heroin.

When it comes to poppy seeds there's quite wide variability in worldwide production and the amounts of opiates that are found in those seeds.  This lists those for you, and you see the word "thebaine" there.  It's one of the opiates that's found in poppy seeds that at one point was thought to be maybe that's the magic bullet.  We can use thebaine detection to differentiate poppy seed eaters from heroin users.

There's also wide variability in how much you eat when you're eating a poppy seed food product.  You don't know whether the seeds were washed or not washed, the latex on the outside of the seeds may or may not, depending on the literature you read, contain the opiates that are in question.  You don't know what the difference might be from cooked and uncooked foods.  The problem is you don't really know where these things come from when you're eating them.

But poppy seeds follow toxicological rules -- whatever goes in one end comes out the other end.  Those ends could be of course on the top of your head or various other places.  And the concentrations in urine are generally described here:  Less than 4,000 nanograms per mil of morphine, less than 200 nanograms per mil for codeine, thebaine, generally below 150 nanograms per mil.  All in urine and all peaking around four hours after ingestion.

It can be extended if you have a lot of butter or rum in some of the products that you're eating.  Hopefully you do when you're eating poppy seeds, because otherwise they're not that delicious.  There is a good deal of high subject intervariability.

In order to rule out poppy seeds, in the late '80s, Mahmood ElSohly and Skip Jones in Mississippi, described in a paper that there may be a strategy using toxicologically placed analytical cutoffs on the confirmation.  If morphine was greater than 5,000, codeine greater than 300, and morphine to codeine ratio less than two, they described that that should not be from poppy seed eating.   This was around the time the military laboratory system was putting in place opiate testing.  They also said 6-acetylmorphine, if you find it, it's heroin.

Well, we went in our laboratory.  We wanted to study the effects.  We thought, well, if there's an analytical approach, we may also be able to join them in the toxicological ingenuity and come up with this good plan for the military.  What we found is that the reasonable toxicological ingenuity does not spell good policy.

The reason for that is in my own study with poppy seed ingestion, we found levels of morphine 11,500 nanograms per mil at peak concentration in one individual; codeine greater than 4,000 nanograms per mil.  The ratios usually work but the cutoffs that were being applied at 5,000 and 200 were not working.

There is very rapid 6-acetylmorphine clearance.  Generally, within three to four hours, you're not going to have 6-acetylmorphine in urine at detectable levels.  Thebaine is not ubiquitous among those that eat poppy seeds, and there's always your idiosyncratic ingester.  So, you may have all the components of a positive, but MROs are routinely not reporting it as a positive.  That's the problem.

There have been bouncing cutoffs.  The military, we started with 300 nanograms per mil for codeine and morphine.  We went to 4,000 morphine, 2,000 codeine.  It's now back down to 2,000 morphine and codeine, the Medical Review Office making the majority of calls in this.

HHS' cutoff changes are shown.  DOT's cutoff changes are shown.  In their Notice of Proposed Rulemaking under new part 40, going in for DOT testing, their cutoffs suggest that at 15,000 nanograms per mil for morphine, there is no possibility of seeds.  It's another hand wasted toxicological ingenuity.  I'm sure we can eat enough -- if there's enough rum and butter, I'm sure we can get enough in our body with the right seeds to clear that 15,000 hurdle.  That's not the point.  Let's detect 6-acetylmorphine.

Now, taking the pieces apart and designing a good program, part one is even the federal government when it was reviewing this issue and looking at changing cutoffs recognized -- this is Dr. Autry's testimony in front of subcommittee of the House in 1998.  Essentially, among 1.1 million tests for drug use, opiate positives at 0.66 percent.  That's 7,294 samples.  Eighty-one percent of them were positive for morphine at a level less than 2,000.

So, only 19 percent of the 0.66 percent had positives greater than 2,000 nanogram per mil cutoff.  And of those, only two percent of the 0.66 percent were positive for 6-acetylmorphine.  So, essentially, if you use 6-acetylmorphine or a cutoff of 2,000, you're going to miss the predominance of positives in this population of 1.1 million people who might be using heroin.

But the policy problem is really twofold. For one point, if labs are doing lots of work and spending lots of money and time to find morphine-positive results in urine and MROs are routinely not reporting them to the company, well, we've wasted a lot of time and money, and that's not right.  So, if you raise the cutoff, you lower the laboratory burden and the cost to the system.

Unfortunately, higher cutoffs also reduce your likelihood of deterring heroin use, and that ultimately is the problem we're after.  We're trying to differentiate poppy seeds from heroin abuse.  Heroin on the street has hydrochloride, or tar heroin, in a smokable form.  We want to distinguish, in effect, in this model, here's your heroin abuser.  You want to a program that has reasonable detection of heroin abuse that will build a good deterrence.  Without good detection you don't have deterrence.

The other policy consensus that should be mentioned is cutoffs should reflect drug use while minimizing potential endogenous, exogenous, and chemical and electronic interferences.  We all recognize that when you find a piece of a tiger, you still have to differentiate the tiger.

(Laughter.)

Another piece of consensus to leave you with is positive opiate results are conceded by all toxicologists poppy seeds can lead to positive urinalysis findings at 300, at 2,000, maybe not at 15,000.  You'll find some consensus at levels below that.

So, heroin abuse is definitely increasing.  My crime laboratory continues to receive advancing numbers of submissions of heroin on a yearly basis.  Poppy seeds creates an evidentiary false positive in urinalysis; therefore, urine tests are not effective in approaching the heroin problem.  Safety is compromised, because if you can't detect a heroin problem, then the safety of the workplace or other tested population is not good and with basically elevated cutoffs, lowering deterrence at the very time that heroin abuse is increasing.

When you hear hoofbeats, you think horses.  Well, the good thing is only 6-acetylmorphine proves heroin ingestion.  One piece of bad news is -- or good news, depending on how you take it -- urine does not routinely accumulate 6-acetylmorphine in levels that are detectable, but hair does.  Six-acetylmorphine readily accumulates at detectable levels of hair.  You've seen a lot of demonstration of that in the data provided today.

The last part of this is what the Hair Testing Working Group provided to DTAB was we suggested good reporting guidelines and maintaining MRO discretion in the review of findings from hair.  We believe hair provides the complimentary matrix that is really needed in the program of deterring heroin abuse.  We should choose unique analytes, like 6-acetylmorphine, and use the lowest cutoffs we can while maintaining reasonable safety.

This is the information that the workplace, the employer, and all of us deserve.  This is the information that urine provides.  We need to do a better job, and hair can do that for us.  Thinking outside the box it was nice to know that the Drug Testing Advisory Board has taken our advice.  The guidelines that will end up in a Notice of Proposed Rulemaking are under development now.  The second draft is already up; they're working actively on the third.  We're proud to have been a part of that process.  It's a very important time for hair drug testing and for heroin abuse.

I appreciate the opportunity to be here today.  I'll slide out of the way now and let you get to a break.  Thanks very much.

DR. CAIRNS:  Mr. Chairman, thank you for the extension of time.  That now concludes our formal presentation.

DR. KROLL:  All right.  Thank you very much.

We're going to take a short break now and report back to give the FDA presentation in 15 minutes.

(Whereupon, the foregoing matter went off

the record at 11:12 a.m. and went back on

the record at 11:26 a.m.)

DR. KROLL:  If the panel will take their seats, please.

Okay, at this time, we are going to continue with the FDA presentations.  And the next speaker is Dr. Albert Peacock.

DR. PEACOCK:  Let me see if I can get some of your panel back.

DR. KROLL:  All right.

DR. PEACOCK:  I didn't want to make a presentation without my Division Director here.

Good morning.  I'm Al Peacock, a scientific reviewer in the Chemistry and Toxicology Branch of the Division of Clinical Laboratory Devices and a member of the team reviewing this device today.

The Psychemedics RIA assay for opiates and hair that is presented here today is a first-of-a-kind submission for the FDA for drug abuse testing in hair.  Currently, there are a number of review issues that we are working with Psychemedics to resolve.

Today, I will be presenting to you several new matrix-related issues that we would like your input on.

I'll summarize for you the study submitted by Psychemedics that correspond to these new matrix review issues, and I'll ask you if these studies are appropriate.

This device is a radioimmunoassay intended for use in the detection of heroin in human hair.  This device produces qualitative as well as semi-quantitative results for morphine.  For a positive report, two nanograms of morphine per ten milligrams of hair are required.  All positive results must be confirmed by a more specific method, such as GC, mass spec, or tandem mass spec or LC/MS/MS.

Now, heroin is metabolized in the body, obviously, to morphine at 6-acetylmorphine, or MAM, 6-MAM.  Consequently, in order to report a positive heroin result, this device has to identify morphine at a concentration of two nanograms per ten milligrams of hair as well as identifying 6-MAM in the mass spectrometry confirmation.

This is an assay flow chart for the analytical process.  There are two pathways that Psychemedics has presented to us for processing presumptive positives after the initial digest and RIA screen.

The blue pathway here has an RIA assay who emits -- get a negative report and the positives go through the mass spectrometry confirmation.  The red pathway emits the second RIA procedure and goes directly to the mass spectrometry confirmation.  One of our concerns with this approach is that we don't think it's been fully demonstrated that these two assay pathways will produce identical results when assaying the same sample.

A total of 11 field studies were submitted by Psychemedics.  Most of these field studies were government funded and were designed by independent scientists not associated with Psychemedics.  Psychemedics was contracted by these independent scientists to perform hair analysis for their studies.

The following study design features are common to all of the clinical field studies submitted by Psychemedics.  The hair cutoff used was two nanograms of morphine per ten milligrams of hair.  The urine cutoff for the opiates assay was 300 nanograms per mil.  Of course, you'll note that currently the opiate cutoff recommended by Sampson is 2,000.  And clinical truth was defined by Psychemedics as self-reporting in the amount of drug ingested, self-reporting of the time of ingestion, and verification of positive drug screen results, urine drug screen results -- excuse me, I'm sorry, verification by a positive drug screen, urine drug screen.  And these were not confirmed.

There were two field studies which were used to determine the minimum detectable dose.  Biofield studies were used to determine clinical sensitivity or positive percent agreement, one field study was used to determine the clinical specificity or negative percent agreement, and three field studies were used to determine the clinical usefulness of hair testing.

The clinical data from research reports and data was collected -- other data was collected from other diverse sources.  The self-report and urinalysis data from these studies were seen by Psychemedics only upon receiving prepublication copies of the papers for the final peer review copies of these papers.

We have the following concerns related to the field studies.  These field studies were uncontrolled, unmonitored, and not validated by Psychemedics.  They had no control of the hair sample collection process.  Psychemedics did not approve any of the study protocols used.  There was no demographic data submitted for these studies, so we couldn't evaluate or make any kinds of conclusions of the potential of race, age, sex, hair color or other individual differences and what their effects could have been on the results.

The field studies lacked the gold standard method of determining true drug use status.  Therefore, we feel that in place of sensitivity and specificity, clinical performance should be estimated in terms of percent agreement with the available or chosen method of determining drug use status.

We also have a concern that the positive percent agreement has been determined primarily in clinical studies of heroin users, while the primary use of this assay will be in a workplace setting.

Additionally, self-reporting has several deficiencies that can influence clinical study results.  These deficiencies include inaccuracy in remembering amount of drug used and remembering the time of ingestion, the truthfulness of reporting, unknown drug purity, varying efficiencies of drug administrations, such as by injection or by smoking or by inhalation.  Validity of self-reporting may differ by the drug being used, and the validity of self-reports of recent drug use may be less at follow-up than at intake.

The environmental studies submitted by Psychemedics utilized ten hair specimens.  The studies were based on their three-domain properties of hair, which are the accessible region, semi-accessible region, and inaccessible domain.  Excessive washing, perming, dyeing, and relaxing hair treatments were evaluated.

The contamination studies consisted of three separate studies:  decontamination of drug pre-hair specimens that have been surface contaminated with morphine, decontamination of hair specimens from heroin users with morphine concentrations above the cutoff of two nanograms of morphine per ten milligrams of hair, and potential of contamination of hair -- drug-free hair by simulation of sweating of hair that has been treated -- surface treated with drug.

One of the problems we have in evaluating this information is there's conflicting evidence in the literature regarding some of these issues, especially the potential for bias of testing results due to hair color.  Several human and animal studies have shown differences in drug accumulation between pigmented and non-pigmented hair.  Other studies have found that hair pigmentation contributed negligibly to variations in drug concentration values.

However, most studies found in the literature melanin concentrations in the hair are not generally determined and correlated to the amount of drug incorporated in the hair, making the evaluation of some of these studies very difficult.

Psychemedics states that their procedure spins out the melanin fraction from the dissolved hair digest.  Consequently, any melanin-bound drugs are excluded from the analytical procedure, and therefore no hair color bias can be present if the Psychemedic methodology is used.

Now, I'll proceed to the questions.  Lunch is approaching.

Okay.  Due to the conflicting literature concerning these new matrix-related issues, we need input from the panel to help us resolve the following concerns.  Question 1 for the panel:  The clinical data in this application is from research reports and data collected from diverse sources and not from perspective controlled clinical trials that evaluate heroin use.  Therefore, a steady hypothesis inclusion/exclusion criteria associated endpoints and a plan of statistical analysis were not provided.  First part of this question:  Can assay performance be established with these types of data?  Why or why not?

DR. PEACOCK:  And I'll go ahead and finish the question:  Do the data presented provide adequate characterization of assay performance?

PARTICIPANT:  Read them all.

DR. PEACOCK:  Okay, sorry.  Question 2:  With respect to making claims for clinical sensitivity and specificity, is a single negative urine analysis plus a negative self-reported drug use a sufficient, unbiased standard for establishing true drug-free status?

Is a positive urinalysis that is not confirmed plus a positive self-report of drug use a sufficient, unbiased standard for establishing true drug-free status?

Question 3:  Should the minimum dose required to produce a positive result be determined?

Question 4:  Should the relationship of the pharmacokinetics of drug use and the incorporation of drug into the hair, that is single dose, multiple dose, and chronic use be determined?

Question 5:  Should the potential for bias by race, age, sex, hair color or other individual differences in the incorporation and retention of drug in the hair be evaluated?  If yes, what additional studies should be requested?

Question 6:  Is the information provided by the sponsor adequate to address the issue of retention of drug in the hair from environmental exposure?  If not, what additional information should be requested?

Seven, Question 7:  Has the sponsor adequately demonstrated the effects of various washing or hair treatment procedures on the internally incorporated bound drug?  If not, what additional studies should be requested?

DR. KROLL:  All right.  Thank you.

At this time, I'd like the panel members, if they have any questions, to ask Dr. Peacock.

DR. KURT:  Tom Kurt.  I'm asking about the precedence within the FDA of a unique extraction method, which apparently is unique here -- not the RIA technique itself, but an extraction method -- and comparing it to a few years ago when Dr. Ostrea, that's O-S-T-R-E-A, came to FDA with a unique extraction method for extracting meconium, or newborn infant stools, for drugs of abuse in a PMA, and was that approved, and under what precedence does that set for this?

DR. PEACOCK:  I might have to refer to Dr. Gutman, that was before my time.

DR. GUTMAN:  There has been -- actually, I think it was clear -- but there has been a meconium assay that was cleared, and the precedent is whether there's reasonable science to support the submission.  It's not anything unique or special.  It's a question of having the right science to make you comfortable that this works in the way it's intended.

DR. EVERETT:  James Everett.  Is the data clear as it relates to race and incorporation of drugs into hair?  The sponsors suggest that there is no difference.  Is that true for the literature?

DR. PEACOCK:  No, it's not.  Hold on one sec.  There are several studies that -- Cone, et al. in 1996 found that cocaine concentrations in hair of African-American subjects were significantly higher than those found in Caucasian subjects.

Glyxner, in 1996, observed that drug residues for clembuturol were lower in light hair, blond and gray, than dark hair, black or brown.

Henderson, in 1998, demonstrated that non-Caucasian subjects incorporate approximately 2.9 times more duterated cocaine in their hair than did Caucasian subjects in equivalent experimental conditions.

Kronstrand, in 1999, observed after a single oral administration of 100 milligrams of codeine to non-subjects, the incorporation of codeine into hair is affected by its melanin content, and that relationship is exponential.

As Psychemedics has mentioned, it's very difficult sometimes to compare some of this data, because people use different methods, different treatment methods, pre-treatment methods before analysis, and makes comparing some of these results very difficult.  But that's one of the reasons we brought this question to the panel today to get your input, because there's such conflicting information out there from Psychemedics as well as some of the articles in the literature.  We were sort of at an impasse on what to believe, and that's why we wanted your input.

DR. KROLL:  Yes, Dr. Lasky.

DR. LASKY:  So, based on what you just said, the uncertainty deals with the drug that was measured and the extraction technique.  Is that a correct interpretation?

DR. PEACOCK:  Well, it's the point and counterpoint.  Psychemedics' position is that their enzymatic digestion process is much better and especially since they spin the melanin out that they don't have an issue with this.  Whereas some of these solvent extraction processes might lead to a problem.

DR. LASKY:  All right.  So, then that's really the issue, the fact that in the literature it says that since that's variable if there is a better technique, then demonstrate that that technique is better.  And I guess that's the question that in fact that we're addressing right now, not what's in the literature but the data that has been submitted, whether or not that's valid.

MS. PINKOS:  This is Arleen Pinkos.  I'm also a member on the review team.  I think part of our dilemma is that the studies provided to us did not have any demographic information associated with it.  So, it was not possible for us to, in an objective manner, evaluate whether there was a bias or not.  And the information that we do have is the anecdotal from other people's studies or literature.  So, we're just asking you is the information that is available adequate?

DR. PEACOCK:  One of the problems is that the studies that were submitted to us came to Psychemedics from their contacts of people doing studies that were independent of them setting up a study to answer some of these analytical questions.  And, so some of this demographic information is not available and will never be available.

DR. KURT:  Tom Kurt, another question.

MS. FLANNERY:  A procedural point:  Can Psychemedics address these questions as they come or are they going to have a chance afterward?

DR. KROLL:  Right now we're discussing these things with the FDA panel.

MS. FLANNERY:  All right.  So, they can address each of these issues afterwards.

DR. KROLL:  We can allow that.  There is a time for open public hearing scheduled at three o'clock.  We'll provide you some time to go ahead and try to answer some of the questions.  Why don't you write them down and we can do it in an organized fashion?

DR. CLEMENT:  I have a question.  You mentioned the lack of a perspective controlled study that's appropriately monitored by the Company, particularly there's no controlled studies of actually ingestion of heroin.  That's a real problem.  How would you, in a perfect world, ethically do a study or how would you envision it be designed to actually look at some of those issues?

DR. PEACOCK:  No, I agree with you that it would be a very difficult study to do, but people are given morphine and codeine under controlled conditions all the time in hospital settings.  And you could set up something.  People take morphine and codeine under controlled conditions for pain management and they take constant chronic doses, and you know what those doses are.  You don't have self-reporting.  I mean those can be administered quantitatively with clinical drugs as well as monitoring the time.

One of the problems we have with some of these studies is that because Psychemedics did not design any of these protocols, even though they were done monitoring people in criminal justice populations and parole, some of the questions they could have answered if they had -- if they were going to do it this way could have been asked but they weren't, because they didn't have control over the studies.  This is data coming to them after the fact, and they didn't -- yes?

DR. CLEMENT:  From looking at the results of the data, they cited four or five studies where the sensitivity of people that they categorized self-reported or urine positive testers sensitivities in the range of like 70 at the low end.  It was as high as 85 percent.  I would think if there's an error that occurred, it would err on the side of less sensitivity.  But from the data that's presented, they did their own in-house study that showed zero false positives in their own workers.

DR. PEACOCK:  Yes.  It's not that I'm arguing with what they did.  I'm just asking your input, because this is the first hair study that's coming in, and if you're giving the advice that this is acceptable to you, it will become less of an issue for us.  And we understand that controlled perspective clinical trials on drugs like heroin, obviously, would be very problematic.  But there are ways to address some of these issues, we feel, at least for morphine for sure.

DR. CLEMENT:  So, if I can rephrase what you're trying to get an opinion to us what is the minimum standard acceptable for an approval of this type of product?

DR. GUTMAN:  We're actually bringing this to panel so that you can in fact provide us some perspective and advice, and I would remind the panel, since Phil Phillips wasn't here to remind you this morning, that we are under an obligation to seek a least burdensome threshold.  So, we're not looking to sell the product short, but we want to do what's reasonable, and we want to make sure we aren't putting either incorrect or artificial hurdles up.  On the other hand, we want to have an honest product with honest performance and honest labeling.

DR. PEACOCK:  In other words, if you would agree or you feel comfortable with the fact that a negative self-report, a negative urine, and a negative hair test is good enough for a gold standard for a negative, and a positive urine and a positive self-report and a positive hair test is good enough gold standard for a positive, well, that question is answered.  But that's why we're asking it, to find out if you feel that's true.

And if you feel that the data submitted to us from studies that they had no control over, as far as I know, if data that came in secondarily is good enough to answer some questions that we feel that could have been answered with a little more of a controlled study that they designed versus somebody else designing.  Is that sort of where we're going?

DR. CLEMENT:  Yes, yes, helpful.

DR. KROLL:  Dr. Manno.

DR. MANNO:  Barbara Manno.  You said that, and I read in the materials sent to us, that they are claiming they spin the melanin out.  Have there been -- is there adequate data comparing one or the other extraction methods simultaneously against their procedure?

DR. PEACOCK:  If there is, I'm not aware of it.

DR. MANNO:  What I'm getting at is melanin-based or lack of melanin specifically, is the data there for that comparison?

DR. PEACOCK:  Not that I know of.  Psychemedics might be aware of it.  They might be able to answer that question better than myself.

DR. KROLL:  Dr. Henderson.

DR. HENDERSON:  I asked Dr. Newel and perhaps in the break someone from Psychemedics, along with Dr. Newel may be able to find some of the answers.  The data that he presented from Glasgow and London, to use that as indicating that there was no effect on hair color -- by hair color and hair type, I think that those are fairly homogenous populations, and it's unclear what groups were analyzed in those populations.  How were they different?  Were their hair textures, hair colors different?

I think there are enough populations that we are fairly confident regularly used drugs -- incarcerated populations, patients in substance abuse programs -- where if there was some interest in looking at whether the test was different in different populations, that could be done.

I'm concerned that the specificities and negativities are reported in populations that have minimal or -- well, they're identified as admitting to drug use, and since we're looking at using this in a general population, I've not seen and I'm unaware of what the predictive negative and positives are for general populations that I would assume may have one and two percent prevalence of substance abuses.

So, some of that information I would like if you could find some of that.

DR. KROLL:  I wanted to ask a question now to try to clarify for me and maybe for the rest of the panel, it appears that Psychemedics has a method that's pretty inclusive.  If someone sends them a sample, they do the sample preparation that uses RIA procedure, and then they have a bunch of confirmatory procedures.

Now, what I'd like to clarify for us exactly what is it that is presented to the FDA for approval, what portions of that?  Because there's like three major sections, each of which are fairly extensive.

DR. PEACOCK:  Right.  The submission itself is for the RIA assay, radioimmunoassay for opiates.  But it's sort of like a package.  I mean the up front processing of the hair is critical to the assay as well as the downstream processing, which is the washing as well as the mass spec.  And especially for heroin, I mean, you have to see 6-MAM in the mass spectrometry.  So, the application itself for the RIA assay, but it's such an integrated system that we can't really exclude any piece.

DR. KROLL:  All right.

DR. GUTMAN:  And we would recognize that, I mean, for most drugs of the assay we would recognize some reference methodology as an acceptable framework for reference.  So, this isn't -- the matrix is unusual, but the idea of using a reference method to try and characterize performance is not.

DR. KROLL:  That's fine.  I mean, I just wanted to clarify that, because it's a rather -- I mean there's the whole back-end, which is the confirmatory portion, isn't really part of the submission.  And it appears like the front end, the sample preparation, is that part of the submission or not?

DR. PEACOCK:  Yes.

DR. KROLL:  That is part of the submission, okay.

MS. PINKOS:  Yes, but the confirmatory -- this is Arleen Pinkos -- the confirmatory method for a urine test would also not be part of our review either.

DR. KROLL:  Right, okay.

Okay, and just my intent, what we're going to do eventually is go through each question, and everybody speak.  So, this is to ask Dr. Peacock.  Do you have any other questions, just to clarify issues in what he presented?

DR. KURT:  Were any attempts made to go to a detox unit in combination with a methadone clinic to take samples in that kind of a situation where you actually have patients in that flow path as opposed to giving them drugs as test subjects?

DR. PEACOCK:  Psychemedics would have to answer that question.

DR. KROLL:  Okay.  That's a question we may actually ask the Company, and what we could probably do is when we come back after lunch is if the Company can succinctly answer some of the questions as we get back into the discussion so we can clarify those issues.

Any other questions from the panel?

Dr. Manno.

DR. MANNO:  Manno again.  You asked the question would we be satisfied with a one negative urine, one negative hair, I believe it was, and a self-report.  Are you talking about hair and urine collection simultaneously in time or are you separating them in time?

DR. PEACOCK:  Psychemedics, their studies collected urine and I think hair samples simultaneously.

DR. MANNO:  Okay.

DR. LASKY:  Okay.  Any other questions from the Committee for Dr. --

MR. REYNOLDS:  Stan Reynolds, consumer rep, and I have just one very simple question:  What is synthetic sweat?  I mean is there a standard formula for this?

DR. PEACOCK:  I think you'll have to ask Psychemedics.

DR. KROLL:  Okay, Dr. Rosenbloom.

DR. ROSENBLOOM:  I was a bit confused by the data that seemed to show that there was a lot of false negatives among expected drug users.  The percentages were, as indicated, 75 to 90 percent, but that still leaves a lot of false negatives.  And I wondered maybe that needs to be reviewed by the Company, those slides, because they're not -- those tables didn't come through in the published -- in the handouts.  And I'm not sure what that represents.  Maybe you can clarify it and whether if we know how many RIA negative individuals are subsequently positive by the GCMS or LCMS testing; in other words, what the false negativity rate is with just the RIA.  Do we have that data?  Did I miss it?

DR. PEACOCK:  I don't think I have that data, no.

DR. ROSENBLOOM:  Well, we need that.

DR. KROLL:  Dr. Everett?

DR. EVERETT:  James Everett.  The studies that they used to support the approval of this device, are the methodologies the same or are they different?

DR. PEACOCK:  The testing methodologies?  Each individual study was designed independently by --

DR. EVERETT:  I know they didn't have control over how they were done, but are the methodologies the same or are they different?

DR. PEACOCK:  You mean testing.

DR. EVERETT:  Right.

DR. PEACOCK:  As far as I know, they're the same.

DR. EVERETT:  Okay.

DR. PEACOCK:  The testing.  And once the hair sample came in to Psychemedics, they were all treated identically.

DR. EVERETT:  But from the different studies or the different groups that analyzed the hair samples, initially, are those the same methodologies or are those different methodologies?

DR. PEACOCK:  I guess I'm not hearing what your question is.

DR. WILKINS:  I might be able to clarify his question for you.  I think what Dr. Everett's referring to is in the published statistical evaluation that was in the Forensic Science International 2000 paper that looked at a series of eight studies that were termed small-scale studies and that were retrospectively analyzed, the data that were presented were analyzed and that paper obtained using the same methodologies employed by the company or by the sponsor.

DR. PEACOCK:  I think some of these statistical studies were done on existing Psychemedics data.  They were pulled out of the Psychemedics database is my understanding.  Like I say, it's not totally clear to us either, so maybe Psychemedics can give us a little clearer answer on that.

It's my understanding, when I was reading I think the same article you were, is that it's sort of like Psychemedics has a database of hair sample results, and then --

DR. EVERETT:  But these are secondary data sets, right?

DR. PEACOCK:  Say it again.

DR. EVERETT:  These are secondary data sets.

DR. PEACOCK:  Yes.

DR. HENDERSON:  May I just ask one question?

DR. KROLL:  Yes.

DR. HENDERSON:  Is the FDA aware of any false positives that have been reported for heroin using this methodology?

DR. PEACOCK:  I'm not, no.

DR. KROLL:  Okay.  Does the panel have any other questions for Dr. Peacock?

DR. WILKINS:  I do have one more, and I may have missed it in the material, so I apologize if I did.  But has the sponsor proposed a monitoring program or a follow-up to track reports of either -- or inconsistencies between perhaps what the MRO might perceive to be this is a drug user, and I keep getting negative results or this is someone who will never in  a million years I would have predicted to be positive, and yet I'm getting positive hair tests.  Is there some type of monitoring proposed to kind of assess for that once it were to be introduced?

MS. PINKOS:  We have not discussed that with the -- this is Arleen Pinkos -- we have not discussed that with the sponsor, but that certainly is something you can give us input on.

DR. LASKY:  Can I --

DR. KROLL:  Yes, Dr. Lasky.

DR. LASKY:  Fred Lasky.  I'd like to comment on that.  Presuming that this device will be cleared for market, the device is subject to quality system regulations, good manufacturing practices.  Part of that requirement is to monitor all complaints and to address all complaints and investigate what the source of those complaints -- what the source is.  So, your question is a requirement once a medical device hits the market.  I'm a little surprised that this device, which will be used only within this laboratory is in fact going through this process, but that's the sponsor's option.

And I think also, based on Dr. Kroll's question that I think is part of what I'm having also some difficulty in segmenting but the clarification was helpful, in that we're looking at a collection device, which has not really been questioned from what I've heard or seen in any of the data.  An internal RIA method that in every case will be confirmed with an accepted, definitive method.  And to me that seems like almost like the device is the whole package, and the reference method, the confirmatory test, I guess, has met other standards that everybody is comfortable with.

So, from my standpoint, the only real question is not the presumptive positives of the RIA tests, because they will be verified by the definitive method, but the potential for a false negative and how that might be addressed and whether or not the data support that.  That's what I'm waiting to get some clarify on.

I had the mike, so I just went on and on.  I'm sorry.

(Laughter.)

DR. KROLL:  Any other questions from the panel?

Okay, it's about 12 o'clock now.  What I'd like to do is we'll take a break for lunch.  We'll come back exactly at one o'clock.  At that time, I think Psychemedics has written down some of the questions that the panel had.  We'll give them ten minutes to try to clarify those issues, and then the panel will go in and start trying to address each question.

We'll adjourn till then.

(Whereupon, the foregoing matter went off

the record at 12:06 p.m. and went back on

the record at 1:07 p.m.)

A-F-T-E-R-N-O-O-N  S-E-S-S-I-O-N

(1:07 p.m.)

DR. KROLL:  I'd like the rest of the panel members to take their seats, please.

All right.  I'd like to resume after lunch.

And now I believe Psychemedics was going to answer some of the questions that were queried before and any other additional questions that the panel may have.  And what we can do is we go and discuss the questions.  If other issues arise, we then can -- we'll ask Psychemedics to clarify those issues as they occur.

DR. CAIRNS:  Thank you, Mr. Chairman.

Ten minutes, or whatever time we've been allotted, is a little short, so we're going to try and make our answers as scientifically brief as possible.

Let's deal with the studies that were submitted, in other words, the clinical studies from the methadone or heroin user clinics.  We find that more appropriate than trying to use a hospital study with critically ill people just on morphine sulfate.  The ingestion of heroin is a different issue from in fact the deliberate giving of morphine sulfate to a chronically ill patient.  So, we would reject that suggestion.

However, let me emphasize that the five studies, A through E, that were presented to you all used the same RIA assay method as is being applied for approval.  And we use the same hair test and the same cutoffs.  So, in essence, the five studies, while independent of Psychemedics, used all the same methodology.

Secondly, there is demographic data, and I'm sorry the panel member who asked the question is absent, but if you look at some of the attachments where in fact the results of studies A through E were published, you will demographic information on ethnicity, race, gender, color.  Those are all embedded in those publications.

The next thing that came up, basically, was the issue of bias.  And you heard me say in my presentation there are no conflicting studies.  It's the issue of using different methodologies.  But let me turn this over for a more succinct analysis by Professor Newel.

MR. NEWEL:  Thank you.  First of all, I'd just like to reiterate a couple of major points.  First, on all of the large-n studies that we looked at, there were no -- there was no evidence of any conflict between them.  We all had the same results.  And what's important in the large-n studies is that with the exceptions of the data from Glasgow and Las Vegas were basically run under different methodologies.  All the other studies, the large-n studies, were run under the Psychemedics methodology.  That's actually a plus, because what we've seen as independent evaluators in this issue is that we can't find any bias even with other methodologies used.

In regards to the small-n studies, Dr. Peacock referred to a number of them as saying that there seemed to be some conflict in the literature, and actually referred to most of the studies that I originally showed.  I'm glad that he did that, because this allows me to emphasize an important point.

First of all, none of those small-n studies used the Psychemedics methodology.  All of these original authors only showed mean concentrations.  In one study, they had three African-Americans, and they reported a mean concentration for African-Americans and for the four or five, I can't remember, Caucasians.

Well, it would be nice if the world were as simple as simply looking at the averages of two small means like that and being able to draw a conclusion that would have any validity in a population.  It would be nice.  As Dr. Peacock correctly pointed out, one author actually did use the word "significant."  He said that there is a significant difference in African-American populations.

Unfortunately, that was probably a misuse of the term.  What we're concerned with is statistical significance.  And, so small-n studies simply don't have the statistical power to be able to show those kinds of differences when we ran the appropriate and standard statistical test on their data, on the data that was originally published by all the same authors that Dr. Peacock referred to.  None of those studies showed statistical significance except for the Kronstrand study.

The Kronstrand study had nine people.  When we look carefully at that data with a linear regression, one data point is considered an outlier.  In other words, one person seemed to throw the results into a significant result.  Now, we can't say at this moment whether or not that's an indicator that there is a problem with codeine or not.  I think it's too early.  But remember it's based on nine people and one person in particular that had an extreme score.

Finally, make a comment about the racial diversity in some of the papers that we talked about.  I mentioned very briefly a paper by Ben Hoffman that appeared in Journal of Clinical Occupational Medicine, I believe.  It's a recent paper; he had 1,800 people.  And I gave the reference to the one member who hasn't joined us yet over the break.  That had a very large and diverse population.  It was a group of police officers.  So, it was workplace testing.  It was not a user population.  Drug use was found in only a small percentage of these people, yet again he found no significant hair color, race or ethnic bias.

And also I want to point out that of all the studies that we reported you have copies, I believe, in the journals that you have.  And our experiences with this have included criminal justice populations, probationers, and arrestee populations.  I spent a lot of time in jail cutting hair.  A very ethnic and racial diverse group.  So, we certainly have a lot of diversity in our data.

I think that's it.

DR. CAIRNS:  Thank you, Richard.

We've heard some panel questions concerning false negatives, and I'd like to turn this over for a compare and contrast of the false negative issue, urine versus hair.

Dr. Selavka.

DR. SELAVKA:  Thank you.  I'd like to say for the record I was never in prison cutting hair.

(Laughter.)

I think on the issue of false negative it requires review of the nature of determining the true value of use history for a person who's having their hair collected.  There's been no question, and the question was actually posed and answered correctly, of a false positive with hair testing.  Hair tests accurately identify heroin users.  They have for 13 years with this particular sponsor.  Frankly, there are many other laboratories that routinely detect the presence of 6-mono acetylmorphine and morphine in the hair of heroin users.

The predictive value for the test therefore demonstrates that you'll have a significantly higher positive rate among tested individuals for heroin when you use hair compared to urine -- eight percent versus 0.6 percent in one of the tables that's provided for you.

If you were to turn this question on its head and say that if the question for false negatives was let's use the gold standard of self-reported use of heroin and a positive hair test, urine would have a very low analytical sensitivity, because the false negative rate is so high.

So, I think if we do look at the clinically significant, the procedurally significant, and the programmatic significant question, which is false negatives, you actually have a much better program with hair than you would with urine.  I think we need to remember that.

If you add to that the fact that there's a medical review officer at the receiving point of a laboratory's output and that they're routinely not going to forward or verify those results from a urine test to a company, therefore no action will be taken on a positive, on the other hand you'll have a much more effective and far greater safety measure with hair testing for opiate positives, because MROs will forward hair positive results.

So, I wanted to put the false negatives into context that way.  Thank you.

DR. KROLL:  All right, thank you.

Any other questions from the panel, from the group right now?  We can ask some questions later on too.

Dr. Lasky?

DR. WILKINS:  Can I ask a question?

DR. KROLL:  Well, Dr. Lasky is first.

DR. WILKINS:  Oh, I'm sorry, Dr. Lasky.

DR. LASKY:  Okay.  I have a general question about your history with the use of this particular test.  And as we discussed earlier, there were issues about the complaints in handling and how you resolve them.  And with 13 years experience, can you just summarize for us any issues that you have received from people who use your facilities and how you've resolved them with regard, of course, to the accuracy of the tests?

MR. THISTLE:  Regarding heroin-positive results, in 13 years of testing, with several million samples being submitted, I had looked for an opiate or a heroin claim that I could have showed you when I showed you the cases when I was speaking.  I couldn't find any court cases on heroin.

I looked back to see if we had claims in that regard, and I couldn't find any until I was reminded that we do have one in several million tests of a high user that's claiming as part of a union issue, first of all, the right of the company to test but also challenging his result.  And, so I have one claim in several million tests.

DR. KROLL:  Okay.  Dr. Wilkins?

DR. WILKINS:  Yes.  I have just one other question.  And I guess I may have missed this, and I want to make sure I understand.  But is the intent of the submission to determine heroin use only, because that's what I keep hearing for determining heroin users, heroine users, and -- let me finish my question.

The reason I ask that is because reading through the documents that I have, which may not be perhaps everything that's necessarily been submitted, it appears that the product would be used for "opiates in general" is how it's termed.  And, so therefore morphine due to codeine ingestion, morphine due to morphine ingestion, morphine due to heroin ingestion.

And I just want to make sure that I clearly understand the intent that you -- the setting or the specific application you want to use it for, just so that I comment later appropriately.  Thanks.

MR. THISTLE:  The intended use has been narrowed to opiates due to heroin use.

DR. WILKINS:  Okay.  So, it's only heroin use.

DR. KROLL:  Dr. Manno?

DR. MANNO:  Could we go back to the negative bias again?  I just want to ask a simple question:  Is it safe for me to make the conclusion that the bias in the negative direction -- let me state that again -- the bias is towards the negative side when it comes to hair color?  The less hair color the more chance you're going to have of having a negative.

MR. THISTLE:  I'm not sure I can answer that.

DR. MANNO:  I think what I'm doing is trying to look at the other side of the coin.  I'm thinking if you have a very blonde person, a White person, you'd have less of a chance with a low concentration of picking up a negative or more of a chance of getting a negative than a positive.

DR. CAIRNS:  I think the bottom line answer is there is no bias, first of all, even on the negative side, because, remember, the melanin is spun out of the assay for the confirmation step.

DR. MANNO:  Okay.

DR. CAIRNS:  And for the RIA step.

DR. MANNO:  Did you present studies in here, I may have missed them, where at some point to show the difference between having the melanin present and the melanin absent?

MR. THISTLE:  Yes.  As part of the studies that were submitted, there were results from Psychemedics where the melanin is spun out, but there were also evaluations of laboratories that don't do that.  And you're not seeing statistically significant deviations in any of those studies.

DR. MANNO:  Okay.

DR. WILKINS:  I'm going to ask another one, just because it's following up on the same issue.  Can I do that or should we wait and come around?

DR. KROLL:  Well, Dr. Clement wondered if he could ask one.

DR. WILKINS:  Okay, go ahead.

DR. CLEMENT:  Well, I can defer to that.

DR. WILKINS:  Okay.  We'll bounce back and forth.

The one question I have and perhaps -- and I'm not a statistician, so I'm going to refer this one just for the statisticians in the room.  But when you do a retrospective examination of literature, peer-reviewed literature to support a particular hypothesis or to test a specific statistical hypothesis, it seems to me that there are several really important things that you need to look at.

One is which you've already pointed out, the small size of the studies, is whether or not these studies have adequate numbers of subjects.  Well, you can't have it both ways.  Either they have too small of a subject to be able to rely on the data and use it or it does.  So, it seems to me like if these are very small studies and have low statistical power, they couldn't have determined a difference anyway.

So, they were not capable of answering that question, plus were the studies originally designed to look at that issue?  And I think restricting it only to the opiate papers, not to cocaine, not to MDMA, not to MDE or any other assay, because to me, just as a reviewer, those, in my mind, are not pertinent to your application per se, because you want to do this for opiates.  So, I'm going to restrict it just to the opiates.

And if you go back and look at those studies that were not originally designed to do that, the population that was selected for the two studies in particular weren't selected for this purpose.  So, I don't know if additional -- there may be some additional variables coupled with a low sample size that we don't even know.  So, relying on two negative studies seems to me be as doubtful or inclusive as the one or two human subject, controlled dose subjects that are.

So, for example, the other -- one more issue, is that, again for the statisticians, if you go back and retrospectively examine data that was not originally designed to look at this question at all, does not use the same analytical techniques -- one of the two papers I'm referring to just was a methanol wash, I wouldn't expect it to agree with your data anyway -- and then do a one-way ANOVA only looking at hair color and a single time point, I might not expect to see a difference either.

I mean I'm not -- as I say, I'm not a statistician.  If you had that data and you were doing a multiway ANOVA with dose, for example, or age or hair color or gender, whatever, it doesn't really matter to me what the issue was.  But it seems like you sort of need to control for that in some way, whether it's analysis of covariance if you're looking at dose and hair color or something.

And I'm just trying to understand how you can use it to support the fact that there's no hair color bias if you don't know all those things.  And maybe I'm too simplistic, and that's a statistician question.  I just don't know.

MR. NEWEL:  I would agree with you, everything that you've said.  I think one of the third or fourth slides that I had up said why is this likely that we see four of the five studies when we reanalyze these data using appropriate statistical techniques?  Why do those initial --

DR. WILKINS:  Is that for codeine studies or opiate studies?  I'm just referring to just the opiate studies, because I'm trying to define.  I don't want to confuse the issue in my mind, because you haven't asked me to look at cocaine.

MR. NEWEL:  Okay.  With whatever studies we reviewed there the essential points were that these are small-n studies.  They do not have the sufficient statistical power to be able to show us that there are or are not racial bias issues.

The point that I was trying to make is to draw your attention to the fact that the original authors on papers that said there may be some relation to melanin content or hair and drug incorporation, those original authors did not use statistics.  What they said was, "Here's a big mean, here's a little smaller mean, and therefore, maybe, who knows, there's some relationship."  What we do is take their same data --

DR. WILKINS:  So, those two opiate publications did that?  I wasn't aware that that's what they were trying to say.  My understanding was is the first one was used to evaluate a method development.  And the second one was to develop whether or not you could even detect heroin and that kind of a thing in hair, which is a different question.  That's why I'm asking it.

MR. NEWEL:  Right.

DR. WILKINS:  I want to really understand this.

MR. NEWEL:  That's the second part of the question, is how do you go back and use what are basically secondary data analysis for these issues?  Because in some cases, even with the large-n studies, we don't have the original authors.  On some of the articles, like the Glasgow study, were never intended to collect data for the purpose of comparing hair color or melanin content and drug concentrations.  That just happened to be in their data set.  We became aware of that and utilized that data on a secondary data analysis.

Now, that's not an illegitimate statistical technique.  It's done -- secondary data analysis is done all the time.  That is not the associated with their A priority hypothesis testing at all.  We're basically using an existing data set and reanalyzing it, not with any particular hypothesis test in mind, other than to say is there some association between these two variables?

DR. WILKINS:  Okay.

DR. CAIRNS:  In some cases, I believe it's a mischaracterization of a piece of research work that was put together to answer an entirely different question.  So, I think the data we presented is we've got one Psychemedics opiate, statistically valid study.  The others really ought to be removed from the table, because they don't even measure up to it as regards.  And the study design originally intended to address the issue whereas the other studies are mischaracterized as conflicting but small-n and other experimental difficulties.

MR. THISTLE:  Some of the studies that were provided to you on other drugs were actually provided because you were given studies originally on other drugs.  We would agree that perhaps a cocaine study has no relevance here, and if you were to not --

DR. WILKINS:  I didn't mean to imply that.  I just meant I was trying to narrowly focus my examination to not bring up irrelevant issues.  I didn't mean to imply that it wasn't good to have looked at that material.

MR. THISTLE:  But I agree.  For the purposes of this application, for looking at an opiate assay submitted by Psychemedics, the important studies would be the results obtained by Psychemedics with their opiate assay and not necessarily the results obtained by someone else with another methodology for another drug.  And that's what I think you should be focusing on.

DR. KROLL:  Okay.  Other questions?

Dr. Clement?

DR. CLEMENT:  I have one question.  I guess it's more of a comment on one the slides on the assay performance for the clinical specificity on the handouts we have.  It's on page 8.  This is regarding panel question 1.  It was also on the Powerpoint presentation as well.  Would you like to comment?  Did you actually mean to put results true positives or is that -- I mean --

DR. CAIRNS:  Is this the slide saying, effectiveness of urine versus -- of hair versus urine?

DR. CLEMENT:  No, no.  This is the clinical specificity test that was done on the in-house employees.  N equals 81 --

DR. CAIRNS:  Oh, yes, I have it now.

DR. CLEMENT:  You had at least written on here, and you may want to look at it, it says the results are true positives for N equal 81s.  Does that mean all your patients were positive for heroin use?

DR. CAIRNS:  No, no.  I'm sure we meant that they were all negative.  I apologize.

DR. CLEMENT:  Okay.  I just wanted to verify that.  That would be very happy employment.

DR. SELAVKA:  This wasn't PCP.  This was a --

DR. CLEMENT:  Okay.  I just wanted to clarify that.  So, if we swap those, the correct numbers -- those true negatives were 81; false positives were zero.  Is that correct?

DR. CAIRNS:  That's correct.

DR. CLEMENT:  And then your formula obviously is correct.  So, your specificity is 100 percent.

DR. CAIRNS:  That's correct, sir.

DR. CLEMENT:  Thank you.

DR. KROLL:  Dr. Rosenbloom, did you have a question?

DR. ROSENBLOOM:  Well, I'm back to my confusion about negatives.  Do we know how many in those who you expect to have -- those who are in treatment programs and so forth, how many negatives are obtained in that population and how many negatives in the RIA are retested and found positives in the definitive test?  Does that make sense?

DR. CAIRNS:  I think you're asking the question is were all of the positives via the RIA morphine assay, were they confirmed as positive heroin users?

DR. ROSENBLOOM:  Yes.

DR. CAIRNS:  The answer would be yes.

DR. ROSENBLOOM:  Okay.  So, there would be then no reason to look for false negatives in the RIA.

DR. CAIRNS:  The cutoff level is determined so, a, the poppy seed ingestion is not detected as a false positive.

DR. ROSENBLOOM:  Right.

DR. CAIRNS:  So, the cutoff is a protection mechanism so that the poppy syndrome does not enter into the equation in the determination of false positives.  The two nanogram cutoff is still there to be able to detect a minimal detectable dose of 173 milligrams per month, which is very little concern with the consumption of the average heroin addict on a monthly basis.

DR. KROLL:  Dr. Kurt has a question.

DR. KURT:  Dr. Cairns, I'd still like to clarify, because you've explained to us how in separate scientific articles the extraction method has shown so many results as positive.  And then in other articles not using the extraction methods there are so many positives.  If you'd explain to me in the same scientific article that samples were divided, some were extracted and some weren't, were the same number of positives above the two nanogram level received or is the extractor group fewer than past the two nanogram bar for positive?

DR. CAIRNS:  Yes.  Let me go back to a fundamental procedure again.  That first RIA sample is an unwashed sample.  The second sample, should it be presumptive positive, is washed.  So, the question came up was the comparison of the two different flow charts within the lab.  And, yes, we've done comparisons.  They're in the submission, volume 3, page 1 through 23.  And there was no significant results between the two mechanisms.

DR. KROLL:  Okay.  Thank you.  Are there any other questions before we proceed to try to answer the questions.

Okay, Dr. Everett?

DR. EVERETT:  Most of my questions have been answered, but I'm still unsure about the issue of the FDA indicating that the literature says one thing and then you guys are saying something different; that is, whether or not the literature supports the differences between the biases where you say there is no difference, and race is only one of those.  In your opinion, is it the same or is it different?

MR. THISTLE:  The differences don't exist in the large population studies regarding bias.  And the differences don't exist in any study utilizing the technology that we're putting forward here.  There's no study that utilizes this technology that is in conflict at all.

With regards to the other issue of external contamination, there is no conflict in the studies.  We do a wash of three hours and 45 minutes and analytically extend it another five hours.  Some of the literature that you have has a wash mechanism of 30 seconds.  There's no conflict there.  We know the 30-second wash will not remove external contamination if it's present.  That's not a conflict, that's just the differentiation in procedures.

DR. EVERETT:  Then you're sure that when you say it's removing contamination from the external surface it truly isn't leaching from the internal surface.

DR. CAIRNS:  Yes, I think we demonstrated that on the wash profile for a typical heroin user where you saw the third, fourth, and fifth wash approach a plateau value.  So, you could imagine the external being removed and then reaching a plateau.  If you continued to wash, that would diminish, diminish, diminish, and then suddenly you have this 30 nanograms of morphine as a skyscraper against the profile of the wash.  So, there's a clear way visually to distinguish a user from someone who's environmentally contaminated.

DR. EVERETT:  Well, the reason I mention that, because in perm and bleaching hairs, frequently there is damage.  And in looking at the studies as retrospective studies, as a prospective study, that doesn't seem to bear out.

DR. CAIRNS:  Yes.  Obviously, when we quoted to you that certain cosmetic treatments would in fact reduce the levels, that is because some of the outer regions of the hair are in fact structurally affected.  But please remember that what goes in easily comes back out easily on the wash kinetic profile that we use.  So, there's the level playing field.  What goes in easily comes back out easily if in fact you're dealing with treated hair.

DR. EVERETT:  It doesn't appear the playing field is really level, because the retrospective studies were not -- many of them, as I look at them, weren't designed to do that.

MR. THISTLE:  I know what you're saying.

DR. EVERETT:  But the data is presented by using that fashion, but they really were never set out to do that.

MR. THISTLE:  I know what you're saying, but the element of consistency between studies A through E and the large population studies that were deliberately designed to look at biases, there's a consistency thread in the methodology running through these.  So, you can do direct comparisons.  Our problem is comparisons where methodologies such as 30-second wash with methanol, that is no way can be compared to a result from an extensive wash procedures that's previously demonstrated to show the differentiation.  You must first of all differentiate external from ingested before you make any statements as regards bias.

DR. KROLL:  Can I suggest we continue now?  We can try to clarify these issues as we proceed.

DR. EVERETT:  I have one other question, and that deals with the essence of deciding to do this as a retrospective study, and I think many of these questions could have been most appropriately addressed as a prospective study.  I'm interested in your reason for choosing retrospective data as opposed to prospective data.

MR. THISTLE:  The contamination studies that -- the question I think started out with contamination issues and leaching of drug out of hair.  The contamination studies weren't done with retrospective data.  We submitted data on hair that we contaminated.

And I think we're confusing two different issues here.  The retrospective data was done for the sensitivity and specificity.  And there's some retrospective data on studies that look at hair color and race and denoted those features.  And I'm not sure what the difference between -- those features don't really change, whether you're looking at the data retrospectively or prospectively.  Those are standard features that you can look at regardless of when you get that data.

But I think we're mixing two things.  We're mixing contamination studies with -- actually, I think we're mixing three things here.  We're mixing contamination with bias studies, which were separate from the retrospective, clinical sensitivity/specificity studies.

DR. KROLL:  Okay.  Why don't we proceed now, and thank you very much.

DR. CAIRNS:  Thank you, Dr. Kroll.

DR. KROLL:  Does Dr. Peacock want to come back up and ask the first question for us?  Is he here?

DR. PEACOCK:  Can I just make one statement before we put the questions up?

One of the reasons we're here today asking these questions is for the very reason you're asking the questions of Psychemedics.  The literature is conflicted, and we had trouble deciding what is correct and what is not.  Psychemedics has their position.  People in literature have different positions with different methods, and that's one of the reasons we're asking these questions of you today, to help us get a feel from the experts that aren't directly involved in the review.

I'm just supposed to read the questions now.

(Laughter.)

Sorry.

Question 1:  The clinical data in this application is from research reports and data collected from diverse sources and not from perspective controlled clinical trials that evaluated heroin.  Therefore a study hypothesis, inclusion/exclusion criteria, associated end points, and a plan of statistical analysis were not provided.

A, can assay performance be established with these types of data?  Why or why not?  And, b, do the data presented provide adequate characterization of assay performance?

DR. KROLL:  Okay.  What I'd like to do is start with Dr. Everett, and then we'll work our way around counterclockwise.

DR. EVERETT:  Okay.  For question 1, can the assay performance be established with these types of data?  Certainly they can, but the question would be whether they were reliable or not, and I don't think they really would be reliable considering the data that we've already looked at.

And then the second part:  Do the data presented provide adequate characterization of the assay performance?  Clearly not, and that's because the intent in the studies are different than what I've seen at this point.  So, I just tend to disagree with this.

DR. WILKINS:  I agree with Dr. Everett's comments, and I think for me I do agree that in some cases, again, that certainly this type of data analysis can be very useful and helpful.  However, I think that especially for a new or first-time application or first time it's been seen, we don't have a lot to compare to, to go on.  That unless we've controlled for some other -- for some critical factors, and I'm not necessarily intimating that means hair color, for example, or anything like that, but unless the study is, in my mind, adequately controlled for those potential variables that might have impacted on the quantitative data that's being used to support the application, I just don't feel like it's sufficient to go and try and pull that out retrospectively from someone else's data in some cases.

The reason I think that's important is not so much related to the specific techniques used but rather to the package insert and interpretive guidelines that are going to be provided for the consumer.  And for me that's the bottom line in my mind is do the studies that are incorporated to make the particular points, are they adequate to support whatever we're going to tell the consumer for interpretation.  And I do not feel comfortable in doing that unless the study has additional controls.  And right now I don't feel that that's there.

DR. KURT:  Tom Kurt.  As a Medical Toxicologist, I would like to say that I'm quite enthusiastic about the prospects of hair testing versus the problematic existence of easy cheating that has arisen under the urine drug testing.  However, I cannot say that my enthusiasm buoys me over the problematic issues that both Dr. Everett and Dr. Wilkins have explained, with which I agree.

DR. KROLL:  Martin Kroll.  I tend to be in agreement with what previously has been said.  I mean I tend to see that retrospective studies can be used if they all are basically in agreement and they sort of cover most of the possibilities.  And from the data I've seen submitted I'm not seeing that here.  It's not there, and it can't be collected, but it's not necessarily apparent that it's all here. Even a small prospective study with very clear hypotheses and set goals perhaps can clarify some of the issues.  I think the fact that so many people here on the panel are confused about what information is presented and where things stand points out one of the problems.

There's also other issues that we keep coming back to hair and hair color.  I guess I get concerned with hair treatments.  There's a lot of variation.  Certainly, there could be some small studies there where you take an equivalent amount of hair and treat it many different ways so that you can look at all the many different types of treatments.  I think those are things that at least I feel are real important for the FDA to consider.  And I would like to see that in a submission.

DR. MANNO:  I can't add anything more to what's already been said.  I agree with all that's come before.

DR. LEWIS:  Sherwood Lewis.  I, likewise, am in agreement with the foregone statements by the panel.

DR. CLEMENT:  Steve Clement.  Well, I never have troubles disagreeing with my colleagues on some issues, and I think on this one, looking at the practical issues of prospectively doing a prospective study of challenging folks with heroin, which is the actual agent to be analyzed, is very problematic, not only in this country but anywhere in the world in terms of inducing tolerance, inducing addicts, inducing all kinds of dependency.  The studies are never going to get done.

And the sponsor did submit these four or five studies that did user assay.  Actually, I think there's some benefit of not having the sponsor involved in the data analysis, because clearly that alleviates some of the bias that could be done in-house by the sponsor, and the studies were pretty uniform.

So, I think even though it's imperfect data in terms of looking at sensitivity, it's the best there is, and I don't think I could come up with any recommendations to improve on that based on the difficulties of treating this disease.

And the issues on specificity look like we cleared up.  It looked like most of the employees are very happy but not on drugs, so it makes it 100 percent specific, which is fine with me.

DR. HENDERSON:  Well, I actually gave my comments earlier, I think, inappropriately.  I certainly agree that there's problems with retrospective data.  But I too agree that I don't quite know how else you could do that.  But I do think that there are opportunities to obtain retrospective data or perhaps even prospective data in diverse populations.

I mean the data that we were submitted from the studies I think without seeing the demographics, because I gather that's not available, but just from the location of the studies, I think they're probably fairly homogenous populations, and that concerns me.

Again, I think most all of these were in populations that were known to use substances illicitly, and some of my concern in the workplace is what are the predicted values for groups of people who do not have a history of using substances?

DR. ROSENBLOOM:  Yes.  I agree that it would be nice to have populations defined prospectively and with the ability to answer all these questions without rejecting the retrospective data, which is convincing.  But I think that there are more studies that need to be done.  So, I guess I'm halfway between what's been said by everybody else.

DR. KROLL:  Okay, thank you, Dr. Rosenbloom.

DR. LASKY:  Fred Lasky.  First, I think that the small studies should just be dismissed, and I think that was the intent of the presentation; that they're not statistically valid, no conclusions can be brought from them, so we should just forget about them.  We should consider the larger studies that were done to demonstrate the effectiveness or ineffectiveness of the method that's under consideration.  And I think that's what we need to concentrate on.

I think that retrospective data is okay, but you have to be very careful with it.  So, I wouldn't dismiss the fact that this is retrospective data.  And it also brings to light the fact that there's over a dozen years of experience that has been had with this assay and the fact that hair testing is based on what we've heard earlier in the day is really an important matrix to look at.

And with that in mind, I think there are two issues, at least, that I'm trying to pull apart in my mind.  One is, is hair as good as people who are using the test believe it is?  And I think that's one of the issues that is in question.  But I think that is true regardless of which test is being done.  So, I don't think that the test -- that the submission before us is in a position to either answer that question one way or another if it's going to be an accepted practice.

I believe that the sponsor has done studies that demonstrate in reasonable studies that in fact hair is a reasonable matrix, that they've done adequate studies to demonstrate that they've taken the precaution to exclude as much as possible any contamination that may occur from the outside.  The washing steps I think are -- and the data are convincing in my mind.  And the way they handle those data demonstrate to me that they are making an effort to reduce the possibility of a false positive.

So, I think if there is an issue to be had here it's whether or not the sample that is collected provides useful information with all of the steps that is in the submission.  And I believe that the sponsor has demonstrated that the sample mix and based on what's in the literature demonstrates hair is a reasonably effective sample, and that they do a good job in trying to measure what's actually ingested through hair analysis.  So, I think it's reasonably convincing.

MR. REYNOLDS:  Stan Reynolds.  I agree with the point that Dr. Clement made that there's really no way you're going to set up a controlled, clinical study for hair.  You're just not going to do that.  But at the same time, as Dr. Henderson and some others pointed out, that doesn't preclude the possibility of gathering the demographic information that you want.

I mean there's no reason that you can't go to methadone clinics and treatment groups and get samples from people who are known heroin addicts and at the same time have all the demographic information that you need:  race, age, ethnicity, hair type.  All that information would be available.

So, at the same time, you can design a study that would give you all the demographic information to have and have the confirmation of the standard urine test, confirmation by GC/mass spec or anything else that you need to show a positive is a positive and a negative is a negative, at the same time getting all the demographic data.

And we just don't have the demographic data here.  You know, we're comparing apples and oranges.  You have study done one way, one study done another way.  I believe the data that they have presented for the test is compelling.  It shows that the test does work.  But we don't know about the population demographics.  That's just the information we don't have at this point.

DR. KROLL:  All right.  Thank you.

Why don't we go now to question 2.

DR. PEACOCK:  Question 2:  With respect to making claims for clinical sensitivity and specificity, is a single negative urinalysis plus a negative self-report of drug use a sufficient, unbiased standard for the establishment -- for establishing true drug-free status?

And part B:  Is a positive urinalysis that is not confirmed plus a positive self-report of drug use a sufficient, unbiased standard for establishing true drug-free status?

DR. KROLL:  All right, thank you.  And this time we'll start with Dr. Clement, and we'll work around clockwise.

DR. CLEMENT:  The first question, I think it may be the best we have, a single urine analysis, although we know the inherent problems with that.  And particularly with self-reporting, there's incredible problems with that.  It's not the gold standard, but it's the best we have.  But I think in this case it's really not applicable, because in the data that was presented by the company it was 100 percent negative, and they're showing not only data of a single urinalysis but multiple urinalyses as well in a fairly contained population.  I think that's probably the best that can be potentially done in any type of situation.

So, I think it may not be the best standard, but it's probably the best that we can ever hope for.

And on the second question, I was trying to think in my own mind why would someone self-report positive taking drugs for some other reason besides he's really taking drugs, unless it's to try to get room and board in some rehab center, and he's not taking drugs.

So, clearly, there could be some underestimation on that.  But even if that's the case, that would underestimate the performance of the assay as opposed to overestimating the performance of the assay.  And in the numbers that are shown here, the performance is in the range of 70 to 80 to 90 percent.  So, even given that problem, if anything, because it would potentially underestimate the performance, it's still not bad at numbers of 80, 85 percent.

So, I say, yes, it's probably true, but it's probably too true and unrelated in this case.

DR. KROLL:  Dr. Lewis.

DR. LEWIS:  I'm of the school with regard to A and B, which says you don't believe anything you hear and only half of what you see.  And as far as self-reporting, I put it in that first category of what is it you hear from the individual as far as claiming use or non-use.  So, I have to sort of reject both of those on that basis.

DR. MANNO:  Manno.  I'm a little bit with Dr. Lewis in terms of dependability of self-reporting.  My experience has been that the percentage of people that are honest in these situations is not too great, so I have a problem there.

But, again, we don't have too much more to go on here.  So, I guess a negative urinalysis and a self-report if you're going to be able to do -- and I'm taking this to mean in the -- at the end of the overall scheme of things to make an interpretation that you may perhaps see this person back later.  I guess I can go along kind of tentatively with A.

I've never been one to rest easy not confirming a positive urinalysis, so I'm not too comfortable with that.  I would much rather see it linked up with a hair sample confirmation, but then again there's a problem with when is it going to be seen in the hair.  We haven't even discussed that here today.  What is the delay from use to the first time you can detect it in hair?  So, I'm just uncomfortable with both A and B, as the way it's presented.  I'll just leave it at that.

DR. KROLL:  Martin Kroll.  I tend to agree with Dr. Manno.  I'm looking at this, and from the data that was presented a single negative urinalysis will a lot of times -- there could be a lot of false negatives there.  So, I think you probably need something a little bit more than that to establish a negative.

DR. KURT:  Tom Kurt.  A single negative urinalysis is too easy, as we know, to cheat on nowadays, and so it's not necessarily a reliable test, but that's what we rely upon under the DOT standards.  Nonetheless, we know the self-reported histories are totally unreliable, and I'll cite the Australia study, the same Australia that I referred to earlier in pediatrics in 1993 who studied the meconium stools of 3,000 infants in Detroit and found that 44 percent of them were positives for opiates, cocaine or marijuana, and questioned the same mothers and found that 11 percent, or one out of four, of the mothers admitted they were using drugs of abuse during pregnancy.  So, one out of four admitted.

A similar Fendrich study reported in the American Journal of Epidemiology last year on 320 samples of hair found only one out of our told the truth about their drug abuse that was verified by hair sampling.

So, self-reported history is generally markedly underestimates or is reasonably wrong.  And it's much more reasonably under two, or B.  The B portion is if the positive urinalysis is reviewed by a reliable medical review officer in questioning the person to find out whether or not they've been using other drugs that could possibly render a test positive under the circumstances.

I would say that a secondary sample, such as the hair sampling test once it becomes more accepted into the Department of Transportation review, should be acceptable in a medical review officer's situation also.  But I can't rely upon the self-reported history.  I reliable upon a reliable laboratory test.

DR. WILKINS:  This is Dr. Wilkins again.  I just want to comment.  I don't have anything in addition to add to the previous speakers.  I agree with actually aspects of all of the previous comments.  And my significant reservation is the issue of relying on self-report and the issue of a single urinalysis result coupled with the timing issue of collection.  And I'm not -- I'm using that simply for -- the study design is where I'm concerned about that.  So, that's it.

DR. EVERETT:  James Everett.  My concern here is that the initial test is considered to be a screening test.  It's not a guarantee that you're going to be drug-free.  But given the limits of science and the sociological impacts, part A where it indicates that the negative urinalysis plus a negative report is okay with me.  That part is okay.  You can't retest everything.

But when you develop a positive test at this point, the general rules for screening tests is to select out those people who most likely have what you're looking for, and they are cheaper, they're easier to perform, and the idea is that you will follow-up on a positive screening test to determine if it is a true positive or if it's a false positive.

So, in reality, performing a screening test without some method for determining if it's truly positive is almost irresponsible.  It's like why would you screen somebody for a disease and then you find the disease on a screening test but you don't look to make sure that it's there.  So, it's not reasonable to perform this test, then get a positive test, and then rely on the person's report now to determine if you're going to do the confirmatory test.

The rules have to be consistent with any scientific test.  If this is how you're going to do it, then it can't be based on what the person says.  It has to be based on the results of the screening test, not necessarily the report of the person, because if you do, it opens up again this ram for abuse where perhaps the screening test was tampered with.  And now you induce a false situation where a person admits to something they really weren't doing or they did but it wasn't in that time frame.

So, in reality, you must be consistent with the rules.  If the screening is positive, then the confirmatory test must be done regardless of what the person's history.

MR. REYNOLDS:  Stan Reynolds.  And I pretty much agree with everything the folks in the Committee have already said.  I don't have any additional comments.

DR. LASKY:  Fred Lasky.  I'm not completely comfortable with the use of these methods, but as Dr. Clement mentioned, I don't know what else to do, quite frankly.

I have two comments with regard to what other people have said and also just an editorial note, that many years ago, probably about 15 years ago, Ray Gambino reported in his lab report to physicians that he had done a study, actually he was looking at sensitivity of lab methods for alcoholism, and he determined that in fact a self-report was more sensitive than a serum alcohol level for determining alcoholism.  And it seemed to me, that was quite enlightening for me as a laboratorian, that just asking a patient would provide more information than I could provide from the laboratory was a tremendous blow to my ego.

But it seems that these sort of patients or subjects would not want to self-report unless there was -- I would think there would be bias to say that they weren't drug users and they were.  And the intent of these samples is to provide a challenge to the method that's being evaluated to have a high probability that indeed these samples did have drugs in them and that that is the only reason that these samples were selected.

If in fact the subject did lie, it seems to me that that would be a greater challenge to the method that is being evaluated because of the reliance of being truthful on this, quote, unquote, "gold standard" as tarnished as it is.

So, my bottom line assessment is it's far from perfect, but I don't know how else to do this.

DR. ROSENBLOOM:  Rosenbloom.  Yes, I have with looking at either A or B as a gold standard, because we're talking about evaluating a test that is measuring something quite different than what a urine test measures.  And we're talking about true drug-free status if looking at B a positive urinalysis not confirmed, that's not confirmed on a repeat urine I take it, not on the same urine -- on the same urine, okay.  But a positive self-report could relate to drug use in the past but not in the past few days and why that would establish a true drug-free status I can't imagine.  So, I have trouble with these, because I have trouble with the concept of a gold standard for sensitivity and specificity for a test that could be itself the gold standard that's replacing or saying something else about drug-free status.  But I don't like either A or B.  I guess that's what I'm saying.

DR. HENDERSON:  Well, I certainly, as others have mentioned, and there are problems with both of those, but having taken care of a lot of pregnant drug users in my special population, I think the self-reporting is a good edition, although new mothers may not confess to having used drugs that their babies are obviously suffering the exposure.  But certainly pregnant women are very likely to admit it primarily because they're afraid that not admitting it will jeopardize the health of their fetus.  So, I tend to believe women when they tell me that.

So, it may just be a special population, but I don't really have objections to either of these, although obviously they're not perfect.

DR. ROSENBLOOM:  I think we have to be careful about interpreting admission data for alcohol as people will readily admit that they've had a few drinks, particularly if they know that they're going to be tested, and they know the reliability of the testing, and it's a legal drug.  So, people will admit to smoking too, but I don't think that relates at all to illegal drug use.

DR. KROLL:  Okay.  Good.  Thank you.

Why don't we proceed and go to question 3.

DR. PEACOCK:  Question 3:  Should the minimum does required to produce a positive result be determined?

DR. KROLL:  Thank you.  And this time I'd like to start with Dr. Wilkins, and we'll go counterclockwise.

DR. WILKINS:  Thank you.  I think my answer to this is truthfully it depends in that if the intent of the product is to say that heroin use, understanding that it's specifically for heroin, heroin use can be detected within the last 90 days, then I think you do have to answer this question, because you need to know whether or not you really would have been able to detect heroin use if it's a single use within the last 90 days.  I mean I'm fairly open on this issue.  To me it depends on what do you want to do with that result and what are you going to say about that result?

So, again, to me, it's an interpretation issue or an ability to provide adequate interpretation issue for a positive or a negative result, whichever way it goes.  Was there -- you know, if they use a lot, will you be able to detect it?  If they use only once, will you be able to detect it?  I'm not even suggesting that you, and I'm certainly not suggesting that we, go out and administer heroin to non-drug using volunteers and give them heroin or even drug-using volunteers and give them heroin.

I think that they're perhaps getting back to our -- this directly relates for me to the other question that you can design a perspective study that is well enough statistically controlled, I think, in the data analysis to do this without having to administer necessarily controlled doses to non-using drug volunteers to answer this question.

So, if the intent is to be able to make statements about your ability to detect individuals as positive or negative, and I think you need to know the minimum dose required to product a positive result.

DR. KURT:  Tom Kurt.  I agree in part with that, but I think this is best left up to an advisory board that reviews on this particular topic, such as the DOT cutoffs were set by an advisory board specifically set for such and Dr. Selavka who was here earlier with his Hair Testing Working Group could very easily set up a kind of a working group in that kind of a situation to review this exact topic based upon the existing scientific studies, taking into account what would be considered a rather substantial use, which I think in part has been explained in the sponsor's application.

DR. KROLL:  Martin Kroll.  I have a tendency to like to see that the minimum dose and detection of any system is determined sort of independent of the matrix.  Now, I understand there's a lot of ethical issues; you don't want to give heroin to people.  But you do appear to have information that can relate the amount of morphine.  That way you could detect versus the amount of the MAM in hair.

And there are certainly plenty of people who take morphine for medicinal purposes where it could actually be fairly well controlled.  You know what the exact dosage would be.  So, it might be appropriate to use something like that as a surrogate to get at least an idea so that if somebody was interested in what the minimum dose was, that there would be some reliable information.

DR. MANNO:  Barbara Manno.  I think I heard something earlier today concerning the fact that this test is really a test to detect chronic use of heroin versus it's not so great for acute use.  And looking at the pharmacokinetics, assuming that this is -- this is a big assumption, I know -- but that the first test, the RIA test, as proposed here, would be confirmed knowing that you have such a short window of detection of 6-MAM in urine, I'm wondering if we don't need to know when that's going to show up in hair, which then reflexes back.  >From that standpoint, we're going to need to know when an approximate minimum dose to give us a positive.

There's a second issue that I don't see -- I did not remember seeing, but here again is addressing just the assay itself, and that is what's the LOD and the LOQ of the procedure itself, both within day and between day?  So, there's really two issues to determine whether -- you don't know what you're looking for to start with without those two things.  And -- yes, that's all I have to say.

DR. LEWIS:  Sherwood Lewis.  I don't think that it's at all reasonable to even talk about minimal doses when we're dealing with heroin.  This is not like conducting a controlled clinical trial where you give the particular drug or medication under known conditions, quantitatively, and then look for whatever the results might be, if their analytical results, by drawing blood samples, what have you.

So, to me, it doesn't have any meaning.  Knowing the notorious nature of the material itself and the roots of administration with regard to heroin, I don't know how you could even begin to talk about minimal doses or any other doses in that respect.

DR. CLEMENT:  I concur with Dr. Lewis.  I think the sponsor showed a negative threshold between their sensitivity of doing the test, the two nanograms, to pick up chronic heroin users, but still is high enough so it does not pick up poppy seed users.

I guess my ignorance of not being familiar with the toxicology is whether or not cough syrup and codeine and cough syrup and other forms of codeine can counteract particularly on a MAM test, but I'll be interested in the sponsor's comments or any of the other panel's comments on medicinal amounts of codeine can be found positive both on the urine test and the confirmatory.

DR. KROLL:  We can try to clarify that now.

MR. IRVING:  Hi, I'm John Irving.  Codeine does not produce 6-MAM in either hair or urine.  It's a specific metabolic product from the use of heroin.  There's no way the body can produce 6-MAM other than that.

DR. CLEMENT:  Okay.  Thank you.  Well, in that case, I think the sponsor's done a sufficient job on clarifying the sensitivity based on the studies available and it's within a sufficient range.

DR. HENDERSON:  Cassandra Henderson.  From the materials submitted, it was clear that MAM was only produced if a person ingested or used heroin in some manner.  So, as a clinician, I don't really care what the minimum dosage is.  If it's there, it's there, and that's all I need to know.

DR. ROSENBLOOM:  Yes, I agree.  This is an illegal drug.  If it's there, it's there.  I don't think -- a minimum dose required to get a positive result might be interesting and someone might want to do a study in legally administered morphine, but I don't think -- oh, it's just in heroin, yes, so there is no legal heroin.  So, I don't know how it could be done or why it would need to be done.

DR. LASKY:  I agree.  It's the obligation of the sponsor to demonstrate that their cutoff point is analytical supportable, and that's the limit of what this test is designed to do, and which that also addresses question number 4, in my opinion.  Trying to move ahead.

(Laughter.)

MR. REYNOLDS:  And I agree with Dr. Rosenbloom and Dr. Henderson that this is an illegal drug.  There's no level that's acceptable.  It doesn't really matter whether you're talking about a low dosage over a period of days or weeks or a large dosage over the course of a weekend.  If it gives you a positive, it gives you a positive, and it doesn't matter how the positive got there.  Whether it was chronic low dosage or short-term high dosage, a positive is a positive.  So, I don't think it's really that critical.  I think that they've established their detection limit, and that's the only thing that's really pertinent in this area.

DR. EVERETT:  James Everett.  Certainly, I think the minimum dose should be detected or determined, which they have.  And that's very important for any test that you do.  You must know what the detection limits of that test really is.  Because this is test performance we're talking about, and this is the real reason for detecting things that might cross-react with your test.

It's when your lab results fall outside of that range, you have to question whether the test is accurate or not.  And without those detection limits, you have no real recourse for knowing whether you're dealing truly with an instrument failure, procedure failure, a technician failure or in this particular case something the patient or the client may have done to their hair.

And in a real-world situation -- I am family practice, and I see tests all the time that are incorrect.  They don't necessarily agree with what the patient says.  but in essence the question becomes who do I believe the test or the patient?  And the first thing I look at is the parameters of the test.  What did the manufacturer say the test is capable of doing, and then I move on to evaluate whether or not the test truly is positive or whether it's a false positive.

But, in essence, without looking at the patients, if you're going to have a test, that test must have performance parameters, and you must know what that test is capable of doing, if for no other reason than to determine if the test is capable of doing what the manufacturer says it can do.

DR. KROLL:  All right, thank you.  Let's go to question 4.

DR. PEACOCK:  Question 4:  Should the relationship of the pharmacokinetics of drug use and the incorporation of drug into the hair, that is single dose, multiple doses, and chronic use, be determined?

DR. KROLL:  Okay.  Let's start with Dr. Lewis, and we'll go --

DR. LEWIS:  Yes, I'm still mulling over my response to question 3.  I read that as meaning minimum dose, literally dose, not detection limits, as Dr. Everett was speaking.  I certainly agree with him on that.

And I would say that for question 4, to get on with it, it would be nice, but I don't think that it's necessary this purpose to have that pharmacokinetics information.

DR. KROLL:  Okay, and let's move clockwise.  Dr. Manno.

DR. HENDERSON:  Should we move back just briefly to 3 to comment on Dr. Everett's interpretation of 3 that it's the parameter of the test?  And I think Dr. Lewis and certainly I thought that it was how much drug does a patient have to take in order to get a positive result.  And I don't care.  I mean whatever they take is -- but Dr. Everett's point is that the parameters of the test need to be identified and specific that you can go to the packet insert and everybody knows what the test measures and what the cutoffs are.

DR. KROLL:  Okay.  Maybe Dr. Peacock can clarify this question and whether we've answered it adequately.

DR. PEACOCK:  We were thinking about the dose required, especially with the 6-MAM and the morphine.  You know, you might -- it goes back to the variation, you know, we were worried about for hair, bias to hair color as well, that it was related to the dose of drug being taken, not the performance of LOD.

DR. ROSENBLOOM:  I think that's the way most of us read it.

DR. KROLL:  Dr. Manno?

DR. MANNO:  Barbara Manno.  Like the physicians, I don't care if a person took a single dose, multiple dose or they're chronic users, but what I do concern myself with from the laboratory standpoint is did the sample get collected in the right time frame relative to drug use, and that gives me -- I'm more concerned about that aspect of the pharmacokinetics at this point than I am anything else.  So, a full pharmacokinetic profile I don't think is necessary in order to say whether the test is working or not.

DR. KROLL:  I tend to agree with Dr. Manno.  I think if there's information in literature or other studies, if they can be put together in a very nice review and they can be modeled some sense to show how different types of patterns or drug use over different periods of time how that would get incorporated in here and what you expect to see, that that would be adequate.

DR. KURT:  Tom Kurt.  Because this is a testing for chronic drug use, I think the pharmacokinetics is somewhat of a moot point, except as Dr. Manno pointed out, the sample needs to be collected proximate in time.  It's nice to know that there are other metabolites that are captured by the hair, such as the 6-MAM and are there, but I think it's important to rate this as a chronic test, and perhaps in the future it could be used in companion with the an acute test, such as a saliva test.

DR. WILKINS:  I agree with the previous speakers, Dr. Kroll and Dr. Manno and Dr. Kurt, what they just said.  And I'd like to say that I don't think full-scale pharmacokinetic studies are necessary for chronic use.  However, again, to me, to my mind, just as a panel member, it's going to totally depend on what the claims are that are going to be made.  And in my mind, if the claims for the consumer or the clinician is going to rely on this test to assist them in making a determination of drug use or whether, a treatment intervention is necessary or whatever.

And as it's been pointed out by the panel member, what they really want to know is, is this -- if I've got a positive, I'm fine.  Well, great.  So, are you always going to find if somebody takes heroin, how is the test going to be reliable enough to determine that this person took heroin and I can detect it?  And am I going to pick up every positive, and am I going to accurately determine the negatives?

And I think to some degree you can't really -- I don't think that the data that I've seen has really answered that question for me, that if you get a negative result, for example, that that means the person has not used heroin, okay, totally.  And I realize that's an issue of false negatives and false positives, but for most products you're looking at sensitivity and specificity, and I need more information for that, I think.

The other issue that hasn't come up yet, but really to me relates directly to questions 3 and 4 in a way, is that the one thing that I haven't seen yet, and again the sponsors may have the information, is for 6-MAM, if this test is intended to be used for heroin use.

Just a simple question:  If you have a hair sample or hair from a donor and you analyze that hair sample 20 times from the same donor, do you always get the same -- how does that relate to your cutoff and the positive result, not standards or controls but a donor whose hair contains drug, and you analyze that over and over again?  Can I rely on the fact that if I get a positive result this time, I'm going to get a positive result the next time and the next time?

That's the reproducibility of the test system that I don't think we've talked about at all, and I may have missed that.  But to me that relates to questions 3 and 4 and the ability to interpret a positive result of the test.

MR. IRVING:  My name is John Irving again.  Let me put this in perspective.  If I'm a drug user, a heroin user, and I get collected now and I'm positive, I get collected four hours from now, I'm negative.

DR. WILKINS:  That wasn't the question.

MR. IRVING:  I know.  If I take a hair sample now and I repeat the sample numerous times, I'm going to get a positive.  If I collect the sample tomorrow and an individual was right off the cutoff, I'm going to continue to get a positive.  So, we are going to get a much more -- I'm not sure exactly what your question is.

DR. WILKINS:  I don't think that that answers -- I'm just saying if I take a heroin user's hair sample and take a sufficient sample that I can analyze that same pool of their hair or specimen that is received by the laboratory multiple times --

MR. IRVING:  I think that that is in the submission.

DR. WILKINS:  -- on the same sample -- well, I haven't seen any data.  That's why I'm asking the question, on hair, not --

MR. IRVING:  It's in the submission we passed on to the FDA.

DR. WILKINS:  Okay.  Because that was one of the questions I had is for 6-MAM how reproducible is that?  And if you have a sample that's around your cutoff -- it's defining the issue of positive.

MR. IRVING:  We had data in there that provided that information, and --

DR. KROLL:  Excuse me, Dr. Gutman wants to make a comment.

DR. GUTMAN:  Yes, that's not a review issue.  We have that information and didn't think it was a problem.

DR. WILKINS:  It may not, but for me to answer these questions I felt that that was something I needed to know.

DR. GUTMAN:  Diana, it's in the submission under physician data.  We went even as far as taking the sample, testing it a year later, the same sample, and got comparable results.  That's in the submission.  We took the sample and --

DR. WILKINS:  No, I thought some of those decreased over time, didn't they?  But that's a different issue.

DR. GUTMAN:  A slight decrease.

DR. WILKINS:  Okay.

DR. GUTMAN:  It's within experimental range, but that was also part of that submission.

DR. WILKINS:  Okay.

DR. GUTMAN:  So, yes, we will continue to get a positive on that sample.

DR. EVERETT:  James Everett.  I don't think the pharmacokinetics under these conditions is truly necessary.  But I didn't read anything in the papers that we got that described whether the hair was incorporated -- whether the drug, rather, was incorporated into the hair uniformly in its distribution.  And that is did they have a particular site where they thought the hair sample should be taken from; that is, is it incorporated in the back of the head the same rate as it is in the front or the side?  I didn't really see that.

DR. HENDERSON:  It's there.

DR. EVERETT:  It's there?

DR. HENDERSON:  It's there.

DR. EVERETT:  Okay.

DR. KROLL:  Stan Reynolds?

MR. REYNOLDS:  Stan Reynolds.  And I pretty much agree with the rest of the panel that basically I don't think you need to do pharmacokinetics.  Again, there's the problem with how would you do it if you wanted to unless you could compare it to morphine or something else.

Basically, what I need to know and what people using this need to know is if an addict took heroin yesterday, can I pick it up today?  It's that simple.  Or can I pick it up tomorrow.  What's the time frame from the time of use that you can pick it up?  I think that's what we all need to know.  It goes back to the sample collection.  When I can be sure that the sample from a user actually is going to become a positive?  Is six hours too soon?  Is eight hours enough?  I think that's the basic information that you want.

DR. HENDERSON:  That's in the submission about it depends on the rate of hair growth.  It's in the submission, and it depends on the rate of hair growth, that I think it was 0.6 millimeters to, I think, four millimeters was the length that I read, and from 14 days to 90 days.  So, that's in there.  And the reproducibility I think is in there where they stored it for a year and then tested it again, and that was similar.

DR. LASKY:  Fred Lasky.  This is not necessary for the submission, in my opinion.

DR. ROSENBLOOM:  Arlan Rosenbloom.  I would like to know -- I think it would be interesting to know if a single-dose exposure -- say a kid goes to a rave and two months later if it's in his hair or her hair.  I think that would be interesting.  I'm not sure how important it is for workplace testing and all the other applications.

It would seem to me that -- and this is not a pharmacokinetic issue as much as how much exposure you need to get a positive test -- I would think that the differences in usage would be worked out by the MRO who's reviewing this data with the user and if they say, "Well, I'm not using it anymore" or "That happened six months ago," is there a possibility of looking at the hair segmentally to get them by the short hairs, as it were, and determine whether they've had drug use in the last month or so?  Perhaps the Company can comment on that.

DR. CAIRNS:  I regretted, perhaps, to educate you to the science of hair testing.  Yes, hair grows half an inch per month.  And as regards the ingestion to the hair appearing at the surface of the scalp for cutting, that would take between five and seven days to grow from base of follicle to surface of skin for cutting.  So, your question is that if someone took heroin today, yes, it would take seven or perhaps a few days longer for that section of the hair that the heroin is incorporated in to be cut from the scalp and tested.

For the issue of segmental analysis, that is correct, you can in fact look at, say, that first half-inch of growth from the scalp which will detect the drug use within that 30-day approximate time window.  The minimum detectable dose, we've outlined in the submission as a mean detectable dose of 173 milligrams per month.  And that's an infinitesimally small amount compared with, say, a chronic user up at 800 milligrams per month.

DR. ROSENBLOOM:  These doses mean nothing to me.  That's not a --

DR. CAIRNS:  But you're correct, it can be detected in a shorter length of hair.

DR. ROSENBLOOM:  Does a dose like that, as a single dose, I mean does that make sense as a single dose, number one?  Number two, if it does make sense as a single dose, then it would appear, I presume, in a segment.

DR. CAIRNS:  Well, a single dose, Dr. Rosenbloom, may well be only somewhere of the order of 16 or 20 milligrams, if you're talking street dose.  In that case, the individual may have to consume several little doses to go just above the cutoff.

DR. ROSENBLOOM:  So, the problem of picking up a single dose is you likely wouldn't.

DR. CAIRNS:  Yes.  We stress the fact that the reason for not doing the single-dose study was an analytical problem, that a single dose would not challenge the assay or the cutoff.

DR. ROSENBLOOM:  Okay.

DR. HENDERSON:  I have nothing to add.  Cassandra Henderson.

DR. CLEMENT:  I have nothing to add.

DR. KROLL:  Okay.  Thank you.  Let's go to question 5.

DR. PEACOCK:  Question 5:  Should the potential for bias by race, age, sex, hair color or other individual differences in the incorporation or retention of drug in the hair be evaluated?  If yes, what additional studies should be requested?

DR. KROLL:  Let's start with Dr. Kurt.  Then we'll move counterclockwise.

DR. KURT:  I think that it should to avoid discrimination problems under the Americans with Disabilities Act and other factors in the federal legislation.  And I think that a point could be made in the hair testing that has been presented in the larger-n groups that there was no division by sex.  As we know in the male population, most of us don't do anything to lighten our hair except when it gets gray.  Women tend to lighten their hair, which, in general, I think probably removes some of their melanin, so that population might be somewhat skewed if included in the population as a whole.  So, I think further studies need to be done in this regard.

DR. KROLL:  I tend to agree with Dr. Kurt.  I don't think they have to be tremendously thorough studies, but there are a lot of women who dye their hair or change their hair color or do other things to it.

And as with respect to hair color, maybe do enough studies or put together the studies in such a way so that things related to the amount of melanin that's in the hair since that seems to be the main culprit.  But even though I think that you probably have some evidence that the melanin's removed, but to make certain that that's very clearly stated what's going on and how that works.

DR. MANNO:  I agree with Dr. Kroll.  Nothing else to say.

DR. LEWIS:  Sherwood Lewis.  I have nothing to add.

DR. CLEMENT:  Steve Clement.  I think some of these should be done but possibly as post-marketing.  From looking at it as best I can it looks like, if anything, any of these variables would, if anything, decrease the sensitivity of the assay, so the bias would be against picking it up instead of for picking it up.

So, from that standpoint, if you still get a positive test, the person took heroin, period.  There's no questions about it.  It's just a matter of whether it picked me up or it picked up someone else.  But if it is positive, I didn't see anything in the materials presented that showed that there would be a bias on causing a false positive, which is the biggest concern I had.

DR. HENDERSON:  I think there's no question that studies need to be done to assess the influence of any of these demographic factors.  And my concern is that it would discriminate against populations where perhaps the test was more sensitive; that is, that it's not that they didn't take it, it's just that their friends who did take it may be at an advantage of being able to treat their hair, wash their hair, do things, because their level was lower than when they started than other populations.  And, so I think there's no question it has to be studied.

I agree with Dr. Clement, it certainly could be post-marketing, perhaps, but definitely that data needs to be gathered.

To Mr. Reynolds' suggestion a couple of questions ago, there are populations that could easily be accessed to collect populations to look at these issues.

DR. ROSENBLOOM:  I have nothing to add.

DR. LASKY:  Fred Lasky.  I agree that these are important studies that have sociological impact.  I believe that the sponsor has shown, at least with reasonable assurance, that this is not a major issue, but I don't think that the specific questions which are important here have been addressed to the specifics.  I would think that these studies would be done because of liability issues, and that would probably be the primary incentive.  But I think it would be an important thing to do post-market.

MR. REYNOLDS:  I pretty much agree with what everyone else has said, that the dyeing and treating and things of this nature seem to, if anything, decrease the sensitivity, and you're not going to get false positives as a result.  The demographic data should be collected.  Once again, it can be collected post-market.  This can be done fairly simply, and I don't think it would be too burdensome for the manufacturer to do that.

DR. EVERETT:  James Everett.  Certainly these items are necessary to be done, if for nothing other than to help prevent fraud and abuse.  This is a test that stands at high potential for fraud and abuse as well as general accusations about who's used a drug, who hasn't used a drug.

Without doing these tests you will not know who the test is not suited for.  If the test is suitable for young kids, the only way you're going to know is you have to check.  If it's suitable or not suitable for elderly people because their chemistry changes quite a bit, and if you don't do the test, then you won't know.  And these are just the fundamentals of all tests that should be done.

DR. WILKINS:  Diana Wilkins again.  I agree with the comments of Dr. Everett just a moment ago.  I think that the studies do need to be done.  And the thing that I'm focusing in on, I think, is the individual differences in retention of drug in hair.  And I think -- because I'm thinking of this as a potential application in a clinical setting.

The average hair growth rate, at least of what I've seen reported in the literature, is a huge range.  And we've talked about earlier with the problems of using a mean data, a mean data point or a mean value.  That can be problematic at times.  I've seen anywhere in the literature from 0.6 to 1.3 centimeters per month.  I don't know which number is valid, to tell you the truth, because it's from a wide range of studies.  So, hair growth rate is important.

So, my question then is, well, if I'm going to rely on this to determine whether someone has stopped using heroin in a clinic or something like that, when am I going to be able to determine this?  I would think hair growth rate would be somewhat important in that issue.  When will a formally positive user become negative?  Or if I keep testing them for six months, are they going to be positive for six months whether they continue to use or not?  I can't answer that question.  I'm not saying that this test isn't sensitive enough to do that, but I don't know that I have the data to convince me that we can answer that type of question.

DR. KROLL:  Okay.  Let's go to question 6.

DR. PEACOCK:  Question 6:  Is the information provided by the sponsor adequate to address the issue of retention of drug in the hair from environmental exposure?  If not, what additional information should be requested?

DR. KROLL:  Let's start with Dr. Manno and move clockwise.

DR. MANNO:  Barbara Manno.  May I ask the sponsor a question --

DR. KROLL:  Sure.

DR. MANNO:  -- before I answer this?  In your procedure, in processing the hair prior to the -- during the wash process, who determines or what determines what you're calling the short process and the long process?  That wasn't exactly clear to me.

MR. IRVING:  Right now, our standard process is for all the samples to go through the entire wash procedure.  We don't do a retest of the sample.  We do an initial screening test.  We do the extensive wash, digest our sample, do a confirmation, and apply the wash kinetics.  And that's across the board for all our samples right now.

DR. MANNO:  Then why do you have -- well, never mind.

MR. IRVING:  Part of the reason that's in there is because some of the initial studies had some of that data in there also.

DR. MANNO:  Oh, okay.  But you are proposing it with that long process.

MR. IRVING:  We're proposing that the single immunoassay followed by the extensive wash and the confirmation.

DR. MANNO:  Okay, thank you.

DR. KROLL:  Can you just give us your name, so we can --

MR. IRVING:  Oh, John Irving, I'm sorry.

DR. MANNO:  Barbara Manno again.  I think that I'm satisfied with that wash procedure that they've presented.

DR. KROLL:  Martin Kroll.  I'm basically satisfied.  I would like to see when they do the analysis, the short analysis, against what they consider typical drug users, it would be nice to manipulate the data and show it against somebody who's right near the cutoff and what the effect is.  I mean, again, when you're working with real problems, it's not the typical user.  Where you get in trouble is somebody who's right close to a cutoff or whatever, and that presents issues and problems.

DR. KURT:  Tom Kurt.  I agree with both Dr. Manno and Dr. Kroll, and yes.

DR. WILKINS:  Dr. Wilkins.  I'm going to say in general yes with one exception.  The only thing, and I'm not necessarily -- I'm just saying this as a possibility to be considered -- is that the effectiveness of the wash procedure and retention of drug in hair, the data looks very good what presented for -- the data that was in the submission looked very convincing.

But from my perspective is that most users or heroin users it was soaked drug, and most heroin users are not soaking their hair in heroin solutions.  And that's always been brought up as a limitation in the past for doing these type of studies is that that's not a really good -- it's not a realistic model for assessing environmental exposure.

Having said that, I think, if anything, it probably sort of pushes it to the one extreme where you would probably, if anything, be getting a lot of drug in there and challenging your system.  But in reality and practice, I think people are being exposed to smoke and exposed to powder on hair, and I think that's a very different -- I think it's a different route or a different means of environmental contamination that the wash procedure would need to be addressed.

So, while the data that is submitted here using the soaked procedure does suggest to me that the washing procedure is quite effective, it would have been significantly enhanced in my mind and very convincing had there also been some data with hair that had been exposed to smoke or perhaps powder or what have you.  That would have really clinched it for me.

DR. EVERETT:  James Everett.  I particularly like the washing data, and that's kind of rare.  But in reality, it does demonstrate that you can remove a consistent amount of drug from the hair before the actual specific test is performed at the end.  And I think, again, it's very difficult to go through all of the possible environmental exposures that a person's hair could go through, particularly one who's using heroin.  So, I'm satisfied with the data that they've already presented.

MR. REYNOLDS:  Stan Reynolds.  In addition to the experimental soaking, they also had the group of police officer, undercover police officers, who did have exposure to the environment in the testing that they did on them.

DR. WILKINS:  That was a different drug.  Wasn't that cocaine?  That was a different drug.  That's why -- I think that last study that the presenter showed this morning --

MR. THISTLE:  Yes, Bill Thistle.  That study with the undercover narcotics officers was with cocaine, because that, in the literature, has always been the drug of issue.  There really has been no literature expressing the fact that heroin is floating around or that people are soaking their heads in heroin.  We did those as extreme contamination scenarios.

The real contamination scenario that has been brought has been brought with regards to cocaine, and that's why -- in fact, the articles that you've gotten about removal external contamination, both from the FDA and from Psychemedics, deal primarily with cocaine.  The exception is the study that we did specifically for this assay where we soaked the hair and where we did the sweat contamination experiment.

DR. KROLL:  Let's move on.  Anymore comments?

MR. REYNOLDS:  But just to finish up that point, in my limited experience in this area, if you're talking about environmental exposure, people who tend to spend a lot of time in shooting galleries, generally are not there as observers.  So, I think exterior contamination of hair is really not going to be a major issue.

DR. LASKY:  Fred Lasky.  I thought the data, as presented, were convincing, but there were a couple of things that I think are worth looking at with the data that's probably in the submission.  And that is we saw the washing steps, and it was based on average data, based on my understanding of what I saw.

And I was intrigued by this extrapolation technique, which I thought was also very convincing.  But the question that I would have is whether or not that extraction technique -- sorry, extrapolation technique compensates for the wash-to-wash variability that might be seen from sample to sample.  I think the data probably contained in the submission.

And my suggestion would be that if it doesn't, that perhaps they extend the extrapolation technique out a little bit further if that is how to deal with that problem.  But based on what I saw and the way it was presented, I thought the data were convincing.

DR. ROSENBLOOM:  I agree.

DR. HENDERSON:  I'd also agree.

DR. CLEMENT:  I think the study was adequate.  We hear all kinds of stories of people saying they use something they're caught with for various reasons.  I remember in Newsweek, they said after the Olympics in Australia, the coaches found, what, 20 vials of growth hormone and they asked him what he was doing with it, and he said he uses it for his bald spot on his head, and he rubs it in his hair.  So, I guess you never can be completely immune from what people will creatively come up with.  So, I think it was a good study, and it proved a point, and there's no other further studies that need to be done.

DR. KROLL:  Dr. Lewis.

DR. LEWIS:  I'm sorry.  I have nothing to add.

DR. KROLL:  Okay.  Let's go to question 7.

DR. PEACOCK:  Question 7:  Has the sponsor adequately demonstrated the effect of various washing or hair treatment procedures on the internally incorporated or bound drug?  If not, what additional studies should be requested?

DR. KROLL:  Okay, I'd like to start with Dr. Henderson, and then we'll go counterclockwise.

DR. HENDERSON:  I was very impressed with the washing.  And as a clinician, I think that I would be hard pressed to argue that there's an environmental risk to claiming someone's positive.

DR. ROSENBLOOM:  Yes, I think that's probably true.  There may be hair treatment modalities of which we are unaware, and I would expect some post-marketing further studies on whether one might be able to get a wash-out from very badly damaged hair.  But I wonder if that kind of thing couldn't even be done in the laboratory without having to do it from a person's head.

DR. LASKY:  Fred Lasky.  I agree with Dr. Rosenbloom's comment that if this does become an issue, and I'm not convinced it will, but if it does, it could be easily handled in a post-market situation.

MR. REYNOLDS:  Stan Reynolds.  I have nothing further to add.

DR. EVERETT:  James Everett.  I agree.  That's probably one of the better parts of the entire study, I thought.

DR. WILKINS:  I agree.

DR. KURT:  Tom Kurt.  I agree, in general.  I would like to point out an issue that this is a kind of an ongoing process that makes this washing technique unique.  The RIA technique is not necessarily unique, because it's been used in the drug testing industry, the laboratory testing industry for decades at this point.  So, that's not the unique part of this process that's being reviewed.

I would like to point out as a footnote one of my concerns that I didn't mention this morning is the sodium azide reagent, which is 20 percent sodium azide, which is metabolized as cyanide in the body.  It's quite a dangerous substance, and I've seen it used in our institution once by a pharmacologists on the faculty for suicide.  And familiar with reports of it in the literature.  It cannot be treated with a cyanide antidote kit, so it's a very dangerous substance and should be appropriately labeled.

DR. KROLL:  Martin Kroll.  I have no other additional comments.

DR. MANNO:  Barbara Manno.  I have no additional comments, but I have an additional question.  I just happened to notice, as I was looking at the presentation of the wash on the -- it was question 6 in your presentation.  I was looking at the numbers for your deriving the 28.4 milligrams.  It's page -- I can't make out which page it is.  It looks like page 11.  As you've derived this final answer of 28.4 nanograms per ten milligrams of hair, I noticed that the test reports out on ten milligrams of hair, but you're weighing in eight.  Do you do a massaging of the data there?

DR. CAIRNS:  No.  Again --

DR. KROLL:  Please identify yourself.

DR. CAIRNS:  Dr. Cairns again.  The original screen is eight milligrams.  Then if it is screened positive, it would move forward for washing and then digesting and confirmation.  We do that on a weight basis.

DR. MANNO:  So, how do you report it then?

DR. CAIRNS:  Per ten milligrams.

DR. MANNO:  But you started with eight milligrams of hair in the analysis.

DR. CAIRNS:  No.  The first screen is eight.  That's gone.

DR. MANNO:  Okay.

DR. CAIRNS:  Then we go back to the envelope and weight out a new amount, and then it's prorated per weight.

DR. MANNO:  Gotcha.  Thank you.  That wasn't clear to me.  Thank you.

DR. KROLL:  Any comments?  Dr. Lewis?

DR. LEWIS:  Yes, to answer the question, I think that there has been an adequate demonstration of the effect of the washings.

DR. CLEMENT:  I agree.

DR. KROLL:  Okay, good.

Now, let me ask the FDA, do they think that we've commented sufficiently on these questions?

DR. GUTMAN:  Yes.

DR. KROLL:  Okay.  Good.

DR. CAIRNS:  Mr. Chairman, may I have your indulgence just for one moment?

DR. KROLL:  Well, right before that, I just want to ask if anybody else in the panel had any other short comments they wanted to make, and then we'll let you speak.

DR. WILKINS:  Comments or questions?

DR. KROLL:  Comments or questions but please keep them brief.

DR. WILKINS:  A short one?

DR. KROLL:  Yes.

DR. WILKINS:  Just another clarification, just so I understand.  On the wash procedure, I'm just a little confused.  Again, this is all for me determining a positive or a negative and how that's occurring.  Is the results of the final wash buffer subtracted from the quantitative mass spec result to determine a value that is then compared to your cutoff?  And if so, is the wash analyzed by the semi-quantitative RIA or by the MS?  I'm just not clear on it, and it doesn't really matter.  I just wanted to make sure I understood.

MR. IRVING:  When we do -- this is John Irving again -- when we do the wash, we do it by radioimmunoassay; however, we do a full curve on that wash, and we do multiply that final value times five, subtracted from the digest, and compare it against the cutoff.

DR. KROLL:  Okay.  Any other additional comments?

You wanted to make a few additional statements?

DR. CAIRNS:  It's very short.

DR. KROLL:  At three o'clock, we do need to do the open public forum

DR. CAIRNS:  It's Dr. Cairns again.  For the record, I just feel it necessary to address a few points made during the discussion of the panel.

First of all, as regards additional information, such as regarding bias, et cetera, that's holding hair to a different standard than urine has been held traditionally.

And as regards to the demographics, I was sorry you had left the room when we addressed the issue, but it is in the attachment 21, 22, 23, as regards the individual studies A through E.  There is demographic information contained in there.

And as regards the clinical standards for exclusion/inclusion, I repeat that Psychemedics just did not use that two parameter.  We also used the fact that there was a negative urine and a negative hair test in some of those inclusion/exclusion criteria.  So, we went beyond the standard as presented in the questions.  Thank you.

DR. KROLL:  All right.  Thank you, everybody, very much.

We need to do the open public hearing at three o'clock, and now it's about five minutes to three.  I suggest we forego the break and do the open public hearing now.

Okay, so we're going to hear from Rosemary Mumm.

MR. MUMM:  Yes, thank you for inviting me and letting myself be invited.

I'm presenting the practioners' point of view for hair testing, and we've had some experience, obviously, with heroin.  I'd like to also, in general, though, address my comments to technology.

I see there's a note here that if we have any financial obligations or indebtedness to Psychemedics that they need to be stated.  We do not.  As the DA's office, we in fact have a high accounts payable with them, and we feel it's well worth it.

I'm the Director of the Diversionary Programs, and I have been for eight years with the New Orleans District Attorney's Office.  The Diversion Program is an alternative to prosecution program for new offenders.  These are both juvenile and adult folks who've been arrested on both narcotics and non-narcotics charges.  They're all non-violent offenders.

I've been using hair testing for eight years to monitor the people in our program to determine whether they need to be in a counseling track for substance abuse or a counseling track for some other mental health or clinical issue.

We assess them, if they are in a drug-free track, to become -- in a drug counseling track, to become drug-free.  That is our intent and our motive.  I'm not trying to find reasons to put people in prison.  I'm trying to find reasons to assist them to deal with their drug problem so that they don't repeat their criminal offenses.  Now, we are very effective, by the way, in reducing criminal recidivism.

Our population -- I know that's been a big issue today -- ranges from the ages of 12 to 66.  We are primarily an African-American city by about 65 percent, and the population in our program reflects that demographics data for the city.  Primarily two-thirds of the people are male, and complimentary, a third are female.

My background has been 20 years in the field of substance abuse treatment.  In the last ten years, that has been the context of criminal justice.  Someone asked what the BCSAC is behind my name, and that's Louisiana credential, a Board-Certified Substance Abuse Counselor.

In addition to the Diversionary Program role that I have, I've been assisting with my boss, Harry Connick, the District Attorney, in helping to implement high school drug testing programs, and I'll talk a little bit about that.

We began our program eight years ago through funding from the Department of Justice, specifically to look at hair testing to monitor offenders in a supervisory program.  Was it effective?  Did it work?  How did we feel about it?  We've used Psychemedics technology during that whole time.

I believe that the program is so -- the technology, hair testing, is so vital in our success that should I go to another program and continue in substance abuse treatment, I cannot imagine myself doing treatment anymore without hair testing.  For us, a piece of hard evidence or as hard of evidence as one can get in an imperfect world about whether someone is using drugs or not.

We also used urinalysis.  We want the benefits of both technologies -- the longer-term window of detection as well as determining whether someone has recently used drugs.  We use hair testing when someone is first assessed when they get into our program to determine, as I said before, what kind of counseling track they go into.

We have a lot of people who are arrested on non-narcotics charges who deny drug use.  And perhaps about half of those, when we administer a hair test, come back with a positive result.  Now, had we not done that and only used urine, I would predict that that would be much lower, because we have scheduled appointments with people.  They know when they're coming in, and they would know then that if they were to be drug tested that many of those could refrain from use for a day or two.

Again, we also use the hair testing to monitor their three to 15 months in our program.  It allows us to decrease the number of urine tests that we do and thus the cost of urine testing.  We really are monitoring whether someone is remaining drug-free as a part of their treatment program.  We would have to be urine testing maybe three times a week, which would get enormously expensive for us.

With the kind of 24-hour tape recording that the hair offers for us, we can reduce that frequency and still periodically do a hair test to see if they're remaining drug-free.  But I wouldn't abandon urine testing.  It's really important to know whether someone has used today or yesterday or some recent -- had some recent use.

In addition, we like the advantages of the collectionees.  Certainly, my staff dislike collecting urine samples.  They don't mind doing hair sample collections.  We like the retest capabilities when a result is challenged.  And again for other reasons, observing urine collections, the frequency of being able to do a hair test rather than the high volume manual labor-wise of doing random urines is greatly beneficial to our program.

We have a very high level of confidence in its detection abilities.  Our staff rates heroin detection in hair as excellent, in light of the self-report of the person being tested, their arrest history, and their urinalysis results.  Fortunately, in our 12 to 16 age population, which is our juvenile program, we have had no self-reported heroine users nor any positive hair tests, nor positive urine tests, I might add.  In my experience, though, beginning with some of our arrestees at age 17, we've seen some very, very serious heroin users who apparently just hadn't been arrested before their 17th birthday.

For us, and this is not related to heroin use as this hearing has directed, but the recent addition of MDMA in the panel has been very welcome because of the expensive urine testing for MDMA, which we would not be able to afford otherwise.  And in New Orleans there's been quite a number of young people who have either died or have had medical emergencies, severe medical emergencies for which we now feel we have a tool to assist with that intervention.

As far as the high school drug testing goes in New Orleans, there are now ten schools in New Orleans that have adopted drug testing programs using hair.  A number of other principals, because of the success of these schools, have stepped forward and are looking for money so that they can adopt similar programs.

The principals have said, those who've used hair testing in their schools, that they've seen disciplinary and behavioral problems drop.  The school milieu has changed, the student attitudes have changed, and they feel that this has been a real positive addition to what's happening in the schools.

Last week, one of the principals who's been doing testing for over three years now, reported that they had tried urine testing.  They knew they had a big drug problem in their school -- I don't think she'd mind if I mentioned the school, called De La Salle High School, known in the community as De La Drug High School prior to beginning hair testing -- and found that urine testing was not effective in turning their school around with their drug use problems.

They gave a 90-day warning to their students and the parents.  This is a private school, so they had the constitutional right to test all their students.  The first year they had a 3.4 percent positive rate.  The second year went to 2.1.  And year three was less than one percent.  I think their population is about 850 students.

They feel it has a very strong deterrent effect in not only discouraging kids from using drugs but also in identifying those few kids who did have a positive detection to get them into counseling and treatment.

All these high school programs are intended as student assistant programs and not punitive programs.

And last week, finally, my boss and I testified before the -- well, spoke before the mayor's Subcommittee on Violence in the Schools.  We had a shooting, like many other high schools in the country.  As a result of that, this Subcommittee was formed.  And the Subcommittee agreed that it would propose to the mayor and city council to seek advocacy from the top administration to look to the public schools to adopt more testing in the programs and endorsing hair testing as the method.

All I can say in summary is that as a clinician and someone whose motivation is to get people into the proper services to turn their lives around, I can't say enough about hair testing.  It's not a perfect technology.  Certainly, we've had people who have challenged the results in our program, people who care not to admit that they're using but when confronted with results 90 percent of the time or higher admit to doing that.

We have power over the situation, unlike a pregnant mother who would want to come forward with acknowledgement of drug use as a step and where the DA's office -- these are people who are not motivated to reveal their drug use, knowing full well that if they did it may jeopardize their court case and potentially land them in jail.

That's all I have.

DR. KROLL:  All right.  Thank you.

All right.  What did you want to do now?  Do you think it's pertinent, Dr. Gutman, to ask her questions from the panel?

DR. GUTMAN:  Certainly.  There's no reason you couldn't.

DR. KROLL:  Okay.

DR. LEWIS:  Sherwood Lewis.  I had one quick question for you.  And I'd like to know what the rationale and the ramifications for doing drug testing in those individuals who are arrested or apprehended for non-drug related offenses?

MR. MUMM:  Well, it's my belief that if someone is arrested on a theft charge, for example, there may in fact be an underlying substance abuse disorder.  We know that people who are needing money for drugs resort to thefts and other criminal activities to get money for their drugs.

We want to screen -- when we put someone in an appropriate counseling program, the priority is first to get them drug-free if they are abusing drugs.  Then if there are other family issues, mental health issues, that can be worked in accordance with the drug treatment program.  But unless we serve to help that person get off of drugs, I personally would expect that that criminal activity would continue.

DR. LEWIS:  I guess my question goes to the point of whether an individual voluntarily subjects himself or herself to that --

MR. MUMM:  Yes, it's a voluntary program, yes.

DR. LEWIS:  -- testing, even though they've been arrested for something not connected with the crime.

MR. MUMM:  Right.

DR. LEWIS:  So, the volunteer to --

MR. MUMM:  Yes.

DR. LEWIS:  -- do this.

MR. MUMM:  Yes, they can enter the program or they can go to court and face the charges there.  So, it has some coercive elements to it.

DR. LEWIS:  They enter the program, but they enter the program subsequent to their having been tested positive.

MR. MUMM:  No, they enter the program whether they're using drugs or not, if they choose to, do so.  Once they sign in and we take a hair test, then that determines what type of program we would ask them to participate in.

DR. LEWIS:  So, they sign on to the program, then you folks do the testing.

MR. MUMM:  Right, and we tell them, though, when they sign in they know that they will be hair tested.

DR. LEWIS:  Thank you.

DR. KURT:  Could I ask a quick question?  Is there a consent form that they sign when they give the specimen or is it an order from the court that they give a specimen?

MR. MUMM:  No, it's a consent form.  There's no court order involved.

DR. KROLL:  All right.  Does anybody else from the public wish to speak?  Okay.

Dr. Gutman, do you have any questions or concerns or any comments you want to make?

DR. GUTMAN:  No, I don't.

DR. KROLL:  All right.  I was talking to our executive secretary, and she said, well, I could summarize what our comments were.  I think it's rather difficult to summarize them, because for many of the questions I think you heard divergent opinions.  And some of them you heard a lot of agreement.  But I'd certainly open it up to the panel if anybody wants to make a closing comment or something else pertinent to the questions that they have to ask.

DR. HENDERSON:  I think in summary we all are excited about the technology and would certainly urge that it be marketed soon, although we do have reservations and perhaps many of those, if not all of those, could be addressed in post-market study.

DR. KURT:  As a medical review officer, I'm certainly looking forward to hair testing that can be performed under direct observation rather than having urine testing performed, for it's so easy to cheat.  However, I think that the problematic areas that we pointed out today should be remedied, and some of it can be done on a post-marketing basis.

DR. KROLL:  Anybody else have any other comments?

MS. CALVIN:  I just wanted to make some closing remarks.  Thank you, Psychemedics, FDA staff, and the panel, of course, for all of your recommendations.  And the next tentative meeting for the Clinical Chemistry and Clinical Toxicology -- how could I forget Toxicology -- Devices Panel will January 17, 2001.

DR. KROLL:  I'd like to also thank the members of the panel, the FDA staff, both from Psychemedics and other people who have made comments for today.

Did you want to make any closing comments, Dr. Gutman?

DR. GUTMAN:  No.  I'll add my thanks to yours.

DR. KROLL:  Okay.  So, we now can adjourn the meeting.

(Whereupon, the FDA Meeting was concluded at 3:18 p.m.)