Atmospheric Mercury Monitoring Services

Statement of Work

Title:  	Monitoring services to operate Tekran automated mercury speciation units and related equipment

Summary

In March 2005, EPA promulgated the Clean Air Mercury Rule (CAMR) to reduce mercury emissions from power plants by 2010.  In conjunction with the Clean Air Interstate Rule (CAIR), mercury emission reductions under CAMR are expected to reduce atmospheric concentrations and deposition of mercury.  Cutting mercury emissions would also translate to reductions in methylmercury contamination in fish, particularly in mercury-sensitive watersheds throughout many parts in the U.S.  This would lead to improved health of people and wildlife that consume fish from these waters.  These reductions will be achieved through implementation of a combination of independently operated state programs plus an interstate emissions cap and trade program modeled after the Acid Rain SO2 Allowance Trading Program.

In order to assess the efficacy of these emerging regulatory programs, scientific information is needed, including a more complete understanding of the fate of mercury emissions with respect to total (i.e., wet and dry) deposition, a quantifiable assessment of the sources (especially coal-fired power plant emissions) contributing to mercury deposition, and an assessment as to whether or not mercury “hotspots” exist or may develop over the course of implementing the mercury rules.  At present, EPA lacks the mercury ambient concentration and deposition data, particularly on dry deposition in source impacted areas, to adequately assess the atmospheric mercury changes anticipated from the regulatory programs.

To meet some of these data needs, EPA is collaborating with the National Atmospheric Deposition Program (NADP) membership of Federal agencies, states, tribes, academic institutions, industry, and other organizations to establish a new, coordinated network for monitoring atmospheric mercury species.  EPA, working in partnership with NADP, is launching this network by initially leveraging existing sites measuring concentrations of atmospheric mercury to form an embryonic network with data from these sites collected and managed in a central archive, network-wide operating protocols for comparison of data across sites, and a centralized quality assurance program.  EPA envisions using up to 10 existing sites that are currently making speciated atmospheric mercury measurements to participate initially in the network. The objective is to harmonize and assure the quality of the mercury measurements in a standardized and coordinated manner that is transparent, involves many partners, and is scientifically credible.

Background

US EPA’s Office of Air and Radiation (OAR) is collaborating with the National Atmospheric Deposition Program (NADP) membership to establish a new, national, coordinated network to monitor mercury in the atmosphere.  The network will measure air concentrations of mercury in its gaseous and particulate forms, event-based mercury wet deposition, and meteorological and land-cover variables needed for estimating dry deposition fluxes.  When implemented, the network will serve many functions:

- facilitate the calculation of wet, dry, and total deposition;

- provide data for evaluating predictive and diagnostic models and for assessing source-receptor relationships; and

- build a data set for analyzing spatial and temporal trends.

The network will consist of monitoring stations with a broad range of classifications including: rural, suburban, and urban; near-source/high-emissions; sensitive ecosystems; and regionally representative.  Stations will adopt standard operational procedures based on methods developed from EPA and other research efforts.  Data shall be quality-assured and accessible online.  The initial network will consist of existing sites (i.e., places where Tekran speciation systems are currently operating, with an accessible measurement record of atmospheric, speciated mercury concentrations).  When quality assurance programs are established for the new network, including standard protocols for automated methods, EPA will request the Contractor to participate in testing protocols, and ultimately adopt these programs.  These programs are currently under development and will necessitate continued input and buy-in from the broader mercury measurement community.

The purpose of this project is for a Contractor(s) to provide monitoring services for the operation of automated Tekran speciation systems.  The objective is to acquire a measurement record of atmospheric mercury concentrations in elemental, reactive-gaseous, and particulate forms from different sites across the U.S.  OAR requires an accessible database of long-term mercury measurements for temporally assessing the impacts of mercury emitting sources affected by federal and state air rules (e.g., CAIR, CAMR) and to support evaluation and improvement atmospheric mercury models that provide estimates of dry and total mercury deposition.

Tasks

TASK 1: 	Quality Assurance and Standard Operating Procedures

The Contractor shall be responsible for a quality assurance plan (QAP) and associated documentation for all aspects of data collection performed under this contract.  The Contractor shall submit the QAP to EPA for approval as required by EPA Order 5360.1 A2, Policy and Program Requirements for the Mandatory Agency-Wide Quality System.  The Contractor may either revise existing quality assurance documents as necessary or prepare new documents to meet quality system requirements.  The Contractor shall provide EPA with detailed standard operating procedures (SOPs) for all aspects of the contract. The SOPs shall describe the various aspects of the instrumentation, the training of site operator(s), operational maintenance schedules, calibration procedures, instrument specifications (i.e., consumable requirements), repair/replacement of parts, and chain-of-custody protocols for any modifications and service performed on the instrument systems, data handling, data reduction and validation.  The Contractor shall also prepare a monthly narrative focusing on deviations from the QAP and SOPs. Other things to report, when appropriate, include significant activities/events, problems detected, replacement of major parts pertaining to the project.  The reports shall include a breakdown of repair/replacement problems (i.e., down time of the instruments), and the number of visits to the sites by the Contractor or support organization.

TASK 2: 	Site Characterization

The Contractor shall prepare a report that describes the existing monitoring site location, including the following information:

- Latitude, Longitude (DMS or decimal degrees)

- Elevation (meters)

- Site sketch (scale drawing of site using NADP template)

- Digital photos of automated speciated mercury system location in 8 directions (N, NE, E, SE, S, SW, W, NW).

- Standardized land use/land cover classification in a circular area of 0.5-km radius (~0.3 miles) around the site. Indicate the fraction of surface covered by forest, desert, uncultivated grassland, cultivated crops, grazed pasture, swamp or wetland, open water, residences, industries, commercial activities, farm buildings, etc.

- Leaf-area index (LAI) quantified directly and/or indirectly and information used to determine LAI values within a .5-km radius (e.g., instruments, methods, remote sensing), if available

- Break down of vegetation species within a .5 km radius (e.g., sugar maple, black birch)

- All objects (vegetation, posts, instruments, etc.) greater than 1 m tall and their description, direction and distance from the automated speciation system within 50 meters

- Proximity of site from sources of emissions within 100 km (mercury and other)

TASK 3:	 Site Operations Support

The Contractor shall be responsible for the calibration, operation, and maintenance of automated mercury instrumentation and supporting equipment/materials.  This task shall include supporting the operation of instrumentation and associated equipment/materials/supplies including, but not limited to:

- Semi-continuous automated ambient mercury Tekran speciation units (2537a, 1130, 1135) Hg0 (5-minute average), RGM, (2-hours), HgP (2-hours)

- Meteorological equipment (hourly)

The Contractor shall deliver to EPA or make available by FTP all raw, original speciated Tekran output on a monthly schedule, as defined below.  The Contractor shall also deliver to EPA or make available by FTP validated Tekran speciated data    and validated continuous meteorology.  Validated reactive gaseous mercury and particulate mercury delivered on hourly aggregation.  Data shall be presented in ASCII CSV files with standardized, consistent file format, as specified in the QAP.  The Contractor shall provide to EPA the computer program code (e.g., SAS) or specific procedures used to determine the validity of the Tekran speciated datasets.

Payment for Tekran speciated data shall be performance-based, based on the percentage of valid data submitted to EPA for a given time period (monthly).  The Contractor shall receive full monthly payment if 70% or more data are valid for the time period.  Thirty percent of the data monthly are allowed to be invalid for the time period to receive full monthly payment.  Payment for valid Tekran data is described in the chart below.

| - 70% or greater valid | - = 100% monthly payment |
| --- | --- |
| - 50 - 69% valid | - = 80% monthly payment |
| - 49% - 11% valid | - = 50% monthly payment |
| - 10% or less valid | - = 20% monthly payment |

Schedule of Deliverables and Timeline

| Task | Deliverable | Due Date |
| --- | --- | --- |
| Task 1 | Quality assurance plan, including specific procedures or logic used to validate the semi-continuous automated Tekran speciated mercury data | Delivered electronically to EPA within 60 days of award of contract |
|  | Monthly narrative focusing on deviations from the QAP and SOPs. Other things to report, when appropriate, include significant activities/events, problems detected, replacement of major parts pertaining to the project. The reports shall include a breakdown of repair/replacement problems (i.e., down time of the instruments), and the number of visits to the sites by the Contractor or support organization. | Delivered electronically to EPA within 10 days after the end of the month |
| Task 2 | Site characterization report | Delivered to EPA electronically within 60 days of award of contract |
| Task 3 | Raw, original semi-continuous speciated (Tekran) data (elemental, reactive gaseous, particulate), including mercury readings, calibration, instrument set-up, and methods data. Data shall be presented in standardized, consistent file format (e.g, .CSV) as specified in the QAP | Delivered electronically to EPA or made available by FTP within 10 days after the end of the month |
|  | Validated data sets: semi-continuous (Tekran) data (elemental, reactive gaseous, particulate) and continuous meteorology Validated reactive gaseous mercury and particulate mercury delivered on hourly aggregation. Data shall be presented in ASCII CSV with standardized, consistent file format, as specified in the QAP | Delivered to EPA or made available by FTP quarterly, by 30 days after the end of the month |