UNITED STATES OF AMERICA

+ + + + +

FOOD AND DRUG ADMINISTRATION

+ + + + +

CENTER FOR DEVICES AND RADIOLOGICAL HEALTH

+ + + + +

MEDICAL DEVICES ADVISORY COMMITTEE

CIRCULATORY SYSTEM DEVICES PANEL

+ + + + +

OPEN SESSION

+ + + + +

Thursday

March 6, 2003

+ + + + +

The meeting came to order in Salons A, B, and C of the Gaithersburg Marriott, 9751 Washingtonian Boulevard., Gaithersburg, Maryland, at 10:30 a.m. Warren K. Laskey, M.D., Temporary Chair, presiding.

PRESENT:

WARREN K. LASKEY, M..D. 	Temporary Chair

SALIM AZIZ, M.D.	Voting Member

CYNTHIA TRACY, M.D.	Voting Member

MARK C. HAIGNEY, M.D.	Voting Member

CHRISTOPHER J. WHITE, M.D.	Voting Member

GEORGE W. VETROVEC, M.D.	Voting Member

RICHARD LEIGHTON PAGE, M.D.	Voting Member

ALBERT WALDO, M.D.	Voting Member

FRANCIS R. GILLIAM, III, M.D.	Voting Member

MITCHELL W. KRUCOFF	Voting Member

MERCEDES DULLUM, M.D.	Voting Member

KENT R. BAILEY, Ph.D.	Voting Member

GERETTA WOOD	Executive Secretary

MICHAEL MORTON	Industry Representative

ALLEN A. HUGHES, Ph.D.	Consumer Representative

BRIAN ZUCKERMAN	Director, Division of Cardiovascular Devices

FDA PARTICIPANTS:

JAMES CHENG	Lead Reviewer

LESLIE EWING, M.D.	Clinical and Animal Review

LILLY YUE, Ph.D.	Statistical Review

CINDY DEMIAN	Biocompatability Review

ELAINE MAYHALL	Sterilization Review

KEVIN HOPSON	Bioresearch Monitoring Review

SPONSORS:

JEAN-PIERRE DESMARAIS, M.D.	Vice President, Scientific Affairs, CryoCath Technologies, Inc.

JEREMY RUSKIN, M.D.	Massachusetts General Hospital

MARC DUBUC, M.D.	Montreal Heart Institute

PETER FRIEDMAN, M.D., Ph.D.	Harvard Medical School

MARWAN ABBOUD	Director of Engineering, CCT

PATRICK CHAUVET	Pre-clinical Scientist, CCT

SUSAN BONDY	Consultant, Cato Research

ANDREW SKRYLOV	Consultant, Cato Research

RICHARD HOLCOMB, Ph.D.	Consulting Statistician

JOHN LEHMANN, M.D.	Consultant, CCT Medical Director

JOSE NAZARI, M.D.	Illinois Masonic Hospital, Chicago

MARK NIEBAUER, M.D.	University of Nebraska, Omaha

DAVID KEANE, M.D.	MGH, Boston

C O N T E N T S

Call to Order	6

Warren K. Laskey, M.C.

Introduction of Members	9

Open Public Session	12

Sponsor Presentation:	13

CryoCath Technologies, Inc.,

P020045, 7F Freezer Cardiac

Cryoablation Catheter

Jean-Pierre Desmarais, M.D.	13, 24, 55

Jeremy Ruskin, M.D.	17

Marc Dubuc, M.D.	27

Peter Friedman, M.D.	31

FDA Presentation:	57

James Cheng	57

Leslie Ewing, M.D.	63

Lilly Yue, Ph.D.	70

Question-and-Answer Period	84

Open Committee Discussion	129

C O N T E N T S

Open Committee Discussion	188

Sponsor Comments	259

Industry and Consumer Representatives'	259

Comments

Open Public Session	262

Recommendations and Vote	262

P R O C E E D I N G S

10:41 a.m.

CHAIRMAN LASKEY:  Good morning.  I would like to call us to order.

The topic before us is a PMA for the CryoCath Technologies' 7 French Freezor Cardiac Cryoablation Catheter, P020045.

I would like to ask Ms. Wood to read the paragraph regarding the Chair, Committee, and conflict-of-interest statement.  Ms. Wood, please.

MS. WOOD:  Just a couple of reminders:  If you haven't signed in at the table outside, please do so.  Also, please turn your cell phones to silent or off during the meeting.  Thank you for your cooperation.

Dr. Cynthia Tracy has been excluded from chairing recent meetings as a result of the regulation governing covered appearance relationships.  Since it is expected that for most future Panel meetings Dr. Tracy's covered relationship will continue to preclude her from functioning fully as the Panel Chair, she has graciously stepped down as Chairperson, but will continue as a voting member of the Panel until the expiration of her term.  Dr. Warren Laskey will be the Temporary Voting Chair for this meeting.

The following announcement addresses conflict-of-interest issues associated with this meeting and is made part of the record to preclude even the appearance of an impropriety.  To determine if any conflict existed, the agency reviewed the submitted agenda for this meeting and all financial interests reported by the Committee participants.

The conflict-of-interest statutes prohibit special government employees from participating in matters that could affect their or their employer's financial interest.  The agency has determined, however, that the participation of certain members and consultants, the need for whose services outweighs the potential conflict-of-interest involved, is in the best interest of the government.  Therefore, waivers have been granted for Drs. George Vetrovec, Kent Bailey, Mercedes Dullum, and Albert Waldo for their interest in firms that could be affected by the Panel's recommendations.

Dr. Vetrovec's waiver involves stock in the parent of a competitor.  The stock is valued from $50,001 to $100,000.

Dr. Bailey's waiver involves a grant to his institution for the sponsor's product study in which he had no involvement and for which funding was between $100,001 and $300,000 per year.

Dr. Dullum's waiver involves stock in a competitor.  The stock is valued from $25,001 to $50.000.

Dr. Waldo's waiver involves consulting for a competitor's unrelated product, for which he receives an annual fee of less than $10,001.

Copies of these waivers may be obtained from the agency's Freedom of Information Office, Room 12A-15 of the Parklawn Building.

We would like to note for the record that the agency took into consideration other matters involving Drs. Vetrovec, Dullum, Mark Haigney, Mitchell Krucoff, and F. Roosevelt Gilliam.  Each of these panelists reported interests in firms at issue, but in matters that are not related to today's agenda.

The agency has determined, therefore, that they may participate fully in all discussions.  In the event that the discussions involve any other products or firms not already on the agenda for which an FDA participant has a financial interest, the participant should excuse him or herself from such involvement, and the exclusion will be noted for the record.

With respect to all other participants, we ask in the interest of fairness that all persons making statements or presentations disclose any current or previous financial involvement with any firm whose products they may wish to comment upon.

CHAIRMAN LASKEY:  Thanks, Geretta.

At this point I would like to have the Panel members introduce themselves, starting to my left, please.

DR. ZUCKERMAN:  Brian Zuckerman, FDA, Division Director, Cardiovascular Devices.

DR. GILLIAM:  Roosevelt Gilliam with the Virginia Cardiovascular Specialists in private practice of electrophysiology, Richmond, Virginia.

DR. PAGE:  Rick Page, Cardiac Electrophysiologist, now at the University of Washington in Seattle.

DR. WALDO:  I'm Al Waldo, Cardiac Electrophysiologist, Case Western Reserve University in Cleveland.

DR. BAILEY:  Kent Bailey, Biostatistician at Mayo Clinic.

MS. WOOD:  Geretta Wood, Exec. Sec. of the Circulatory System Devices Panel.

CHAIRMAN LASKEY:  Warren Laskey.  I'm an Interventional Cardiologist at the National Naval Medical Center.

DR. WHITE:  Chris White.  I'm an Interventional Cardiologist in the Ochsner Clinic Foundation in New Orleans, Louisiana.

DR. VETROVEC:  George Vetrovec, Chairman, Division of Cardiology, Virginia Commonwealth University.

DR. DULLUM:  Mercedes Dullum, Cardiac Surgeon, cryo practice, Washington Hospital Center.

DR. TRACY:  I'm Cindy Tracy.  I'm at Georgetown University Hospital, and I not so graciously did try to get my brother fired, but it didn't work.

(Laughter.)

DR. HAIGNEY:  I'm Mark Haigney.  I'm a Cardiac Electrophysiologist at the National Naval Medical Center and Director of Cardiology at Uniformed Services University.

DR. HUGHES:  Alan Hughes, Assistant Professor of Decision Sciences and MIS at George Mason University, and I'm the consumer representative.

MR. MORTON:  I'm Michael Morton.  I'm an employee of Sorin-COBE Cardiovascular, and I'm the industry representative to the Panel.

CHAIRMAN LASKEY:  Thank you, colleagues.

Geretta, if you could read the voting status statement, please?

MS. WOOD:  "Pursuant to the authority granted under the Medical Devices Advisory Committee charter, dated October the 27th, 1990, and as amended August the 19th, 1999, I appoint the following individuals as voting members of the Circulatory System Devices Panel for this meeting on March the 6th, 2003:

"Mark C. Haigney, M.D.; Christopher J. White, M.D.; George W. Vetrovec, M.D.; Kent R. Bailey, Ph.D.; Mitchell W. Krucoff, M.D.; Richard L. Page, M.D.; Albert L. Waldo, M.D.; Francis R. Gilliam, M.D., and Mercedes K. Dullum, M.D.

"For the record, these individuals are special government employees and are consultants to this Panel under the Medical Devices Advisory Committee.  They have undergone the customary conflict-of-interest review and have reviewed the material to be considered at this meeting.

"In addition, I appoint Warren K. Laskey, M.D., to act as Temporary Chairperson for the duration of this meeting."

And it is signed by David W. Feigal, Jr., M.D., MPH, Director, Center for Devices and Radiological Health, on February 26th, 2003.

CHAIRMAN LASKEY:  I would like to open the public hearing portion of this morning's program.  Is there anyone in the audience who wishes to address the Panel on today's topic or any reasonably-germane topic?

(No response.)

If not, then I will close the open public hearing session.

I would like to begin with the sponsor's presentation.

MS. WOOD:  I would also like to remind the speakers to introduce themselves and to state your conflict of interest before speaking.

DR. DESMARAIS:  Good morning, ladies and gentleman, Mr. Chairman, and Panel members.

My name is Jean-Pierre Desmarais.  I am CryoCath Vice President of Scientific Affairs.  It is my pleasure to present our technology to this Panel.  As well, we will present the results of our pivotal study.  Here is the agenda of the meeting:

I'm introducing the group.  Then we will present the evolution of technology, the details of the freezor device, a brief synopsis of the pre-clinical testing, the pivotal clinical study, and a few concluding remarks.

The presenters will be myself, Dr. Jeremy Ruskin, Director, Cardiac Arrhythmia Service, Massachusetts General Hospital.

Dr. Marc Dubuc, Co-Principal Investigator, Chief of Electrophysiology Service, Montreal Heart Institute.

Dr. Peter Friedman, Co-Principal Investigator, also Professor of Medicine, Harvard Medical School, Brigham and Women's Hospital.

Also available for questions will be from CryoCath:  Marwan Abboud, Director of Engineering; Patrick Chauvet, Pre-Clinical Scientist.  Consultants:  Susan Bondy, VCRA, Cato Research; Andrew Skrylov, Study Statistician, Cato; Richard Holcomb, Consulting Statistician; John Lehmann, Consultant, CryoCath Medical Director.

Investigators present with us as well are:  Dr. Jose Nazari from Illinois Masonic Hospital in Chicago; Dr. Mark Neibauer, University of Nebraska, Omaha, and Dr. David Keane from Massachusetts General Hospital in Boston.

CryoCath is a public-traded company that was founded in 1995 and which counts over 150 employees now.  CryoCath designs, develops, manufactures, and distributes products from its location in Montreal, Canada, as well as the freezor catheter which is the subject of this PMA.

CryoCath offers products to interventionalists and cardiac surgeons for the treatment of arrhythmia.

Our catheter and probe-based products include Freezor and Freezor Xtra for the treatment of cardiac arrhythmia, Arctic Circler for the treatment of AF originating from pulmonary veins, and SurgiFrost, a cryosurgical probe for the surgical treatment of cardiac arrhythmia.

Our products have been on the market since 2001 in Europe, Canada, Australia, and other countries.  Several clinical trials are being conducted in Europe and elsewhere with the Freezor Catheter, and CryoCath has put in place an international patient registry to collect data from the use of its various products, including Freezor.

In the USA, a feasibility trial is underway with the Freezor Catheter to assess the feasibility of treating posterseptal pathways through the coronary sinus.

Today we will present the results of the FROSTY trial, as well as other pre-clinical and clinical information related to this PMA.

CryoCath has proposed the following indications for use:  The CryoCath Freezor is indicated for the cryoablation of the conducting tissues of the heart in the treatment of patients with atrioventricular node reentrant tachycardia, AVNRT, and for identification of aberrant tissue responsible for supraventricular tachycardia, using reversible electrophysiological cryomapping near the AV node to minimize AV block.

We will now show you the evidence that we believe supports the following conclusion:  Cryoablation with Freezor has an excellent safety profile.  No permanent AV block is clinically effective in AVNRT patients, has excellent long-term success, and has reversible cryomapping.

I would like to introduce Jeremy Ruskin, who will present the evolution of cryoablation.

DR. RUSKIN:  Thank you, and good morning.  Mr. Chairman, Panel members, ladies and gentlemen, I am Jeremy Ruskin.  I am a paid consultant to and own a small equity position in CryoCath, and my travel expenses are paid by CryoCath.

The purpose of my talk is to address very briefly with you some of the history and specific technologic aspects of cryoablation therapy.  This form of ablation is not new.  In fact, it dates back over a century and currently is widely used in a number of areas of clinical therapeutics, including gynecology, urology, tumor surgery, and dermatology.

There's also extensive experience with cardiac surgery using cryoablation for the treatment of cardiac arrhythmias, and this goes back more than 30 years.

While they are working on the slides, in the interest of time, I will try to do the best I can from memory.  There are several areas of cardiac surgery in which cryoablation therapy has been applied.  The first was intentional creation of AV block in patients with disabling or life-threatening drug-resistant supraventricular tachycardias.  It became obvious, when the technique was applied in this setting, that it was both safe and highly effective in producing cold injury that would result in the induction of AV block.

It subsequently was applied to treatment of both AV nodal reentrant tachycardias and Wolf Parkinson-like syndrome and, in fact, was the mainstay, the only nonpharmacological option available for the curative treatment of AV node reentry and accessory AV connections until catheter ablation emerged, particularly that using radiofrequency current in the late 1980s.

Cryoablation has also been used in the treatment of ventricular tachycardia, primarily as an adjunct to endocardial resection, but also as a primary therapy in some situations.

That's it.  Thank you.  And I will note in this particular indication, it became evident that, because of the ability of cryoablation to preserve the underlying tissue architecture while destroying myocytes, one was left with intact tissue with good tensile strength, and surgeons became comfortable applying this therapy to the papillary muscles in the intraventricular septum, sites at which incisional therapy and heat-based therapies could not be used because of concerns about perforation.

More recently, cryotherapy has been applied for the treatment of atrial fibrillation by surgeons using both open and minimally-invasive techniques.

Next, please.  What I would like to do in the next few minutes is speak briefly about the mechanisms of tissue destruction, make a few comments about the histologic characteristics of cryolesions, and then comment briefly on two unique properties of cryoablation, that of adhesion and reversibility.

Next, please.  When applied to the myocardium, cold energy produces a very well-demarcated area of freezing that results in a zone of irreversible cell injury.  This, then, progresses to replacement fibrosis.  Interestingly, as mentioned previously, despite the destruction of myocytes, the integrity of fibrous stroma is preserved, so that the underlying tissue architecture and the tensile strength of the tissue is maintained, even during the acute application of cryoablation.

Next, please.  This is an artist's rendition of application of a cryocatheter to the endocardium of the heart, resulting in a zone of freezing that produces irreversible cell injury and results in the long-term formation of a dense, fibrous scar.  There is an outer halo which is cooled to lower temperatures and at which the effects of the cryoapplication are reversible.

Next slide, please.  This is a typical histologic section.  This is taken from a canine heart one week after application of cryo energy.  The key features are illustrated here.

First, there is an extremely homogeneous dense fibrous scar with very well-demarcated margins, and there is preservation of the endocardial surface.  These features are relatively unique to cryotherapy.

Next slide, please.  This is just a photograph from our institution in the late 1970s illustrating the use of a surgical cryoprobe in a patient with ventricular tachycardia, and you can see the frost forming on the probe at this point.  As mentioned, the surgeons were particularly comfortable using this form of energy because of their ability to treat tachycardias over the papillary muscles and also the septum without major concerns about the risk of perforation.

Next slide, please.  Cryo has a number of unique properties, two of the most important of which are illustrated in the next two slides.

The first is that, when a cryoablation device is applied to the myocardium, it adheres firmly to that tissue during the freeze period.  Loss of adhesion occurs immediately upon rewarming.  This is a unique attribute of cryoablation, and it allows application of cold energy in close proximity to critical structures like the AV node without concern about catheter migration or slippage.  This property is distinctly different from what is seen with heat-based energy sources.

Next, please.  This slide is a movie of a cryocatheter, outlined in blue, introduced through the neck and into an anterioseptal position on the tricuspid annulus.

I hope that this continues to play.  If we can run that one more time, what you will see during this is remarkable stability of this catheter on the tricuspid annulus in an anterioseptal position.  You see termination of the tachycardia, restoration of sinus rhythm with literally no movement of the catheter except in conjunction with the cardiac cycle.  In that position, which is a very delicate position, that kind of stability is almost impossible to achieve with radiofrequency devices.  That is a result of adherence during the freeze.

Next, please.  The other unique property is that reversibility, which is defined as using warmer temperatures or shorter durations of application to cause reversible alterations in cardiac conduction.  This is currently referred to as cryomapping and historically was known as ice mapping, something that was used extensively by the surgeons, and in some cases still is.

It is used most commonly when working in close proximity to the AV conduction system, and is a unique attribute which allows reversible interruption of cardiac conduction without creating permanent electrophysiologic effects.

Next, please.  This is another artist's rendition, just to emphasize the fact that cryomapping and cryoablation are part of a spectrum and that cryomapping, which is defined as a reversible change in the electrophysiology of the tissue, occurs at warmer temperatures and with shorter durations of exposure, whereas true ablation occurs at colder temperatures and with longer durations of exposure.

Next, please.  In conclusion, then, cryoablation is a unique technology that has been used safely and effectively in the surgical treatment of cardiac arrhythmias for more than 30 years.  So this is not a new form of ablation.  What is unique about today's presentation is the availability of this energy source in a catheter-based platform.

Cryoablation has different properties from heat-based energy and results in the following unique attributes:  One, it produces well-demarcated, homogeneous, dense fibrous lesions with minimal to no disruption of the endocardium and a low propensity to thrombus formation.

Cryoablation devices adhere to cardiac tissue, providing an inherent stability to catheters during the ablation itself, and cryoablation provides the ability to cause reversible alterations in cardiac conduction, which provide operators with the potential for targeting desirable sites for ablation while avoiding undesirable sites, such as the AV conduction system.

Thank you for your attention, and I would like to turn the podium back to Jean-Pierre Desmarais.

DR. DESMARAIS:  Thank you, Dr. Ruskin.

I would like to present a technical aspect of this device.  First, I will begin with some definitions of cryoablation terms.  Cryocatheter, an ablation catheter which causes necrosis by application of low temperature.  Cryoapplication, cooling a specific portion of the heart with either diagnostic mapping or therapeutic ablative intent.  Cryoablation, cryoapplication at the coldest attainable temperature for a period sufficient to cause tissue necrosis.  In the FROSTY study, cryoablation consisted of one four-minute cryoapplication at a set temperature of minus 68 degrees Centigrade or colder.

Cryoadhesion, the adherence of the tip of a cryocatheter to tissue during a cryoapplication.  Cryolesions, lesions created by cryoablation.  And, finally, cryomapping, a limited cryoapplication intended to cause reversible alteration in cardiac conduction, achieved by using shorter times or warmer temperatures.

This table shows the differences between cryoablation and RF ablation catheters, and I would like to point to the main difference.  The energy source is nitrous oxide for cryocatheters, which is widely used as anesthesiology gas.  The mode of action and ablation mechanism cryofreezes tissue, leading to ice formation and disruption of cellular membranes while RF heats tissues and denatures protein.  Finally, cryo energy has the ability to provide reversible mapping, a unique feature, unlike RF current energy.

The system is comprised of a console which contains a refrigerate, the user interface, and the safety-monitoring system, a 7 French catheter with a 4-millimeter deflectable tip, electrical and mechanical umbilical to transport the refrigerant and electrical signals between the console and the catheter.

The major safety features include a vacuum system that maintains the catheter shaft under negative pressure to prevent any potential release of refrigerant gas into the bloodstream, flow detection monitoring and vacuum sensors to monitor key operating parameters to prevent catheter pressurization.

A leak detection system which disables refrigerant injection, in the unlikely event of a catheter breach; a blood detector in the handle to prevent the aspiration of blood through the catheter.  We believe these were well-designed and tested features and will offer unparalleled safety for this device.

This slide presents the extensive design verification testing done on the safety system and the device attribute.  All the testing conformed to FDA guidance and industry standards.

I would like now to introduce Dr. Marc Dubuc from the Montreal Heart Institute, who will summarize the pre-clinical experimentation.

DR. DUBUC:  Thank you.  Ladies and gentlemen, Mr. Chairman, and Panel members, my name is Marc Dubuc.  I am Co-Principal Investigator of the FROSTY trial and I am also a paid consultant to CryoCath.  I do not have an equity position in CryoCath.  My travel expenses for this meeting have been paid by the company.

I would like to present a pre-clinical testing.  Extensive animal testing in over 250 animals confirmed cryocatheters' handling characteristics, safety features, and ability to make cryolesions.

In addition, we demonstrated the ability to cryomap, and we compared cryolesions to RF lesions, demonstrating reduced thrombogenicity compared to RF lesions.

We also have demonstrated the ability to ablate safely adjacent to coronary vessels.  These studies provided sufficient evidence of feasibility and safety to proceed to IDE clinical study.

On the left is a typical cryolesion, homogeneous, sharply-demarcated, and dense.  This lesion was made with the four-minute cryoapplications at minus 70 degrees Centigrade.  Note the intact endocardium.

On the right is representative RF lesions for comparison made with temperature-controlled RF at 70 degrees Centigrade for 60 seconds.  Note the adherent thrombus.

Next slide.  A study randomizing RF and cryolesions in the canine heart demonstrated that the median depth for cryoablation of 5 millimeters and RF lesions 5.3 millimeters, shown in the first column, were the same.  However, lesion volumes for RF ablation, 95 cubic millimeters, were roughly twice as large as cryoablation lesions at 49 cubic millimeters.  This resulted directly from the larger surface area of RF lesions in comparison to cryolesions.

Note that the interquartile range for cryolesions are much smaller than those for RF ablation.  Cryolesions are more focused at equal lesion depths with less tissue destruction and have much less dimensional variability than RF ablation lesions.

Next.  In this same study, lesion-associated thrombus was compared using computerized morphometry.  Only 13 percent of cryolesions had lesion-associated thrombus compared with 76 percent of the RF lesions.  This difference was highly significant.

In addition, the median thrombus volume was zero for cryolesions and 2.8 cubic millimeters for RF ablation lesions.  This difference was also highly significant.

This slide shows the lesion-associated thrombus in this study.  On the left you see cryoablation lesions of 23 lesions with no detected thrombus and three lesions with thrombus volumes of 0.4, 0.6, and 5.5 cubic millimeters.

On the right you see that the RF lesions at over 75 percent of the lesions with associated thrombus ranging in size to several thrombi lighter than 20 cubic millimeters.  The association of RF lesions with increased frequency and size of thrombus remain even after all other variables were controlled for.

Next.  The use of the freezor catheter to ablate accessory pathways via the coronary sinus system was also studied in a canine model.  Ablation was performed within 2 millimeters of the adjacent coronary artery, using RF at 30 to 50 watts for 60 seconds.  Cryoablation was administered at minus 75 degrees Centigrade for either one or two four-minute cycles.  Sacrifice was performed at one week for RF and one week and three months for cryo.  Assessments were made for transmurality of the coronary sinus lesions and for stenosis by IVUS and geography and histology.

These slides depict the placement of the ablation catheters in the coronary venous system .  On the top, three angiograms of RF animals, we see the development of a significant stenosis in the left coronary artery one week after RF ablation in the CS adjacent to the artery.  Compare this to the lower panel in which no stenosis developed in cryolesion animals one week and three months after cryoablation.

This study showed that both RF and cryoablation led to transmural coronary sinus lesions.  At all subsequent time points, measures of stenosis were much worse in the RF-ablated hearts as compared to cryoablated hearts with no cryoablated lesions showing atrial stenosis.  Histology at three months showed healed and endothelialized with no stenosis, only minor intimal proliferation, two to three cell layers.

In summary, our animal experimentation with the CryoCath cryoablation catheter demonstrated that cryolesions are dense, homogeneous, and well-demarcated.  Cryolesions are focused as deep as RF lesions with a smaller volume of tissue destruction.  Cryolesions are associated with significantly less frequent and significantly smaller thrombi, even when lesion dimensions are controlled for.

Cryoablation in proximity to coronaries is significantly less likely to produce coronary artery stenosis.  These studies provided sufficient evidence of feasibility and safety to proceed to IDE clinical study.

Now I would like to introduce Dr. Peter Friedman.

DR. FRIEDMAN:  Thank you, Marc.  Ladies and gentlemen, Mr. Chairman, and Panel members, my name is Peter Friedman.  I am Co-Principal Investigator for the FROSTY trial.  I'm also a paid consultant to CryoCath.  I do not have an equity position in CryoCath.  My travel expenses for this meeting have been paid by the company.

I would like to present the FROSTY study.  Next, please.  The FROSTY study was a non-randomized, single-arm, multicenter study.  It utilized FDA's previously-determined objective performance criteria, or OPCs, that were developed during RF catheter testing.

The primary safety endpoint was acute major complications, or AMCs, and the primary effectiveness endpoint, acute procedural success, or APS.  The safety secondary endpoint was the occurrence of long-term major complications and the effectiveness secondary endpoint, freedom from recurrence of the targeted arrhythmia.  Furthermore, this study was designed to investigate the cryomapping effect and the reversibility of cryomapping.

Next, please.  The inclusion criteria for FROSTY were a documented history of supraventricular tachycardia, either AV nodal reentrant SVT, AV reentrant tachycardia utilizing a bypass track, or atrial fibrillation requiring AV node ablation for rate control.

Patients had been referred for radiofrequency ablation.  The ejection fraction had to be greater than or equal to 35 percent, age greater than or equal to 18 years.

Finally, the diagnostic electrophysiologic study had to document inducible sustained AVNRT or AVRT in the baseline state or the presence of a rapid ventricular response in the setting of AF, requiring AV nodal ablation.

Next slide, please.  These are the exclusion criteria in the trial, typical for what one would expect.  Note that atrial tachycardia was an exclusion criterion, and the others are ones that you would predict.

Next, please.  These are the FROSTY study hypotheses for safety:  the 95 percent, double-sided upper confidence bound of the proportion of safety subjects with AMCs was greater than 7 percent for the null hypothesis and less than or equal to 7 percent for the alternative hypothesis.

For effectiveness, acutely, the 95 percent double-sided lower confidence bound of acute procedural success, or APS, for intent-to-treat subjects was less than 85 percent for the null hypothesis and greater than or equal to 85 percent for the alternative hypothesis.  Long-term clinical success, freedom from recurrence at six months would be greater than or equal to 85 percent.

We estimated that a sample size of 165 evaluable safety subjects was required to establish a 95 percent upper confidence bound of less than or equal to 7 percent, when the acute major complication did not exceed a target value of 3 percent.

We also estimated that a sample size of 150 subjects intent to treat was required to establish a 95 percent lower confidence bound of greater than or equal to 85 percent, when the acute procedural success exceeded a target value of 91 percent.

Next, please.   One hundred and sixty-six subjects were enrolled at 14 sites.  Eleven of these were in the United States; three were Canadian.

The first subject enrolled on March 23rd of 2001.  The last six-month followup was on October 23rd of 2002.  Protocol compliance supported the scientific validity of the study conclusions.

Next, please.  Overall, 166 patients were enrolled with a mean age of 48 years.  Patients with AVNRT and patients with AVRT tended to be older than the atrial fibrillation patients.

As one can see, the mean of body surface indices were typical for what one would expect in North America, and there was a female preponderance in each of the three groups in a ratio of approximately two to one.

Next, please.  One hundred and sixty-six subjects were enrolled.  All of these are considered safety subjects, and they were comprised of 103 patients with AVNRT, 51 patients with AVRT, and 12 with atrial fibrillation.  This distribution of diagnoses closely matches that seen in routine clinical practice.

Two patients with AVRT were not treated because of equipment failures, which left a total of 164 intention-to-treat patients.  One of the subjects initially diagnosed as AVNRT turned out to have atrial tachycardia, which was an excluded arrhythmia in this trial, and, therefore, cryo treatment was abandoned in that individual.  Six other subjects did not have a protocol-specified, qualifying cryoablation; namely, one lasting for four minutes at less than or equal to minus 68 degrees.  Excluding these seven patients yields 157 per-protocol patients who were correctly diagnosed and had at least one qualifying cryoablation.

Next slide, please.  Basic study procedures were nearly identical to standard diagnostic and therapeutic RF ablation procedures with the addition of protocol-mandated assessments.  All of the preparation of these patients and the diagnostic portion of the EP study were identical.

When the cryocatheter had been introduced into the body and correctly positioned, the console was activated, either for cooling to achieve ablation or for mapping at the ablation catheter tip.  A physician monitored the response at the tip, noting particularly temperature and time of application.  When the cryoapplication was complete, the console stopped the flow of refrigerant gas, the catheter was allowed to thaw, and the catheter was then moved, if desired.

Next, please.  This slide shows a typical panel from the cryo console during the cryoablation.  On the righthand side are the controls to set the pre-determined tip temperature and the time of application.  In the mid-panel note the orange tracing which depicts the pre-determined tip temperature that was set.  The white tracing represents the actual tip temperature in the catheter, recorded by a thermocouple.  We can see how closely it mirrors the pre-determined desired temperature during the cryoapplication.

At the end of 240 seconds, the injection is halted and very quickly one sees a return of the tip temperature to body temperature; namely, 37 degrees.  The total application here was 240 seconds.

Next, please.  This slide illustrates the procedural details for all the patients.  One hundred and sixty-four patients underwent cryoablation.  There were 1,230 cryoapplications for ablation with a mean of 7.5, ranging between 1 and 36.  One hundred and thirty-five subjects had cryomapping with a total of 820 cryomapping attempts, a mean of 6.1 per patient, ranging from 1 to 37.  Mean fluoroscopy time was 25 minutes, mean procedural duration 265 minutes.  The mean cryoablation temperature, minus 71.2 degrees Centigrade, and mean cryomapping temperature, minus 26.6 degrees Centigrade.

Next, please.  I would like to now present the safety results.  I will discuss major safety outcomes; namely, acute major complications, or AMCs, deaths and serious adverse events, SAEs, and other device-related adverse events.  I would then like to present the safety results for the AV nodal reentry tachycardia subjects specifically.

I would like to add that all adverse events were reviewed by the investigators, study monitors, and the medical monitor, and were then reviewed and adjudicated for final characterization by an independent adverse event adjudication committee which convened on three separate occasions.

Next, please.  Seven subjects experienced a total of eight acute major complications, for an AMC rate of 4.2 percent with a 95 percent confidence interval ranging from 1.7 to 8.5 percent.

No AMC was device-related, as adjudicated by the independent committee.  All subjects recovered completely.  However, the primary safety hypotheses based on OPC with a 95 percent upper confidence bound of less than or equal to 7 percent was, therefore, not met.

Next, please.  This slide illustrates the AMC in each subject.  One subject had a minor pulmonary embolism requiring brief hospitalization.  One subject developed prostatitis after traumatic placement of a Foley catheter.

One subject, after a lengthy procedure utilizing cryoablation which failed and then RF ablation, was found the day after the procedure by echocardiography to have adherent thrombus in the inferior vena cava.  This patient's procedure began at 9:00 in the morning, and heparin was not administered until three or four o'clock in the afternoon.  The patient was treated with anticoagulation, and a follow-up echocardiogram revealed resolution of the thrombus.

One subject with a left lateral accessory pathway failed cryoablation and then was treated with radiofrequency catheter ablation.  Several hours after the procedure he developed evidence of acute myocardial infarction.  Emergency angiography revealed occlusion of the right coronary artery, a site distant from the site of ablation.  This was treated by angioplasty and stenting, with complete resolution of the symptoms and of the evidence of infarction.  He was discharged some days later with a normal ejection fraction.

One subject was noted to have a partially-sheared introducer sheath during removal of a diagnostic catheter.  One subject had perforation of the right ventricle during manipulation of the RV apex diagnostic catheter.  This patient required pericardial centesis.  The same subject was also noted to have deep venous thrombophlebitis on day two after the procedure

Finally, one subject with failed cryoablation -- this was the subject with the atrial tachycardia -- had the cryoablation catheter removed from the body, and preparations were then initiated for a transseptal catheterization because the focus was located in the left atrium.  An intracardiac echo was performed that revealed thrombus that was adherent to a diagnostic catheter extending down the IVC.  That catheter and the thrombus were removed.  A small, residual thrombus remained on the right atrial wall.  The patient was anticoagulated and had no permanent sequelae.

Next, please.  This slide shows deaths and serious adverse events.  One unrelated death occurred in an 89-year-old subject who was hospitalized for cholecystectomy and who had a stroke.  This was four months after successful cryoablation.

Twenty subjects, or 12 percent of the total, had a total of 26 serious adverse events.  None of these was device-related, as adjudicated by the independent committee.  Seventeen of 20 subjects recovered completely and three did not:  the subject who died, another with unrelated compression fractures, and a third unrelated depression.

Next, please.  All device-related adverse events were various kinds of AV conduction disturbances.  The duration of these was a median of 45 seconds, ranging between 10 seconds and 24 hours.  All of the longer-duration AV conduction disturbances were examples of incomplete or complete right bundle branch block.  No patient developed permanent AV block.

The etiology of these conduction disturbances was felt to be related to cryoablation in seven, related to cryomapping in one, due to mechanical trauma in one patient during pacing and was noted post procedure in one patient.

Next, please.  This table shows you all 14 instances of AV conduction disturbance in the 11 patients.  You see that the longest duration abnormalities were those of incomplete or complete right bundle branch block.  AV block, ranging from first degree to complete AV block, generally was very, very short-lasting, often resolving within a matter of seconds.  I would like to emphasize how dramatically different this is compared to radiofrequency catheter ablation.

During RF ablation, an accelerated junctional rhythm is induced which impairs the operator's ability to assess anterograde AV conduction.  When transient AV block occurs during RF ablation, 30 to 50 percent of those patients go on to have permanent conduction abnormalities that require a permanent pacemaker.  The chance of that developing ranges from 1 to 4 percent, depending on the site of ablation.

Cryoablation is very, very different.  There is no junctional rhythm.  One can assess anterograde conduction during the ablation, and if AV block develops, it resolves very quickly.  Indeed, the investigators in this trial quickly obtained a high degree of confidence with this technology in using it very close to the AV node.

Next, please.  This slide shows the procedure-related acute major complications in the AV nodal reentry track of cardio subjects.  Specifically, three of the 103 patients, or 2.9 percent, had procedure-related acute major complications.  One was the subject with the pulmonary embolism; another with prostatitis due to a traumatic Foley placement, and the third, the patient we have mentioned already, with atrial tachycardia who developed thrombus on a diagnostic catheter.

There were, of course, again no permanent AV blocks in this group.  This is a clinically-excellent safety profile in this patient group.  You can see that the 95 percent confidence interval for the AMC rate was .6 to 8.3 percent, and with the Bonferroni correction for multiple comparisons, it ranged from .4 to 9.6 percent.

Next, please.  To summarize the safety results in FROSTY, it is true that the overall statistical endpoint was not met.  However, as an experienced electrophysiologist, I can assure the Panel that this outcome represents an excellent safety profile overall, particularly for the AVNRT patients.

I should emphasize again there were no device-related AMCs.  There were no device-related SAEs, and there were no instances of permanent AV block in 151 patients at risk.

Next, please.  Now I would like to present the effectiveness data.  I'll present the acute procedural success for all intent-to-treat subjects and for the AVNRT subjects specifically.  Then I will discuss those subjects who were acute cryoablation failures who were treated with radiofrequency ablation.  Finally, I will discuss long-term clinical success.

With regard to acute procedural success in the intent-to-treat subjects, there were 164 such subjects with AVNRT, including the one misdiagnosed who had atrial tachycardia, AVRT, or AF.  One hundred and thirty-six subjects had acute procedural success, yielding a point estimate of 83 percent with 95 percent confidence intervals ranging from 76 to 88 percent.  The primary effectiveness endpoint, a 95 percent lower confidence greater than or equal to 85 percent, was, therefore, not met.

Acute procedural success, however, was significantly different among the three diagnostic groups.  This slide shows overall procedural success of 83 percent in the 164 patients overall.  Note that the acute procedural success in the patients with AVNRT was 91 percent compared with 69 percent in the AVRT patients and 67 percent in the AF subjects.

Next, please.  A number of subjects were also treated with RF ablation when cryoablation had not been successful.  Twenty-five of 28 subjects with acute procedural failure during cryoablation went on to have RF ablation; two subjects did not.  Sorry, three subjects did not.  Twenty-three of these 25 of the RF-treated subjects had acute clinical success.

Considering both cryoablation and RF ablation, acute clinical success was achieved in 161 of 166 patients.  This was a point estimate of 97 percent with confidence intervals ranging from 93 to 99 percent.  Adverse event measures in the cryoablation plus RF ablation patients did not differ from those in the cryoablation-alone subjects.

Next, please.  I would like now to discuss long-term clinical success.  There are two ways of assessing long-term outcome.  One is by life table analysis, another by Kaplan-Meier survival analysis.

This slide shows the life table analysis for all 136 patients who had acute procedural success.  We see that the long-term clinical success rate in this group was 91 percent with 95 percent confidence intervals ranging from 86 to 96 percent.

This outcome is clinically excellent, exceeds the protocol-specified rate of 85 percent and exceeds the subsequently-published FDA RF ablation guidance of a conditional long-term clinical success rate of 90 percent with a 95 percent lower confidence bound of 80 percent.

This slide shows the Kaplan-Meier analysis with the same 136 patients, all of whom had acute procedural success.  Using this method of analysis, we get essentially the same long-term, chronic success rate of 91 percent with 95 percent confidence intervals ranging from 86 to 96 percent.

Note that there was no recurrence of the targeted arrhythmia after the third month of followup.  This durable response is exactly what we are accustomed to seeing with radiofrequency ablation.

This is the life table analysis for the AVNRT subjects and demonstrates that the long-term clinical success rate was 94 percent with a corrected 98.3 percent confidence interval ranging between 87 to 100 percent.  Again, this outcome exceeds the FDA generic RF ablation guidance of a conditional long-term success rate of 90 percent with 95 percent lower confidence bound of 80 percent.

Here is the Kaplan-Meier survival analysis for the AVNRT patients.  This results in a very similar outcome as the life table analysis with a long-term clinical success rate of 94 percent, corrected 98.3 percent, with a confidence interval ranging from 87 to 100 percent.

Next slide.  Thank you.  To summarize the FROSTY effectiveness outcomes, for all 164 intent-to-treat subjects, the primary endpoint was not met with an overall acute procedural success rate of 83 percent.  The secondary effectiveness endpoint, however, was met with a long-term clinical success rate of 91 percent at six months of followup.  For the 103 AVNRT subjects, acute procedural success was 91 percent with a long-term clinical success rate of 94 percent at six months of followup.

Post-hoc analyses need to be carefully assessed.  The analysis of the AVNRT group has substantial clinical and statistical validity for a number of reasons.

To begin with, this is a clinically-relevant group.  Furthermore, the group was defined a priori in the protocol.  There was a highly significant correlation between diagnosis and acute procedural success.

The univariate relationship between diagnosis and acute procedural success was for AVNRT 91 percent; AVRT, 69 percent, and AF, 67 percent.  By chi square, this was highly significant with a p-value of .002.

The multivariate relationship demonstrated that diagnosis was a highly-significant factor in predicting acute procedural success.  With a stepwise logistic regression, the p was .022.

Finally, there is a significant amount of additional clinical information, much of it filed with the PMA, that supports the FROSTY AVNRT safety and effectiveness results, as we have reported here.  For these reasons, we believe that this analysis of the AVNRT subjects is valid and is clinically useful.

Next, please.  I would like now to discuss cryomapping, which I believe is one of the most unique and potentially most important features of this new technology.  Cryomapping is defined as a limited cryoapplication intended to cause a reversible alteration in cardiac conduction.

In the FROSTY study, cryomapping was done at minus 30 degrees Centigrade for up to 80 seconds.  This allowed for a reversible assessment of the potential effects of cryoablation at a chosen target site.

Next, please.  I would like to show you just one example, and for the non-electrophysiologists on the Panel whose eyes may glaze over, please bear with me.  We see two panels here during cryomapping.  In the top panel and in each panel you see surface electrocardiographic leads one, two, three, V1 and V6, and then high-right atrial electrograms here from the high-right atrium, from the ablation catheter, and from the his bundle electrogram.

In the baseline state, this patient had sustained AV nodal reentrant supraventricular tachycardia of the typical variety.  These atrial electrograms represent retrograde conduction to the atrium over the fast AV nodal pathway.  The impulse then turns anterograde and conducts through the AV node over the slow AV nodal pathway to complete the reentrant circuit.  The cycle length here is 390 milliseconds.

During cryomapping one sees progressive slowing of the tachycardia cycle length here and here, and then, finally, termination of the tachycardia.  Note that the termination occurs in the anterograde limb, the slow pathway, which is exactly where the catheter had been placed.  When the catheter was then rewarmed, this patient again had inducible, sustained AVNRT.

Next, please.  One hundred and thirty-five intent-to-treat subjects had a total of 820 cryomap attempts.  The remaining 29 intent-to-treat subjects had no attempted cryomaps.

On a per-subject basis, 87 of the 135 subjects, or 64 percent, with an attempted cryomap, had one or more effective cryomaps.  On a per-cryomap basis, 164 of 820 attempts, or 20 percent of the attempts, resulted in effective cryomaps.  So a majority of study subjects and about two-thirds of those in whom it was attempted had one or more effective cryomaps.  Any given cryomap attempt had a 20 percent chance of being positive.

With regard to cryomapping reversibility, of the 164 effective cryomaps, 102 were converted immediately into cryoablation without warming to assess reversibility; 62 of the effective cryomaps were warmed after mapping.  Forty-nine of these reversed immediately and completely; nine reversed in one to six minutes after rewarming, and four went on to cryoablation from four to twenty-one minutes after cryomapping but before complete reversal had been observed.  Ninety-four percent, 58 out of 62, of the effective cryomaps definitely and rapidly reversed.

A subject with one or more effective cryomaps was significantly more likely to have acute procedural success than patients with only ineffective cryomaps or patients in whom cryomapping was not attempted at all.  Looking at the total group, for patients with effective cryomaps, the acute procedural success rate was 94 percent as compared to only 67 percent in patients with ineffective maps or 76 percent in patients with no cryomapping attempted.

Thus, considered as a diagnostic tool for catheter location, having one or more effective cryomaps predicts a group significantly more likely to have acute procedural success.  This significant difference is present in the AVR patients, but the trend in the AVNRT patients and the AF patients was not significant.

I suspect that the effect in the AVRT patients represents the much greater localization challenge inherent in treating accessory pathways and the corresponding benefit that cryomapping offers in that group.

To summarize cryomapping, the FROSTY study demonstrates that cryomapping can reversibly alter cardiac conduction in the majority of subjects in whom it is attempted.  Effective cryomaps are associated with a higher procedural success rate, but this trend did not reach statistical significance in the AVNRT diagnostic group.

Cryomapping utility has many interesting possibilities.  Clinical applications at the moment include the ability to map in order to confirm a desirable site for ablation, but also to identify sites that are undesirable where ablation would likely result in an unintended outcome.

To summarize the FROSTY study with regard to safety and effectiveness, cryoablation with the Freezor catheter is a safe modality in patients with supraventricular tachycardia.  There were no instances of permanent AV block in 151 ablated AVNRT and AVRT subjects, despite frequent ablations near the AV node.  AV block can be monitored and is quickly reversible.  There were no device-related AMCs or serious adverse events.

Cryoablation with Freezor is a clinically-effective treatment for AVNRT subjects, resulting in an acute procedural success rate of 91 percent and a long-term clinical success rate of 94 percent.

Cryomapping with Freezor was observed in 64 percent of subjects attempted.  Cryomapping was quickly reversible in almost every case, and subjects with effective cryomaps had a significantly higher acute procedural success rate compared to subjects without an effective cryomap, 94 percent versus 67 percent.

From FROSTY, I think we can conclude that Freezor is safe for the treatment of patients with supraventricular tachycardia.  Freezor is a clinically-effective treatment for patients with AVNRT, and cryomapping with Freezor may offer advantages for safely confirming desirable ablation sites and also avoiding undesirable ablation sites.

Thank you for your attention.  I will now turn the podium over to Mr. Desmarais.

DR. DESMARAIS:  Thank you, Dr. Friedman.

Here are my concluding remarks:  Cryoablation for arrhythmia has a long history of safety.  Cryoablation has been shown to be safe and effective for the treatment of cardiac arrhythmias in surgery over 30 years.

Freezor has many proven safety systems included in the catheter and console to ensure patient safety.  In over two years of worldwide commercial distribution and over a thousand procedures, there is no reported instance of permanent AV block.  In studies of almost 600 cryoablated SVTs there have been no instances of permanent AV block as well.

Freezor cryoablation, as demonstrated, an excellent safety profile with no permanent AV block.  It is clinically effective in AVNRT patients, has excellent long-term success and reversible cryomapping.

Ladies and gentlemen, Mr. Chairman, Panel members, thank you for your attention.

CHAIRMAN LASKEY:  Thank you very much, and congratulations on adhering to schedule.  It was much appreciated up here.

Barring any burning questions from the Panel, which I would like to hold until after lunch, I would like to proceed with the FDA's presentation.  Does anybody have a question for the sponsor?  If not, we will just hold until after lunch.

MR. CHENG:  Good morning.  I'm James Cheng, and I am the Lead FDA Reviewer for the CryoCath Cryoablation System submitted under PMA P020045.

The FDA review team was comprised of myself, our Medical Officer, Dr. Leslie Ewing, who will present the FDA clinical review summary; our Statistician, Dr. Lilly Yue, who will present the FDA's statistical review summary; Cindy Demian, who performed the biocapability review; Elaine Mayhall, who performed the sterilization review, and Kevin Hopson, who performed the bioresearch monitoring review.

The sponsor-proposed indications for use as seen earlier, the basic components of cryoablation system, the Freezor catheter, the cryoconsole, and the umbilicals and accessories:

The Freezor catheter is a 7 French single-use catheter with a 4-millimeter gold-plated metal tip, 3 EGE ring electrodes, a thermocouple sensor, and a flexible, maneuverable shaft.  The catheter lumen contains a refrigerant injection tube, ECG wires, a leak detection wire, and a thermocouple wire.  The catheter handle contains a deflection mechanism.

The cryoconsole provides refrigerant delivery and recovery.  It maintains a vacuum condition inside the catheter lumen.  It controls the refrigerant pressure and flow rate to achieve the target temperature range.  It contains a device safety system.  It monitors the integrity of umbilical connections, and injection control is implemented in dedicated hardware and has a manual override function for shutting down injection.

The umbilicals and accessories consist of the coaxial umbilical which delivers liquid nitrous oxide under pressure to the catheter and evacuates the nitrous oxide gas; the electrical umbilical, which carries the catheter electrical signals to the auto connection box; the auto connection box, which connects the electrical umbilical and ECG cable to the console, and the ECG, which carries the catheter ECG signals to an external monitor.

The basic principles of operation of the system, cryogenic temperatures are generated only at the catheter tip.  Pre-cooled liquid nitrous oxide is injected under pressure to the tip, where the liquid nitrous oxide expands to a gas.  The face change that the nitrous oxide undergoes is an endothermic reaction which generates the cryogenic temperatures at the catheter tip.

Cryoablation involves achieving a catheter tip target temperature of between minus 68 to minus 75 degrees Centigrade and maintaining that temperature for up to 240 seconds.

Cryomapping involves achieving a catheter tip target temperature of between minus 25 to minus 30 degrees Centigrade and maintaining the temperature for 60 seconds.

The FDA pre-clinical review goals were to ensure the safety and reliability of the device.  For safety, we want to ensure that the device has been appropriately designed and tested, that potential device hazards have been analyzed and mitigated, and that the device safety features have been qualified for use.  For reliability, we want to ensure that the design and manufacture of the device provide us with assurance of consistency with performance specifications.

Testing of the catheter included biocompatibility testing of the catheter materials, reliability testing of the catheter design, mechanical and electrical testing of catheter performance, and qualification of a sterilization process.

Qualification of the console included software and hardware qualification.  For software, we assess the sponsor's design and development methodology, the device software hazards analysis, and the software verification and validation process.  For hardware, we assess the design of the nitrous oxide injection and recovery systems, the temperature controller performance, the device risk analysis, and the design and performance of all device safety features.

The major hazard posed in the device design is the potential for causing a gas embolism.  The manufacturer addressed this hazard with several mitigation features.  One mitigation features involves the design and qualification of the catheter, which included burst and leak testing.  Another mitigation is that the catheter lumen is kept under continuous vacuum during the procedure, which prevents the release of the refrigerant gas into the patient's bloodstream if there is a catheter breach.  There is also a catheter safety interlock which prevents device operation until all the catheter connections have been properly configured.

Additional mitigation features include redundant blood and fluid detector systems that would detect the presence of blood or fluid in various parts of the catheter as a result of a catheter breach.

The flow profile of recovered injectant gas is also monitored to detect any unusual catheter performance.  The pressure relief valve helps ensure that the catheter doesn't become pressurized, and any loss of vacuum will immediately disable the injection.  To minimize the risk of exsanguination, again, you have the blood and fluid leak detectors and the catheter design and qualification.

Another hazard that was addressed by the sponsor was the risk of experiencing freezing temperatures along the catheter shaft, instead of only at the tip.  Catheter qualification testing demonstrated that a break in the injection tube inside the catheter did not allow external shaft temperatures to approach freezing, and the injection flow profile monitoring also helps monitor for catheter failure conditions.

One last device hazard is the risk of software controller failure.  This hazard was addressed by the use of a dedicated, hardware-based injection controller with a manual override for stopping injection delivery and by the use of hardware-based watchdog circuitry to monitor the software for failure.

In conclusion, based on the documentation provided to the FDA by the sponsor, the pre-clinical testing performed by the sponsor is appropriate and acceptable.  Specific hazards posed by the device have been appropriately analyzed and addressed by the sponsor's device design and qualification testing.  Overall testing shows that the device is reliable for human use.

Now to present the FDA clinical review summary, it is Dr. Leslie Ewing.

DR. EWING:  Good morning.  I am Leslie Ewing, and I have no conflicts to report.  I will be presenting the results of the study, and following me will be Dr. Lilly Yue, who will talk about the statistical analysis of those results.

The initial purpose of the study, as you have heard, was to study the safety and effectiveness of this cryoablation system to treat the two types of SVT AV node reentry SVT and SVT due to an accessory pathway, and to treat patients with atrial fibrillation who have rapid ventricular response.

The study was a single-arm, randomized, multicenter study using OPCs, as previously described.  These OPCs were based on the medical literature on RF ablation and designed to be used for the entire pooled study population.  These OPCs have been used in previous ablation clinical trials reviewed by the FDA.

As also was mentioned, since the beginning of this clinical trial, the FDA has put out a guidance document which was issued on July 1st, 2002 entitled, "Cardia Ablation Catheters Generic Arrhythmia Indications for Use, Guidance for Industry."  It can be found at the web address as listed.  These recommendations were put out and intended for radiofrequency ablation catheters and were based on medical literature for RF catheters.

This is a table from that guidance document, and I will point out, as has been previously mentioned, the chronic success or freedom from recurrence was 80 percent, as is listed here.

In this study there were three patient populations included:  AV node reentry, AVRT, patients with accessory pathway, SVT, and patients with atrial fibrillation.

The inclusion criteria is as stated here and has been previously discussed, and the exclusion criteria have been previously discussed.

The primary effectiveness endpoint for the study was acute procedural success.  For the patients with the AV node reentry and accessory pathway SVT it was absence of spontaneous or inducible sustained SVT at the end of the procedure, and for the patients with atrial fibrillation it was absence of AV node conduction.

The lower-bound OPC for the acute procedural success, as you can see here, is 85 percent.  For a secondary effectiveness endpoint or long-term success for the patients with SVT, there was to be no recurrence of sustained SVT by the time of their three-month followup, and for the patients with atrial fibrillation, there was to be no evidence of AV node conduction at the three-month followup.

For the study, the chronic success lower bound was to be 85 percent with the asterisk pointed out to the recently-published FDA guidance, which would be 80 percent.

The safety endpoint for the study is that the number of major complications following the use of the device should have 95 percent upper bound of less than 7 percent.

The definition of major complications used for the study is the standard for FDA, the definition of major complications, and it is any adverse event which occurs within the first week following the use of the investigational device and is life-threatening or results in permanent impairment of a body function or permanent damage to a body structure, or necessitates significant intervention such as major surgery to prevent permanent impairment of a body function or permanent damage to the body structure or requires hospitalization or an extended hospital stay, results in moderate transient impairment of a body function or transient damage to a body structure, or requires intervention such as medication or cardioversion to prevent permanent impairment of a body function or damage to the body structure.  In previous ablation studies the definition has been consistently applied by the FDA in a fairly strict fashion.

As has been described, the patients all had screening and enrollment into the study and all had diagnostic electrophysiology studies.  Some patients had cryomapping, and the determination of which patients were to have cryomapping was at the investigators' discretion.

Cryoablation occurred in all patients that passed the diagnostic electrophysiology study, and the followup occurred at seven days, one, three, and six months.  The followup at seven days was by telephone and also at six months was by telephone.

A hundred and sixty-six patients were enrolled and 164 patients received cryoablation lesions.  This was at 14 study sites, 11 in the U.S. and three in Canada.  The diagnosis after the EP study showed that 61 percent of the patients had AV nodal reentry; 31, accessory pathway; 7 percent, atrial fibrillation, and as was discussed, one patient was diagnosed after a short cryoablation lesion to have -- or cryoablation application to have atrial tachycardia.

Of the 166 patients enrolled, as I said, 164 had cryoablation.  There were 157 that had per-protocol lesions, per-protocol cryoablation application, and that is a full duration, 240-second ablation application.  Six had less than full duration lesions and were, therefore, qualified as acute failure of the device.  All those patients went on to have radiofrequency ablation.

There were, of the total of the patients who received cryoablation, 28 failures, procedural failures, and there were 136 acute successes.  There were 122 of that patient group that had chronic success, and one of those patients, as has been described, is the one patient who had at four months post ablation.

So, to reiterate what has been described by the sponsor, there were 91 percent of the AV node reentry patients that had acute success, 69 percent of the accessory pathway tachycardia patients, 67 percent of the atrial fibrillation, and as the total there were 83 percent.  As has been described, this did not achieve -- so they did not achieve their acute procedural success or their primary effectiveness endpoint.

The chronic results are the patients who had no recurrence of tachycardia or AV node function at three months.  Also previously discussed, of the total groups, 90 percent and 91 percent for AV node, for the accessory pathway patients 88 percent, and 75 percent of the atrial fibrillation -- so of the patients who had acute procedural success, the ones that remained long-term success.

There were seven patients with acute major complications within seven days of the procedure, which also has been stated by the company exceeds the safety endpoint of the study.  Three of these patients had the AV node reentry ablation procedure and four with an accessory pathway procedure.  The company has very clearly identified and discussed these patients, so I will not go into detail with these patients.

Cryomapping was performed on a subset of the entire patient population.  As you have heard, this mapping was performed by using a reversible cryo effect on the conduction system.  The use of cryomapping was decided on a per-case basis by the investigator, and the decisionmaking process was not collected as part of the study.

The criteria for effective cryomaps were pre-determined per tachycardia.  Cryomapping was not part of the pre-determined endpoints of the trial in terms of effectiveness and safety.

A hundred and thirty-five patients out of 164 had cryomapping attempts.  Of those patients, 65 percent had effective cryomaps and 35 percent had only ineffective cryomaps.  The total number of cryomaps attempts was 812 with 20 percent of those being effective cryomaps, but also negative cryomaps may have helped the investigator determine unsuccessful cryoablation with patients.

In the data-collection form process for this study, the investigator could mark "reversible," "not reversible," and that they had gone immediately to cryoablation attempt without attempting rewarming.  There were seven AV node patients who had "not reversible" marked on their data-collection form.  All of these seven patients had successful cryoablation procedures with no adverse events reported and, as has been previously stated, no patients had unintentional AV block.

Dr. Yue will present the statistical analysis.

DR. YUE:  Good morning.  My name is Lilly Yue, Statistician at FDA.  Following Dr. Ewing, I will speak on the study results and give clinical and statistical conclusions.

As specified in the protocol, the primary effectiveness endpoint:  acute procedure success.  The primary safety endpoint:  major complication occurrence.  The secondary effectiveness endpoint:  long-term clinical success at the three-month followup conditional on acute procedural success, evaluated for the entire SVT patient population.

Protocol indicated that the study's, like I said, criterion for the acute procedure success was that the lower 95 percent two-sided confidence bound of the acute success rate for all intent-to-treat patients should be larger than 85 percent.

Please note that the two-sided confidence interval here is necessary when the acute success rate could be larger or smaller than 85 percent, as it was agreed upon by the sponsor and the agency at the design stage.  We will see this, indeed, necessary in a few minutes.

The intent-to-treat patients are those who have a prior ablation catheter activated.  For the major complication occurrence, the study success criterion was the upper 95 percent two-sided confidence bound of the major complication incidence rate for all safety patients should be less than 7 percent.

Here the 15 patients are those who have a cryoablation catheter inserted.

For the conditional long-term success there were statistical hypotheses specified in the protocol.  The alternative hypothesis said the conditional long-term success rate should be larger than or equal to 85 percent.  It is the basic statistical concept that the success rate indicated in the statistical hypotheses are in the population parameter, not an observed point estimate. 		So the study success criterion for this endpoint is the lower 95 percent two-sided confidence bound should be larger than 85 percent for all intent-to-treat patients.

Let's look at the study results.  As we can see, the point estimate of the acute success rate is 83 percent, less than the OPC 85 percent.  So the two-sided confidence interval is, indeed, necessary, and there is no way for the lower confidence bound to be larger than 85 percent.

The lower confidence bound is 76 percent.  Therefore, the study has failed to meet the OPC 85 percent for the acute success for planned patient population.

Also, the upper confidence bound of major complication incidence rate, 8.5 percent, exceeded the OPC 7 percent.  Therefore, the study has failed to meet the primary safety OPC for planned patient population.

For the conditional long-term success, using the protocol specified exact confidence interval approach, if we consider the crude probability of success, that is, classified the two lost-to-followup patients as failures, the lower confidence bound is 83 percent, less than the OPC 85 percent.

If we use a more liberal method and exclude the two lost-to-followup patients from the dataset, the lower confidence bound is 85 percent, just on the border line.

The agency also suggested survival analysis, either Kaplan-Meier or life table cohort analysis, which should give a lower confidence bound between 83 percent and 85 percent, and closer to 85 percent in this case.

Then after looking at the data, the sponsor performed two types of post-hoc subgroup analysis for endpoint statistical hypotheses.  The first one is a retrospective subgroup analysis on the ablation safety and effectiveness endpoints for the three individual subpopulations, using the OPCs originally developed for the entire patient population.

The second one is the retrospective subgroup analysis on the impact of effective cryomapping on ablation acute success.  However, in the protocol, there were no statistical hypotheses and no study success criteria and claims generated for all cryomapping at all.

According to the post-hoc subgroup analysis results, the sponsor made two new claims to support the new indications for use.   No. 1, the AVNRT subgroup met the FDA OPCs for the safety, acute procedure success, and the long-term success.  No. 2, there was a significant association between effective cryomapping and ablation acute success to support these two new indications for use.

Question:  Can the post-hoc subgroup analysis results be used as valid evidence for the new claims?  First of all, when should we perform subgroup analysis?

Generally speaking, subgroup analysis is to determine if there is a reportable subset for which the treatment effect is either significantly more effective or harmful.  So, when the study has succeeded in pre-specified overall analysis, subgroup analysis may be useful in suggesting hypotheses to be tested in future studies or help refine labeling.  If the data shows a significant overall treatment effect, but non-significant treatment effect in subgroups, then the device still could be approved for general use.

However, subgroup analysis is not intended to be used to rescue a study with non-significant overall treatment effect.  So, when the study has failed in pre-specified overall analysis, generally, we do not perform subgroup analysis because the risk of false positive results from subgroup analysis increases.

Particularly, here are some criteria to check if a subgroup analysis is appropriate:

No. 1, is the hypothesis for subgroup analysis pre-specified?  Pre-specified is more believable than post-hoc-specified.

No. 2, is the subgroup classification clinically relevant?

No. 3, is there significant treatment effect in overall analysis?  If yes, may do subgroup analysis; otherwise, generally, we do not do.

No. 4, is there significant interaction of treatment with subgroup variable?  This question generally refers to a two-arm trial, but here we just have one arm.

Let's check if the subgroup analysis of ablation safety and effectiveness for the three individual patient subpopulations is appropriate:

No. 1, in the protocol, no statistical hypothesis was generated for the three individual subpopulations; no study success criteria and claims were developed for the subpopulations.  And the sample size estimation was not based on the three individual subpopulations.

No. 2, yes, the three subpopulations are generally referred to SVT patients.

No. 3, no, there is no significant treatment effect in pre-planned overall analysis.

No. 4, not applicable, but, indeed, there is a big difference in the performance of the device between the three patient subpopulations.  That is why the OPCs developed for the entire patient population cannot be directly used for subpopulations.  It is just our second concern on the next slide.

OPCS developed for the entire patient population may be wrong for all subpopulations.  So it cannot be directly used in a subpopulation.  In fact, from medical literature, the acute success rate for AVNRT patients is significantly higher than those for AVRT and AF patients.  Therefore, the acute success OPC for AVNRT should be higher than 85 percent.

Our other concern is, if the subgroup analysis and the data analysis had been planned in the protocol, a multiplicity adjustment for a significance level, such as Bonferroni adjustment, should have been performed.  Otherwise, the overall Type I error rate of the study, that is, the probability of incorrectly approving the device, would be inflated and could be close to 15 percent.

Let's back up one step and let's suppose the original OPCs were appropriate for the subgroups.  The point estimate of acute success rate is 91, 69 and 67 percent for the three subpopulations, respectively; 91 percent for AVNRT is significantly better than 69 percent and 67 percent.  You can see a huge difference here.

Here the 95 percent confidence interval is the analysis result without multiplicity adjustment.  We also suggested Bonferroni-adjusted confidence intervals.  The highlighted numbers here are the lower confidence bounds, but none of them match the OPC 85 percent in a group.

Similarly, for the major complication incidence rate, none of these upper confidence bounds reached the OPC 7 percent, no matter the confidence intervals are multiplicity-adjusted or not adjusted.

For the conditional long-term success, the second column in the table gives the results of crude probability of success for AVNRT subgroup.  None of these lower confidence bounds reached the OPC 85 percent.  The last column gives the result of excluding the two lost-to-followup patients from the data analysis.  The lower bounds, 86 percent and 85 percent, they are omitted for the AVNRT subgroup.

So suppose the OPCs were appropriate for the subgroups.  For any patient subgroup, with or without multiplicity adjustment, the study has failed to meet the primary safety and effectiveness OPCs.

In conclusion, none of the three subgroups met the OPCs for either primary safety or primary effectiveness.

Let's look at a second subgroup analysis on the association of effective cryomapping and ablation acute success.  Of 164 patients with cryoablations, 135 had cryomap attempts and 29 didn't.  Of the 135 patients with cryomapping attempts, 87 had effective cryomaps and 48 had ineffective cryomaps.  So the point estimate of the effective cryomapping rate is 64 percent.  Of 87 patients with effective cryomapping, 80 had reversibility so the rate is 92 percent.

Now we have three patient subgroups according to cryomapping:  effective, ineffective, and no attempts.  The ablation acute success rate is 94 percent for effective, 65 percent for ineffective, and 79 percent for no attempts.

The sponsor groups "ineffective" with "no attempts" and called it "without effective," then compared "effective" with "without effective" in terms of ablation acute success, and claimed that "effective" was significantly better than "without effective" in ablation acute success.

These claim was supported by a p-value less than .05 from exact test for the overall intent-to-treat patient population.  However, the significant result was driven by 49 AVRT patients.  You can see the p-values in the last column.  However, the AVRT subgroup is not the one the sponsor is currently claiming for.

Instead, the AVNRT is the only group indicated in the new indication for use.  However, there is no significant difference detected in the ablation acute success between "effective" and "without effective" patients for this subgroup, with a p-value of 0.45 from exact test.

For this post-hoc comparison, we have the following concerns:

No. 1, what is the meaning of this comparison?

No. 2, why grouping "ineffective" with "no attempts"?

No. 3, is the subgroup classification "effective" versus "without effective" clinically-relevant?

It seems that if we try to attach the impact of the effective cryomapping on ablation acute success, we could compare the "effective" group with  the "ineffective" group and use "no attempts" as a control.

The ablation acute success rate, 94 percent for the "effective" group, is significantly better than 65 percent for the "ineffective" group, but a patient had only 64 percent of a chance to have a successful cryomapping when he had cryo attempts.

On the other hand, if we try to test the impact of attempted cryomapping on ablation acute success, we could compare "attempt" with "no attempts."  Here "attempt" includes "effective" and "ineffective."

We performed the comparison and found that there was no significant difference detected in the ablation acute success between "attempt" and "no attempt" groups, with a p-value of .59 from the exact test.  The point is made that the ablation acute success rate is 84 percent and 79 percent for the two subgroups, respectively.

From these post-hoc subgroup analyses, we can see that there are many ways to generate a subgroup hypothesis after fact.  Different post-hoc hypotheses and data analyses could lead to different results, either significant or non-significant.

This is the situation we always try to avoid, in which subgroups are defined by the data and the unplanned statistical hypotheses are generated by study results.  Often, a treatment effect is so suggested and then confirmed with statistical significance on that same dataset.

Clearly, the subgroup hypotheses and the data analysis on the association of effective cryomapping and ablation acute success were not pre-defined.  We are not sure if the subgroup classification "effective" versus "without effective" is clinically-relevant.  Also, we do not have information on why those 29 patients did not have cryomapping attempts, purely by chance or due to some patient characteristics.

No. 3, there is no significant treatment effect in overall analysis of ablation acute success.  So, the sponsor's significant subgroup analysis result used as evidence of device performance is questionable.

Clinical and statistical conclusions:  The device did not meet the primary effectiveness of 50 OPCs for either the overall study population or any patient subgroup.  No patient had unintentional, permanent AV node block at the end of the procedure or during followup.  There were a low number of recurrences after successful cryoablation.  Cryoablation adherence appears to have a durable effect.

The post-hoc assessment of cryomapping effectiveness is questionable.  There was no significant association detected between effective cryomapping and ablation acute success for the AVNRT subgroup.  There was no adverse event reported due to cryomapping.

Thank you.

CHAIRMAN LASKEY:  Thank you, Lilly.

At this point, for the record, we will read the questions.  We are still, fortunately, a tad early here for the one o'clock break.  So are the questions from the Panel to the sponsor or to the FDA, for that matter, the FDA presenters?

I had one question for the sponsor.  I was very interested in some of the acute and subacute and chronic changes in the subjacent coronary artery in your pre-clinical studies.  Do you have any further information on the pathology or histopathology of these lesions?  What is it that is acutely constrictive there?  Is it dynamic?  Is it fixed?  Is it something that we need to worry about, if the latter?

DR. DESMARAIS:  Mr. Chairman, I will ask Patrick Chauvet, our pre-clinical scientist, to answer your question.

MR. CHAUVET:  Thank you.  My name is Patrick Chauvet.  I am a pre-clinical scientist for CryoCath, a full-time employee, and I own equity.  My trip was also paid for by CryoCath.

The question was dealing with the cryoablation lesions in the coronary sinus and the effect on the adjacent left circumflex coronary artery:  What were the histological changes acutely, subacutely, and chronically?

In the early trials that we performed we had noticed histological changes both with RF and cryo that were significantly different.  In the RF lesions, the coronary artery acute constructs, and the hypothesis there is because of heating of the collagen fibers in the medial area of the coronary artery, this effect is quite dramatic, and acute constriction is evident; whereas, with cryoablation there is no heating of the collagen fibers, and our hypothesis is that prevents acute constriction.

In the subacute and chronic cases, we have noticed in both cases medial necrosis of the coronary artery in similar amounts.  However, due to the differences in human processes, the cryoablated coronary sinus and adjacent coronary artery healed very well, and there was no subsequent stenosis up to three months in the cryoablation regions.

CHAIRMAN LASKEY:  So is the acute stenosis reversible with nitroglycerin, for example?  You alluded to spasm.  It is spasm?  Did you give nitro and make it go away or some other anti-spasm?

MR. CHAUVET:  Nitroglycerin was given during the study.  As you saw in the presentation, stenosis was still present one week after the procedures.  So it is not just an acute phenomena.

CHAIRMAN LASKEY:  One question from Dr. Waldo and then Dr. Dullum.

DR. WALDO:  Actually, I have a few, but I will just take one then.  I have a question about the definition of efficacy of the cryomap.  I'm not sure I understand the definition because to me, if you apply adequate cooling, almost anything is effective, the way I look at it.  If you don't get an effect, I think that was implicit in some of the things that presented, if you don't get an effect, then are you sure you know you're in an area that is not useful to ablate?

Do you also know that it is safe to ablate if you want to ablate with something else?  In your instance, your presentation, it wouldn't be like atrial flutter where you might want to ablate and you would want to make sure you're not close to any structure.

I don't see why all the cryoablation shouldn't be considered effective if adequate cooling is applied, because it is telling you something each time.  I am not sure why you have to have reversibility as part of it.

DR. DESMARAIS:  We will ask Dr. Lehmann, our medical monitor, to answer the question with regards to the definition.

In terms of the cryomapping clinical applicability, I would like to ask Dr. Friedman to answer that question.

DR. LEHMANN:  My name is John Lehmann.  I am a paid consultant to CryoCath.  I do not own equity, and my expenses were paid by the company to come.

Dr. Waldo, your question bedeviled us as well.  There is a kind of cryomapping where you are looking for an effect, in which case we ultimately simply defined that as, when you cool to a non-destructive level, did you see a physiologic change or not?  And there is a kind of cryomapping where you wouldn't like to see an effect, or what is loosely called "negative cryomapping," where you are working with it at the node and you cool and you don't see it.  Exactly how to determine those in study was less clear to us a couple of years ago.

But at the moment we use the terminology "cryomapping effect" in a positive sense, in a place we would like to see the change, and we do or don't with a non-destructive cryomapping application, and cryomapping reversibility just for those situations when we do see change in physiology and with warming that change goes away.

DR. WALDO:  So then, with Dr. Yue presenting the statistical analysis, if I understand it, when you were ineffective, it was because you didn't see anything.  There was nothing reversed.  But I would suggest to you, why is that ineffective?  It doesn't strike me that that is ineffective.  It might be very effective.  You know you're not in a place you don't want to be, for instance.  That is the most obvious example.  I don't know why that is ineffective.  		If you look at just the statistics, that would clearly affect this.

DR. FRIEDMAN:  Peter Friedman.

I think that is a very good point, Dr. Waldo.  As John Lehmann alluded to just a moment ago, we struggled with how to gauge and how to measure mapping, because it is not something that has been available with any other technology and had never been done before.

You're right in the sense that, if you do cryomapping in a site and you don't see the intended effect, that's a negative cryomap, but, on the other hand, it tells you that you are in a place where it is actually probably safe to ablate.

DR. WALDO:  Or not the place you want to ablate.

DR. FRIEDMAN:  Or not the place that you want to be.  So it is another way of analyzing the data.  Of course, we could include those all as effective, and then the numbers would be very different.

DR. WALDO:  It would certainly affect the statistical analysis we just heard, which was suggesting that it was useless, and I think that statistical analysis doesn't make any sense to me in terms of how we understand it.

DR. FRIEDMAN:  Yes, I think as the study went on, we learned more about what the true utility of this technology is.  It is true that, if during cryomapping one sees block in an accessory pathway or block in the slow pathway, that indicates that it is a good place to do ablation.

But we learned about what we called "negative cryomaps" or "safety cryomaps," also, that if you're applying cryo energy to an area and you see an unintended effect, that is an area not to ablate and it is another utility for mapping that was not addressed when the protocol was written.

DR. WALDO:  Am I wrong in that, do you think?  If you apply it to an area where nothing happens, why would you continue?  Did you find that, if you don't see any effect, you're sure that that is an area that is not desirable to ablate?  For instance, if you don't see any effect on a slow pathway, which would be critical for your presentation, did it prove in your data that this was an area clearly that you should not ablate, that you were wasting your time to ablate that?

DR. FRIEDMAN:  No, and it is difficult to analyze those data from the way they were collected.  But, for example, there were investigators who did cryomapping specifically to localize the slow pathway and were unable to show a reversible block in the slow pathway.

Now one could conclude from that either you're too far away or that the temperature you achieved during the mapping procedure was not adequate to reach the slow pathway which may have been further below the endocardial surface.  So there were actually instances in the study where people, after mapping and being unable to find an effective, quote, "cryomap area," actually went back and ablated at those areas, and some of those turned out to be successful.

CHAIRMAN LASKEY:  We do that in interventional cardiology as well.  So it is not an uncommon precedent.

Dr. Dullum, you had one question?

DR. DULLUM:  I just wanted to ask a followup to your question about the coronary occlusion.  I noticed the intravascular ozone catheter was in the heart in one of your pictures with the RF ablation but not on the cryo.  Did you have it in during the ablation procedures for both of them or did that just happen to be that one picture that we saw?

MR. CHAUVET:  Patrick Chauvet.

No, the AFIS catheter was positioned in every single animal whether it was for RF or cryoablation.

DR. DULLUM:  So during the ablation procedure the AFIS was in there, in the animals?

MR. CHAUVET:  That's correct, yes.

CHAIRMAN LASKEY:  Dr. Krucoff?

DR. KRUCOFF:  You mentioned that you have an experience in Europe.  Is the catheter in clinical use outside of the United States?  Are there any data available that would support any of these discussions from your U.S. activity?

DR. DESMARAIS:  Jean-Pierre Desmarais, CryoCath.

F7, the presentation, the catheter is in commercial use in Europe, in various countries in Europe, and in Australia and other countries in the world.

We found with the PMA there's two separate studies which are not sponsored by us which are an individual initiative from the German Heart in Munich and from the Rotterdam in The Netherlands which compared the NRT to RF, and we do have data on that.  We have submitted those data to FDA during the PMA.

We are trying just to plug the computer and present the data to you in a minute.

DR. ZUCKERMAN:  Okay, I think I have to take issue as to whether the agency has seen these data.  I don't believe so.  And the Panel needs to recognize that any data not contained in the PMA Panel pack have not been thoroughly reviewed by the agency and need to apply the appropriate caveats to such data.

DR. DESMARAIS:  Thank you, Dr. Zuckerman.

CHAIRMAN LASKEY:  I am sorry, are we looking for summary of the European data here?

DR. DESMARAIS:  In about 30 seconds.

CHAIRMAN LASKEY:  Sure, and then at that point I would move to adjourn for lunch.  So we will end on that response.

Oh, I'm sorry.  I'm sorry, Dr. Waldo.

DR. WALDO:  Just a very quick followup to what I was asking before:  Have you any instances where you looked for parahisian pathways?  I mean, this is one of the areas where --

DR. DESMARAIS:  I would like to have Dr. Nazari to answer that question.  Excuse me, I will let Dr. Chauvet to answer that question.

DR. DUBUC:  Mark Dubuc.

As you know, the product is commercially available in Canada, and we have used that to do about 17 cases with parahisian.

DR. WALDO:  Seventy or 17?

DR. DUBUC:  Seventeen, 1-7.  Actually, we were successful in 15 of these cases, and we were able to make what we call a negative cryomap or safety cryomap, and we had clearly a his electrogram on the ablation catheter when we did that.

I can say that maybe in two cases we had transient AV block and we had to cancel the procedure.  This is why we have the 15 out of 17 were successful.

DR. WALDO:  So you also used the mapping, the cryomapping technique?

DR. DUBUC:  Yes.  Well, mainly, the negative cryomapping.  I want to be sure not to ablate at the same time that I'm doing --

DR. WALDO:  Precisely.

DR. DUBUC:  So another thing interesting with cryo is that during the application what you cannot do with RF, you can do program stimulation; you can pace the atrium, pace the ventricle, to see if the AV conduction is still there during the application.  The catheter is very, very stable.

CHAIRMAN LASKEY:  Dr. Bailey had a question.

DR. BAILEY:  Yes, I was wondering if you have compiled data on all initially-enrolled patients as far as their rhythm status at the end of the -- I know there was combining the ones who were successful after radiofrequency ablation.

DR. DESMARAIS:  Dr. Lehmann will answer that question.

DR. LEHMANN:  Could I clarify that you want the overall clinical success at the end of the procedure?

DR. BAILEY:  With the strategy of cryoablation, followed perhaps by radiofrequency.

DR. LEHMANN:  Right.  That was in the presentation.  There was a slide that shows that, of 166 enrolled subjects, 161 had the endpoint success at the end of the procedure.

DR. BAILEY:  And then what about at three months or six months?  Did you follow the radiofrequency cases for their status at three months?

DR. LEHMANN:  We don't have that data here today.

CHAIRMAN LASKEY:  Do you have the non-U.S. --

DR. DESMARAIS:  This is partly the non-U.S.  As explained earlier, that was part of the PMA that FDA maybe did not have the chance to review.

There is a study by Dr. Zrenner at the German Heart.  At the time of filing of the PMA, that was the data we had available showing comparable success rate fluoroscopy induration in terms of AV study.

And there is a second study by the group of Dr. Jordaens in Rotterdam, The Netherlands.  In this case, he was doing it at the NRT and right septal pathways.  Again, the success rate is comparable at this point in time, and that was what we had at the time of filing.  Obviously, these studies are continuing, but we don't have formal updated data at this point.

Does that answer your question?

DR. KRUCOFF:  Yes.  I guess you had references to thousands of uses.  Obviously, these are smaller trials then.

DR. DESMARAIS:  Correct.

DR. KRUCOFF:  You were talking about your commercial use?

DR. DESMARAIS:  Correct.  We do have registry data.

DR. LEHMANN:  We have four IDE-like studies to either Canada or the U.S., which comprise roughly 300 subjects with AVNRT and AVRT.  Those had no permanent block.  There's a total of 600 patients in the registry collection, and then there's between 1,000 and 2,000 commercial uses for which there are no event-reporting systems of any permanent AV block.

DR. KRUCOFF:  Can you just flip back to the first data slide?

DR. LEHMANN:  Yes.

DR. KRUCOFF:  So on the cryo side, does that four third-degree heart blocks only transient?  So "no block," you mean permanent --

DR. LEHMANN:  No permanent AV block.

DR. KRUCOFF:  Correct.

CHAIRMAN LASKEY:  Okay, thank you.  On that note, we're at the magic hour, and I think we'll have plenty of opportunity for each panel member to ask additional material after lunch.  Let's break for lunch and let's meet again at 1:55 sharp.  Thank you.

(Whereupon, the foregoing matter went off the record for lunch at 12:56 p.m. and went back on the record at 2:05 p.m.)

A F T E R N O O N   S E S S I O N

2:05 p.m.

CHAIRMAN LASKEY:  I would like to begin this afternoon's portion of the proceedings by having our Lead Reviewer, Dr. Cynthia Tracy, present her review and ask questions.

I also want to make a bold, and perhaps indefensible request, for Panel members to limit their exegesis to no more than 10 to 15 minutes, so we can get through the afternoon, if that is possible.

I'm assuming we actually got through all the panelists' questions before the lunch break.  Is that a fair assumption?  That is not?  Can you hold them until it is on your round?  Would that be okay?

Great.  So, Dr. Tracy, the floor is yours.

DR. TRACY:  I think it is appropriate that this is room is very frosty for this discussion.

(Laughter.)

With that in mind, I would like to congratulate you on your presentation.  You covered a lot of data and did it very well.  One thing that I thought was remarkably well done was the pre-clinical testing.  I think this device has a lot of potential safety issues that have all been addressed very carefully with excellent safety features built into the device.

It strikes me that both the strength and the weakness of this device is the small, uniform lesion that is created.  I think that is why effective in many instances, but ineffective -- it is effective and safe in many instances, but ineffective in places where you might want to create a larger lesion.

Just a brief question relative to that:  Is there any way that a lesion or any development that might be done could increase the size of the lesion to make the device more applicable for a greater number of arrhythmia substrates?  Is that something that is potentially planned?  Would a larger catheter tip be useful or some other modification?

DR. DESMARAIS:  At the current time, we believe that the catheter that we have is adequate for the treatment of AVNRT, but in our design strategies, you know, to pursue other indications, we have a lot of design strategies and looking and addressing other arrhythmia.  We will surely approach FDA in due course on how best to bring those products to market.

DR. TRACY:  Okay.  I think the cryolesion has the nice feature of not being very thrombogenic, and that might have potential implication for left-sided structures in the future.  So it was just a passing thought.  But I think at this point you are correct that AVNRT is probably your most approachable lesion with the current device setup.

In terms of the device itself, you mentioned that there were equipment failures in two cases.  In looking through the packet, I think one was a console failure, and it looked like on the second one multiple catheters were connected to the console and it was failure to deliver the nitrous oxide, I think.

Can you explain, if I am doing math right, about a 1.2 percent device failure?  Are you happy with that?  What exactly were the problems?

DR. DESMARAIS:  I believe that the catheter consumption throughout the study was about 1.14 catheter on average, I believe.  The console failure was a technical failure where a valve did not open and did not let refrigerant into the catheter.  So that case never took place with trial as such.

In the other failure there was a catheter that kinked on two occasions, and one the injection tube was blocked, so we could not inject.  Unfortunately, the site ran out of catheter, so we could not complete the case with trial.

DR. TRACY:  Okay.

DR. DESMARAIS:  And just to answer your last question, in general, we are very happy with the catheter as it is, and in terms of its reliability and the way it is used.

DR. TRACY:  Okay.  I'm only going to touch on this because I assume Dr. Bailey will be asking much more detailed questions regarding the whole statistical analysis issue, but if you had designed this study for AVNRT alone, what endpoints would you have had to achieve in order to prove statistically that the device was safe and effective?  Have you put thought into that?

DR. DESMARAIS:  I will ask Dr. Lehmann to answer that.

DR. LEHMANN:  The short answer is that we have not at this point attempted to design just a trial for AVNRT.  In this trial we used the OPCs, which, as Dr. Yue stated, were for a cooled population, and those were the numbers that we used.

DR. TRACY:  Okay.  I suspect this issue will get much more discussion as we go around the room, but I will leave it at that for the moment.

A couple of procedural questions:  The procedure time average was 265 minutes.  Does that include from the moment the patient is put on the table until the moment they are pulled off the table?

DR. DESMARAIS:  Yes, it is skin to skin.

DR. TRACY:  Skin to skin?  Okay.  So that is probably a bit long, I think.  Is that related to device issues, people having to set things up?  Is it hard to set up to deliver refrigerant?

DR. FRIEDMAN:  I think there are a number of factors that go into that procedure time.  I think it is true that any investigational study by its very nature, because of the data collection that occurs, takes longer than a routine clinical case.

Some of the time was certainly getting familiar with hook-up procedures, the umbilicals.  Mapping I think took some time.  There was a new technology, and people were interested in exploring how it could be used.  That added to the procedure time.  I think all those things together would explain that.

I think in routine clinical practice, not in the context of a clinical trial, procedure times would not be dramatically different than RF.  In fact, in the data that you saw just before the lunch break, where patients had been randomized RF versus cryo in Europe, there was really no significant difference in overall procedure duration.

Maybe Dr. Dubuc from Montreal would care to comment on that because he probably has the largest experience and has been using it a lot since it is commercially available in Canada.

DR. DUBUC:  Well, I had a chance to use this product because it is commercially available.  I did in my institution more than 150 cases with this technology.  Actually, even our group decided to do all AV nodal reentries with this technology since last August.

The time is comparable, and I mean the time of the procedure is comparable, and also the fluoroscopy time.  Even the fluoroscopy time has a tendency to be a little bit lower because you don't have to watch or monitor the position of the catheter during cryoapplication because of the adherence of the catheter.

DR. TRACY:  You are comfortable enough that that thing is really going to stick on there, that you don't wash --

DR. DUBUC:  Oh, sure, no doubt about it.

DR. TRACY:  Okay.

DR. DUBUC:  No doubt.

DR. TRACY:  Okay.  A procedural question that relates to one of the major adverse events:  I'm assuming that that right corner infarct either was a catheter inadvertently positioned down the right coronary artery during a retrograde approach or was a completely unrelated event.  Do you have a comment on that particular complication?

DR. DESMARAIS:  Dr. Friedman will answer that.

DR. FRIEDMAN:  It's difficult to know precisely.  That, I think, was a case from Canada.  My understanding of that procedure was that cryoablation for the left lateral accessory pathway was done via a retrogradic aortic approach, and there were no technical difficulties encountered crossing the aortic valve.  That was not difficult.  So I think it is unlikely, based on what was reported, that that catheter went down a coronary.

The RF catheter ablation procedure was actually done by a transeptal approach.  That also turned out to be unsuccessful.  So this was a very lengthy procedure in a patient with pre-existing coronary disease, and the occluded vessel, as you alluded to, was actually the opposite side of the heart.

DR. TRACY:  Okay, thanks.

I wanted to ask a few questions about the cryomapping.  The whole concept of it is a new concept that we have to try to figure out exactly what it means.

I think that part of the idea of doing cryomapping is that it is a way of testing where you are without creating permanent damage, and yet I note that there's a little bit of contradictory information that is being given.  In one part of the packet it mentions that 100 percent of cryomapping was reversible.  Yet, we talked this morning about seven not reversible cryomaps.  Are those cryomaps, were they actually cryoablations, or what's the difference there?

DR. DESMARAIS:  I think what I would like to do is to answer that question twofold:  first, with Dr. Friedman, to really explain cryomapping and make an attempt in re-explaining that, and, secondly, for Dr. Lehmann to discuss exactly the reversibility in terms of the data collection.

Dr. Friedman?

DR. FRIEDMAN:  I can answer that question with some assurance because two of those seven patients were patients that were done at our institution, and I remember the details.  Those were patient who actually had cryomapping, not cryoablation, of the slow pathway during AVNRT.

In the baseline state, both patients had inducible, sustained SVT, and during cryomapping we showed that the slow pathway was completely blocked and SVT was no longer inducible.  With cryomapping turned off, the slow pathway recovered function and there were single, occasionally two, echoes, but not sustained SVT.  We waited a while, and still there was no sustained SVT, although the slow pathway had recovered.

We marked that as non-reversible because it did not return to baseline, although I think it was a function of time.  Had we waited 20 or 30 minutes, in all likelihood it would have.

DR. DESMARAIS:  Dr. Lehmann?

DR. LEHMANN:  Part of your question related to data that doesn't seem to match, and that is really the record of our good collaboration with the FDA, to clarify what we felt the true situation was.  So the report in the FDA presentation was that 80 out of 87 cryomap subjects with an effective cryomap had "reversible" marked on their form and seven had the "not reversible."

We went back and looked at that.  Part of the issue was, what was "reversible," quote/unquote?  Is it seconds?  Is it minutes?  So we did the additional analysis of actually looking at the cryoapplication data to see how long it was before the next application occurred either with normal conduction restored or not.  That is the data that we showed.  Either way, you get a better than 90 percent reversibility assessment in short notice and no adverse events related to cryomapping.

DR. TRACY:  Okay.  In terms of the clinical utility of cryomapping, how useful is this?  There was, I think it was, 84 percent attempted and successful, 79 percent not attempted, and yet the ablation was successful.  So it is 79 percent versus 84 percent.  Is that hugely different?

DR. DESMARAIS:  Dr. Friedman will answer.

DR. FRIEDMAN:  Well, I'm a believer in cryomapping.  I think that it is useful.  Having used this technology, in my mind, it is very useful.

I think there are differences among the different patient groups, and that is partly why we presented the data that we did.  For AVRT, if you think about patients with an accessory pathway, these tend to be very discrete structures and they are not the sorts of things that one can localize simply by looking at catheter position fluoroscopically.  It requires very careful mapping technique.  It needs to be very precise.

In my mind, the clearly-demonstrated effect of this, of cryomapping, in that group relates to the fact that, with cryomapping, you're affecting a relatively small area with the tip of that catheter.  You need to be very close to the accessory pathway in order to interrupt conduction in that pathway temporarily.  If you see that effect, then that predicts effective cryoablation because you're right on the spot.

With AV nodal reentry, it may be different because in many cases patients are undergoing AV node r-entry ablation based on anatomy fluoroscopically.  You can position the catheter in what you think is the slow pathway position without regard to mapping -- this is typical of radiofrequency ablation -- and often get a successful ablation.

So the area of ablation for the AVNRT may be larger and mapping precisely may be less critical, but, conversely, the negative cryomap in that group I think is a very important and useful thing clinically, because to map during the slow pathway ablation a spot and find transient AV block identifies a place that you do not want to do ablation.  So it is useful in a different kind of way.

DR. DESMARAIS:  I would like also to ask Dr. Keane from MGH comment on that.

DR. KEANE:  My name is David Keane.  I'm an investigator, and I practice at the Mass General Hospital in Boston.  I have no equity, am not a consultant, but my travel trip expenses, my flights, were paid for by CryoCath.

The reason I came down here today is because I'm strongly motivated to see it be introduced to the clinical practice that we have in Boston, and it has implications just not imminent to mapping but also ablation alone.

I get a call once every two to three weeks from New England physicians with patients who have -- typically, younger patients than older patients with parahisian pathways -- and I have over 10 patients whom I have sort of held off over the last five years on the basis that I was involved in the animal work with this system and have always told them that it is only a matter of time before you will see this thing come through.

They have been holding out for a number of years with recurrent SVT, treated on medication, and a lot of these are teenagers who are treated with beta blocker and anti-arrhythmic drugs.  Some of them are young females who wish to become pregnant and they continue to take their anti-arrhythmic drugs.  That is the real downside for them.  I think it is a shame that they have to wait, in particular, for this system to come along and kind of eradicate their arrhythmia on a permanent basis.

For us, it is really key to see an approval for parahisian pathways because they are the folks that suffer the most in that they have had enough SVT to come for ablation; they have turned down.  They have been brought all the way, almost like being a horse to the water, but they haven't been allowed to drink.  We have mapped these very accurately and found them to be truly parahisian.

There are a lot of patients out there waiting in the wings, not only directly on our own waiting list, but also on the referring docs.  So the answer today has been for AVNRT.  I think the real value for cryomapping, actually the relevant importance of cryomapping greatly exceeds the AVNRT group for people with parahisian pathways.

As Dr. Waldo referred to it, it is a negative predictive value of a cryomap that has been more important than a positive predictive value for these people with parahisian pathways.  The issue is that you have total adhesion to the spot.  It is the same with the mahine fiber.  If you have mahine fiber, if you do a bump map, there is a good possibility that by the time you go on and your tachycardia terminates, again, it is like Dr. Ruskin's video this morning:  If the tachycardia terminates, the catheter moves, you never get it back there again.

With the adhesion, the ability to do a map down to 30 degrees, go on at that spot, if you have elimination of your target physiological endpoint, to be able to go straight on from minus 30 to minus 69 or minus 70 without having to rewarm is the biggest plus of this system, because you are guaranteed that your spot is exactly where you mapped.  If you do it with an RF, either by bump mapping or even a low temperature, 10-watt output, you are still getting smudge lesion.  With every other system, you are constantly smudging, the same way with microwave and ultrasound.

With this cryo, because it is a pinpoint lesion, as you point out, they are very small lesions, and that's why it is critical for these accessory pathways that they are precisely mapped.  Perhaps that was underappreciated in the trial, the critical importance because you get so much collateral damage with an RF.  If you're not exactly on the spot but close to it, you have success.  With cryo, you have to be absolutely on the spot.  I think that may have been underappreciated when we were performing this trial.

But, for me, the biggest issue about mapping is the ability to map a parahisian pathway, particularly for these teenagers and young people who have been waiting now for several years to see this system come through.

Thanks.

DR. TRACY:  I appreciate that comment, but I think that the device isn't being considered for parahisian pathways at this point.  I think our discussion is just related to AVNRT.

DR. DESMARAIS:  I think Dr. Lehmann has a final comment on that.

DR. LEHMANN:  I have just redisplayed this slide from this morning's presentation that I think relates to your question.

First, I would like to point out that this display doesn't relate to our indications for use, but we thought it was a fascinating finding from the study.  In the effective cryomap column, at the bottom you see the 94 percent success rate for those 87 subjects who had an effective cryomap.  For the 48 subjects who had cryomapping attempted but didn't think demonstrated, they had a 67 percent complete procedural success.  For subjects, the 29 who had no cryomapping attempted, it was 76 percent.  The comparison of the first to the second and third columns is clearly significant.

So there is this positive predictive value.  That is really all we were saying, that if you do happen to cryomap and get an effective cryomap, your certainty about ultimate success is high, and so that is really the remark.

DR. TRACY:  Thank you.

Moving on to the cryoablation itself, it strikes me that 11 of 103 patients had transient heart block, and that is high for even -- I would not expect that high of a level of transient heart block with RF energy.  What explains that, and what explains the development of a right bundle branch block when you're working in a slow pathway zone?  I can't quite figure that piece out.

DR. DESMARAIS:  I will ask Peter, Dr. Friedman, to answer that.

DR. FRIEDMAN:  You have touched on a number of important issues, and I think, Dr. Lehmann, if we could have that one slide that I had asked you to show, we will talk about this.

With regard to the right bundle branch block, there is a spectrum of catheter stiffness or flexibility among the catheters that are available in the marketplace, with EPT catheters being probably the most flexible and on the other end of the spectrum maybe a Biosense Webster being less flexible.  Different catheters are used to different degrees in different institutions, and people are accustomed to certain kinds of handling characteristics.

This catheter is probably in the middle part of that spectrum, but tending toward the Biosense Webster.  I think if someone is used to an EPT catheter, which is very flexible, then all of a sudden picks this catheter up, it is a little bit stiffer and it takes a little time to get accustomed to.

I think some of the right bundle branch blocks that you are seeing there are not necessarily related to ablation or mapping.  It is just the mechanical handling characteristic of the catheter.  After all, we see right bundle branch block just with diagnostic his catheters on occasion.  So I think those are my comments about the right bundle branch block.

With regard to the high incidence of transient AV block, I think this relates to the fact again that this was an investigational procedure with a new technology, and the investigators really were set on giving this technology a good test.  So, for example, this is a case actually that we did at our institution.  A woman with AVNRT -- I don't have a laser pointer -- but a woman with AVNRT who had, I don't remember the exact number, a few cryoablations, and after the last of those, the sustained SVT was no longer inducible.  There were still single echo beeps that were inducible.

Now in radiofrequency ablation, that would be an endpoint.  Most people would stop at that point and not go on.  But this is a new technology, and we didn't know whether that was an adequate endpoint or not.  So we persisted, and we did ablations closer and closer to the compact AV node.

So here's an ablation actually in the high to mid-portion of the septum, fairly close to the compact AV node.  You see the artifact on the ablation catheter because of the ice formation around the catheter, and you see here the surface ECG, and here's an atrial electrogram.  That's conducted; that's conducted.  Here's a blocked P wave.

The first blocked P wave is right there, and you can see within one or two seconds the ablation was turned off and you see disappearance of the ice ball this quickly.  Now this is a continuous strip, and you can see that this transient AV block now is gone within four or five seconds.

So this may relate to why we saw transient AV block perhaps more frequently than you would have expected, because people were using this technology aggressively, knowing that if AV block occurred, it would disappear.

I would conclude by just drawing your attention to the Calkins data in a previous submission with a different kind of catheter, where if one looks at the -- yes, and here we have the data for you.

So this is the previously-submitted Calkins data that was reviewed prior to that catheter's approval.  This shows you the incidence of any block during ablation of AVNRT or AVRT.  The incidence was 3.4 percent, so maybe a little bit higher than people are accustomed to seeing.  Ours was 7.2 percent.

But here's where the major difference resides, and I alluded to this in my comments earlier.  Of this 3.4 percent, you know, nearly a third or a half went on to have persistent AV block and needed a pacemaker.  Here is where our difference is:  that AV block is transient.

DR. TRACY:  It is interesting, the electrophysiologic definition of the substrate is not the traditional thing that we look at for success.  It talks about the success being measured as not inducible to 15 seconds or more of SVT.  I would never stop with 15 seconds of AVNRT residual.  Why was that chosen, and how often did you see an effect on both slow and fast in the cryo?

DR. FRIEDMAN:  That endpoint was chosen not for its own value, but we struggled with what were the appropriate inclusion criteria.  We didn't want to do ablation in people who had non-sustained arrhythmia at baseline because it would be so difficult to judge the effect of the intervention.

So we arbitrarily chose greater than 15 seconds' duration as a definition for sustained SVT at baseline.  Once we established that, then the endpoint of a successful ablation was anything less than that.  But I would argue that the success rate long term attested to the fact that that was a clinically-useful endpoint.

DR. TRACY:  And the 94 percent long-term success rate, that's of the cryo patients, not of those cryo-plus-rescue?  That is cryo alone?

DR. FRIEDMAN:  That's correct, that is just the cryoablation alone.

DR. TRACY:  Okay.  In patients with recurrence, at what point was that seen?  Is that an early recurrence, within a day or two, or is that -- it's early?

DR. FRIEDMAN:  It is usually within a day or two, certainly within the first month.  No recurrences after three months.

DR. TRACY:  Okay.  How many patients had first re-AV block following ablation?

DR. FRIEDMAN:  I think we have the slide.  None of them was permanent.  I think the table that I showed you there -- go back.  So there was one, two, three, four, and all of them resolved with 24 hours.

DR. TRACY:  None were permanent?  And none developed first re-AV block in the first month or at a point later?

DR. FRIEDMAN:  No, no.

DR. TRACY:  Okay.  There was sort of a -- we didn't talk about it in your presentation this morning, but the learning curve seems to be -- it seems to take quite a bit of learning to get to use this catheter correctly.  It looked like there was an acute success rate of 85 percent early versus 97 percent late.

How hard is it to learn how to use this catheter and what does that mean in terms of how you would train physicians to use the system?

DR. DESMARAIS:  I think I'll have a multiple answer to that question.  I'd like to start with Dr. Friedman to discuss how to use the device.  Secondly, there is a real number where we have Dr. Lehmann that can answer some of that.

Then, in real practice, Dr. Ruskin.  In the real practice, then we can talk, Dr. Dubuc can, who has been using the catheter now for over 150 cases, as he reported earlier.

DR. FRIEDMAN:  As a clinician who has been around longer than I care to recall, I think it is fair to say that there is a learning curve with any new technology, and that was true of radiofrequency ablation.  If one goes back and looks at acute procedural success in some of the older trials, you know, 85 percent was acceptable.  Now we use an OPC of 92 or 93 percent.  So, seemingly, there is an evolution in the field as people do more and more and get more comfortable, but that is true of also the cryoablation catheter.

It is difficult to demonstrate that statistically within the confines of a small study.  So we actually didn't demonstrate statistically a learning curve.  Maybe Dr. Lehmann can comment on that further.

DR. LEHMANN:  I think it is worth just keeping it very simple.  There was a very minor trend.  There were a couple of difficulties.

One is there were a number of subinvestigators, and when you take 166 or 164 cases and divide them amongst 13 sites with multiple subinvestigators, it is just impossible to do any -- you don't get enough case accrual.  That was really a major issue.  When we did it on a per-site basis, there was nothing significant.

DR. DESMARAIS:  And Dr. Dubuc?

DR. DUBUC:  As I said before, we did more than 150 cases in our institution.  I did myself 100 cases of these.  When we switched to -- we decided to go for cryo for all our AV node reentry cases afterwards, I mean since last August; my colleagues started doing it, but they knew about the technology; they knew about the result of the studies we performed in Canada previously.  And they quickly got it.  Right after three or four cases, they were able to do these things by themselves.

You have to realize that actually the thing you learn from the technology is that you don't expect the same response from the energy source.  Like when you do RF ablation, you expect irradiated junctional rhythm during the ablate.  You don't have that.

The catheter adheres to the under-cardial surface.  So it is different.  You don't have to monitor on the fluoroscopy.  This is what you have to learn.

More than my colleagues, also the personnel working in the lab, they know what to expect and they make the connection quickly, and we know what to expect from the technology.  So it comes very, very quickly, if you have somebody -- not everybody has to go through the same learning curve, I would say, you know, like I did.  My colleagues near me, they did it very quickly.

In closing, your question about the complexity of the system itself, you have a catheter; you have, compared to RF, you have two connections instead of one, and you have an on-and-off button.  So from a complexity standpoint, it is not very complex.

DR. TRACY:  Okay, thanks.  Just to return to some of the labeling issues, I don't know if you want to look at your labeling section, but I think that, based on what you have presented today, I agree these are the appropriate indications that you should be seeking approval for.

Under the precautions, just a couple of statements:  The one, two, three, four, fifth bullet down, "Consider periprocedural coagulation therapy for patients undergoing left-sided and transeptal catheter procedure and for selected patients undergoing right-sided procedure."  That doesn't quite fit with the indication, but I understand why that might be there.

DR. DESMARAIS:  Well, obviously, when we wrote the indication originally, we understood that indications and instructions for use is something that has to be worked on at a later date, obviously.

DR. TRACY:  Okay.  And down further, the third-from-the-bottom precaution:  If patients need to be defibrillated during the procedure, disconnect the catheter's electrical connection part to do defibrillation.  Why?

DR. DESMARAIS:  I would have Marwan Abboud to answer that question.  He is our Director of Engineering.

DR. TRACY:  Okay

MR. ABBOUD:  Ladies and gentlemen, Mr. Chairman and Panel members, I am Director of Engineering.

The different relation issue is mainly --

DR. TRACY:  Excuse me.  Could you please tell what your conflict of interest is?

DR. DESMARAIS:  Marwan Abboud is an employee of CryoCath.

DR. TRACY:  Okay.

MR. ABBOUD:  Director of Engineering at CryoCath.

DR. TRACY:  Okay.  I'm sorry I didn't understand.

MR. ABBOUD:  The defibrillator issue is mainly to protect the console.  As any equipment, when you defibrillate, since we have a thermocouple measurement,  in order to prevent destroying the temperature measurement circuit, we recommended to disconnect the catheter.

DR. TRACY:  Okay.  That's fine.  All right, that's basically all the questions I have at this point, but I sort of reserve the issue of the statistical questions, might come back to that.  I am hoping others will do that, too.

CHAIRMAN LASKEY:  Great.  Thank you, Cynthia.

While Dr. Tracy had the luxury and prerogative of spending 15 minutes, you can appreciate the number of people up here who need their moment in the sun.  So I would like to ask the Panel members to share their thoughts and direct their questions to the sponsor and the principal investigators in less than five minutes, five minutes or less, if they can summarize the key issues that they have before us.  That way, everyone can have a fair shake.

MS. WOOD:  Actually, at this time it would be general questions for either the FDA or the sponsor.  We would ask the sponsor to vacate the table at this time, please.

CHAIRMAN LASKEY:  I would like to begin the Panel's questions and comments with Dr. Gilliam.

DR. GILLIAM:  I have a few questions related.  I think Cindy touched on several, but I am looking at your labeling.  You do say, "The foot switch is available only in Europe."  Is that the plan?  I am just wondering, why was that set up that way, your foot switch for your console?

DR. DESMARAIS:  Obviously, every device has design evolution.  When we designed this trial, we designed it with the current product that we had.  At that time there was no foot switch that existed with the design of the console that we have, but the design evolution in Europe is moving faster.  So it is available there.  But, in due course, when it is time to file amendments to the current product, we will do so within the boundaries of the FDA regulations.

DR. GILLIAM:  Another thing, looking at the first panel of your labeling, I guess that is on the first page on the righthand side, the third point down:  "Do not connect CryoCath to radiofrequency generations as it may result in patient electrocution."  It seems pretty drastic.

Is there a real possibility of that?  I guess I would direct that to your engineering.

DR. DESMARAIS:  I will ask now our Director of Engineering to answer that.

MR. ABBOUD:  I think, as any RF, you, even for the eruption rater, they do recommend not to put a diagnostic catheter.  Since our catheter printout on the connection are different from regular RF catheter, we recommend not to do it at this time.  Thus, we are using different thermocouple.  Since we have a cold temperature, we use a different thermocouple.

DR. GILLIAM:  I'm not opposed to not hooking up to an RF.  It was just quite drastic when you saw electrocution.  Have you all had some experience with that perhaps?  You don't have to answer.

(Laughter.)

While you are there, I do have one question.

MR. ABBOUD:  Yes.

DR. GILLIAM:  Your panel that you viewed the temperature that is right on the console, is that exportable to any type of monitor that an operator could see if they are not directly looking at the console?

DR. DESMARAIS:  At the current time, it is not exportable.  It is not designed to be in such a way.  However, again, in Europe we are looking to introducing that concept.  Whenever it is ready for us to introduce in the USA, we will do so again, within the boundaries of the FDA regulations.

DR. GILLIAM:  This may be one of your investigators could answer this.  Looking at your AV node reentry population, is it typical that you would be doing the slow pathway ablation while the patient would be in tachycardia insofar as you will not have the junctional rhythm that we typically see with RF?  How would you, other than looking at a negative map, cryomap, how would you know you have achieved some degree of success with slow pathway modification during an ablation run?

DR. FRIEDMAN:  There are a number of ways that one can do that.  The example I showed you this morning was cryomapping, just to illustrate how the tachycardia would terminate, and in the anterograde slope pathway direction.

During ablation, most typically, the ablation is actually during a sinus rhythm.  The advantage to that is that the patient's in sinus; you can monitor the PR interval.  Indeed, you can even do atrial pacing or programmed atrial stimulation during the ablation application.

That catheter tip is fixed to the myocardium and will not move, and that allows you to monitor during sinus rhythm or during atrial pacing, while the ablation is going on, when the slow pathway disappears.

DR. GILLIAM:  As far as the shipping and storage of the catheters, are there any special precautions of this catheter, given it is sort of a multi-lumen catheter?  I notice the kinking is a concern, obviously, for reasons, but how does the catheter come in storage?

DR. DESMARAIS:  The catheter is packaged into a tray system.  It is packaged flat and it is not coiled on itself, and the storage conditions for catheters and console is standard of any ablation catheter out there.

DR. GILLIAM:  Those are all the questions that I have right now.

CHAIRMAN LASKEY:  Thank you, sir.  Thank you twice.  Mitch?

DR. KRUCOFF:  I also will try to be brief.  I guess I am most interested in the use of the OPC-based trial design.  Obviously, that is consistent with the precedence in this area of industry.

But I think it has been remarkable to me, in listening to your presentation, the difference between the fascinating findings from this study, which is to me the investigator's view of what happens in a patient and how it works, versus what you have welded yourself to with an OPC-based trial design that I think Dr. Yue did a very eloquent job of illuminating.  It is really a population kind of statistic.

So the observed behavior of the catheter in individual patients, the PI sort of view, and how it will behave in the population, and whether or not you qualify based on boundaries for your endpoints as successful or not are two very different worlds.

I guess what I would like to hear is, why do you think you failed to show the boundary outcomes?

DR. FRIEDMAN:  I think there are a couple of points to mention in answering that question.  When the protocol was being designed, we were faced with a choice of either using an OPC, an historically-derived OPC, or doing a randomized comparison between cryo and RF.  The randomized comparison would not have allowed us to look at cryomapping because there is no way to map with radiofrequency.  So we were, basically, stuck with the OPC comparator for the purpose of this trial.

I would point out that that OPC is derived from results recently, and RF is a very mature technology.  It has been around for nearly 20 years now.  This is a very young technology.  So we are comparing ourselves to a very high bar.

Beyond that, I think that the technology was demonstrated to be clinically effective in the group for which an indication is being sought:  Ninety-one percent in the AVNRT patients is a clinically-valuable result.

The reasons why the OPC weren't met I think relate to the fact that some of the ablations were in AVRT, left-sided accessory pathways.  This technology is very different than radiofrequency.  I think you saw the acute effectiveness results were less dramatic in the AVRT group than the AV nodal reentry group.

I think you have to think about the biophysics of ablation and the differences between these two technologies to understand that.  With radiofrequency ablation of the left free-wall pathway, typically done along the mitral valve annulus, this is a high blood-flow area.  That high blood-flow actually serves to cool the radiofrequency tip, and in a way it is almost akin to a saline-irrigated RF ablation.

From Dr. Dubuc's animal studies, we know that radiofrequency ablation causes a larger area of damage on the endocardial surface.  That relates to the fact that the catheter is moving.  It is not stuck to a certain point.  It paints back and forth across the endocardial.  So that one doesn't need to be exactly on the accessory pathway to achieve a successful ablation.

It is very different for cryoablation.  Cryoablation adheres to the point and does not move from that point.  That high blood-flow along the mitral valve actually acts as a heat source that minimizes the efficiency of the cryoablation.  I think that relates to why there may be a difference in effectiveness in that group, which affected the overall results.

But, conversely, I would look at it a different way, because in high-flow areas where cryoablation may not have an advantage, in contrast to low-flow areas, it may have a very real advantage, and specifically in post-receptal accessory pathways within the coronary sinus, for example, where radiofrequency often can't be done because of high impedance and low-energy delivery, and in the slow pathway position, which is not as high a blood-flow area.  It results, I think, in the fact that in that circumstance it is highly effective.

DR. KRUCOFF:  Well, I mean, I hear you, but I'm honestly not sure to which side of my question this -- you know, again, to me, what I hear is an enormous amount of conceptual and intellectual fascination in why and where this technology really might be an advance beyond RF.

But how much of that is reflected in this clinical trial or can be deduced, or even in the 91 percent in AVNRT, is the actual observation?  It is the boundaries around that that become even relevant to talk about with an OPC.  As was described, it is not the boundaries around the whole population.  It would be the boundaries around the RF population and AVNRT.  Again, at least my understanding of the data, unless I missed something, is that even the boundaries around the 91 percent for AVNRT do not compare favorably to the RF historical boundaries that you would pre-define if you had prospectively pre-defined for that population.

So, I mean, I get that there are a lot of theoretical both safety and effectiveness possible subpopulations that this instrument might be terrific for.  On the other hand, I also get that, like so many things that we do, it is possible that the more damaging milia that RF creates may have areas where it is actually more effective.

The fact that I saw two different sets of numbers, but whether it is 23 out of 25 or 25 out of 27 of the cryo failures that were successfully done with RF, I don't need an explanation of that.  To me, it just says that one is more histologically pleasing to look at than the other.  It ultimately begs the question of when you start applying this to a population of human beings, where is the data that gives you confidence that this is a safe and effective approach?

And with an OPC set of boundaries that, even in the group who you are asking for an indication, if you drill down retrospectively to your point control population, you still don't make the boundaries for the AVNRT.

So I am just missing a link as to, if you bought into this trial design from the beginning, which obviously you must have at some level, how do you come to see the data supporting the fascination?  The fascination is self-evident?  The data that this is safe and effective compared to a well-established, as you say, mature technology is, I think, what the patient or user side question is.

I am really interested in why you think, both in the safety and in the effectiveness, you failed to make your boundaries?

DR. LEHMANN:  I will just briefly comment on the design.  As you do know, sponsors don't have a free hand in choosing a clinical study design.  It is negotiation with the agency.

We either had the choice of a --

DR. KRUCOFF:  Well, you could have done a randomized trial.

DR. LEHMANN:  Yes, we could have done a randomized trial, but there are a number of problems with that, including standardizing RF therapy procedures and equipment, which is actually rather hard to do -- it will cause a large study -- and the difficulty with cryomapping.

As to the implications of what we have demonstrated, I think Dr. Friedman will just have a few remarks.

DR. FRIEDMAN:  I will try to answer your question.  I understand what you are wrestling with, and it is a difficult issue.  I will try to answer it by putting it in clinical perspective.

I think it is very helpful to step back and try to look at things from the perspective of the patient.  If I'm a 20-year-old patient with AV nodal reentry that is interfering with my lifestyle and a physician gave me two choices:  One choice, I can have radiofrequency ablation with a 94 percent chance of acute success; however, there is a 1 to 2 percent chance that I am going to wind up with a permanent pacemaker.  Or I could have a cryoablation with a 91 percent acute procedural success rate and no chance -- no chance -- for a pacemaker.  It is very clear in my mind what I would choose, and I would venture to say that most individuals faced with that decision would make the same decision that I have just made.

CHAIRMAN LASKEY:  I must say I need to object to that last statement.  It is not zero.  The confidence intervals of zero events observed in 150 patients go out on my back-of-the-envelope, amateur stat. program here to 2 percent.  It's not zero.  I think you do need to be rather intellectually honest to some of these questions.

I think we are on the threshold of perseverating as well.  So I'm not sure we are going to get to what you really need to know.

DR. KRUCOFF:  I think a question is on the table, and I think, you know, again, I am just going to move on because I only have a couple of other brief questions.

In your European experience, are you aware of any instances where RF ablation has failed and cryoablation has succeeded?

DR. DESMARAIS:  Well, in fact, every time that there has been ablation in AVNRT and do defibrillation near the aveno where the physician either pulled back, because of fear of I think an AV block and subsequently tried trial, there is a lot of data that concurs that we have successful cryoablation after RF.

DR. KRUCOFF:  Okay.  Thank you.  So what I am hearing is, in situations that the operator sort of judged to be too risky to actually fire RF energy --

DR. DESMARAIS:  Or has fired RF and failed in that, as well I think there's a current trial in the U.S.  Again, this was filed initially by the PMA.  The data we have, we know that these are patients that failed RF, that the study conducted by Dr. Chapman right now, which is a small subset study, but we don't have any data to present on that.

DR. KRUCOFF:  Okay, and are you aware of any experience with cryo patients that were not durable, that came back, that were redone with cryoablation?  Can you go back and freeze the same site twice.  Do you have any experience with that?

DR. DESMARAIS:  Yes, and I think Dr. Dubuc can answer that.  He has done some of those patients.

DR. DUBUC:  I think I have two parts of this question, if we have patients who have failed RF and they went to cryo or if they had two sessions of cryo?

DR. KRUCOFF:  Yes.

DR. DUBUC:  Okay.  Well, yes, with no problem.  No problem at all.

DR. KRUCOFF:  Have you looked at that in animals, if you repeatedly cryoablate the same tissue?

DR. DUBUC:  Nothing will happen because the process is already in evolution.  Actually, I could say that, even if you stand there, and why we picked that time of four minutes is because we know that after three minutes you reach a plateau that the ice ball inside the tissue will not grow anymore.  It will stop there, and that has been proven by cardia ultrasound, because you can see the ice ball with ultrasound, with intracardiac ultrasound.  Actually, we can monitor that.  Actually, the correlation coefficient between the measurement done by echo and the histology was .95.

DR. KRUCOFF:  Okay, I have no other questions.

CHAIRMAN LASKEY:  Thank you, Mitch.  Dr. Page?

DR. PAGE:  Thank you.  I will try to keep this as short as possible.

The first question is -- actually, the first two questions have to do with the footprint of this device.  Peter, maybe you could answer this, Dr. Friedman.

In terms of looking for biological plausibility for why you found what was, I guess, an unexpected result, that AVNRT was successful and AVRT was not, as I wrestled with the biological plausibility of that finding, as opposed to just chance, which is what we always have to be concerned about with a post-hoc analysis.  Would you say it has to do with the footprint of this device being smaller and being adherent at the time of ablation lesion?

DR. FRIEDMAN:  I think those are the answers, as I alluded to earlier, yes.

DR. PAGE:  Okay, fair enough.  Do you have an idea, how long is it before you get adherence of this catheter at the ablation site?

DR. FRIEDMAN:  It occurs within a matter of 10 or 15 seconds.  It depends a little bit on how stable the catheter is to begin with and what kind of contact one has with the tissue.

But by the time the artifact on the catheter that I showed you there indicating freezing appears, that catheter is stuck.

DR. PAGE:  And, likewise, would you say that it detaches at about the same interval that we saw on that electrogram?  So a few seconds, but pretty promptly?

DR. FRIEDMAN:  Within three or four seconds.

DR. PAGE:  Within three or four seconds, good.

And as someone who performs ablation, I recognize the advantage of having an adherent catheter at the spot you are burning or freezing in this case, because, especially if you are burning during tachycardia --

DR. FRIEDMAN:  Yes.

DR. PAGE:  -- then it will stay put when the geometry and the motion of the heart changes as you break the tachycardia.

Let me flip that around.  I have spent the last 10 years doing ablations in Texas, and every once in a while a patient would get up off the table in the middle of a burn or grab or move a leg.  What data do we have in terms of, if a catheter is pulled when adherent to the tissue -- you understand my question?

DR. FRIEDMAN:  I do.

DR. PAGE:  We've all been through this, being in a busy lab, and suddenly a patient moves, and your catheter might move with him.  With the standard RF catheter, it may perforate or it may more likely pull away from where you are burning.

But if you are attached and someone took a tug at that catheter, what happens to the tissue?

DR. FRIEDMAN:  Well, there are some animal studies that we could cite, but I would follow up just by saying that, if the patient moved, injection can be halted within a matter of seconds.

DR. PAGE:  I understand that, but this happens so fast, and we have already agreed that it sticks for a couple of seconds.  This happens in less than half a second, as we all -- anybody who has been in an EP lab recognizes sometimes the patient will suddenly come up off the table.

DR. FRIEDMAN:  Yes.

DR. PAGE:  So if that does occur with adherence, what happens to the underlying tissue?

DR. DUBUC:  I mean, if the range of motion you are talking about is rather small, I mean I don't think that the patient would start running from the table.

DR. PAGE:  No, but imagine -- let me put it clearly because it happens.

DR. DUBUC:  Yes, I know.

DR. PAGE:  Imagine someone in the middle of the burn tugs it forcefully.  I mean, what comes with the tip?

DR. DUBUC:  Okay.  If the patient moves, nothing will --

DR. PAGE:  That's not my question.

DR. DUBUC:  Okay.  If somebody tugs really just a little --

DR. PAGE:  No, I'm saying a good, strong tug.

DR. DUBUC:  Well, you can do damage to the tissue underlying the --

DR. PAGE:  The underlying tissue?

DR. DUBUC:  Yes.

DR. PAGE:  Okay.  And my only caution being, you know, we have had physician experts talking about looking forward -- I should just mention, first of all, that at the AV node slow pathway area, this probably would not be a major catastrophe.  If someone is taking this to other places -- we have talked about parahisian.  What if we wanted to outflow track ablations, something like this, where the tissue is thin, atrial tachycardias where the tissue is thin, do we have any data on what would happen then?

DR. DUBUC:  Well, something wrong can happen because, you know, we made animal studies on that or on purpose we pulled --

DR. PAGE:  Sure.

DR. DUBUC:  -- you know, strongly on the catheter.  Actually, if you pull a lot, you can bring part of the heart inside the IVC and even have avulsion of the tissue.  Really, if you have two people on the -- we did that testing.

DR. PAGE:  Sure.  It doesn't happen often, but it does happen rarely.

DR. DUBUC:  Yes.

DR. PAGE:  Okay, thank you.

One question I had, and I just have one other question after this:  We seem to agree that there were no catheter-related complications, but there were a number of complications that I think related to procedure time and DVT, atrial clots, and bladder infection related to a Foley that you don't need for an hour-and-a-half procedure, but you might need for a 260 average minute procedure, the average number of burns is 7.5, and those were four-minute burns?  Is that right?

DR. FRIEDMAN:  Four-minute freezes, yes.

DR. PAGE:  Thank you.

(Laughter.)

I am so old-fashioned.

So four-minute ablations, if you will.  So that is about 30 minutes of actual ablating time.  So this is, indeed, a longer procedure, related to whether it is the experimental condition or not, it is a longer procedure.  I think the length of the procedure related to the experimental protocol, indeed, was responsible for increased complications.

I would mention that the European data you showed us, which I never had a chance to see before today, was reassuring in terms of number of lesions as well as duration of the procedure.  So I could presume that those wouldn't be ongoing complications, and the procedure time would, indeed, shorten.

My last question was, in terms of the comment on parahisian pacing, which, as I mentioned, concerns me because we are talking about a very limited indication here, and with approval, I am already hearing that people are looking forward to using off-label use of the device and potentially in tissues where we have already discussed could be a problem if the catheter weren't handled properly.

But if I saw the data properly, when you have a failed cryomap, that still translated to a 67 burn success at that site?

DR. KRUCOFF:  Ablation.

DR. PAGE:  Ablation.  The slide, it was like 90 percent ablation when you have a successful cryomap, but even a failed cryomap, you ended up getting successful ablation 67 percent of the time.  So my only caution is, with these 20-year-olds that are waiting to have their parahisian pathways ablated, a negative test pulse, cryomap, will not necessarily translate to failure to ablate the his bundle.  Am I interpreting the data correctly?

DR. LEHMANN:  I would make one minor correction.  The figures that you are quoting that were recently up there were on a per-patient basis.  The distinction between a subject with one or more effective cryomaps versus subjects who were attempted but had not effective cryomaps, versus the group that was never attempted.

So that isn't on a --

DR. PAGE:  So on a per-burn basis, it is lower than that?

DR. LEHMANN:  Well, the overall is -- we would have to go back and look.  I don't have that on the tip of my tongue.

DR. PAGE:  I think you're right, it's significantly lower.

DR. LEHMANN:  Yes.

DR. PAGE:  My only caution being that I don't know what comfort I have that the test pulse, indeed, will not result in true ablation of the normal conduction system.

And my last question is, the reversibility of the heart block that you have seen, that wasn't just with cryomapping; that was also during an attempted ablation, is that right?  And then you turned off in time and, indeed, conduction returned?

DR. LEHMANN:  There are two aspects to this.  One is the straight cryomapping where we had 162 effective cryomaps, of which only around 64 of them, 62 of them were warmed for reassessment.  That is where we got the 90-plus percent reversibility of an individual intended cryomap effect.

Then we had 11 adverse events, these device-related transient AV block, of which one was related to cryomapping and led to a death.  It was intended.  It wasn't the therapy, they contend.  They were trying to go after one of the slow pathways, and that subject had unintended AV block that lasted 20 seconds.  So that was an adverse event.

DR. PAGE:  But in terms of when you are therapeutically ablating, those do also, if you see heart block, they tend to get better?

DR. LEHMANN:  Those, we had the 14 instances in 11 subjects all reverted.

DR. PAGE:  And that is reassuring.  Thank you very much.

CHAIRMAN LASKEY:  Nobody in this room is more sensitized to the off-label use of the material that we're discussing.  So could we please restrict our comments to the material at hand and not encourage off-label use?  But thank you very much.

Dr. Aziz?

DR. AZIZ:  I, too, enjoyed the presentation, and I must say I learned a little more electrophysiology today.

Let me just address a few questions.  Looking at the number of patients who actually developed either arterial or a venous clot, I guess part of that may be related to the time that the catheters were in place, and it seems that the times with experience do decrease.

Do you put these patients on any anti-platelet agents or anything that you suggest that that should be done?

DR. FRIEDMAN:  In the study protocol, decisions about anticoagulation during and after the procedure were left to the investigator based on the practice in that laboratory and, as you might imagine, varied widely from one institution to another.

For example, the one case of a pulmonary embolism that I mentioned as an AMC was a woman who actually had a fairly short procedure and did not get heparin during her procedure, but she went home and spent two days in bed and developed a deep venous thrombitis that embolized, undoubtedly related to the catheter insertion site, but in a sedentary, bed-ridden person.

These other procedures that I showed you, one of them I mentioned specifically was a very, very long procedure where no heparin had been given for some eight hours.

So our own practice at our institution is that patients get anti-coagulation during right- and left-sided ablations, and they all get sent home on aspirin, but I think that that is going to vary from investigator to investigator.

DR. AZIZ:  And you have seen the histological slides of the difference in damage done with cryoablation versus RF.  In clinical usage, did you see that translate into better or less CPK leakage with cryoablation versus RF, or you didn't think about that?

DR. LEHMANN:  We did see some CPK rise and we didn't compare it to any RF experience, but there is some CPK rise with --

DR. AZIZ:  In somebody who had a prior, let's say, tricuspid valve repair or a tricuspid valve replacement, could you still use this technology, depending on, obviously, the type of --

DR. LEHMANN:  These subjects were excluded from this trial.

DR. AZIZ:  I guess, when it does come out, could you see it being used in somebody who has had prior tricuspid valve surgical procedures done?

DR. FRIEDMAN:  That's a problem even with radiofrequency ablation.  I don't know the answer to that.  As I think about the biophysics of this technology, nothing occurs to me right away that would make me think that it would be contraindicated.  If there is a sewing ring in the tricuspid valve and you are trying to ablate beneath that, I am not sure how effectively the cryo energy would be transmitted.  It might not be that effective, but I don't think there would be any particular safety concern.

DR. AZIZ:  Let me just go over -- there were one or two patients.  On page 85, Subject 0917, could you just have a look at that?  This was actually a 76-year-old male who had the procedure performed.  Echo post-procedure showed that his EF had increased and he had quite significant mitral regurgitation, and that the left atrial size had actually also increased.  The dimensions had increased by 9 millimeters.

DR. LEHMANN:  Which document?

DR. AZIZ:  Page 85, Subject 0917.

CHAIRMAN LASKEY:  In the general packet itself, the patient descriptions.

DR. LEHMANN:  I can't find the reference.

DR. AZIZ:  The patient was 0917.

DR. FRIEDMAN:  I'm reading here from the report.  I don't know this patient, but let me just read the details of what we know.

DR. AZIZ:  Okay.

DR. FRIEDMAN:  An 89-year-old female had a baseline echo which showed aortic stenosis, mild tricuspid, and mitral regurgitation and inject --

DR. AZIZ:  Actually, it is the next one.

DR. FRIEDMAN:  Okay.  Seventy-six-year-old male, baseline echo showed thickened aortic valve leaflets, mild MR, mild TR, and an injection fracture of 63 percent, underwent successful AV node ablation for rate control and permanent pacemaker without adverse event.  Post-procedure echo demonstrated dilated left atrium; the pacemaker in the right ventricle.

I think it is difficult to compare those two echoes for a couple of reasons.  No. 1, baseline, the patient had a very rapid ventricular response and, presumably, a normal QRS because of rapid conduction through the AV conduction system.

Post-ablation, clearly, had a much slower heart rate, which by itself would increase cardiac dimensions by any measure, and, in addition, had a right ventricular pacemaker, because there was no intrinsic AV conduction.  Pacemaker implantation alone in someone with AV block could lead to ventricular dilation and even some degree of mitral regurgitation.

My guess is that that is related to the pacemaker implant and the presence of AV block post-procedure as compared to pre-, not a function specifically of cryoablation.

DR. AZIZ:  Thanks.

DR. WALDO:  I just had a few short questions.  The histology that you showed us, the iced lesions looked like ventricle.  Am I correct?  Have you done that at atrium?  Is there any more of the same, because --

DR. DUBUC:  For the purpose of presentation, we showed mainly ventricle lesions.  Those lesions are all ventricular, the lesions you saw this morning in the presentation, but we did the same in the atrium.  Naturally, all these lesions were transmural.

DR. WALDO:  That was my point; they are transmural?

DR. DUBUC:  Yes.

DR. WALDO:  And no other problems?  Okay.  How do you decide on 4 millimeters being necessary duration of the application of freezing?  For minutes?  I'm sorry.

DR. DUBUC:  The four minutes, yes.  Well, we -- and this work is already published, but we demonstrated with the ultrasound we can monitor the ice ball growth within the tissue, and we can monitor the size and it is growing for about three minutes, and then you reach a plateau after three minutes.

So we decided that four minutes was okay.  This, as I said previously, the size of the ice ball correlated very well with the coefficient, the correlation coefficient of .95, when we compare histology to the size we found by ultrasound.

DR. WALDO:  Okay, two other very brief questions:  Peter, when you answered an earlier question, you didn't address the issue of the relatively surprising to me poor success rate for his cryoablation.

DR. FRIEDMAN:  Yes, that does stand out, Dr. Waldo, though I am not sure why.  The numbers are small, and those confidence intervals are fairly wide.  I think it is hard to make any definite conclusions.

DR. WALDO:  Okay, now just a little heresy from me:  I don't know, I think this is a philosophical point I would like to make in considering this.  I think the statistics here have been very well presented.  I am not a statistician.  We are going to hear a lot of more, and I think if you just took a statistic, there are some real questions, and Mitch addressed some of them.

But, for me, that is why it is heresy, because I hope -- I have been a scientist in my time, and I think data are very important, but I look at this, do look at this in part as a clinician.  I think one of the things I look upon this as another option.  There are not a lot of secrets about AV nodal ablation.  We know the mechanism.  We know how to go about it.  There are not a lot of secrets about hisp ablation, even about getting accessory AV connections.

I just see that there are times when one would want to have an alternative technique.  And what becomes important there is safety.  And if I understand the data, I think we have seen that it is remarkably safe.

I think efficacy is always important, but I haven't seen a terrible efficacy really.  I have seen quite a good efficacy, maybe not as good as some of us would like to see, especially statistically or if you look at ITT and a lot of other things.  I understand that thing.

But I think that we shouldn't lose sight of the fact that this provides the clinician with an alternative therapy that may be very important sometimes, and it is safe.  I would just like to leave it at that.

CHAIRMAN LASKEY:  Not heretical at all, Dr. Waldo.

Dr. Bailey?

DR. BAILEY:  Well, I feel like everyone's waiting for my remarks here.

(Laughter.)

As a statistician, I guess I belong to the group that we don't allow second chances; we don't allow rescue procedures for analysis.  If you miss it on your first shot, we don't bail you out, although I think the sponsors were well motivated to try because it certainly looks like a very interesting device.

However, as a patient or as a consumer, I am certainly interested in rescue procedures.  I guess most of my remarks tend to be on the philosophical.

I agree, if you accept the design of the study, the OPC pretty much ropes you in and you don't even need to think about it.  I agree, I thought both the sponsor and the FDA made excellent, clear presentations of the data, with obviously some differences in interpretation.

Of course, the subgroups, you know, we don't let you look at subgroups, but, of course, you have to look at subgroups.  You have to; they are interesting.

I guess, what does it mean to have an overall OPC if, indeed, the efficacy is intimately linked to the composition of the study population?  Now I don't know how much of the variability would have been known ahead of time, but I understand in the radiofrequency literature there's also variability of efficacy of ablating these three different entities.

So I guess from a philosophical point of view, how do you come up with an overall OPC when you have heterogeneity that may influence what the result is in any given population?

There was a little bit of vagueness in the way the protocol was written, but I think, to be hard-nosed, you have to accept that the overall results are what count.  However, this thing about the rescue procedures, I mean, as a consumer, you know, if I am being offered a procedure that has much better safety in some sense, but I agree with the comment that that's not necessarily been shown, then I might be willing to have a lower efficacy, initial efficacy rate, if I can also, then, get the RF procedure as a backup.

I guess it has sort of surprised me that we weren't looking at the overall success rate of the strategy.  Now the RF procedure, is it done in the same catheterization or do you have to come back later for it?

DR. FRIEDMAN:  It was done at the time of the same procedure.

DR. BAILEY:  I presume it adds quite a bit of time though to the overall procedure.

DR. FRIEDMAN:  It depended on the patient.  Some patients who failed cryo also failed RF.  They, obviously, tended to be longer procedures.  Some had a successful RF very quickly, and it was a short procedure.

DR. BAILEY:  I had thought that the complications weren't related, but, as I understand, they could be related to the length, overall length, of the procedure.  So it would be important to know whether the complication rate was, indeed, related to the length of the procedure and, indeed, to the performance of the back-up procedure or two procedures rather than just one.

But, subject to that, I would think that, from the point of view of seeing where this fits in or if it fits in, it is not really fair to compare the initial efficacy to the efficacy of radiofrequency ablation, unless you are proposing this procedure as a substitute for it without the opportunity to perform RF ablation.

But if you are proposing a strategy, then now that may imply that we need a higher limit than 85 percent, but I think that, as a consumer, not a statistician but a consumer who happens to have a Ph.D. in statistics, that is what I would be looking at, is, you know, what's the efficacy, given that I am going to be able to have a rescue procedure, if necessary?  Not rescue, a back-up procedure.

That is why I would have been interested in -- well, we were given that initial success rate, and it was well over 90 percent.  I would have been interested in the success rate at three months, and so forth.

In terms of the safety issue, it is also amazing how different, again, the OPC is from the way the data seem to be that are most relevant, in that although -- again, that was probably from naivety -- I didn't realize that it is possible that it may be somewhat actually difficult to determine which complications are related to the procedure versus just the device.

But the one thing that stands out is that, in order to achieve your OPC, you had to have four events -- I'm sorry -- you needed to have a rate such that you were expecting four events, four complications, five was the limit, and then you actually observed seven.

I guess what is striking is how little tolerance there is between those numbers.  So I think this gets at the point that this is a very small study on which to base safety, given the vagaries of what causes these complications, and the extent to which they are dependent on the devices is unknown.  That is why either large numbers or a comparison group or something.

In terms of the randomized trial, I didn't follow the argument why the cryomapping prevents you from doing a randomization.  You can still look at the efficacy and safety of the strategy of cryoablation.  You could even randomize the use of cryomapping or not, if you wanted to, versus radiofrequency without the alternative of -- of course, on the cryoablation side, you would have the opportunity to use radiofrequency, if necessary.  Then you would be comparing efficacy immediately at three months and, of course, complications.

But I didn't follow the -- I heard you say that the cryomapping made it impossible, but I didn't follow the logic of that.

And I agree with all the comments that have been said, that the cryomapping really, in order to understand it, you really have to compare the groups that get cryomapping versus groups that don't, if you are trying to look at the efficacy of it from a strategy point of view, and probably even in this study, even doing that analysis is very fraught with the usual observational biases.

So those are my questions.

CHAIRMAN LASKEY:  If they were questions, do you want comments or questions or more thoughts?

DR. BAILEY:  They are comments that are questions.

CHAIRMAN LASKEY:  Oh, I see.

DR. BAILEY:  We should give the sponsors an opportunity to respond then, starting with --

CHAIRMAN LASKEY:  Why couldn't you do a trial?

DR. LEHMANN:  Well, it could be possible to do that.  I rose to just respond to a few of the numbers, if I could address those.

Just to the overall strategy of cryo and then RF, if cryo failed, resulted in a rate of 97 percent for every study subject, and in each instance RF was undertaken in the same procedure.

Any measure of adverse events was not correlated with either the number of cryoapplications, the number of cryomaps, the number of cryoablations, nor procedure duration.

In descending order of certainty, we have 151 IDE subjects in this trial with no AV block.  In a prior IDE trial of a 9 French device, otherwise essentially identical to this one in its mode of action, and two trials done in Canada under essentially the identical protocol, using the same monitoring group and submitted to the TPT, and filed with the FDA in the PMA.

So we have almost 300 very-carefully-observed patients followed for many months, and the number AV block in that population was zero.  Zero out of 300 is, back of the envelope, is about 1 percent upper limit.

Now with lesser degree of certainty, there's another 300 AV node reentrant tachycardia, AV reentrant tachycardia subjects in a European registry, where they do get more attention than normal.  It is not carefully monitored in a classic sense, and that is about 600 no AV block.  Then, of course, there is the commercial experience, which, as we all know, has its issues on reporting.  So those are some numbers in relationship to some of the remarks you make.

CHAIRMAN LASKEY:  Thank you.  Dr. White?

DR. WHITE:  You're not going to take a turn?

CHAIRMAN LASKEY:  I'm not allowed to speak.

(Laughter.)

DR. WHITE:  I would like to echo, without repeating the words, what Dr. Krucoff said.  I think we hold very similar opinions about understanding the magic or the theoretic nice pieces of this procedure, but wondering why the device didn't actually perform.

Could I just ask a specific question about the IFU, about the instructions for use?  Could you clarify for me whether it is possible for this device to be connected to an RF generator?

DR. DESMARAIS:  The answer is no.

DR. WHITE:  Then why did you take the time to write it in here?

DR. DESMARAIS:  I think, you know, you want to address from when we do a risk hazard analysis, which is a technique used in the industry when you look at all the potential hazards.  By design we've mitigated that, but it's always there and a good practice to put all kinds of warnings and precautions to address the risk analysis that you conduct.

DR. WHITE:  Is the criteria for LV ejection fraction being greater than 35 percent, I think, in your protocol, is that necessary for clinical practice?

DR. DESMARAIS:  I will ask Dr. Friedman.

DR. FRIEDMAN:  As a practicing electrophysiologist, I wouldn't see that as an exclusion.  I think for the purposes of the trial, the intent was to identify healthy patients who were less likely to have co-morbid conditions.

DR. WHITE:  How do you feel about IVC filters and working through them?

DR. DESMARAIS:  Dr. Friedman?

DR. WHITE:  Is that a problem or no problem for this device?

DR. FRIEDMAN:  With other catheters, radiofrequency catheters, we traverse IVC filters without difficulty.  The average dimension in this catheter is really 7 French, the same size.  I don't foresee a problem.

DR. WHITE:  And for aortic valves retrograde, do you mean mechanical valves or do you mean a pig valve as well?

DR. FRIEDMAN:  No, we don't think we would advocate traversing retrograde a mechanical aortic valve.  Those patients would have to be done transeptally.  But for a porcine prosthetic valve or a native valve, this catheter could be used retrograde.

DR. WHITE:  Because the contraindication simply says any aortic valve replacement, you may want to be more specific about that.

In the demographics, in slide No. 46 that you showed, it struck me that you had excess of women in the trial.  Is it that women have more of these arrhythmias than men?

DR. FRIEDMAN:  That's a very good question actually, but that mirrors clinical practice.  Women present with AVNRT and AVRT more commonly than men.  I don't know whether it is a genetic difference or there is a difference in diagnosis.  Maybe men don't complain of palpitation or are less troubled by the palpitation.  There could be a host of explanations, but --

DR. WHITE:  But you would agree that percentage would be representative of the clinical practice.

DR. FRIEDMAN:  I think that is representative of clinical practice.

DR. WHITE:  And the minorities were not represented in an adequate fashion here.  I mean, is there a reason why American Indians, Blacks, Hispanics were not represented in percentages that are reflected in our population?

DR. FRIEDMAN:  I don't have an explanation for that, but I would point out that overall it was a fairly small trial, and you might not expect to see it reflected in those small numbers.

DR. WHITE:  We have heard today, I think, a lot of reasons to think that, because the device attaches and because the device has a smaller footprint, that there may be some theoretical safety benefits, but I am not sure I am convinced that the device has been proven to be safer because it has not been directly compared.

As I think Dr. Page mentioned, I think some of the complications which are not catheter-related certainly appear to be procedure-related.  And I think that if, as you are comparing the patient who is going to have one or the other of these procedures, those complications need to be taken into account.

The other issue is that I am not an electrophysiologist, so I am hearing so much about AV block and how terrible this is with RF, but in the paper that you provided us in the Panel pack, the Calkins paper that looked at the thousand patients with radiofrequency ablation, the incidence of heart block was only 1 percent in that population.

So I guess I am not being struck with -- I am not convinced that in 151 patients that you have actually convinced me that AV block will happen in less than 1 percent of your patients.

DR. FRIEDMAN:  I think there are a couple of points to mention there.  First of all, if one actually does a careful literature search, the incidence ranges from 1 percent even to as high as 14 percent in some series actually reported from fairly busy laboratories with experienced investigators.  So 1 percent is probably the lower bound, if you will.  It is the best that you would expect.

The other point I would mention is that it does not include patients who were referred for AVNRT ablation, who may get one or two ablation attempts, and then because of proximity to the compact AV node, the procedure is stopped.  I have seen patients, like every one of us who does these procedures have patients like that.  They don't develop AV block, but they also don't have a successful procedure, and they're not reflected in those data.

DR. WHITE:  The other thing I would like to ask you about is in slide 64 and 66, which were the Kaplan-Meier curves for long-term success.  It appears that the patients drop out of your Kaplan-Meier curve, and it is not clear to me why you haven't retained the patients in the Kaplan-Meier curve.

You've got at six months, for the long-term success of all subjects, 77 patients being measured in the Kaplan-Meier curve, whereas your flow sheet at six months you've got, I believe, 119 patients who were available.  It is slide 64.

DR. FRIEDMAN:  Well, I will see if we can find the slide.  Sixty-four.

DR. WHITE:  Actually, 64 and 66 both show it.

DR. FRIEDMAN:  Right.  So just bring that one up.

DR. WHITE:  You see at six months you have 77 patients when in your flow sheet here you have accounted for 119.

DR. LEHMANN:  First of all, here's the three-month point, starting here, and the effect is stable.  This six-month telephone followup, here is a dispersion, because there was a window of following the subjects.  The software, of course, takes the exact number as of the exact point.  When you do a live table, you take the inferred endpoints.  So you've got all of these subjects with the nominal six months, but in this subject attrition line it stops dead right there, losing -- so that is why you do both of the assessments.

But really with this curve you can see that there has been absolutely no change, and all of the six-month follows are represented, although they don't come to that number.

Have I addressed the question?  No?

DR. WHITE:  Am I dumb as a rock or what?

(Laughter.)

I mean, how many patients got measured at six months?

DR. LEHMANN:  A hundred and twenty something.

DR. WHITE:  Then how come it says 77 at the bottom?

DR. LEHMANN:  Because if the phone call occurred at six months and one day, then it doesn't show up in that six-month number in the Kaplan-Meier; in the survival analysis it measures every day as a distinct event.  In the life table analysis it measures the increment, the normal increment.

DR. WHITE:  All right.

DR. BAILEY:  So, technically, it shows the status of that patient.

DR. WHITE:  That is not the way I do my Kaplan-Meier curves.  I don't understand them, I guess.  Are you telling me that's right?

CHAIRMAN LASKEY:  Dr. Bailey?

DR. BAILEY:  No, in a way, I think you might want to consider interval censoring there.  If you assume that you don't get recurrence that goes away, if you just assume that the later time point is representative of the earlier time point, you could sort of back it --

DR. WHITE:  I guess I would like to be assured that all of the patients have been followed and that their events are being accounted for.  So when I see 40 percent of the patients not being accounted for, I am concerned that you are not seeing them.

DR. LEHMANN:  No, I forget the exact number, but over 95 percent of the subjects with acute procedural success were followed.

DR. WHITE:  Right.  Well, no, they are here in your flow sheet.

DR. LEHMANN:  Yes.

DR. WHITE:  And they are identified here.  So I am just wondering why they are not showing up.

DR. LEHMANN:  Well, go back to the live table.  Go up to the live table.  I think if you look at the life table, here you see we start with 136, and these are the subjects that are lost.  So by month six we have lost three out of 136 from analysis.  There's 122 who remain successful --

DR. WHITE:  So does it strike you as strange that your graph doesn't show 122 at six months?

DR. LEHMANN:  It doesn't strike me as strange at all.  That is just the way it works in the two different analyses.

DR. WHITE:  Okay.  What do I know?  I'm dumb.

CHAIRMAN LASKEY:  No, Chris, you're just rebelling against how computers constrain our thinking.

(Laughter.)

That's the output of the tests.  Thanks.

DR. WHITE:  Yes.

CHAIRMAN LASKEY:  George?

DR. VETROVEC:  Very briefly, was there any evidence that procedure time got shorter by the number of procedures done?  In other words, in the second half of the study, is the average procedure length less?

DR. LEHMANN:  We haven't assessed that.

DR. VETROVEC:  And then in terms of training, the observation was made that, "I learned to do a hundred of them and then I could teach the other fellows in four or five procedures."  What is it that you have to teach them that is unique about this, because this is going to be important, if you come down to, what do you specify as a training issue?  What is it that's different?

Even if you couldn't document a learning curve, there was something, it sounds like, you had to train the other personnel in doing.  What was it?

DR. DESMARAIS:  I will let Dr. Dubuc answer that, specifically what has to be done for training of the person.

DR. DUBUC:  I think you can call it training or teaching.  People doing this procedure with this new technology, they have to acknowledge that it is different from RF, and that is the only point here.

The thing is, like I said previously, when you do AV nodal reentry ablation and you are going into a slow pathway area, you expect the technology to produce, when you are successful, to do rapid junctional rhythm.  This you don't have with cryo.  So this is different.  So the operator has to know more than that; their catheter will stick to the wall, adheres to the endocardial surface.  So during that time you can do pacing or program stimulation, which you cannot do with radiofrequency.  So this is different.

And, third, you have the attachments.  So you cannot move the catheter during the position, and the catheter is very, very stable.  It will stay there, not like RF, when you have this slippage from the area.

So those are the three main reasons, and the people doing the procedures, they have to acknowledge that it is different from RF from the beginning.  So it takes time because you change the way you ablate and you make your procedure.

DR. VETROVEC:  I don't disagree with you, but it just is an issue in terms of, how has this been disseminated into the practicing population?

DR. DESMARAIS:  In that respect, I will put up just a few slides.  I don't even need that.

But we believe that through instruction for use, operator manual, and we believe that there is a need for didactic training, as Dr. Dubuc just alluded.

We will utilize skilled physicians experienced with Freezor to train and initiate at new sites.  We have learned that from our experience in Europe, and we will provide ongoing support with clinical support specialists, which we are doing in Europe as well.  We are convinced that this is a need.

DR. VETROVEC:  Are you talking about physicians proctoring this or are you talking about --

DR. DESMARAIS:  Yes, yes.

DR. VETROVEC:  And how many procedures do you anticipate as necessary?

DR. DESMARAIS:  That is very difficult.  It depends on the site in terms of the volume, the number of physicians.  So it is a very difficult to estimate at this point in time.

DR. VETROVEC:  I don't have anything else.

DR. DULLUM:  You said that in the ice ball you saw after three minutes stayed stable, so that is why you chose four minutes.  In surgery there is usually a dry field in this aspect that a little bit was talked about, that there is going to be warm blood going by.  So that was monitored, and you still know that at three minutes with warm blood that it stays the same?

DR. DUBUC:  Well, we know in a beating heart, which could be different from surgery sometimes, in a beating heart we know that it is very stable after two minutes.  Even if you keep the catheter there for six, seven, or eight minutes, there will be no change.  But if you go in the preparation, like a fine muscle preparation, you will see that the ice ball will still grow after three minutes.  The longer you will keep the catheter there, the bigger will get the lesion, which does not happen, actually, in a beating heart.

DR. DULLUM:  So you basically have a cool  sink.

So did I understand that it took an average of seven-and-a-half cryoablation procedure times with your catheter?  Is that the same with RF?

DR. DUBUC:  I would say it is about the same, yes.  It is the same range.

DR. DULLUM:  So are you going to recommend that in clinical practice then, when do you know when to stop?  Do you say, okay, after seven-and-a-half or do you just keep going until you get tired or the patient gets tired?

DR. DUBUC:  I think it is like our effort goes with the clinical judgment.  I mean, you are doing the procedure and sometimes you know that on the average it will take, four, five, six lesions, but sometimes it takes more.  I mean, even with RF, I have seen cases you do 30, 40 applications.  So that can happen also with cryo, I imagine.  I mean, that goes with the clinical judgment.

DR. DULLUM:  And my last question:  Are you going to recommend that if you go with cryo and you don't ablate, then you switch to RF, or vice versa?  I think, as someone said, there was an RF that was not successful and then you used cryo.

DR. DUBUC:  I think we had no safety problems with that in this clinical trial, doing RF after cryo.

DR. DULLUM:  And vice versa?

DR. DUBUC:  Well, we did not -- just to correct, there were no subjects in the study who got RF first.  Everyone, out of 166, 164 had cryo first.  Some failed; some succeeded, and some of the failures were RF.  Two of the subjects did not have a cryoapplication; they had RF alone without any cryo following.

We did look at the complications, any measure of adverse events in the group of cryo and followed with RF, and there was no difference between those two groups.

Furthermore, in terms of the number of cryoapplications, either mapping or ablation, they had no relationship with measures of adverse events as well.  So that bears a little bit on some of your questions.

CHAIRMAN LASKEY:  Thank you.  Mark?

DR. HAIGNEY:  Okay, I will be brief.

I am still confused about your indication.  You are asking for an indication for parahisian pathways?  You're not?  Just AV node?

DR. DESMARAIS:  As presented today, the indications for it --

DR. HAIGNEY:  AV node --

DR. DESMARAIS:  Correct, AVRT indication.

DR. HAIGNEY:  Okay, so not mid-septal pathways?

DR. DESMARAIS:  Correct.

DR. HAIGNEY:  Okay.  On page 62 of the final report, there is something that really confused me.  It said, under adverse events, it says, "Thirty-one of 42 Canadian subjects and eight of 124 United States subjects were reported with these AEs."  Is that what I think it is saying?

DR. LEHMANN:  We believe that the three Canadian sites had a very different view of what constituted an adverse event.  No, we don't feel there was any -- other than sequelae that would indicate they had had a much different means of accessing the circulation, and I know it was quite consistent across the Canadian sites.  So we think it is a reporting anomaly.

DR. HAIGNEY:  Well, it is a reporting anomaly with somebody.  Whether it is with the Canadians or the Americans, they're not going to say.

Finally, I have one more weird comment to make.  Have you considered as a possible contraindication patients with known cryoglobulinemia anemia?  It just occurred to me over breakfast, and it seems to me that that would probably be -- if you had a patient who had known cryoglobulinemia, that it is probably not the technique you would want to use.  So I would suggest putting that in as a contraindication.

On the whole, I think the device represents an important contribution potentially to the progress of EP.  It's a niche device, and it doesn't replace RF ablation.  So I am not as bothered by the fact that you didn't make your OPCs.  I agree that there are no clear device-associated complications.  So given the fact that this is not replacing RF ablation, I think I'm inclined to be a little more lenient than perhaps Dr. Krucoff.  I usually am.

(Laughter.)

CHAIRMAN LASKEY:  You're not supposed to spill the beans until after the break.

(Laughter.)

We're about to take a short break, following which we will reconvene and do the voting.  However, I would like to raise one point here in my prerogative.

I think in the future -- and I think it is unfortunate in this trial, but in the future when we have these sorts of interventional trials requiring prolonged fluoroscopy, we should have some measure of dose.  I think all of the agencies around, there's none that I can think of that should be more aware of patient dose.  Procedures that have up to whatever the estimate is, many, many minutes of fluoro time, that is not a surrogate for dose.  I think we should require that of the studies down the road, because we are going to see more and more prolonged laboratory time.  So let's build that into the protocol.  It might even be a measure of safety as well.

Thank you.  So let's take a 10-minute break.  Let's reconvene.  Then we will go over the questions to the Panel and the Panel's preferences.

Thank you again, Panelists, for sticking to the schedule.

(Whereupon, the foregoing matter went off the record at 3:48 p.m. and went back on the record at 4:05 p.m.)

CHAIRMAN LASKEY:  Good, thank you.

At this point I would like to proceed with reviewing the questions put to the Panel, if we can have them up.

As we go through these, I will do my best to summarize points of agreement and disagreement amongst the Panel members.

"The results of this clinical trial were compared to objective performance criteria, OPCs, established for the study for both safety and effectiveness.  The OPCs were determined from the radiofrequency ablation medical literature."

The first question to the Panel with respect to safety:  "The safety endpoint was the occurrence of major complications, as defined in the study protocol.  The FDA interprets the definition of major complications to include all adverse events requiring treatment which occurred within seven days of the procedure.  The upper 95 percent confidence bound for the major complication rate was 8.5 percent.

"This exceeded the safety OPC, which specified an upper 95 percent confidence bound of less than 7 percent.  Please comment on the following:

"a.  Please discuss whether the results of the clinical study provide a reasonable assurance of device safety for the intended patient population."

To paraphrase Dr. Bailey, I think that you certainly were roped into these constraints by the OPCs in that there is not a lot of wiggle room in a small study.  One patient either way could have made the difference, and, unfortunately, it didn't work in your favor.

"b.  Please discuss the applicability of a safety OPC for cryoablation which was based on reported clinical experience with radiofrequency ablation."

I think that the Panel alluded to this in some of their comments.  That is basically all we have to fall back on as a benchmark.  It may or may not be appropriate, but that serves as a benchmark and it is a not unreasonable benchmark.

Question 2, with respect to effectiveness of ablation:  "The device did not meet the effectiveness OPC for the overall study population or for any patient subgroup.  The lower 95 percent confidence bound for acute success for the entire study population was 76 percent.  The OPC for acute success specified a lower 95 percent confidence bound greater than 85 percent.

"a.  Please discuss whether the results of the clinical study provide a reasonable assurance of effectiveness in (a) the overall patient population or (b) in any individual patient subgroup."

I think that both statisticians have spoken cogently to this point.  Again, you are constrained by your boundaries.  One patient either way could have made the difference.  Again, it's an awfully small study to allow for that kind of non-leeway.

With respect to the individual patient subgroups, I think that the Panel has done a good job, as well as the FDA statistician, at pointing out the foibles of doing post-hoc subgroup analysis.  It's really treacherous, and you really live or die by the limitations of this technique.

"b.  If the clinical trial does not provide enough evidence of effectiveness, please discuss what would be needed."

I think we need to hear really how people are voting before we can give you an answer to that.

DR. ZUCKERMAN:  Could we, for Question 2, which really gets to the heart of the matter, could we hear from several other Panel members?  Is there a consensus on this summary as you presented it?

CHAIRMAN LASKEY:  Let's go back to a and b with -- both parts, a and b?

DR. ZUCKERMAN:  Yes.

CHAIRMAN LASKEY:  Okay, well, feel free to chime in, Panelists, if I didn't quote you correctly.  Cindy?

DR. TRACY:  I would like to jump in, because this is sort of the crux of the struggle with this thing:  What would it take -- can you construct a study that would look at AVNRT or do the data that we have -- how close are we to defining safety and effectiveness in AVNRT?

I have whole issues with the safety, given the way there is such a variance in reporting between the Canadian side, just the variance in reporting in this study.

But, anyway, in terms of effectiveness, what has to be done?  I mean, are we close?

CHAIRMAN LASKEY:  So this is "b"?  This is really 2b?

DR. TRACY:  "Or not to be."

(Laughter.)

CHAIRMAN LASKEY:  Please discuss what would be needed, because we have before us a negative trial.

DR. TRACY:  Dr. Bailey was struggling with this.  Maybe he can struggle some more for me.

DR. BAILEY:  Well, yes, as a true/false question, the answer is false, but if it's an open-ended question, I mean the data do seem to support efficacy, but only if you look at the strategy of cryoablation with the possibility of RF as a backup.  That's to say nothing of the safety, but in terms of efficacy, you get 161 successes out of a 166, which seems pretty reasonable.

DR. WHITE:  Well, but the problem, though, Ken, is that's not what we were being asked to judge because you have no comparer.

DR. BAILEY:  Right, but that's why I say, if it is a true/false question, the answer is false.

DR. WHITE:  The question is, what Cynthia is asking is, what would it take to make us say that they have satisfied efficacy?  Is it because there were so few patients that the margins were so narrow and the lower bounds were not met?  If there were another hundred patients enrolled, would the bounds narrow?

DR. BAILEY:  No, no, I think the OPC as stated doomed them to the result because that was not a quirk of sample size.  You can't get that high an efficacy with this technique.  But in terms of the overall clinical application, it's pretty high.

DR. KRUCOFF:  I think that is a really key point.  I think we have actually heard it stated several times, but it seems like when this trial was first designed, that the catheter behavior and performance was expected to be the same as an RF catheter.  Yet, it sounds more like the actual experience that's evolved out of the trial is that the catheter performance may be a little inferior to an RF catheter.

When it is, you can bail out with an RF catheter, but the other safety elements and the potential to pre-map and to protect more vulnerable areas, that's what gets actually into an area where you're suggesting a trial design that would look at a potentially slightly inferior efficacy from the device, but at least an equivalent behavior in an intention to treat, which includes the RF backup, where overall the safety would be better.  I don't see that it would be very complicated to design that kind of a trial.

The trouble is that that's not how that OPC was designed.  That's, unfortunately, I guess the question that we're going to be asked today in 2a, and then in the 2, bottom "b," I think it would be pretty straightforward, based on this experience, to see a design where you would do an intention-to-treat analysis for efficacy and really concentrate on safety, and whether these are catheter or dwell time or radiation time safety issues or not.

DR. WHITE:  But without a comparison group, I have a hard time understanding the safety --

DR. KRUCOFF:  No, not in a randomized trial.

DR. WHITE:  Because I don't see the safety in 151 patients.

DR. GILLIAM:  I don't think you can see the safety in 150 patients, but I also don't think it would change if they had 650 patients or a thousand patients, because the comparison that we really are talking about is looking at, for instance, use of radiofrequency in an AV nodal reentry group and skilled operators today.  I mean your heart block risk, if you will, is something less than 1 percent.  I mean it is.

If we start looking at the comp. we're seeing, no one is saying that there is something intrinsically not safe with this catheter.  I mean we are getting complications because they only had 150 patients, and if we probably take 150 patients from any of the procedures, we may have like complications.  These are procedure-related but not catheter-related.  It may very well have been RF cases we were looking at to get the same type of complications.

I think the reality is that this is a procedure that potentially may be less likelihood of creating complete heart block in an AV node reentry population, and, ultimately, the overall efficacy of this one alone may not be as good, but taken with the option of using RF, it may provide an additional tool.

That's saying a clinical trial -- I'm not sure what we're going to compare it with.  I think it would be very difficult to provide such a study.  I don't see how it could be done easily.

DR. WHITE:  Well, I don't understand why you say that because there is a standard clinical procedure now to treat these patients.  So you take the experimental procedure, you compare it to the standard clinical procedure, you get a comparison.

DR. GILLIAM:  If we were going to replace the procedure, I would agree with you.  But I don't think that anyone would suggest that cryoablation is intent to replace radiofrequency ablation.  I think that that's where I think it would be difficult because it may very well be that you can go in and just look at RF, you know, first or second.

You may go in with cryo first, and if you find an easy point, you map and you get it right away, then fine, you may just go with it.  If you need a bigger lesion, if you will, or you just use RF, I just don't think that you are going to be able to compare one to the other because they're not going to be used the same.

I don't think you're going to replace RF with cryo.  It's a smaller, more discrete lesion.  That means that you're likely to have less efficacy, I mean all things being even.  But, on the other hand, potentially that one or two people that surprise the heck out of you, that when you step on the pedal, they all of a sudden develop heart block, and you jump off the pedal real quick, you know, if you're lucky, as we most of the time, frankly, are -- I mean that we don't get permanent heart block.

But every now and then you look at it happen, and you say, "Oh, my God, I wish it was reversible."  Maybe this is reversible.  You know, I don't think they've proved that, but I don't think they've proved that it is any more dangerous.

CHAIRMAN LASKEY:  But, again, we're not here to talk about replacement.  We have heard the word "supplement" and "adjunct to."  Whether that lends itself to non-inferiority versus class RCT is, I guess, what Dr. Zuckerman needs to know here in Question 2b.

DR. VETROVEC:  But let me ask you, if the goal, based on what you say, is to look at this as something to prevent it, then it will become your first choice.  Otherwise, you're going to have heart block the whole time.

So now you're suddenly substituting this as a first choice, and RF is becoming second choice, and do you know that that's a fair trade?  That's what you're suggesting, if you're going to use a hundred times to prevent one or two heart blocks.

DR. TRACY:  Do we have data that -- is there a way to look at this data for non-inferiority?

CHAIRMAN LASKEY:  Which data?

DR. TRACY:  The effectiveness data.

CHAIRMAN LASKEY:  The way I look at this trial, this was a non-inferiority trial.  That's the study design to me.

DR. KRUCOFF:  Well, it didn't make its boundaries on --

CHAIRMAN LASKEY:  No, exactly, but --

DR. KRUCOFF:  It is a negative non-inferiority on a safety compared to an historical.  So a classic non-inferiority, this is not a randomized dataset.

DR. TRACY:  Effectiveness, not inferiority.

DR. GILLIAM:  But if we're looking at heart block, let's just look at just heart block.  Then, in effect, this study doesn't --

CHAIRMAN LASKEY:  We need numbers.  We don't have numbers here.  There's no precision of the estimate, and maybe I guess, Dr. Page, you said something during the break about adding some precision to the number.  We know it's not zero.  Nothing is zero.  But could it conceivably be on the same order of magnitude as RF and is it, therefore, as safe from that standpoint?

DR. PAGE:  Yes, I would just like to comment that, first of all, I think in all fairness, as Rosie Gilliam said, in good hands the risk of heart block, and I would say there are a number of good EPs in this room, and please speak up if your heart block is over 1 percent, but most of us are less than that, I believe.

That being said, I believe that the cryo is probably as low or lower in terms of heart block.  I see this as another arrow in the quiver, if you will, as being supplemental.

In addition, one thing that has to be mentioned is we've focused on the specific indication for AV nodal reentry, but, in addition, to have this other tool in terms of cryomapping is going to be valuable in a number of circumstances.

If I understand the indication as it is written, we are not limited to the device being used in the AV node reentry for the cryomapping, is that correct?

DR. ZUCKERMAN:  No.  Dr. Page, your comments are well-founded, but I think you're getting into Question 3.  Perhaps if we could go back to Question 2 and take it in an ordered fashion, it will help us decide what the trial showed and what can be done with this device.

I would like, Dr. Laskey, first of all, to know, for Question 2a, the overall patient population.  Is there a Panel consensus as to whether there's reasonable assurance of effectiveness?

DR. PAGE:  Just so I'm clear, is effectiveness only in successful ablation, having nothing to do with the second indication, effectiveness as a mapping tool?

DR. ZUCKERMAN:  That's correct.  Right.  For acute procedure success; we would all agree that they didn't meet it for the overall patient population.

CHAIRMAN LASKEY:  I think we all agree with that.

DR. ZUCKERMAN:  Good.  Okay.

CHAIRMAN LASKEY:  It would be hard to say otherwise, yes.

DR. ZUCKERMAN:  Now Question 2b asks -- a subgroup analysis was done.  You've heard multiple comments as to whether or not the subgroup analysis is valid for AVNRT.  It's important for the agency to get a consensus, if there is one, or just a response from the Panel as to the validity of the result, of the effectiveness results for the AVNRT population.

DR. BAILEY:  I thought your statistician made the excellent point that the OPC was for the overall mixed population, and we don't know what would have been specified, had it been this one subgroup that was recruited.

CHAIRMAN LASKEY:  As a non-expert in this area, though, it did look as though the AVNRT group is the group most likely to do well of the three groups anyway.  So, therefore, the bar would be even higher.  So that we would expect this procedure to do better in that group which tends to do better.

So I think that you need to sit down with the applicant and go over the applicability of OPCs.  There must be data in the literature on the success rate of AVNRT only rather than all-comers, and that may be helpful, although retrospective still.

DR. ZUCKERMAN:  Okay, that's one way to look at it, but the electrophysiologists here in the Panel have offered the suggestion that these two devices, cryoablation plus RF ablation, could be used in a treatment strategy approach because of a comfort level associated with cryoablation.

So another approach which is similar, Dr. Laskey, to what we do with guide wires to cross-total occlusions is, if we do believe the first device has some intrinsic merit, might be safer, we may accept a lower success rate with a first device, and then if it doesn't work, you go to your next device.

So I'm not sure I have heard from the electrophysiologists that for the subpopulation called AVNRT we necessarily need a higher bar, given the risk/benefit profile.  Can the electrophysiologists comment?

CHAIRMAN LASKEY:  You don't need a higher bar.

DR. TRACY:  I'll speak now as an electrophysiologist, and I would be happy with this level of success, knowing that there are certain patients with very narrow -- there are some people with very narrow anatomic windows between the slow and fast pathway, that this would be a very nice thing, very stable catheters used.

There would be certain circumstances where I would use this as a first line, and I would be very comfortable knowing that I had a backup of RF, and in my lab the chance of heart block is less than one-half of 1 percent.  So it's not high, but in those patients that have the funky windows, those are the ones that, if they are going to get into trouble, those would be the ones that would have the problems with heart block.  So there's a group of people where this device would have a distinct advantage to use.

So if it is a 91 percent success rate with this device, that's not that far off from reported data on other AVNRT studies, although our success rate is 98 percent.  If I took this catheter and put it in a heart, what my success rate would be -- presumably, it would be better after the tenth case than on the first two or three.

To me, it's in shooting range of what would be acceptable for any device.

DR. WHITE:  But the problem I have is that there is one option to approve this device for -- what is it called when there is a very narrow indication for this, like we use for the atrial septal closure devices?

MR. MORTON:  Those are HDEs, and the population is way --

DR. WHITE:  Right.  Well, you're talking about a very small population that you would choose this device for.  So it would be a very narrow window.  But when this device fails, the OPC, the lower bound, then I don't understand how we can say that we're in the neighborhood and we ought to let that go.

I don't see the purpose of the OPC if you're going to ignore the OPC's lower bound.  I mean, once it is set, once it is agreed and we proceed with that, then I find it a specious argument to work around the number, if the device fails the lower bound.

MR. MORTON:  I think that's the hardest thing about this, is that, obviously, there's a patient population who it seems to be a real consensus would benefit from doing cryos first, but if, as a plumber, I'm understanding what I'm hearing, that patient population isn't identifiable upfront.

DR. TRACY:  It is.  In short order it is.

MR. MORTON:  It is?  So when you get into the lab, you can see when the window's --

DR. TRACY:  It is.  Once you have the catheters in place and you've done your baseline study, then you know who you would pull which catheter out for.  And if you had both available in your lab, you would have a very clear distinction which catheter would be appropriate to use.

In terms of the OPC, I think that the OPC is not appropriate.

DR. GILLIAM:  I agree.  I think that the question isn't -- I think the OPC is a high standard, which is good.

DR. WHITE:  A high standard?

DR. GILLIAM:  I think it's very high.

DR. WHITE:  For AVR?  I mean, I don't think you can -- that's not true.  I don't know how you came to that conclusion.

DR. WALDO:  You know, I think if you're introducing a brand-new concept -- ablation is not a new concept.  I mean if you're introducing a brand-new concept, if you're doing a mated one, or maybe if suddenly you're taking a new technology and introducing it to a new treatment mode, you have to be extremely -- well, you should be rigorous always, but this is something very different.  This is ablation.

We have several tools for ablation.  There is not just one RF ablation technique.  There are many RF ablation techniques, from the size of the catheter to cool tip, to, you know, saline irrigation.  There are all kinds of ways to do these sorts of things, but the concept is clear.  This is a concept of ablation.  It's another way to do ablation.

I think we can even get some information from Dr. Dubuc, because what if your laboratory switched to only --

DR. WHITE:  Well, I don't think that's relevant.

DR. WALDO:  I'm going to suggest an answer then.  If you don't think it is relevant, I'm going to think that they decided in their group -- I'm guessing because I haven't talked about it, but they decided in their group that it was worth doing the cryo technique because they found it very, very effective and they found it safe.

DR. WHITE:  I disagree.  He's conflicted.  He's a paid consultant for that company.

DR. KRUCOFF:  I think we've got to recognize that there's a difference between what these numbers are from the study and what these numbers potentially mean when they are applied to a population.  Kent, please correct me if this wrong, but I think that's really what boundaries are about.

The OPC historical control creates a non-randomized clinical trial venue.  What that means is that the 91 percent doesn't mean -- I mean it's 91 percent for this 130 patients; when you get to 136,000 patients, the boundary is what you may see.  You may see that's 97 percent or the upper boundary or you may see that it's 83 percent or the lower boundary, and that 83 percent, if that's what happens in 136,000 patients compared to your 98 percent with radiofrequency ablation, that is my understanding of what an historical boundary branded number is telling us for application of population.

That's, to me, what is so difficult and what we seem to be wrestling with.  Ninety-one percent sounds really good, and certainly if it's safer and there's a population who would benefit from it, you know, that's where the consumer side of a Kent Bailey steps up and says, "Maybe that's what I would like to have used."  But 91 percent is what happens in this 130-ish patients.  The boundaries tell us what might happen at 136,000 patients, and that's why if the trial design pays attention to the boundaries rather than to the number -- and somewhere we've got to figure out how to digest that into whether this is just another neat arrow for the quiver or whether what this trial may be telling us by failing its boundaries both in safety and efficacy is that, when you really get out there and start using this thing, compared to what you're already getting with radiofrequency ablation in your lab, there's going to be a bigger gap than the 91 percent kind of number indicates.  That seems to me to be what we are wrestling with.

So we've come full circle.  We basically all agree that they failed to meet all their pre-specified endpoints, and we're not being terribly helpful here to the agency or to the vendor.  But I would have to say, if we did 136,000 or million, the confidence limits would shrink.  They wouldn't stay that way.  They would be awfully tight.  Is that right?  They wouldn't stay where they are?

DR. BAILEY:  The learning curve.

CHAIRMAN LASKEY:  But to get back to what your needs are, you know, the Panel is torn.  We have a body of scientific evidence before us which meets the null hypothesis, which has not been rejected.  Yet, as clinicians we have a gut feeling that this probably is safe.  If you take out the prosthetist and you take out a few other weird things that can happen, it probably is safe and it probably is effective.

But you're asking us, and we need to qualify that again and again and again, that we vote as clinicians and not as methodologists or statisticians.  If you want a positive study, it needs to redone in some manner with a larger sample.

Does that reflect --

DR. ZUCKERMAN:  That's helpful as a start because it helps us through 2a.  We have a difference of opinion as to what the data show in terms of effectiveness right now.

So then we go to 2b.  Given this difference of opinion, is there any consensus on what would be required in terms of replication of results or new dataset that could help everyone here concerned?

DR. WALDO:  Well, you're not going to get a trial of a thousand patients.  I mean you know that.

DR. ZUCKERMAN:  No, but I don't think the agency would require a trial of a thousand patients for RF ablation.  I think we have to go through the usual panoply.

One is you always start with a randomized trial.  There's a difference of opinion there.  But at the other end of the spectrum, do we need another registry experience where we can replicate a number that was developed in a post-hoc analysis to make everyone feel that we've reached a bar that's acceptable?  I mean there's always a wide range of designs when we're talking about device trials.

CHAIRMAN LASKEY:  And I, for one, would suggest that.  I think that if post-hoc analysis is all about hypothesis-generating, then certainly registry of AVNRT only, which is what the indication here -- it looks like that's the indication they're going for.  If that's the indication, then a registry of some reasonable number of patients, less than a thousand but greater than a hundred, would answer the question as to the safety and efficacy track record.

I think that we're sort of jumping the gun here.

DR. TRACY:  What registry are we talking about?

CHAIRMAN LASKEY:  A registry such as has been suggested to add to the qualifications of a vote, for example.

DR. TRACY:  You're talking about a registry that this catheter would be entered -- or patients would be entered into who had their ablations performed with a cryocatheter versus a registry that we would pull out of the shelf in the library and say, let's compare this data to a bunch of AVNRTs?

CHAIRMAN LASKEY:  Prospective.

DR. TRACY:  We're talking about a prospective though.

DR. ZUCKERMAN:  I would just take a step back.  I'm looking at the screen, and we've dealt with Question 2a, which suggests that there's a difference of opinion as to effectiveness.  So then we go to 2b, which is, how much more might be necessary to go to the goal line?

I think it's important for folks on this Panel to understand that it doesn't imply that needs to be a randomized trial.  It doesn't imply that needs to be a single-arm registry.  We're just looking for ideas here for those who feel uncomfortable with the present dataset, realistic ideas.

DR. KRUCOFF:  I think one thing that has been said a couple of times is that the size of the current dataset plus or minus one or two patients flirts source.  One way to consider going forward, it would seem to me, would be to focus on your indication population and sit down with a calculator and see how many patients enrolled in the identical protocol, but with just that one indication, treated in the same way, could be potentially appended to take away some of the flirtation with the boundaries and really find out whether you're at or better than the boundary or whether you're not.

DR. WALDO:  Can I ask a point of information?  In terms of safety, we heard that the events, when they adjudicated all the cases, that none of these things were device-related.  Isn't that my understanding?

Now I understand that there are procedure-related problems in all things, and the real issue, then, if I understand it correctly, is that, is the device indirectly the cause of some of the problems?  That's what you're struggling with, is that right?

Because I read through.  I pulled it out a couple of times, and I read through all the things.  I don't know how to answer that because I'm not a statistician, but as a clinician I thought a lot of those things, the countings we see anyway, most people don't look that hard.

I recently sat on an events committee where we found some things that were newly-recognized because they were looking very hard.  When we started looking at other patients who were not in this kind of study, they found the same thing these other patients had, a very, very good laboratory.

So I don't know, I would not like to -- I mean, what I hear a couple of things, and I think our Chairman has stated them very well earlier -- I really think there is a comfort level for me, and I'll speak only for myself, about the safety of this device.  I think it has a utility.

Then the question is, do we want to just disregard that at this point because it is a small study and there are -- I don't know if I'm using the right term -- secondary concerns because it's not the device per se that has been associated with the safety concerns; it's the procedure, and we don't know if the procedure of the use of the device or if there is something unique about the device and the procedure.  My suspicion is that it isn't, but I don't know that I can say.

I don't know if I have made myself clear, but I think we're focusing on numbers, and we're not focusing on what I think is the substance of something that we in electrophysiology deal with a whole lot, and that's ablation.  I think the ablation, the results of ablation are clear.

For me, a 91 percent and 97 percent and 96 percent, it's still an option for a physician and there are times you want to use it.  That's why I made that remark earlier.  I think that it's an option that I think is legitimate, and I think the data are sufficient to me to suggest that there's an adequate efficacy, and the safety problems that I see here are not the kind problems -- I haven't seen life-threatening things.  I've seen things -- you've proselytized the most obvious, but someone had a problem with a sheath in a diagnostic catheter.  I mean those things happen from time to time.  There are a bunch of other things like that.

So I think the forest and the trees thing here is what I'm talking about.  I think seeing the forest, I think this is something useful.  If you look at all the trees, then you haven't got a sense of what the forest is about here.  That's what I see.

DR. WHITE:  I guess my problem is that they failed on every single boundary.  They failed every one.  It wasn't like two were good and one was bad, and shouldn't we overlook that?  I mean, we're talking about rescuing the --

DR. WALDO:  But you're focusing on efficacy.  I mean, I don't think 91 percent and 97 percent is to be ignored, but we're not talking about life-threatening rhythms.  We're not talking about something terrible.

I think Dr. Bailey really had a very -- he saw it immediately, and he called it as a consumer, but I think really the approach, if you take the approach if you want to start with cryo, and if it doesn't work, go on to RF, you've got an enormously good success rate.  I think that is how we think as clinicians.  That's how they did it in the study.

You have to ask yourself then, why do OPC trials at all if you're not going to accept the data?

DR. WHITE:  I haven't thought about it forever, but just off the top of my head, I mean if you're breaking new ground, if you bringing in a whole new direction of therapy, if you're introducing your first implantable defibrillator or your first bimitricular pacer, talking in my own area of devices, I think you're breaking new ground.  But we're not really breaking new ground in terms of ablation.  We're playing with the technology of ablation; that's all.  That's how I see it.

CHAIRMAN LASKEY:  Well, but this is not a 510(k) either.  This is a new --

DR. WALDO:  Well, it's not 510(k).

CHAIRMAN LASKEY:  This is a new energy source.

DR. WALDO:  I mean, the safety issues to me are not safety issues that are out of my cage; put it that way.

CHAIRMAN LASKEY:  No, I would agree with you that these are procedure-related complications, but these are the chances you take when you do a study.  You roll the dice and that's it.  If you win, great, but in this case they didn't.

But they're procedure-related and they're often hard to disassociate from any device.  It's in the hands of the operator, and there's seven or so variables that go into procedural complications.  Nevertheless, I tend to agree with you that I think we do see the forest through the trees, but the trees are awfully big.

DR. TRACY:  Does anybody have a rough estimate of what it would take to achieve an effectiveness endpoint?  Is it five patients or is it 5,000 patients?

DR. PAGE:  If I may just ask, I think it troubles the heck out of me that an OPC trial was done and it failed, but in retrospect I don't think it was the right trial.  I think it was too small, and given what I believe the safety is of this device, then I think the bar was set too high in terms of efficacy.

Then we're looking back at the safety issue.  When we dissect out the safety issue, which I think is fair in this case, because one or two going the other direction changes the whole result, when we dissect out the safety, then it doesn't look like it is the catheter that's responsible for it.

So I feel like I'm totally schizophrenic here because, as a statistician, if I were one, I would not accept what I just said, but as a clinician I think we would be wasting our time to do another pre-approval trial because I don't know if this Committee could even figure out what our endpoint was, then much less go through all the time of running a trial.

I think a registry afterwards of this type of patient would be very valuable, the first thousand cases, and really look at it.

DR. ZUCKERMAN:  Right, but the problem is you need the data before approval to label a device for a certain indication.  So to respond to Dr. Tracy's question, though, I think we need to remember that with the approximate number of AVNRT patients that were in this trial, the lower bound was about 82 percent.  So if you were going to do a subsequent trial, where you could accept a lower OPC, for some of the reasons alluded to, you're in the same ball park, not as exactly, as to what was studied, what was retrospectively determined to AVNRT patients, around a hundred or so.

MR. MORTON:  Dr. Laskey, just a quick comment:  I know that we're all looking at the same screen, but one thing that I would emphasize that the agency and the sponsor have come before the Panel to ask is, is there proof of reasonable safety and effectiveness?  And that is exactly what I'm hearing you wrestle with.

DR. GILLIAM:  My concern is, just looking at the OPC, I mean I'm not so certain that if we were to do this study with 5,000 patients and it didn't meet the OPC, does it mean that we should not approve the device?  I guess at the end of the day we have to take a vote.

As Dr. Waldo said, I sort of feel that, if you were to say, "Is this device as effective as RF for ablation," then I would probably say RF is probably a little bit better.  I mean, it's a bigger lesion maybe, and maybe that's why it's better or maybe quicker, whatever.

But is this an effective treatment?  I would have to say, yes, it's effective.  Maybe it's not as good as RF, but it's different.  It's not the same thing.

So I think using the OPC standards we have for RF may not fully be applicable to this type of therapy.  I mean it's an ablation, but it's different.  I mean, maybe the efficacy is not as good, but it doesn't mean that it is still not effective.  I think that's the distinction that I make.

DR. HAIGNEY:  Would it be possible to have a registry that would be aiming at safety?  I think it's quite possible that this isn't as good as RF, but RF is darn good and it has been developed over 10-11 years.  If it turned out, after looking at a thousand patients, it had an 85 percent success rate, I would say fine, that doesn't bother me because there are certain patients who I want to use it in.

It would bother me, however, if we did a registry and we started getting more heart blocks.  Then I would say, well, wait a minute, maybe this isn't the thing I want to try in the area of the AV node.  In a registry, couldn't we design a registry that was a couple hundred patients with the express purpose of being sensitive to detecting AV node block and other -- the sort of device-related complications that we're all --

DR. WHITE:  If we haven't seen heart blocking in over 600 patients I heard, how are you going to find it?  This is going to be very hard to find.

DR. WALDO:  Well, that's terrific.  I don't want to find it.

DR. WHITE:  No, no, no, but I mean for pre-approval.  We already have 600 --

DR. WALDO:  You've only had 150 patients to look at.  You've had 150 European experience.

DR. WHITE:  Well, I would love to see that data.  I would love to see it, but it wasn't presented to us.  So I have 150 patients with a 1 percent incidence.

DR. WALDO:  If the prevalence is zero, the confidence is because the numbers are small, is what we're talking about.

DR. WHITE:  But it would only take one or two patients.  The next two patients have heart block, and all of a sudden --

DR. WALDO:  AVNRT has fewer patients than that, but I think we have to decide what the agenda of the Panel is.  If, ultimately, there's a regulatory process that's asking a question about an indication based on data, then that's one conversation.

I think you could probably put this catheter in the hands of any of the EP people sitting here, and in about five or ten cases they would know 99 percent of what we have been talking about today, but that's not data.

I think ultimately we have to decide whether there is information here sufficient to support a regulatory approval and/or indication or not, and if not, then what would provide that, I think is what I interpret to be the part of this question to allude to.

DR. WHITE:  I may be confused, but what I'm hearing from the EP group here is that this may not be as good as RF, but it has a niche.  But, on the other hand, what the sponsor, I think, asked the FDA was, how could we have this approved that it's at least equivalent to using the guidelines of RF ablation?

We didn't quite hit that, and that is the problem we are struggling with.  I have sat on a number of these panels, and I'm always startled by the fact that somehow the studies really never have the power to answer the question you would really like.  I think when you go back and look at these in the future, you hope that people will do the studies that are going to answers the questions you want or design the question, design the trial to answer the question.

Maybe it should have been designed as a niche device trial in some way rather than the way it is being presented, but somehow the position of the product needs to fit the study that's done, and an adequate study ought to be done to answer that question.

DR. KRUCOFF:  Well, the OPC questions are on the safety side that the boundary established tells us that the use of this catheter is not going to do more harm than radiofrequency.

And on the effectiveness side, the other boundary basically tells us that this catheter is at least as good as, within a range, the point of comparison.  The reality is I think what we are hearing is that clinicians who do these procedures would be happy to have this instrument in their lab, even if it was not as good as.  The trouble is that's not what this trial addresses.

Either we have an approval process based on the trial or we have an approval process based on what, unfortunately, may or may not be some degree of bias, having looked at the data as to where you think, in fact, the right trial design might have been, if you were to go backwards.  It's a little head spin.

CHAIRMAN LASKEY:  Does it help the agency at all for the Panel to at least say with some unanimity, I think, that the OPC criteria, these criteria chosen are not applicable, and, therefore, need to be evaluated in that context?

That gives us the room in which to make a clinically-driven decision rather than the one that we're agonizing over, which is our clinical horse sense says one thing, but our methodologic rigor says another.  But if we were all willing to discount or to freely acknowledge the known limitations of the OPC criteria as to another device, to this device, that would make us all feel better in terms of our final recommendations.

DR. WHITE:  I'm not sure I feel guilty about this.  I mean, I'm not the one who choose to do the OPC trial.

CHAIRMAN LASKEY:  No, but having done it, you know, we're trying to be fair here to everyone, and it is fair, I think, to acknowledge the limitations of somewhat arbitrary criteria that act as a benchmark, which is a moving target, and perhaps if done in 2003, and not in 1990, would be different.

DR. BAILEY:  But I think the considerations that make people think that it has a niche are the heart block issue and things like that.  It seems it might theoretically have a better safety profile, but we don't have data --

CHAIRMAN LASKEY:  There are no numbers to support that.

DR. BAILEY:  We don't have enough precision on the thing that gives it the niche that it needs.  It has to be better at something.

CHAIRMAN LASKEY:  You keep putting us back in the hole.  We need to get out of the hole.

DR. ZUCKERMAN:  Oh, but that's fair.  I think it's fair to summarize there's a difference of opinion on Question 2a and b regarding the effectiveness shown presently.  We'll get into it later with the Panel voting, but I'm not sure there's more to say about this question more.

CHAIRMAN LASKEY:  Thank you.  Agreed.

(Laughter.)

There are so many ways to rephrase it.

Question 3, effectiveness in the cryomapping area:  "The submission describes the use of cryomapping technology and effectiveness evaluation.  Please discuss whether the study results show that the cryomapping technology is effective for use in the intended patient population."

First of all, the intended patient population is now AVNRT.  So to answer that question, we need to delimit the Panel's response to AVNRT, not to AVRT or AF, is that correct?

DR. HAIGNEY:  No, I don't think so.

CHAIRMAN LASKEY:  No?

DR. HAIGNEY:  I believe the indication is for cryomapping around the AV node for accessory pathways, as well as AV node reentry.  Isn't that correct?

DR. WALDO:  No, it says, essentially, says for AV-conducting tissues.  That would include --

DR. ZUCKERMAN:  Dr. Haigney, do you have Section 3 of the label?

DR. HAIGNEY:  Yes.  It says, "Cryomapping of conducting tissue near the AV node."

CHAIRMAN LASKEY:  You're correct.  That's absolutely correct.  That's my misread, uh-hum.

DR. TRACY:  I think that is, in fact, what that says, but when it was asked earlier of the sponsor, maybe that would need to be clarified, but their answer was pertaining to the AV node.  You could say, well, that is anterioseptal pathways conducting tissue near the AV node.  So that it's not -- if you want to limit this specifically to AV nodes, you have AV nodes or peri-AV nodal tissue, something to more specifically state that.

MR. MORTON:  I have a question.  Shouldn't we be looking at the slides that were presented today by the sponsor and the agency, rather than what was in the labeling, because this labeling was presented or actually was what was used to start the study, as I understand?  Is that not correct, when we're talking about the indications for use?

CHAIRMAN LASKEY:  I suppose that's a good point, Mike.  We need to make sure we're on the same page.

DR. TRACY:  It's actually on page 4 of the sponsor's -- the identification of a variant conducting tissue responsible for SVT using reversible electrophysiologic cryomapping of conducting tissue near the AV node.  It doesn't really narrow it down any further.  In fact, it broadens it if you consider aberrant conduction to mean accessory pathway.

DR. HUGHES:  To me, that means that you could use it to cryomap accessory pathways where you think there's danger of causing AV block, but not ablate those accessory pathways with the cryoablation catheter.  Just a little bit odd.  I guess you would pull the catheter out at that point.

CHAIRMAN LASKEY:  Well, that's right.  I mean that's the problem with the semantics here, and it does need to be reworded.  You can't reword it on the fly?  Can we reword it on the fly, because as it is it's not --

DR. GILLIAM:  Did they show that cryomapping was successful at any point for anything?

CHAIRMAN LASKEY:  No, the results were the same with and without.

DR. GILLIAM:  My question, really to be on the other hand, I think I'm kind of liberal with the first area.  As a clinician, I want to go that way.  But the second area, I think I might be the guy with the dagger to put into it because I think at that point -- I mean, I haven't seen anything in any of the study, in any subgroup, in any way that suggests that cryomapping has been effective in any way.

DR. HAIGNEY:  Didn't they show that they had a higher percentage of successful burn for efficacy in the sites with positive cryomaps?  Is that correct, the 95 compared to 67, which they had positive or negative in the seventies if they didn't do a cryomap?

DR. TRACY:  I think the problem was it was more effective than a negative cryomap, but it was particularly more effective than if you had not done a map.  But if you had done a map in the places where you had not done a map, you might have had further data in one direction or another.

So it was only a piece of the data that was collected, and it was not mandatory to collect cryomapping.  So it makes it a little hard to say too much definitive about it.

DR. WALDO:  But having more rigorous and selective -- I mean savvy, it would have been really nice to focus specifically on the patient groups of interest, but I think we all know why we would love to have this, if it really works, which theoretically it ought to.

CHAIRMAN LASKEY:  But just to remind everybody, it was P equals NS.  I think it was .59 or something from the chi square result, that was just not even close, the efficacy with and without mapping.  So they chose to do the analysis and it was way off.

All right, can we move on?

This study did not show it's effective, at least if we talk about the intended patient population, which was the chi square that they gave us.

No. 4 -- oh, sorry.  I will read on.  With respect to the training and learning curve, "Acute success rate varied per institution in this study," albeit the numbers were also variable.  "Acute success rate per institution ranged between zero and 100 percent.

"a.  Please discuss the concept of the site-based and physician-based learning curves."  Can you go back to A?  		Thank you.

DR. WHITE:  You know, in our world there are roll-ins, and that's how we take care of learning curves with plumbing devices, is we decide how many you need to get comfortable with this, and then you have roll-ins and then you start your pivotal trial.  So I think that's the way that you handle a learning curve with any new device, which we all agree is there.

The only question is, how many roll-ins do you need to feel comfortable?  Three?  Five?  Seven?

DR. WALDO:  But maybe to demonstrate this, I thought that they admitted that they didn't demonstrate any learning curve.  The numbers are so small.

DR. WHITE:  But I think it was because they had multiple operators at the sites, and so they didn't have numbers -- it kept changing over time, too.

DR. WALDO:  But I don't think there's any question among us that a new device requires some training and some comfortability, and that you don't want to hit the deck with the first time that it ever touches your hand, impacting the safety and efficacy pivotal trial, and that's the purpose of a roll-in.  So you get to try it and it doesn't count you if you have a problem until you get comfortable.

CHAIRMAN LASKEY:  So the Panel agrees with the concept of physician-based learning curves.  What the number is is up for grabs.  I think we all espouse that.

"b.  All new devices inherently involve a learning curve in their evaluation.  Please discuss whether the concept of a learning curve, either per site or per physician, may be considered in the evaluation of device safety and effectiveness."

Yes, it may be considered.  Some trials actually do allow for that in terms of looking at roll-in patients and not counting them in a randomized but still analyzing them in the same way.  So I think we all would agree that they may be considered in the evaluation, and, in fact, should be.  Roll-in patients should be followed in every detail the same way as the rest of the trial.

"c.  Please discuss whether and/or what type of physician training should be required for this device if approved."

Here I think there was lots of discussion.  So, Cindy, do you want to lead off whatever consensus you took away from the recommendations for training?

DR. TRACY:  I think there has to be some physician training.  I don't know whether that has to be a visitor come to your lab and show you how to use the equipment versus you go and observe the use of the equipment someplace else.

I think it would be preferable to have somebody come and train onsite.  I'm not sure it needs to be a physician.  It could be a well-trained clinical field engineer or something like that.  But I think there should be some specific individual physician training as well as training of the staff in the laboratory on how to use the device correctly.

I think I would hesitate to mandate the specific number of cases that would have to be performed with an observer in the lab with you.  I think it would depend on the individual expertise of the person learning.

CHAIRMAN LASKEY:  It sounds reasonable.

DR. PAGE:  Will you forgive my addressing Question No. 3 one more time, and just to point out, because I think what we just agreed on wasn't what I remembered and isn't supported by the FDA slide series?

If I can just draw your attention to slide No. 70, which shows a p-value of .0001 in terms of the analysis of cryomap.  The overall group -- so the prospectively-defined, overall, no subanalysis group is .001.

The next slide, 71, points out that it's driven by the AVRT patients, but, in fact, that's where I see the mapping as being useful.  Actually, that's where the proposed indication is.  The second proposed indication reads, to my interpretation, and maybe it needs to be clarified, but reads to my interpretation exactly the issue of mapping AVRT around the AV node.

So it seems to me that, at least my impression was, we dismissed the idea of cryomapping showing any efficacy, and it seems to me we have both an indication of efficacy on slide 70 and a request for indication for that exact purpose on slide four.

CHAIRMAN LASKEY:  The request for the indication calls for the identification of tissue.

DR. PAGE:  Right, for mapping.

CHAIRMAN LASKEY:  Period, right.

DR. PAGE:  But, as I interpret that, aren't they talking about parahisian conduction in and around the AV node?

CHAIRMAN LASKEY:  They didn't specify.

DR. TRACY:  It's not specified, and it's inconceivable that somebody would put in a catheter, map an anterioseptal accessory pathway successfully, pull that catheter out, and put another catheter in and try to get it in the same place.  I mean it's a can of worms that you open there.

DR. PAGE:  Well, I didn't open it.

(Laughter.)

But that's the request that they have.  Indeed, the data are the data.

CHAIRMAN LASKEY:  Lilly, could you help me out here?  What was the results of your analysis in the 2-by-2 table of the outcomes by mapping or no mapping?

DR. YUE:  Okay.  Now please remember we have three subgroups:  effective, ineffective, and no attempts.  The sponsor combined effective -- sorry, the sponsor combined "ineffective" with "no attempts," correlated without cryomapping, then compared this without effective cryomapping with effective cryomapping.

Okay, now here the overall p-value, the p-value for the overall analysis is less than .0001, but you'll see it is driven by 49 AVRT patients.  This subgroup is not the group that the sponsor is, it is occurring to me, claiming for.  Instead, the AVNRT group is the only group they are claiming for.  We have a concern about this grouping.

DR. GILLIAM:  I think you're asking what ablation --

DR. YUE:  What is the meaning of the comparison?  Why are they grouping "ineffective" with "no attempts"?  Is the subgroup classification "effective" versus "without effective" biologically plausible?

It seems if we try to test the impact of effective cryomatching ablation with success, we could compare the effective group with the ineffective group, then use "no attempts" as a control.

Now from here, we can see the acute success rate is 94 percent for effective group.  It is much better than 50-65 percent for the ineffective group, but the chance of having an effective cryomatching is only 64 percent.  On the other side, if we try to test the impact of our attempt cryomatching on ablation with success, we could compare attempts with no attempts.  Attempts include effective and ineffective here.

We performed this study.  Then we found there's no significant difference in ablation acute success between attempt and no attempts.

Now here the p-value is .59 here.

DR. BAILEY:  The point is, if you were doing a randomized trial, it would be use of cryomapping or not using cryomapping.  You can't just take the people that are positive and say, because their success rate is high, that shows that cryomapping is a good thing to do.

DR. WALDO:  You know, this is the problem.  We understand the statistic very clearly, and there's no challenge to that.  But the specific thing that we are talking about is very unique.  What you are trying to do is avoid the his bundle, and I'm not sure they presented that specific data.

We are not talking about ablating -- this is very, very unique, and you can't apply it the same as looking at pathway on the left side with a his bundle.  This is a very unique application where we have nothing at the moment to help us, and, fortunately, it is not the most common type of Wolf Parkinson-wide or accessory AV connection problem, but it's very, very real.

The hope is, and maybe we heard some of it anecdotal, unfortunately, here, but the hope is that you can use this cryomapping technique to find the sweet spot and avoid the his bundle and help the patient effectively and safely.  That is as simple as that.

That's not going to come out of anything we say about the statistics here because it is just not going to.  It's not there.

DR. TRACY:  It's also not necessarily something that needs regulating.  This is something that people over time can find out:  If I attempt, I'm successful, I ablate; what is my success at ablation at that spot?

So I think this is something that you can do with this catheter that is unique.  You can't do it with anything else, but it doesn't seem to be the make or break whether this catheter is useful or not.

DR. KRUCOFF:  I think this is also a good illustration of the difference between the data we actually have to discuss and the questions that go through your mind.  In fact, I think the sponsors during their presentation came up with a discussion that actually the inability to show any effect may be a great way to show that it's safe to ablate in certain patients, but while it's a great concept, you know, this trial doesn't give us the basis to address it with data.  I mean it's a great concept.

CHAIRMAN LASKEY:  Frankly, all bets are off with the cryomapping because there was an arbitrary selection criteria that we were not privy to as to who were mapped and who were not.  It is really difficult to take any lessons home from that.  I just wish it weren't here, but it's the second bullet in their IFU.  So sorry.

Labeling:  "Labeling for a new device should indicate which patients are appropriate for treatment, should identify potential device-related adverse events, and should explain how the device should be used to optimize its risk/benefit profile.  If you recommend device approval, please address the following:

"a.  Please discuss whether the proposed warnings, precautions, and contraindications are acceptable, based on the study results."

Let's be clear about one thing:  We're now back to the AVNRT group, is that right, Panel members?  Right.

Are the warnings, precautions, and contraindications acceptable, based on the study results?

DR. GILLIAM:  I'm going to go back to my electrocution question still.  I still have real issues with that, and maybe I'm not electrocuting a dead horse.  But it is not clear to me how this device can electrocute a patient if it were hooked to an RF generator.  I think if it is possible, then maybe a little sterner warning needs to be connected, because I can't see how you can avoid or prevent me from hooking this device to an RF generator.

DR. TRACY:  I don't think it fits.

DR. WHITE:  I asked them specifically that.

DR. GILLIAM:  You said they can't do it, but if you're going to be able to record, I assume at some point you record an electrogram from the end of this catheter some way.  That means it's got to plug into your recording machinery you have.  So at some point there's a pin that gets to my, if you will, Pruca or whatever system you have.

So that would suggest to me that I could plug it into an RF generator, because all I need is one pin.  If that is a significant risk, it may need to be a little bit higher.

I mean I've never seen anything like this in anything we've done other than plugging the pin directly into the shield, when we went to hooded pins --

DR. TRACY:  Well, but I'm not sure, Rosie.  You might, then, carry it to the extreme of saying:  Don't plug this into the wall because that could be dangerous.

I think the connectors are simply not meant -- they don't fit into an RF generator.

DR. GILLIAM:  Okay.  I think that's how we ended up with the hooded pins, by the way, so people didn't plug them into the wall.  But I'm maybe getting a little bit overboard here.  I just find it very curious.  I have never seen that specific thing written out in any way.

CHAIRMAN LASKEY:  Well, it's certainly a red flag, but can we settle this one more time from the engineer?  Can you explain the physical basis for the inability to make this hazardous connection?

MR. ABBOUD:  Actually, technically, you can connect anything to an RF generator, right?  You can take any wire and connect it to the RF generator.

The catheter has a tip, a metal tip, and has a wire on it.  And like you said, you can take that wire and put an adapter and put it on the RF generator, but no electrocution can happen.  That catheter is not designed to deliver enough energy, and there is no reference electrode first to connect it.  That means someone has to take that catheter, that second end wire, make all the adaptation and put it to the right pin to an RF generator and make it work, right?

CHAIRMAN LASKEY:  Okay.  All right, and we do have a big warning with a box around it, is that right, in the IFU?

Let me just go back to identify potential device-related adverse events.  I think Dr. Waldo, in the lead, explained it is very difficult to pinpoint specific device-related AEs when they are procedure-related.  So I might suggest modifying that language to get us off the hook.  We can't look at device-related AEs unless it's heart block, and that didn't occur.

So other than Rosie's concern --

DR. ZUCKERMAN:  Okay, Dr. Laskey, maybe I can take you and the Panel through the labeling because we do need some comment on the big ticket items.  I'm looking at the label which is in Section 3, "Indications for Use," Mr. Morton, which is the same as on our slide.

This is the key point.  The first bullet regarding cryoablation, is that still acceptable to the Panel, and the second bullet should be removed?

CHAIRMAN LASKEY:  What page are you on?

DR. ZUCKERMAN:  I'm in Section 3, "Labeling of the Panel Pack," and page 2 starts with, "Device description" and then "Indications for Use."

The "Indications for Use" somehow got deleted from Question 5, but is our major question here:  Is the device appropriately labeled?

DR. TRACY:  I think it would save a lot of angst simply to remove that second bullet.

DR. ZUCKERMAN:  Okay.  Part "b" would be:  "Is the Panel satisfied with the current contraindications?  Should some be removed?  Are there others that are important?"

DR. TRACY:  Didn't we decide that we could put it through a bioprosthetic aortic valve, so that that slight change in the wording in the retrograde -- but then again, why are you going retrograde for the AV node?  I suppose you could, if you were ablating it on the left side.

I mean, it's there.  It's not pertaining to this population, but --

DR. VETROVEC:  If I'm correct, if one reads the package insert for a porcine bioprosthetic valve, it says it should not be crossed by a catheter, not to say that we haven't, most of us in this room, done that, but --

MR. MORTON:  I think the valve manufacturers would definitely not want you crossing anything across prosthetic valves.

DR. GILLIAM:  Do we have guidance from other like RF catheters and their ablation systems?  The contraindications, are these substantively different from that in any way?  I guess I'm asking FDA.

DR. EWING:  The indications for all the non-generic-indicated RF catheters include accessory pathway and adrophobe.

DR. GILLIAM:  I was thinking more of the contraindication, not the indications.  I mean, the indications I will grant here a specific arrhythmia, but are there any specific contraindications to RF systems?

DR. DESMARAIS:  I can answer that.  We based the contraindication of this indication for use based on prior RF ablation safety and effectiveness data that we lifted from the FDA website, and this is the model we used to create this impression for use and this labeling.

DR. WALDO:  By eliminating your angst, we eliminated the one potential advantage of this because we haven't had data, and we've made our peace with that, is that what we're saying, unhappily, but that's the way it is?  Is that we're saying?

DR. TRACY:  Yes, that's what that means.

DR. PAGE:  Does everybody agree with that?  I personally think that having this as a mapping device for cryomapping is a potentially useful indication.

DR. HAIGNEY:  I agree with that.  I think it does require us to have certain amount of faith in our fellow man that they will map and then pull the catheter out, but I don't think -- I think it depends on whether you have a Hobbesian view of man or not.

CHAIRMAN LASKEY:  Again, we're straying from the data.  I think it always helps to go back to homeplate, and I think we all understand its attractiveness, but we're straying from the data to support that.  We all know that you're going to do it anyway.

DR. WALDO:  That's an easy thing to recommend, a recommendation on how to focus them on identifying a bunch of patients like this --

CHAIRMAN LASKEY:  Yes.

DR. WALDO:  -- and giving us the data.  That's the one easy thing I think of all the things we talked about so far.

CHAIRMAN LASKEY:  Well, we'll get to that.

Brian, do you want to lead us through the wilderness some more?

DR. ZUCKERMAN:  Okay, now we're dealing with warnings.  Dr. Gilliam has expressed an opinion on one warning.  Are there any other warnings that are of concern here for the precaution section?

DR. HAIGNEY:  The cryoglobulinemia contraindication I think should be included.

CHAIRMAN LASKEY:  I think we're all in agreement about what not to do.  There's confusion about what to do, but I think Part a is pretty easy here for the labeling.  Do you agree?  Yes?

Do we feel the IFU described how the device should be used?  From a technical standpoint, I think so, yes.  Agree?  Good.

Are we awake?

No. 6, from a post-marketing standpoint:  "If you recommend approval, please discuss whether a post-market study should be performed to address any issues that are unresolved, but not essential to the pre-market approval of the device."

DR. KRUCOFF:  Were one to approve the device, then one might append a post-market look at the value of mapping, for instance, if it's built into the same device.  That might be --

DR. WALDO:  Specifically, the parahisian mapping.

CHAIRMAN LASKEY:  One might, agreed.  What else one might recommend?

DR. VETROVEC:  Can you look for things that aren't necessarily indicated?  If we talked about removing one of the indications, then that wouldn't be a legitimate question to ask someone to do.

DR. ZUCKERMAN:  I think Dr. Vetrovec's viewpoint is correct.  The parahisian question is an important one, but that could be studied in a new IDE study with the intent to get that on the label.  We're really talking about some issues for the intended indication which are not show-stoppers for making a decision today, but would be nice subsequently to study because there would be utility for having that additional data.

DR. TRACY:  This is a dreadful suggestion because it just goes full circle to where we were earlier in terms of our registry on effectiveness.

DR. AZIZ:  What about monitoring the instance of heart block that people have spoken about?

CHAIRMAN LASKEY:  Sorry?

DR. AZIZ:  The instance of heart block that people have spoken about?

CHAIRMAN LASKEY:  Yes, I couldn't agree more.  I am very enthusiastic about -- I would go further than one might.  I would strongly suggest a prospectively-designed registry, designed for a safety outcome analysis.  I think the efficacy is believed by most, but safety is the issue, if this is a complementary or an alternative technique.

DR. ZUCKERMAN:  I think, again, in understanding the nature and intent of post-market surveillance for a problem like complete heart block, which has an extremely low incidence, while that information might be interesting to clinicians, it would really require several thousand patients perhaps to make a definitive determination.  That's not the intent usually of post-market surveillance.

What we are asking here is for the intended use indication under study, are there some additional points where we still feel very uncomfortable, not where we still want to dot an "i" rather than get an A-plus-plus on a paper?

DR. GILLIAM:  I think the utilization of cryomapping even in the AVNRT patients could be effectively shown.  I mean, you could show, when you achieve an effective cryomap in a specific place, I mean in the study here it suggests that it wasn't particularly useful, but I think the data are really suspect because we don't know how those patients were selected or, you know, in any way.  I think that is something post-market could be handled within the study population.

DR. BAILEY:  I don't see how you could learn about the efficacy unless you design a study to look at it.  I mean, post-marketing, some people will get cryomapping and some won't, but how do you know, what does that show?

DR. GILLIAM:  I mean specifically design a study to look at that.

DR. BAILEY:  But that's not post-marketing.

DR. ZUCKERMAN:  Right.  Again, that's a traditional IDE study for a new indication on the label.  Perhaps to put this in perspective and allow the Panel to better appreciate what we're getting at here, if this was a traditional RF ablation catheter, there wouldn't be any post-market study required because, as several people have pointed out, this is a tried-and-true technology.

Is there something unusual here that has been indicated in the dataset about a safety issue that would make you really want post-market data?

CHAIRMAN LASKEY:  The consensus seems to be no, but --

DR. VETROVEC:  Heart block seems to be the obvious.

CHAIRMAN LASKEY:  Heart block seems to be the obvious one, yes.

DR. BAILEY:  Yes, I think heart block would be --

CHAIRMAN LASKEY:  With due respect to the precision of the estimate, I would think that a thousand cases could be done across the country in a reasonable period of time.

DR. BAILEY:  And if the incidence is zero, the estimate will be quite precise?

CHAIRMAN LASKEY:  Yes.

DR. KRUCOFF:  The other part of this is the premise of approval based on current data, because while device-related safety looks pretty good, procedure-related safety, if a significant percentage of these procedures end up being bailout RF ablations, you're going to have catheters in the blood -- I mean the safety, the potential safety issues from a procedural point of view, if this was approved for use, might well be worth tracking in the post-market, but that begs the question of whether you see the safety issues as sufficient to approve the device in the first place.

So the way this question is stated, it's a little hard to --

DR. TRACY:  Does it make any sense to look at how the catheter is used, how many times it is a bailout for RF versus how many times RF is used to bail out from a cryo?  Is that something that would be -- I know it would be interesting.  Is it something that's worthwhile?

CHAIRMAN LASKEY:  Or easy to do.  I mean, recording that kind of stuff is just one more thing to record.  Generally, the surveillance studies are just do it and we'll look at the outcomes, but this is yet another iteration.  So it's one more thing to do.  It's getting closer to a study than just an observational kind of thing.  I would think it's a good idea, but it just requires one more thing to do and puts the sponsor on the hook for paying for that information.

DR. KRUCOFF:  I think what we are wrestling are ideas for actually it would just be better studies to do.  Maybe pertinent to the agency's interest, I haven't heard anyone suggest that there's anything about the lesion in animals or the behavior of the catheter, or whatever, that ought to spook us in some way that would make you want to keep watching longer.

I would suggest that the real question here is approval, not sort of unique post-market kind of issues.

CHAIRMAN LASKEY:  So as we get ever closer -- at this point I would like to ask the sponsor -- at this point I would like to ask the agency and then the sponsor if they have any additional comments or questions before the vote.  So the agency first?

DR. ZUCKERMAN:  The agency doesn't have any additional comments.

CHAIRMAN LASKEY:  Again, on behalf of the Panel members, I thank the various contributors who have helped us out significantly.

Any final comments from the sponsor?

DR. DESMARAIS:  Yes.

CHAIRMAN LASKEY:  Brief?

DR. DESMARAIS:  Very brief.

CHAIRMAN LASKEY:  Thank you.

DR. DESMARAIS:  We really appreciate and are fortunate to be here and to present our technology to the Panel.  We are very grateful to the Panel and FDA for their thoughtful comments.

We believe that the risk/benefit profile for this product for this profile is favorable and that the technology will play an important role in the patient care.

Once again, on behalf of CryoCath, I would like to thank the Chairman, Panel members, and ladies and gentlemen for letting us present our technology.  Thank you.

CHAIRMAN LASKEY:  Thank you, sir.

Before we proceed with the motion, I would like to ask our industry and consumer representatives if they have any final thoughts, beginning with Mr. Morton.

MR. MORTON:  Just to thank the sponsor for a clear and good presentation of data and clinical utility of the device, and thank the Panel.  This is exactly the sort of decision that the agency needs Panel input on, new technology and the outcomes not being crystal-clear.  So thanks very much.

CHAIRMAN LASKEY:  Mr. Hughes?

DR. HUGHES:  Thank you Dr. Laskey.  Let's see, I know it's late in the day and I would like to be brief, should be brief.  If it wasn't for spring break, I would have done a three-hour lecture at my institution.

I think that, first of all, I appreciate the presentation by the manufacturer and the FDA, as well as the comments and discussion by my colleagues.  I think that there were a couple of points that I really would like to highlight some that have come around.

That is one with regards to learning curves and being able to translate whatever level of success the device has to other physicians, other surgeons, and certainly a concern when it comes to the consumer.  The consumer I think of as being the patient, as opposed to his or her learned agent; that is, the surgeon or physician.

I also appreciate the comments of Dr. White having to do with labeling and mechanical valves; that is, making sure that the labeling does not say specifically or does not indicate that it might be okay to traverse mechanical valves, as well as representativeness of any studies, minority representation, gender representation, things like that.

I think overall, though, this particular device does need in some form more study, more statistical study to assure its efficacy for the patient.  This particular kind of study or set of studies I would think could take the form of perhaps post-approval studies.  But, otherwise, I think that from the perspective of alternatives, the consumer/the patient should have alternatives, and in that regard the physician, of course, should have alternatives at his or her disposal.

I think that this device conceptually theoretically looks very, very promising, but I think that there are additional studies more specific to the intent as opposed to taking the OPCs and applying them.  Anyway, more specific studies would be nice and they could take the form of post-approval studies.

Thank you.

CHAIRMAN LASKEY:  Thank you, sir.

I would like to quickly, hopefully, open the public hearing.  Anyone in the audience who wishes to address the Panel on today's topic before we vote?

(No response.)

Thank you.  If not, we will close the open public hearing.

I would like to ask Geretta to read the voting options.

MS. WOOD:  "The medical device amendments to the Federal Food, Drug, and Cosmetic Act, as amended by the Safe Medical Devices Act of 1999, allows the Food and Drug Administration to obtain a recommendation from an expert advisory panel on designated medical device Pre-Market Approval Applications, PMAs, that are filed with the agency.

"The PMA must stand on its own merits, and your recommendation must be supported by safety and effectiveness data in the application or by applicable publicly-available information.

"Safety is defined in the Act as reasonable assurance, based on valid, scientific evidence, that the probable benefits to health under conditions on intended use outweigh any probable risks.

"Effectiveness is defined as reasonable assurance that in a significant portion of the population the use of the device for its intended uses and conditions of use, when labeled, will provide clinically-significant results.

"Your recommendation options for the vote are as follows:  approval if there are no conditions attached, approvable with conditions.  The Panel may recommend that the PMA be found approvable subject to specified conditions such as physician or patient education, labeling changes, or a further analysis of existing data.  Prior to voting, all of the conditions should be discussed by the Panel.

"Not approvable.  The Panel may recommend that the PMA is not approvable if the data do not provide a reasonable assurance that the device is safe, or if a reasonable assurance has not been given that the device is effective, under the conditions of use prescribed, recommended, or suggested in the proposed labeling.

"Following the voting, the Chair will ask each Panel member to present a brief statement outlining the reasons for their vote."

CHAIRMAN LASKEY:  I now ask for a motion on the PMA.  Anyone?  I now ask for a motion on the PMA.  May I have a motion?

DR. TRACY:  I am making this motion as an electrophysiologist with the understanding that the reasonable safety and effectiveness of this device has been shown.  I move to approve with conditions.

CHAIRMAN LASKEY:  Is there a second?

DR. GILLIAM:  I second.

CHAIRMAN LASKEY:  It has been moved and seconded that the motion to approve with conditions move forward.  May I have condition one for the PMA?

DR. TRACY:  The first condition would be some minor modification in the labeling to include cryoglobulinemia as a contraindication to cryoablation.

CHAIRMAN LASKEY:  Do we have a second?

DR. GILLIAM:  Second.

CHAIRMAN LASKEY:  Do you need a count for these?  Can we have a hand-raising response to the second, the seconding of the first condition that Cindy just raised?

All members in agreement with cryoglobulinemia to be added to the labeling, please raise hands.

(Show of hands.)

It looks like unanimous.

This is new for me, too, Rosie.

DR. KRUCOFF:  Can I ask a clarification, Mr. Chairman?

CHAIRMAN LASKEY:  Yes, I was just about to invite the Panel discussion on this condition.

DR. KRUCOFF:  Oh, it isn't about the condition.

CHAIRMAN LASKEY:  Oh.

DR. KRUCOFF:  I am trying to understand what we have moved in terms of the indication for ablation and/or the indication for mapping.

DR. TRACY:  That's a condition.  We have moved for, what I was moving to approve was for the condition of ablation in patients with AVNRT.

DR. KRUCOFF:  And the other one was in the --

CHAIRMAN LASKEY:  We had previously agreed, before we got to this point, to strike the second bullet.  Is that correct?

DR. ZUCKERMAN:  Right, but right now we are voting to discuss each condition.  You're going to vote on each condition of approval, and you're going to put the package together.

So you have started with labeling Condition No. 1, which refers back to cryoglobulinemia.  You want to discuss that and have a vote.

DR. TRACY:  We had the vote.  We did vote on that, yes.

CHAIRMAN LASKEY:  We had the vote.

DR. ZUCKERMAN:  Fine.

CHAIRMAN LASKEY:  There didn't seem to be a lot of discussion.

DR. ZUCKERMAN:  Then, Dr. Tracy, Condition No. 2?

DR. TRACY:  Condition No. 2 is that we strike the cryomapping from the indication.

CHAIRMAN LASKEY:  Discuss first or before?  Yes.  Do we have a second?

DR. GILLIAM:  Second.

CHAIRMAN LASKEY:  We do.

DR. PAGE:  I would just like to state that I think it's useful to have that in there.

CHAIRMAN LASKEY:  Okay, there's the discussion.  Okay, let's vote on the second condition then.

All in favor of striking the second bullet from the current IFU?

(Show of hands.)

One, two, three, four, five in favor.

All against?

(Show of hands.)

One, two, three, four, five against.

DR. ZUCKERMAN:  Okay, it's important to state who is in favor and who is against the vote for the record.

CHAIRMAN LASKEY:  I have to name my colleagues in public?

DR. ZUCKERMAN:  Yes.

(Laughter.)

CHAIRMAN LASKEY:  How about we do it henceforth or do you want another vote?  Let's do this procedurally correct then.  Let's just do the hand vote again.

All in favor of Cindy's second condition, which is the elimination of the second bullet?

(Show of hands.)

In favor are Cindy, Dr. Dullum, Kent Bailey, Mitch Krucoff, and Dr. Gilliam.

Okay, and those against this condition?

(Show of hands.)

Drs. Haigney, Vetrovec, Waldo, Aziz, and Page.

I've never had to do that.  All right.  It's five/five.

MS. WOOD:  If you voted against the condition, could you please raise your hand again?  I missed -- okay.

CHAIRMAN LASKEY:  One, two, three, four, five.

MS. WOOD:  Okay.

CHAIRMAN LASKEY:  It's five/five.

Okay, is there a third condition?

DR. TRACY:  Yes, a third condition --

DR. ZUCKERMAN:  Well, we need to decide what's going on with this condition.

CHAIRMAN LASKEY:  Of course.

(Laughter.)

You need a tie-breaker?

DR. ZUCKERMAN:  Yes.

CHAIRMAN LASKEY:  Well, I vote with the group that recommends striking it.

The third?

DR. TRACY:  I move that a formal training program, the details of which can be worked out at a later date, be part of the approval.

CHAIRMAN LASKEY:  A second?

DR. GILLIAM:  Second.

CHAIRMAN LASKEY:  Thank you.  Do we have any discussion on this one?

(No response.)

Let's have a hand vote for all in favor of the training criteria to be elaborated subsequently, articulated subsequently.

(Show of hands.)

Drs. Haigney, Cindy, Dullum, Vetrovec -- we have unanimous agreement on this, and I would agree with that, too.  That's the third condition.

Is there a fourth?

DR. TRACY:  Maybe.  I move that there be post-market surveillance to monitor for the incidence of AV block.

CHAIRMAN LASKEY:  May I have a second?

DR. HAIGNEY:  Second.

CHAIRMAN LASKEY:  And a vote?

DR. DULLUM:  Can we have discussion?

CHAIRMAN LASKEY:  Oh, I'm sorry.

DR. DULLUM:  Yes.  I just want to go back to, is that going to require a thousand patients, do we think, and how is that going to be trackable and relevant, the less than 1 percent incidence most places?

CHAIRMAN LASKEY:  It depends on what happens, I guess.  All right, surveillance for a relatively-rare event, although not astonishingly rare, if we think it is on the order of 1 percent?

Kent, what would be a reasonable --

DR. BAILEY:  Well, if it's really 1 percent, then you can never show that it's less than 1 percent.  If it's really very low, like one in a thousand, then you should be able to show it in a few hundred patients.

CHAIRMAN LASKEY:  And that would be the goal?

DR. BAILEY:  But if it is a percent, then it sounds like what I'm hearing is that it's not necessarily any advantage over RF.  So I think you're really trying to show that it's better, which would lead to, presumably, a lower sample size.

CHAIRMAN LASKEY:  Than the equivalent, yes.

MR. MORTON:  I have a question on this to the agency, Dr. Zuckerman.  Is post-market surveillance actually not an early-warning system, something that we would look for for an event that we suspect could happen, an adverse event that we suspect could happen?  In this case, we actually suspect just the opposite?

DR. ZUCKERMAN:  Right.  Again, I think there are two issues.  One, we want to distinguish the need for post-market surveillance versus the need for addressing interesting scientific questions.  Certainly, if someone wants to go out and organize this registry of "X" thousand patients to definitively prove that the incidence of heart block is less with one form of therapy than the other, that's a very interesting scientific question, but it's not necessarily in the purview of the FDA's post-market authority.

The agency would be more interested in perhaps the following question:  Perhaps if you add a certain delta to that 1 percent level, and proposed a registry and the incidence of heart block is unexpectedly a safety problem because it is coming in at 3, 4, 5 percent, but I think one needs to pose the right post-market question here initially:  What is the question?

DR. HAIGNEY:  Yes, I think the question is not, is it better than RF ablation?  I think it is, is it worse than RF ablation?  So I would be interested if the incidents maybe blocked were 2 percent.  I would consider that an important factor that I would want to know before I used the catheter.

DR. DULLUM:  Well, wouldn't it be a reportable event anyway, MDR, or not, AV block?

CHAIRMAN LASKEY:  But, Mark, doesn't that presuppose some kind of simultaneous or concurrent control or would you resort back to the literature for comparison?

DR. HAIGNEY:  I would resort to the literature.

CHAIRMAN LASKEY:  I mean, I personally don't see what's wrong with a surveillance that looks for the incidence of complete heart block in "X" consecutive patients undergoing this procedure.  I'm not allowed to -- right, but that's my concept of the surveillance; it's not necessarily a study, but it's an observational approach to event-collecting.

DR. HAIGNEY:  We were having trouble trying to find a number of patients that we want in the registry, and I'm saying that I don't think we need to look for the high standard of less than 1 percent incidence.  Do you see what I'm saying?  I'm proposing a more liberal --

CHAIRMAN LASKEY:  Yes, but my point is we can do better than 150 patients.  I think we can come up with a better estimate with more than 150 patients, which is exactly what a post-marketing survey would do for us.

DR. HAIGNEY:  If I have learned anything today, Dr. Laskey, it's that we need more than 150 patients.

DR. WALDO:  Would it be permissible to use people in other countries or would it only be people in the United States?  My question is, in this surveillance, would it only be post-marketing in the United States or could they use worldwide information?

CHAIRMAN LASKEY:  You would have to ask the agency for how they would receive that data.

DR. ZUCKERMAN:  Again, what is the question --

DR. WALDO:  But you get it much faster with worldwide, the numbers anyway.

DR. ZUCKERMAN:  Right, but what is the question being asked?  The question that we would like to ask is, for the label indication, is there an unusual safety problem that produces heart blocks, say, in 3 percent of patients?

So if you were convinced that outside U.S. data was being utilized, was being obtained in that intended patient population, then, yes, a large simple registry, both U.S. and OUS patients, might be applicable.  On the other hand, if the OUS data contains patients who are being treated multiple other indications, it's not going to help us answer the question.

DR. TRACY:  I think that the intent of the condition is to identify an unusually large number of AV blocks, an unanticipated number of AV blocks.  I think that right now I'm not certainly prepared to say what that number of patients would be required, but I think we could statistically come up with something that would be a reasonable number of patients to survey for that phenomenon.

CHAIRMAN LASKEY:  Yes, I agree.  I don't think we should be on the hook for the study design and the sample size, but I think we're simply recommending that this be an additional condition.  I think it is designable.  This is a study which is designable.

DR. WALDO:  So can we move the question?

CHAIRMAN LASKEY:  So, therefore, we need to vote on the fourth condition, which calls for the configuration of a post-marketing surveillance for the event of complete heart block in a sample size to be defined.

May I see all in favor?

(Show of hands.)

Drs. Haigney, Cindy, Dullum, Kent, Waldo, Aziz, Page, and Gilliam.

And those against?

(Show of hands.)

Dr. Vetrovec.  Okay.

DR. VETROVEC:  I guess I am supposed to state why.  My position would be, I think, reflecting the industry comment, which is we're really doing this because of our interest.  We don't think there is a problem with heart block.  We think there's not a problem with heart block, but that's not sort of the spirit of a post-marketing surveillance.

DR. GILLIAM:  But we don't know that there's not, and I think that's why the post-marketing is necessary.  We haven't been presented enough data to show that this is, in fact --

DR. VETROVEC:  I would suggest you not suggest that when you're trying to convince some of us that, because you're already off the beam, that may be one of its advantages.

CHAIRMAN LASKEY:  Okay.  Any further conditions, Cindy?

DR. TRACY:  Not for me.

CHAIRMAN LASKEY:  So we have a motion.  We have conditions.

DR. ZUCKERMAN:  Okay, so now you're going to --

CHAIRMAN LASKEY:  Do we need to raise any other conditions?  Are there any other conditions that any other members of the Panel want to raise besides Cindy?

(No response.)

Good.  All right, and I think at this point we're ready to vote on the motion presented, which is to move to approve with the following five conditions:

To add the contraindication that patients with cryoglobulinemia be excluded from consideration for this treatment.

The second condition being that we strike the second bullet on the IFU pertaining to the mapping.

The third condition being the establishment of a better-articulated and more-specific training criteria, perhaps including the number of cases for a learning curve.

The fourth condition being -- there's only four conditions here -- being the post-market surveillance for heart block as a measure of safety of the procedure.

So the one motion with four conditions.

DR. KRUCOFF:  Do we discuss the motion itself?

DR. ZUCKERMAN:  Sure.

CHAIRMAN LASKEY:  Yes, we can, as is.  That's en bloc now, yes, before we vote.  Sir?

DR. KRUCOFF:  I would just like, before we vote, to encourage everybody to think about what this vote means beyond just this product, but with regard to the process of designing a trial prospectively between a sponsor and the agency, and then having to expend this kind of energy to really salvage what amounts to the wrong trial or a trial whose pivotal data suggests the opposite of what you're intuitively inclined to do.

That may well be the right decision for this device, but what it will stop is any ability to answer that question, and whether it will roll over to then, as a precedent, for future studies to leave it to this body to bail out trial designs that don't answer the questions that actually are brought forward as an indication question, and I'm really concerned about that.

It's really not an issue of whether -- you know, I don't think there's any question about the diligence with which this trial was done, about the interest of the sponsor of bringing this device forward, or about the interest and the passion of the investigators who have had their hands on it.  In fact, we're looking at a very similar situation upcoming with a coronary device.

But I really do think that the process of prospectively designing a clinical trial that actually addresses a question that has to do with an indication is a vital process.  If we decide to vote around that today, that there's a bigger issue that impacts by setting that precedent that ultimately this body would be put in a position to suggest to the FDA or for a sponsor to rely on to bail out having done the wrong trial, or whatever, to bring a device forward.  So I'm really concerned about where we go with this.

CHAIRMAN LASKEY:  I think we all appreciate that.  The reason we're here until six o'clock is because we were giving due diligence to the process and realizing how the serious the task is.

So it being six o'clock and with those reminders, let's vote.  May I see, by a show of hands, all in favor for approval with the conditions as enumerated?

(Show of hands.)

In favor for:  Drs. Haigney, Tracy, Dullum, Waldo, Aziz, Page, and Gilliam.

And those against?

Drs. Vetrovec, Bailey, and Krucoff.

How did Dr. White vote?

DR. ZUCKERMAN:  He didn't vote.

CHAIRMAN LASKEY:  He didn't?  He did, but he didn't.

DR. ZUCKERMAN:  But you could guess.

(Laughter.)

CHAIRMAN LASKEY:  Okay.

DR. ZUCKERMAN:  Dr. White, for the record, is no longer present.

CHAIRMAN LASKEY:  So the motion passes by a vote of six to three.

I would like to spend 60 seconds or less for each member of the Panel to please state your name and the reason for your vote, please.  Dr. Haigney?

DR. HAIGNEY:  Mark Haigney.  I thought the sponsor showed that the device is safe.  I think that the clinical trial, as it was designed, it was entirely appropriate because of the way this device is going to be used.  I think I'll stop at that point.

DR. TRACY:  Cindy Tracy, and I voted for approval because I think the sponsor has shown reasonable safety and effectiveness for the device.

DR. DULLUM:  Mercedes Dullum.  I voted for approval because of showing the reasonable safety and efficacy.

DR. VETROVEC:  I voted against because I think that every time we say something is reasonable but it didn't hit the target that the sponsors said they would hit, we're really leaving ourselves open for question as to why we made that decision.  I'm not comfortable in this circumstance, and would almost like to encourage this as a reason that people really should answer the question that they set out to or set up studies that approach that.

CHAIRMAN LASKEY:  Yes, I think that today this has probably been one of the more difficult sharp edges to sit on.  We didn't have with support what our clinical intuition suggested.

I would strongly recommend, just as a personal opinion, that we relook at the concept of OPCs.  They're moving targets, particularly in the technologic arena, where things move awfully dramatically and they can be better or worse within a six-month period of time as the numbers accumulate.  So I would be wary of those kinds of strawmen basically to design a study, and there's nothing like a good, old, randomized, controlled trial that, hopefully, will answer the question with more statistical rigor.

DR. BAILEY:  Kent Bailey.  I voted against approval based on the evidence that was presented today.  Although I think the sponsor demonstrated reasonable efficacy, that is not -- although the device works in a vacuum, you can't look at that in a vacuum.

It's probably not as efficacious as the other conventional approach.  So, well, maybe it's safer.  Well, it didn't met the safety standard either.  So we can't really say that it's safer.

But, well, it has less heart block, but we don't have the evidence to show that either.  So, potentially it's not as good in any of those three categories.

I also agree with the process question, that this is a bad precedent to set, although I think it is a very promising device and I would have liked to have seen data presented that showed that it was as good or superior in at least one category.

DR. WALDO:  Dr. Waldo.  I voted for.  I would like to just second what Dr. Laskey said.  I think he said it succinctly and well after a long afternoon.

I voted for it because I thought that the efficacy and safety was reasonably presented and reasonably demonstrated.  I really think that it would serve us well to relook at some of the criteria that were the strawmen in this target, the strawmen targets in this study.  Thanks.

DR. AZIZ:  Salim Aziz.  I voted in favor of the device because I think, firstly, it showed relative safety and efficacy.  I think it adds a new way of dealing with a somewhat difficult problem for the EP and people, and, hopefully, with the surveillance the evidence will bear out the fact that it has a low incidence of problem.

DR. PAGE:  Rick Page.  I was convinced by the evidence that there is reasonable assuredness that this device is efficacious and safe, and I think it will be a useful tool to the clinical electrophysiologist.

DR. KRUCOFF:  Mitch Krucoff.  I voted against because, while it is obvious that from the user's point of view, this really does look like a promising device, that the lesson we have is that intuition in medicine is a very dangerous direction to follow exclusively, and that's why we do clinical trials.  In this particular clinical trial, neither the safety or the efficacy data landmarks were hit.

An OPC trial, by and large, we consider to be an easier target to hit than a randomized trial, and that's why I voted against it.

DR. GILLIAM:  Roosevelt Gilliam.  I voted for the motion, aware that it did not meet the OPC trials of radiofrequency device, a device that has an extraordinarily high standard of success and safety that's been proven over years.

I recognize that this device, I felt, demonstrated it has a reasonable degree of safety and efficacy, although both may be close to, or at least falling short of, the accepted radiofrequency, it provides a different approach to a very difficult problem.  So it provides an additional tool.

I think that insofar as this is not to replace radiofrequency ablation, I think the data so we can have a relative appreciation of its safety was there, and I was comfortable with that level of safety.  Thus, I voted for the device.

CHAIRMAN LASKEY:  Mr. Morton, Mr. Hughes, any final words?  No?

DR. HUGHES:  Yes, I have just a quick comment to send out to the manufacturer, and that is, with regards to the post-market surveillance, I feel like in the long run that it will be a good thing for the manufacturer.

In fact, I think we have been, or the Panel has been, somewhat mild with regards to those recommendations with regards to post-market surveillance.  A very thorough kind of post-approval study I think would be in order, including such things as the potential for cryomapping.  I think that, once again, in the long run this would be beneficial not only to the consumer, to the patient, but also to the manufacturer overall.

There are regulatory kinds of issues that have to be taken care of by the FDA, but that does not preclude the manufacturer from taking additional steps voluntarily, and I want to urge the manufacturer to consider that as it goes forward.

Thank you.

CHAIRMAN LASKEY:  Okay, on behalf of the Panel members, I want to thank the sponsor for a superb presentation and for sitting here with us until this late hour.  Thank you again.

This concludes the report and recommendations of the Panel on PMA P020045 from CryoCath Technologies for the Freezor Cardiac Cryoablation Catheter for cryoablation of cardiac tissue to treat patients with AV tachycardia.

Thank you again, and good evening.

(Whereupon, the foregoing matter was concluded at 6:06 p.m.)