Preface

In developing the optical design of the Micro-shutter infrared testbed three-lens collimator we geared our specifications toward commonly available test plates by United States manufactures of Infrared lenses.  We found that the sets of test plates fall into two groups as described in tables A3 and B3.  Therefore, our design specification is duplicated for the two groups of test plates available.  The design is described as [Micro-shutter IR testbed design] A and [Micro-shutter IR testbed design B. Manufacture of the three lens system using either design A or design B will be functionally equivalent.  Quotation need to be based on only one design type.  If other test plate groups are identified additional designs can be considered.

Micro-Shutter IR Testbed Collimator Design A

A1. Introduction………………………………………………………………………………………………2

A1.1 General Description……………………………………………………………………………………	2

A1.2 Environment……………………………………………………………………………………………	2

A1.3 Scope of Work…………………………………………………………………………………………	2

A2. Number of Lens Systems…………….………………………………………………………………….	3

A3. Optical Coatings…………………………………………………………………………………………3

A4. Lens Covers……………………………………………………………………………………………..	3

A5. Tolerance Data…………………………………………………………………………………………..	3

A5.1 General Tolerances…………………………………………………………………………………….	3

A5.1.2 Surface Tolerances…………………………………………………………………………………...	4

5.1.2.1 Surface Tolerances………...………………………………………………………..4	
5.1.2.2 Surface Wedge Tolerances……...………………………………………………….4

5.1.2.3 Index Tolerances……………………………………………………………………4

5.1.2.4 Element Tolerances…………………………………………………………………4

A6. Prescription Data…………………………………………………………………………………………5

A6.1 Surface and Element Data…..…………………………………………………………………………..5

6.1.1 Surface Data…………………………………………………………………………...5

6.1.2 Element Data…………………………………………………………………………..5

A6.2 Index of Refraction Data…………………………………………………………………….………….5

A6.3 Thermal Coefficient of Expansion Data…………………………………………………….…………..5

A7. Element Drawings………………………………………………………………………………………...6

A7.1 ZnSe Lens……….……………………………………………………………………………………….6

A7.2 BaF2 Lens………..………………………………………………………………………………………7

A7.3 LiF Lens…………………………………………………………………………………………………8

A8. Performance Requirements……………………………………………………………………………….8

Tables:

Table A1. General Testbed Collimator Lens Data………………………………………...…………………....3

Table A2. Tolerance Data Summary…………………………………………………….…………….….…….3

Table A3. Surface Tolerances……………………………………………………………….…………..………4

Table A4. Wedge Tolerances…………………………………………………….……………………….……..4

Table A5. Glass Tolerances…………………………………………………….………………………….……4

Table  A6. Element Tolerances…………………………………………………….………………………........4

Table A7.  Surface Data…………………………………………………….……………………………….…..5

Table A8. Element Data

Table A9.  Index Data…………………………………………………….……………………………………..5

Table A10.  TCE Data…………..…………………………………….…………………………………………5

Figures:

Figure A1. IR Testbed Collimator……...…...…………………………………………………….……………..2

Figure A2. ZnSe Lens.………………………...…………………………………………………………………6

Figure A3. BaF2 Lens.…………………………...………………………………………………………………7

Figure A4. LiF Lens...……………………………...……………………………………………………….……8

A1. Introduction

A1.1 General Introduction

This optical system will be used in an infrared testbed designed to test the optical performance micro-shutter (micro electro-mechanical) MEMs devices. The three lens collimator is used to take collimated light and create a diffraction limited point source at the micro-shutter device. After light passes through the shutter, a replica of the previous three lens collimator is used to re-collimate the light. The system was designed to create an f/15 testbed with a full field of view of 0.2500 at the micro-shutter array. The collimator is designed using Zemax to work at wavelengths from 633nm to 5um with no anti-reflection coatings. Although the lens assembly will be mounted in a barrel by the lens manufacturer, a flange is not required on the mount.

A1.2 Environment

This optical system will operate in a lab environment. As such, its design and manufacture shall incorporate appropriate safeguards to ensure that all performance goals and system stability are satisfied within this environment. The lens barrel mounting system shall maintain proper alignment of the optics over the specified survivability temperature range and handling of the system in a lab environment.

A1.3 Scope of Work

Figure A1 illustrates the lens layout of the collimator. Note that two of these lens assemblies will be required. One system will be used to create a point source from collimated light, the other will be used to create collimated light from a point source. The scope of work described in this document includes (1) the lens manufacture requirements, (2) the mounting system design, and (3) performance testing for the completed assembly. The scope of this work does not include incorporating the optical assembly into the infrared testbed set-up.

Although this lens system was designed to perform from 632.8nm to 5 microns, for the purposes of this document all wavelength related tolerance data and the prescription data is given with respect to the test wavelength of 632.8nm. Also, the dewar windows, which were modeled in the development of these lenses, have been removed for the purposes of this document.

Figure A1. IR Testbed Collimator Overview

The lens manufacture, tolerance and mounting system shall conform to the specifications given in Tables A2 – A10. Zemax prescription data is available to aid in the lens mounting and design layout. General lens data are listed in Table A1.

## Table A1. General Testbed Collimator Lens Data

| Surfaces | 6 |
| --- | --- |
| System Aperture (Entrance Pupil Diameter) | 56.26mm |
| Glass Catalog | Infrared |
| Design Temperature | 20ºC |
| Design Pressure | 1 atm |
| Operating Temperature | 20ºC 4ºC |
| Effective Focal Length | 706.181 |
| Back Focal Length | 621.212mm |
| Total Track | 669.852mm |
| Working F/# | 13.9 |
| Image Space N.A. | 0.0359 |
| Stop Diameter | 56.90mm |
| Design Wavelengths | 0.633, 1, 2, & 5 microns |
| Test Wavelength | 632.8 nm |

A2. Number of Lens Systems

Two lens systems will be manufactured as per the requirements set forth in this document.

A3. Optical Coatings

The lens elements shall not be treated with an anti-reflection optical coatings.

A4. Lens Covers

The lens assembly shall be equipped with lens covers only if this does not add additional cost to the lens assembly.

A5. Tolerance Data

A5.1 General Tolerances

The lens manufacture, alignment, and mounting shall conform to the following general tolerances. Tolerances for specific lenses will be defined later in this document.

## Table A2. Tolerance Data Summary

| Surface Figure | 0.1 Waves PV @ 632.8nm |
| --- | --- |
| Surface Quality | 60/40 Scratch Dig |
| Surface Roughness | 15-20 Å RMS |
| Material Index | +/-0.001 |
| Abbe Number | +/-0.5 |

A5.1.2 Surface Tolerances

Each radius has been fitted to test plate indicated. The power tolerance is the percent deviation of the radius of curvature from the nominal value. Irregularity is in double pass fringes at 2 microns.

A5.1.2.1 Surface Tolerances

#### Table A3. Surface Tolerances

| Surf | Radius (mm) | Test Plate # | Radius Tol | Thickness (mm) | Thickness Tol (mm) |
| --- | --- | --- | --- | --- | --- |
| 1 | 107.2134 | 4.2210 | 0.05% | 17.03 | +/-0.1 |
| 2 | 90.805 | 3.5750 | 0.05% | 10.61 | +/-0.1 |
| 3 | 115.8189 | 4.5598 | 0.1% | 8.0 | +/-0.1 |
| 4 | 605.028 | 23.8200 | 0.1% | 5.0 | +/-0.1 |
| 5 | -1412.494 | 55.6100 | 0.1% | 8.0 | +/-0.5 |
| 6 | 863.6 | 34.0000 | 0.1% | -- | -- |

### A5.1.2.2 Surface Wedge Tolerances

#### Table A4. Wedge Tolerances

| Surf | Surface Tilt X (mm) | Surface Tilt Y (mm) |
| --- | --- | --- |
| 1 | 0.005 | 0.005 |
| 2 | 0.005 | 0.005 |
| 3 | 0.015 | 0.015 |
| 4 | 0.015 | 0.015 |
| 5 | 0.015 | 0.015 |
| 6 | 0.015 | 0.015 |

### A5.1.2.3 Index Tolerances

Table A5. Glass Tolerances

| Glass | Index Tol | Abbe Tol |
| --- | --- | --- |
| ZnSe | 0.001 | 0.5 |
| BaF2 | 0.001 | 0.5 |
| LiF | 0.001 | 0.5 |

### A5.1.2.4 Element Tolerances

Table A6. Element Tolerances

| Element # | Srf1 | Srf2 | De-center X (mm) | De-center Y (mm) | Tilt X (degrees) | Tilt Y (degrees) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | 2 | 0.15 | 0.15 | 0.08 | 0.08 |
| 2 | 3 | 4 | 0.15 | 0.15 | 0.08 | 0.08 |
| 3 | 5 | 6 | 0.15 | 0.15 | 0.08 | 0.08 |

# A6. Prescription Data

Manufacture & testing assumed at STP (20ºC and 1 atm)

### A6.1 Surface and Element Data

A6.1.1 Surface Data

Table A7.  Surface Data

| Surf | Radius (mm) | Thickness (mm) | Thickness Label | Glass | Clear Aperature (mm) | Conic |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 107.2134 | 17.03 | t1 | ZnSe | 56.26 |  |
| 2 | 90.805 | 10.61 | t2 |  | 50.82 |  |
| 3 | 115.8189 | 8.0 | t3 | BaF2 | 51.42 |  |
| 4 | 605.028 | 5.0 | t4 |  | 50.86 |  |
| 5 | -1412.494 | 8.0 | t5 | LiF | 50.41 |  |
| 6 | 863.6 | -- |  |  | 49.80 |  |
| IMA | Infinity |  |  |  |  |  |

A6.1.2 Element Data

Table A8. Element Data

| Element # | Srf1 | Srf2 | Lens Diameter (mm) | Edge Thickness (mm) |
| --- | --- | --- | --- | --- |
| 1 | 1 | 2 | 63.50 | 17.82 |
| 2 | 3 | 4 | 57.15 | 5.27 |
| 3 | 5 | 6 | 57.15 | 8.72 |

### A6.2 Index of Refraction Data

Table A9.  Index Data

| Glass | Temp (ºC) | Pressure (atm) | Index for(632.8nm) |
| --- | --- | --- | --- |
| ZnSe | 20.00 | 1.00 | 2.5906 |
| BaF2 | 20.00 | 1.00 | 1.4732 |
| LiF | 20.00 | 1.00 | 1.3913 |

### A6.3 Thermal Coefficient of Expansion Data

Table A10.  TCE Data

| Glass | TCE () |
| --- | --- |
| ZnSe | 7.1 |
| BaF2 | 18.4 |
| LiF | 34.4 |

A7. Element Drawings

A7.1 ZnSe Lens

Figure A2. ZnSe Lens

A7.2 BaF2 Lens

Figure A2. BaF2 Lens

A7.3 LiF Lens

Figure A4. Element #3

A8.  Performance Requirements

Each lens system should be tested to make sure that in single-pass its wavefront error does not exceed 0.30 waves rms.

Micro-Shutter IR Testbed Collimator Design B

B1. Introduction……………………………………………………………………………………….…….10

B1.1 General Description……………………………………………………………………………….…...10

B1.2 Environment………………………………………………………………………………………...…10

B1.3 Scope of Work……………………………………………………………………………………....…10

B2. Number of Lens Systems…………….…………………………………………………………….……11

B3. Optical Coatings…………………………………………………………………………………………11

B4. Lens Covers………………………………………………………………………………….……..……11

B5. Tolerance Data………………………………………………………………………………….………..11

B5.1 General Tolerances…………………………………………………………………………………..…11

B5.1.2 Surface Tolerances……………………………………………………………………………………12

B5.1.2.1 Surface Tolerances………...………………………………………………………12	
B5.1.2.2 Surface Wedge Tolerances……...…………………………………………………12

B5.1.2.3 Index Tolerances……………………………………………………………...……12

B5.1.2.4 Element Tolerances………………………………………………….………..……12

B6. Prescription Data………………………………………………………………………………………….13

B6.1 Surface and Element Data…..……………………………………………………………………....……13

B6.1.1 Surface Data…………………………………………………………………….…….13

B6.1.2 Element Data………………………………………………………………………….13

B6.2 Index of Refraction Data………………………………………………………………..………..………13

B6.3 Thermal Coefficient of Expansion Data…………………………………………………….……………13

B7. Element Drawings…………………………………………………………………………………………13

B7.1 ZnSe Lens……….…………………………………………………………………………………..……13

B7.2 BaF2 Lens………..………………………………………………………………………………….…….14

B7.3 LiF Lens…………………………………………………………………………………………….…….16

B8. Performance Requirements…………………………………………………………………………...……16

Tables:

Table B1. General Testbed Collimator Lens Data………………………………………...…………….……...11

Table B2. Tolerance Data Summary…………………………………………………….………………….…..11

Table B3. Surface Tolerances……………………………………………………………….………………….12

Table B4. Wedge Tolerances…………………………………………………….……………………….…….12

Table B5. Glass Tolerances…………………………………………………….………………………….……12

Table B6. Element Tolerances…………………………………………………….………………………..…...12

Table B7.  Surface Data…………………………………………………….…………………………….……..12

Table B8. Element Data

Table B9.  Index Data…………………………………………………….………………………………….….13

Table B10.  TCE Data…………..…………………………………….……………………………….………...13

Figures:

Figure B1. IR Testbed Collimator……...…...…………………………………………………….….…………10

Figure B2. ZnSe Lens.………………………...……………………………………………………….………..14

Figure B3. BaF2 Lens.…………………………...……………………………………………….…….………..14

Figure B4. LiF Lens...……………………………...…………………………………………….……….……..16

B1. Introduction

B1.1 General Introduction

This optical system will be used in an infrared testbed designed to test the optical performance micro-shutter mems devices. The three lens collimator is used to take collimated light and create a diffraction limited point source at the micro-shutter device. After light passes through the shutter, a replica of the previous three lens collimator is used to re-collimate the light. The system was designed to create an f/15 testbed with a full field of view of 0.2500 at the micro-shutter array. The collimator is designed using Zemax to work at wavelengths from 633nm to 5um with no anti-reflection coatings. Although the lens assembly will be mounted in a barrel by the lens manufacturer, a flange is not required on the mount.

B1.2 Environment

This optical system will operate in a lab environment. As such, its design and manufacture shall incorporate appropriate safeguards to ensure that all performance goals and system stability are satisfied within this environment. The lens barrel mounting system shall maintain proper alignment of the optics over the specified survivability temperature range and handling of the system in a lab environment.

B1.3 Scope of Work

Figure B1 illustrates the lens layout of the collimator. Note that two of these lens assemblies will be required. One system will be used to create a point source from collimated light, the other will be used to create collimated light from a point source. The scope of work described in this document includes (1) the lens manufacture requirements, (2) the mounting system design, and (3) performance testing for the completed assembly. The scope of this work does not include incorporating the optical assembly into the infrared testbed set-up.

Although this lens system was designed to perform from 632.8nm to 5 microns, for the purposes of this document all wavelength related tolerance data and the prescription data is given with respect to the test wavelength of 632.8nm. Also, the dewar windows, which were modeled in the development of these lenses, have been removed for the purposes of this document.

Figure B1. IR Testbed Collimator Overview

The lens manufacture, tolerance and mounting system shall conform to the specifications given in Tables 2 – 10. Zemax prescription data is available to aid in the lens mounting and design layout. General lens data are listed in Table B1.

## Table B1. General Testbed Collimator Lens Data

| Surfaces | 6 |
| --- | --- |
| System Aperture (Entrance Pupil Diameter) | 56.90mm |
| Glass Catalog | Infrared |
| Design Temperature | 20ºC |
| Design Pressure | 1 atm |
| Operating Temperature | 20ºC 4ºC |
| Effective Focal Length | 702.892 mm |
| Back Focal Length | 616.923mm |
| Total Track | 660.593mm |
| Working F/# | 13.8 |
| Image Space N.A. | 0.0361 |
| Stop Diameter | 56.90mm |
| Design Wavelengths | 0.633, 1, 2, & 5 microns |
| Test Wavelength | 632.8 nm |

B2. Number of Lens Systems

Two lens systems will be manufactured as per the requirements set forth in this document.

B3. Optical Coatings

The lens elements shall not be treated with an anti-reflection optical coatings.

B4. Lens Covers

The lens assembly shall be equipped with lens covers only if this does not add additional cost to the lens assembly.

B5. Tolerance Data

B5.1 General Tolerances

The lens manufacture, alignment, and mounting shall conform to the following general tolerances. Tolerances for specific lenses will be defined later in this document.

## Table B2. Tolerance Data Summary

| Surface Figure | 0.1 Waves PV @ 632.8nm |
| --- | --- |
| Surface Quality | 60/40 Scratch Dig |
| Surface Roughness | 15-20 Å RMS |
| Material Index | 0.001 |
| Abbe Number | 0.5 |

B5.1.2 Surface Tolerances

Each radius has been fitted to test plate indicated. The power tolerance is the percent deviation of the radius of curvature from the nominal value. Irregularity is in double pass fringes at 2 microns.

B5.1.2.1 Surface Tolerances

#### Table B3. Surface Tolerances

| Surf | Radius (mm) | Test Plate # | Radius Tol | Thickness (mm) | Thickness Tol (mm) |
| --- | --- | --- | --- | --- | --- |
| 1 | 106.63 | TP000469 | 0.05% | 16.92 | +/-0.1 |
| 2 | 90.28 | TP000429 | 0.05% | 7.63 | +/-0.1 |
| 3 | 111.31 | TP000477 | 0.1% | 6.35 | +/-0.1 |
| 4 | 518.7 | TP000841 | 0.1% | 6.44 | +/-0.1 |
| 5 | -471.17 | TP000823 | 0.1% | 6.35 | +/-0.5 |
| 6 | -3750.2 | TP001015 | 0.1% | -- | -- |

### B5.1.2.2 Surface Wedge Tolerances

#### Table B4. Wedge Tolerances

| Surf | Surface Tilt X (mm) | Surface Tilt Y (mm) |
| --- | --- | --- |
| 1 | 0.009 | 0.009 |
| 2 | 0.009 | 0.009 |
| 3 | 0.018 | 0.018 |
| 4 | 0.018 | 0.018 |
| 5 | 0.018 | 0.018 |
| 6 | 0.018 | 0.018 |

### B5.1.2.3 Index Tolerances

TableB 5. Glass Tolerances

| Glass | Index Tol | Abbe Tol |
| --- | --- | --- |
| ZnSe | 0.001 | 0.5 |
| BaF2 | 0.001 | 0.5 |
| LiF | 0.001 | 0.5 |

### B5.1.2.4 Element Tolerances

Table B6. Element Tolerances

| Element # | Srf1 | Srf2 | De-center X (mm) | De-center Y (mm) | Tilt X (degrees) | Tilt Y (degrees) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | 2 | 0.25 | 0.25 | 0.08 | 0.08 |
| 2 | 3 | 4 | 0.20 | 0.20 | 0.08 | 0.08 |
| 3 | 5 | 6 | 0.25 | 0.25 | 0.08 | 0.08 |

# B6. Prescription Data

Manufacture & testing assumed at STP (20ºC and 1 atm)

### B6.1 Surface and Element Data

B6.1.1 Surface Data

Table B7.  Surface Data

| Surf | Radius (mm) | Thickness (mm) | Thickness Label | Glass | Clear Aperature (mm) | Conic |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 106.63 | 16.92 | t1 | ZnSe | 56.90 |  |
| 2 | 90.28 | 7.63 | t2 |  | 51.45 |  |
| 3 | 111.31 | 6.35 | t3 | BaF2 | 51.89 |  |
| 4 | 518.7 | 6.44 | t4 |  | 51.47 |  |
| 5 | -471.17 | 6.35 | t5 | LiF | 50.95 |  |
| 6 | -3750.2 | -- |  |  | 50.54 |  |
| IMA | Infinity |  |  |  |  |  |

B6.1.2 Element Data

Table B8. Element Data

| Element # | Srf1 | Srf2 | Lens Diameter (mm) | Edge Thickness (mm) |
| --- | --- | --- | --- | --- |
| 1 | 1 | 2 | 63.50 | 16.80 |
| 2 | 3 | 4 | 57.15 | 3.92 |
| 3 | 5 | 6 | 57.15 | 6.95 |

### B6.2 Index of Refraction Data

Table B9.  Index Data

| Glass | Temp (ºC) | Pressure (atm) | Index for(632.8nm) |
| --- | --- | --- | --- |
| ZnSe | 20.00 | 1.00 | 2.5906 |
| BaF2 | 20.00 | 1.00 | 1.4732 |
| LiF | 20.00 | 1.00 | 1.3913 |

### B6.3 Thermal Coefficient of Expansion Data

Table B10.  TCE Data

| Glass | TCE () |
| --- | --- |
| ZnSe | 7.1 |
| BaF2 | 18.4 |
| LiF | 34.4 |

B7. Element Drawings

B7.1 ZnSe Lens

Figure B2. ZnSe Lens

B7.2 BaF2 Lens

Figure B2. BaF2 Lens

B7.3 LiF Lens

Figure B4. Element #3

B8.  Performance Requirements

Each lens system should be tested to make sure that in single-pass its wavefront error does not exceed 0.1 waves rms.

# Text boxes

t1

t2

t3

t4

t5

For t0, t1, t2, t3, t4, and t5 values, see

Thickness Data (see Table A7)

Mounting for 3 lens system to be suggested by the manufacturer.

t1

t2

t3

t4

t5

For t0, t1, t2, t3, t4, and t5 values, see

Thickness Data (see Table B7)

Mounting for 3 lens system to be suggested by the manufacturer.