3.9 GHz Flash Cryostat – Pressure Vessel Spec

Bernd Petersen DESY-MKS

May 7,2007

1.General:

Fermi National Accelerator Laboratory (FNAL) will manufacture a 3.9 GHz Cryostat in the framework of a collaboration contract with DESY (see email and note from Helen Edwards dated April 25,2007.) The 3.9 GHz Cryostat will be installed in the FLASH-linac at DESY Hamburg.

The 3.9 GHz FLASH cryostat has to conform to the European legal regulations. DESY will be the importer into the European Union and has to certificate the conformity to the European regulations.

As part of the cryogenic installations of the FLASH-linac (former TTF2-linac) the 3.9 GHz cryostat will be treated according to the TTF2 risk analysis (Gefahrenanalyse nach Anhang I der Druckgeräterichtlinie DGRL, May 8, 2003). With reference to this analysis, the cryogenic equipment of the 3.9 GHz crystat will be assessed to fall below module I of the DGRL, as a consequence the cryogenic installations require no authorized third party inspections.

According to our risk analysis, the 3.9 GHz cryostat shall be treated in line with the standard 1.3 GHz cryomodules: the niobium material of the cavities can not be certified for the use at low temperatures. The thermal shields and the vacuum vessel of the cryostat act as containment for the inner 2K pressure vessel (see also DESY note D5/591, December 1, 1993.)

The DESY safety rules are fulfilled as soon as the cryostat is conforming to the EU rules.

2. Technical rules

The design and construction of the 3.9 GHz cryostat has to be conforming to the DGRL. According to the DGRL, the manufacturer has to take the responsibility for the application of suited technical rules, which are not defined in detail in the DGRL. In case of the 3.9 GHz cryostat, DESY as the importer takes action for the manufacturer outside the EU.

At DESY we apply the AD2000 technical rules or the European harmonized rules. But according to our risk analysis also the US ASME regulations for pressure vessels can be applied. This is underlined by the fact that since a decade, we have already cryogenic components in use at TTF and FLASH, which were designed and constructed according to ASME regulations (FLASH feed box, test cryostats).

2.1 Choice of materials for low temperature use

The choice of materials for use at low temperatures must accord to the European harmonized rules (see attachmentA1). If you will use the austenitic stainless steels  316L (1.4435 DIN EN 13445-2) or 316LN (1.4429)  you should be perfectly in line not only with the EU rules but also with ASME.

On default, materials may be used, which are allowed by the ASME rules for pressure vessels at low temperatures (but steel 321 must be avoided!). For the Ti-vessel we use titanium  DIN/EN 3.7035, this corresponds to Ti grade 2.

According to the DGRL there is no way to certify niobium for the use at low temperatures.

All ‘pressure vessel parts’ – all parts of the 2K, 4.5K and 40/80K circuits (apart from niobium)- require material certificates. (On default, the content and quality of the certificates should correspond to the ASME code for pressure vessels and use at low temperatures.)

The weld additives used for the pressure vessel parts shall correspond to the ASME code (or related rules), shall be suited for low temperatures and need material certificates.

3. Pressure vessel spec

Apart from the cavity made from niobium, all parts of the 3.9 GHz cryostat shall be designed and manufactured in accordance to the ASME pressure vessel code for the use at low temperatures.

3.1 Spec of cryogenic circuits

3.1.1   40-80K shield circuit

Design temperature: 40K ; design pressure: 20 bar abs

    - 4.5K shield circuit

Design temperature: 4K; design pressure: 20 bar abs

    - 2K circuit

design temperature: 2K; design pressure (*):  2 bar abs

(*) at room temperature an outer pressure not higher than 2 bar abs (cavity under vacuum) can be applied to the cavities. As a consequence, all safety valves of the 2K circuit of FLASH are set to an opening pressure of less than 2 bar abs. At cold the cavities may tolerate pressures up to 4 bar abs. Therefore it is recommended to design all parts of the 2K circuit – apart from the cavities – to a pressure of 4 bar abs, if possible.

4. Standards & Procedures

All standards and procedures according to the ASME pressure vessel code (including test procedures)  shall be applied. In particular:

4.1 Complete certification for all pressure vessel materials is required.

4.2 All welding procedures have to be supervised and documented.

4.3 All test procedures, in particular pressure tests, have to be documented.

4.4 A complete technical documentation has to be delivered (technical drawings, engineering docs etc.)

4.5 An operation description is required.

A 1)

European Harmonized Rules

Stainless steels for the use at low temperatures

DIN EN 13445-2

B.2.2.4  Lowest material temperature for austenitic stainless steels

Apply also for 2K !

Tabelle B.2-2—Niedrigste Werkstofftemperatur TM nichtrostender austenitischer  Stähle

(material spec corresponding to ASTM type AISI may differ  ! On default apply ASME spec)

| Material spec | DIN EN number | ASTM type AISI | TM (in °C) |
| --- | --- | --- | --- |
| X1NiCrMoCu 31-27-4 | 1.4563 |  | -270 |
| X1CrNiMoN 25-22-2 | 1.4466 |  |  |
| X1CrNi 25-21 | 1.4335 | 310 L |  |
| X2CrNiMoN 17-13-3 | 1.4429 | 316 LN |  |
| X2CrNiMoN 17-11-2 | 1.4406 |  |  |
| X2CrNiMoN 18-12-4 | 1.4434 |  |  |
| X2CrNiMo 18-15-4 | 1.4438 | 317 L |  |
| X2CrNiN 18-10 | 1.4311 | 304 LN |  |
| X2CrNiMo 18-14-3 | 1.4435 | 316 L |  |
| X2CrNi 19-11 | 1.4306 | 304 L |  |