| Name(s) of Risk Team Members: R. Colichio, A. Emrick, Setlow, Shea-McCarthy, Lin, Kumaran Completed: 10/4/06 | Point Value → Parameter ↓ | 1 | 2 | 3 | 4 | 5 |
| --- | --- | --- | --- | --- | --- | --- |
| Job Title: General Cryogenic Work Job Number or Job Identifier: DJ-JRA16-Cryogens | Frequency (B) | <once/year | <once/month | <once/week | <once/shift | >once/shift |
| Job Description: General work with non-combustible cryogenic fluids and systems. Does not include liquid hydrogen or liquid oxygen. | Severity (C) | First Aid Only | Medical Treatment | Lost Time | Partial Disability | Death or Permanent Disability |
|  | Likelihood (D) | Extremely Unlikely <<1x/20yrs | Unlikely 1x/10-20yrs | Possible >1x/10-20yrs | Probable 1x/yr | Multiple >1x/yr |
| Training and Procedure List (Optional): |  |  |  |  |  |  |
| Approved by: R. Colichio Date: 10/6/06 Rev. #: 0 |  |  |  |  |  |  |
| Stressors (if applicable, please list all): Moving dewars outdoors | Reason for Revision (if applicable): | Comments: |

|  | Before Controls |  | After Initial Controls |  | After Additional Controls |
| --- | --- | --- | --- | --- | --- |
| Job Step / Task | Hazard | Stressors Y/N | # of People A | Frequency B A | Severity C | Likelihood D | Risk* AxBxCxD | Initial Controls | # of People A | Frequency B | Severity C | Likelihood D | Risk* AxBxCxD | Control(s) Added to Reduce Risk | # of People A | Frequency B | Severity C | Likelihood D | Risk* AxBxCxD | % Risk Reduction |
| Moving storage vessel (dewar) to/from location | Overexertion injuries- excessive lifting, pushing, pulling, holding or carrying | N | 1 | 3 | 3 | 3 | 27 | Wheeled dewar design, Use of carts, Cryogloves or no gloves. | 1 | 3 | 3 | 2 | 18 |  |  |  |  |  |  |  |
|  |  | Y | 1 | 2 | 3 | 4 | 24 |  | 1 | 2 | 3 | 3 | 18 |  |  |  |  |  |  |  |
| Moving storage vessel (dewar) to/from location | Being struck by object (Dewar tipping over) | N | 1 | 3 | 3 | 3 | 27 | Design of dewar, attention to speed, threshold design | 1 | 3 | 3 | 2 | 18 | Removal of bump on 463 loading dock. | 1 | 3 | 3 | 1 | 9 | 50% |
| Transferring liquid nitrogen to storage dewar at fill station. | Contact with temperature extremes (Frost bite) | N | 1 | 4 | 2 | 5 | 40 | Gloves, face shield, proper clothing, training, posting | 1 | 4 | 1 | 2 | 8 |  |  |  |  |  |  |  |
|  | Oxygen deficiency | N | 1 | 4 | 5 | 3 | 60 | ODH alarms (including fill valve shutdown, auto door open and Control Room alarm response at NSLS), postings, training, ODH evaluations | 1 | 4 | 1 | 2 | 8 |  |  |  |  |  |  |  |
| Connecting, transferring and disconnecting storage dewar to/from experimental equipment | Contact with temperature extremes (Frost bite) | N | 1 | 3 | 3 | 3 | 27 | Hand protection, eye protection, proper clothing, training | 1 | 3 | 1 | 2 | 6 |  |  |  |  |  |  |  |
|  | Falls to lower level from ladder or other cryostat access | N | 1 | 2 | 4 | 3 | 24 | Ladder/platform design, equipment inspection, training, Tier I | 1 | 2 | 4 | 2 | 16 |  |  |  |  |  |  |  |
|  | Explosion from plugged dewar | N | 1 | 2 | 5 | 3 | 30 | Training, emergency response procedures, vessel design/certification, pressure relief valves | 1 | 2 | 3 | 2 | 12 |  |  |  |  |  |  |  |
| Hand transfer of cryogenic fluid to system | Contact with temperature extremes (Frost bite) | N | 1 | 4 | 3 | 4 | 48 | Hand protection, eye protection, proper clothing, training | 1 | 4 | 2 | 2 | 16 |  |  |  |  |  |  |  |
| Moving helium dewars | Oxygen deficiency | N | 1 | 1 | 5 | 3 | 15 | Training (No personnel allowed to ride with dewar in elevator) | 1 | 1 | 5 | 2 | 10 |  |  |  |  |  |  |  |
|  | Overexertion injuries- excessive lifting, pushing, pulling, holding or carrying | N | 1 | 1 | 3 | 3 | 9 | Wheeled dewar design | 1 | 1 | 2 | 3 | 6 |  |  |  |  |  |  |  |
|  | Being struck by an object due to pressure release | N | 1 | 1 | 3 | 3 | 9 | Training, eye protection, equipment inspection, vessel design/certification | 1 | 1 | 2 | 2 | 4 |  |  |  |  |  |  |  |
|  | Being struck by object dropped during lift | N | 1 | 1 | 4 | 3 | 12 | Training, certification, hard hat, inspection of lifting devices | 1 | 1 | 4 | 2 | 8 |  |  |  |  |  |  |  |
| Work in close proximity to cryogenic auto fill system | Oxygen deficiency | N | 1 | 5 | 5 | 3 | 75 | Training, ventilation, ODH calculations, engineering design, system verification | 1 | 5 | 5 | 1 | 25 |  |  |  |  |  |  |  |
|  | Contact with temperature extremes Frost bite | N | 1 | 5 | 2 | 3 | 30 | Training, postings on large systems | 1 | 5 | 2 | 2 | 20 |  |  |  |  |  |  |  |
| Manipulating samples in liquid nitrogen bath | Contact with temperature extremes (Frost bite) | N | 1 | 5 | 4 | 3 | 60 | Training, safety glasses | 1 | 5 | 2 | 2 | 20 |  |  |  |  |  |  |  |
| Mounting and manipulating samples within cryostream | Contact with temperature extremes (Frost bite) | N | 1 | 5 | 3 | 3 | 27 | Training, safety glasses | 1 | 5 | 2 | 2 | 20 |  |  |  |  |  |  |  |
| Disposing of waste liquid | If disposed improperly (on floor) --Frostbite, trip hazard if tile cracks, electrical shock | N | 1 | 3 | 3 | 3 | 27 | Glasses, gloves, proper disposal containers or evaporation | 1 | 3 | 3 | 2 | 18 | Proper disposal containers and communication of rqmts | 1 | 3 | 3 | 1 | 9 | 50% |
| Further Description of Controls Added to Reduce Risk: Have researchers try “alligator gloves” (R. Colichio will purchase and distribute for pilot). Conduct ODH calculations on tissue culture room and 9-401 in Bldg 490 and request NSLS conduct for “triangle lab” across from X25. Investigate and purchase proper disposal containers so researchers don’t throw on floor to evaporate. Recommend revision to cryogen postings to include “Do not unplug clogged dewar—call 2222” |
| *Risk: | 0 to 20 Negligible | 21 to 40 Acceptable | 41 to 60 Moderate | 61 to 80 Substantial | 81 or greater Intolerable |