# Iridium Oxide pH Sensor for Nanoscale Electrochemical Cell Monitoring

Erik Douglas, E298A-EE290B, 6 May 2004

# Abstract

A device was developed to record a spatial and temporal map of cell-surface pH on the sub-cellular size scale.  The pH can be correlated with cell metabolism, thereby yielding valuable information about the status and activity of the cell.  An array of 100 nm wide gold electrodes was patterned with the LBL CXRO electron beam writer, and a pH sensitive layer of iridium oxide was electrodeposited on the surface.  The pH response of the electrodeposited IrO2 film sensor was characterized, and found to be linear, with a slightly super-Nernstian response of -65.3 mV/pH unit.

# I.  Introduction

Single cells are the basic unit of biology, and nanofabrication holds great promise for electrochemical interaction with these micron-scale components.  Communication with single cells could lead to biomimetic neuroprosthetics and produce personalized drug screening.  Since many cells are on the size scale of a few microns in diameter, nanoscale devices produced by electron beam lithography could allow for high resolution spatial and temporal interrogation of the surface of a cell.

A device is envisioned which would isolate a target cell on an array of nanoscale electrodes, with surfaces tailored to detect biologically relevant signals and provide electrical stimulation.  For example, because cells acidify their environment in the course of normal metabolism, the extra-cellular pH can be used to measure cell metabolic rate.  This concept has lead to the development of the cytosensor microphysiometer which has been a valuable tool in cell surface receptor pharmacology.  The directed isolation of a single cell also provides a considerable challenge, but the needs of this project could be met by elevating the surface around the array, thereby placing the electrodes in a sunken well where a cell will settle during flow.  It would also be possible to capture the target cell on an electrochemical device by modifying either the device or cell surface.  Figure 1 displays an example of a single cell captured on a microfabricated platinum dot using protein adhesion facilitated by this type of cell-surface modification.

The present work describes the fabrication of an array of 100 nm wide pH sensitive electrodes for use in cell monitoring.  To measure pH using this array, the potential was measured between an electrode and a common, distant Ag/AgCl reference electrode.  The voltage and pH are linearly related by the Nernst equation, which describes the expected equilibrium potential of an electrode: E=E0 + (2.3RT/nF)logAH+.  Calibrating the sensor at two points gives the slope of the line, which allows the conversion from voltage to pH.  The pH response of the device was characterized using standardized pH buffer solutions over the physiologically relevant range of 4 to 10.

II. Experiment

# Fabrication

A two-by-sixteen rectangular array of 100 nm wide gold electrodes with a chromium adhesion layer was patterned on a glass substrate by metal evaporation and the lift-off technique (20 and 80 nm thick, respectively).  Glass was chosen as the substrate because it will facilitate visualization of the cells in subsequent experiments.

To allow for the use of glass, the KRS resist was coated with a layer of the water-soluble conductive polymer AquaSave to prevent charge buildup.  While the results were generally good with this approach, Figure 4b illustrates the fact that the AquaSave was sometimes not effective in preventing distortion of the pattern due to charging.  The AquaSave was readily removed in the development process by rinsing with deioized water after exposure.

The electron beam part of the array pattern is roughly 500 m on a side, and the larger fan-out leads were be patterned by optical lithography in the UC Berkeley Microlab.  The layout was designed to ensure that the leads are all of equal area to prevent differences in resistance.  In order to fabricate the array with one electron beam exposure, a simple array pattern was chosen.  The pattern, shown schematically in Figure 2, is two columns of 100 nm wide lines separated within a column by 400 nm, and between columns by 1 m.

The larger, optically patterned, leads terminated in 100-by-400 m pads which were to be addressed by gold wires with a wire bonder.  This fan-out contact layer was aligned to the features in the first layer, and completed with metal evaporation and lift-off.  The gold layer was poorly adhered, however, and pulled off of the surface when the wire bonder was used.  This resulted in tests of the device being limited to the collective response of each column, with a wire connected across the contact pads.

The active area was isolated by masking the rest with a 15 m thick layer of the rugged negative photoresist SU-8, which will insulate the gold wires, and also form a well into which a single cell will settle during flow in upcoming experiments.  The remaining exposed area of the gold electrodes was electroplated with iridium oxide (IrO2) following the protocol of Yamanaka.  Anodic electrodeposition at 10mA/cm2 for 150 sec using a gold counter electrode produced a 120 nm thick layer of IrO2.  A second “dummy” gold electrode, with an area of 0.15 cm2 was also plated in order to improve control over the total current density.  A bluish film was observed on this electrode after deposition which confirmed the presence of the iridium oxide.

Characterization

The electrode device was characterized using calibrated pH buffer solutions at 4, 7 and 10, and an Ag/AgCl reference electrode.  Initial tests recorded an overload resistance between the nanoelectrodes and the reference, which prevented measurements of voltage.  Instead, the larger electrode which was simultaneously plated with IrO2 was used to characterize the sensitivity of the device.  This may be due to blockage by, or misalignment of, the SU-8 layer, or disconnects in the electrode array.  The results which were obtained for the larger electrode should be very similar to those with a functional nanoelectrode device.

Though it has not yet been accomplished, a single-cell experiment would make use of a PDMS microchannel formed with a mold made from SU-8.  A syringe pump would drive the liquid flow, and as the cells passed over the well, one would eventually fall in and fill it, so the rest would pass by.  The cell surface pH would be monitored, and then the cell would be killed with heat.  If the device was successful, the cell death would register as an increased pH, since it is no longer acidifying its local environment by pumping out H+ ions.  The characterization of the device has so far been limited to its native pH sensing properties, due to the complex measurement compensation which is required when the interrogation volume is made very small by the presence of a cell, as well as the aforementioned complications with fabrication.  As discussed by Slowinska and Majda, the very small volume causes the measurement to use up the ions being sensed and substantially alter the pH being measured.  Single cell experiments are planned in the near future which will take this compensation into account and allow the device to serve its intended purpose of single-cell monitoring.

# III. Results and Discussion

Following fabrication and electrodeposition, the pH response was characterized by measuring calibrated buffer solutions over the physiological range of pH 4 to 10.  Because the array device seemed to be inoperable, the characterization data in Figure 5 were measured with the larger (0.15 cm2) iridium oxide coated gold electrode.  This electrode was seen to have a linear pH response of -65.3 mV/pH units over the range of 4-10, with an accuracy of (0.25 pH units.  This result is in close agreement with the literature, and presents an exciting opportunity for further development of similar nanoscale sensors.

Iridium oxide appears to be a good choice for a pH sensitive layer because of its strong response and straightforward patterning.  Iridium oxide films have the attractive qualities of being stable over a wide pH and temperature range, fast response, and low disturbance by redox interference and stirring.  Marzouk et al. have demonstrated the use of anodic electrodeposited iridium oxide film (AEIROF) for use in physiological measurements.  They reported a super-Nernstian (exceeding the Nernst equation which describes the expected equilibrium potential of an electrode) response of –63.5 mV/pH unit after conditioning, and an accuracy of (0.02 pH units, which is quite close to the sensitivity described in the current work.  They also found that the pH selectivity of the film was not compromised by the presence of other physiological cations such as Na+, K+, Ca2+, and Mg2+.  This property is very encouraging for use in future integrated multi-ion-selective devices, and allowed that group to conduct the simultaneous measurement of pH, K+ and lactate using other selective electrodes.  The probes they used were on the order of 1 mm in diameter, but they state that greater spatial and temporal resolution would allow them to better understand the mechanism and interactions of the biological system in question.

One immediate application of this technology could be a feedback-control system for electroporation of single cells.  Electroporation is often used to introduce foreign material into cells, such as DNA for transfection, but the instrumentation provides only a crude interface with cells.  Huang and Rubinsky described a system which detects the formation of pores in a single cell membrane by measuring current passing through them under applied potential.  Such a system indicates the onset of membrane permeability, but does not provide information about the location or the number of pores.  A nanoscale array of electrodes with some having a pH sensitive layer and others of bare metal could do just that, by measuring the amount and location of the change in pH levels.

# IV. Conclusions

Nanoscale single-cell interface devices promise to elucidate the fundamental workings of the cell, and hold potential for applications such as drug screening and neural prosthetics.  Electron beam lithography allows for the generation of probes which are well below the size scale of a single cell.  The pH sensitive electrode demonstrated herein could be an important step toward this goal.

# V. Acknowledgements

The author would like to thank Alex Liddle, Erik Anderson, and Josei Lee for their enormous help in fabrication, and Nick Toriello for his help with electrodeposition.  The work took place in the CXRO Nanofabrication Laboratory and the UC Berkeley Microlab.

# Endnotes

McConnell, H. et al.  SCIENCE 257 (5078): 1906-1912 SEP 25 1992.

Huang et al.  ANAL CHEM 76 (2): 483-488 JAN 15 2004.

Yamanaka, K.  JAPANESE J APP PHYS 28 (4): 632-637 APR 1998.

Slowinska, K. and Majda, M.  J. SOLID STATE ELEC CHEM In Press.

Fog, A. and Buck, R.P.  SENS ACTUAT 5: 137-146 1984.

Marzouk, S.A.M. et al.  ANAL CHEM 70 (23): 5054-5061 DEC 1 1998.

Huang, Y. and Rubinsky, B.  SENS ACTUAT A 89: 242-249, 2001.

# Text boxes

Figure 3 SEM image of the patterned resist indicating the final electrode design.

Figure 2 Conceptual layout of the device.  The lines are gold, masked by the gray SU-8, and selectively plated with the blue IrO2.  The SU-8 forms a well around the electrodes to capture the cell.

Figure 1 A single cell captured on a Pt dot via protein-based adhesion.  Bar = 8 m.  (Douglas, unpublished results.)

Figure 5 Characterization of the electrodeposited IrO2 pH sensor reveals a linear response of -65.3 mV/pH unit.

Figure 4 Left: Optical micrograph of the final device with contacts and SU-8 insulation removed over the active site.  Right: Distorted pattern due to charging.