December 2005

Performance of Precision Floating Gate Analog Voltage References at Cryogenic Temperatures

Richard Patterson, NASA Glenn Research Center

Ahmad Hammoud, QSS Group, Inc. / NASA GRC

Background

Precision voltage references based on Intersil’s proprietary Floating Gate Analog (FGA)TM technology are reported to feature very low temperature coefficient, excellent long term stability, and high accuracy [1–2].  They consume very little power, exhibit very low drift and noise, and provide excellent line and load regulation.  Their extremely ultra-low supply currents allow these devices to be powered-on continuously as they present almost no load to the power source, especially in battery-powered circuits.  In addition, reference devices achieve their highest accuracy and lowest drift when powered up continuously [2].  They can source and sink currents up to about 10 mA and are specified for operation from –40 (C to +125 (C [1-2]. The performance of two versions of these FGA voltage reference chips at cryogenic temperatures was investigated in this work.

Test Procedure

Two series of Intersil’s precision FGATM voltage reference chips were mounted on printed circuit boards for evaluation in the temperature range between +20 (C to –195 (C. These devices, which comprised of Intersil ISL60002BI and X60008B, were characterized at test temperatures of 20, -50, -100, -150 and -195 (C in a liquid nitrogen cooled environmental chamber.  Each device performance was evaluated in terms of its output voltage and supply current at different input voltage levels as a function of temperature.  Line regulation was also established at four load levels and at different temperatures.  Finally, the effect of temperature on load regulation of these devices was also determined.  Table I lists some of the manufacturer’s device specifications [1-2].

Table I.  Manufacturer’s specifications of FGATM voltage reference chips.

| Parameter (Unit) | ISL60002BI | X60008B |
| --- | --- | --- |
| Output Voltage Temp. Coeff., TC VOUT (ppm/°C) | 20 | 3 |
| Supply Current, IS (nA) | 400 | 500 |
| Source & Sink Current, IO (mA) | 7 | 10 |
| Supply Voltage Range, VS (V) | 2.7 – 5.5 | 4.5 - 6.5 |
| Output Voltage, VO (V) | 2.5 | 2.5 |
| Temperature Range, T (°C) | -40 - +85 | -40 - +85 |
| Package | 8 lead SOIC | 8 lead SOIC |
| Lot # | MZ41301A | MZ41406A |

Results and Discussion

Output Voltage

Figure 1 shows the variation in the output voltage of the ISL60002 device as a function of temperature.  The data was obtained at four different input voltages under no load conditions.  It can be clearly seen that, regardless of the input voltage level, the output voltage of this reference exhibited a gradual but very minute drop with decrease in temperature.  At an input voltage of 4 V for example, the output value had a value of 2.5003 V at room temperature but it decreased to 2.4996 V at -195 (C temperature.  This corresponds to an output voltage temperature coefficient of about 3.25 ppm/(C.  Similar behavior was obtained for the X60008 device in terms of its output variation with temperature, as depicted in Figure 2.

Figure 1. Output voltage of ISL60002 reference as a function of temperature.

Figure 2. Output voltage of X60008 reference as a function of temperature.

Supply Current

The supply current of the ISL60002 voltage reference is shown in Figure 3 as a function of temperature and at various input voltages.  It can be seen that the supply current, in general, remained almost steady with change in temperature.  With an applied input voltage between 5 and 6 volts, the supply current, however, displayed little higher values; notably at room temperature.  The input voltage range for this reference is specified between 2.7 V to 5.5 V, and driving it near or beyond the maximum input voltage is most likely what contributed to the increase in the supply current.

Figure 3.  Supply current of ISL60002 voltage reference versus temperature.

Figure 4.  Supply current of X60008 voltage reference versus temperature.

The supply current of the X60008 voltage reference exhibited excellent stability with temperature as shown in Figure 4.  Regardless of the input voltage applied, the current did not change at all within the test temperature of +20 (C to –195 (C.

Line Regulation

Line regulation characteristics of the ISL60002 voltage reference was found not to be influenced much by the change in test temperature between +20 (C to –195 (C.  These curves, which were obtained at four load levels, are depicted in Figures 5 and 6 at the test temperature of +20 (C and –195 (C, respectively.  The two set of curves were very similar in their trend as they maintained almost the same change in the output voltage, ∆ VOUT, under similar operating conditions.

Fig 5. Line regulation of ISL60002 at +20 °C.

Fig 6. Line regulation of ISL60002 at -195 °C.

The X60008 voltage reference exhibited similar behavior in line regulation as that experienced by the ISL60002 reference.  The line regulation curves of the X60008 voltage reference are depicted at +20 °C and -195 °C in Figures 7 and 8, respectively.

Fig 7. Line regulation of X60008 at +20 °C.

Fig 8. Line regulation of X60008 at -195 °C.

Load Regulation

Figure 9 shows the variation in the output voltage of the ISL60002 device as a function of load current at three different temperatures.  It can be seen that the output voltage dropped very slightly as load current was increased.  Varying the test temperature between +20 °C and -195 °C seemed to produce very little change in the load regulation of this voltage reference, as displayed in Figure 9.

Figure 9. Load regulation of ISL60002 reference at various temperatures.

Similar behavior was observed in the load regulation characteristics of the X60008 voltage reference as that displayed by its ISL60002 counterpart.  Load regulation curves of this voltage reference are shown in Figure 10 for test temperatures of +20 °C, -100 °C, and -195 °C.

Figure 10. Load regulation of X60008 reference at various temperatures.

Conclusions

Preliminary evaluation was carried out on the performance of precision voltage reference integrated circuits at cryogenic temperatures.  Intersil’s ISL60002BI and X60008B precision voltage reference chips, which are based on Floating Gate Analog (FGATM) technology, were tested in the temperature range between +20 (C to –195 (C.  Device performance was obtained in terms of output voltage and supply current at different input voltage levels as a function of temperature.  Line and load regulation characteristics were also established at four load levels and at different temperatures.  While the output voltage of either reference dropped slightly with decrease in temperature, the supply current exhibited minimal change with temperature.  Line and load regulation of both devices were very similar and did not show very strong dependence on temperature.  Although the two references are specified for operation from -40 (C to +85 (C, they sustained very good operation at cryogenic temperatures down to – 195 (C.  Therefore, these devices represent good candidates for potential use in extreme low temperature applications.  Output accuracy and stability under long-term exposure to extreme temperatures and thermal cycling, however, need to be established so that reliability assessment of such devices can be determined for use in space exploration missions.

References

- Intersil ISL60002 Precision 1.25 & 2.5 Low Voltage FGATM References Data Sheet, FN8082.2, September 17, 2004.

- Intersil X60008B-25 Precision 2.5 FGATM Voltage Reference Data Sheet, FN8140.0, March 15, 2005.

Acknowledgments

This work was performed under the NASA Glenn Research Center GESS Contract # NAS3-00145.  Funding was provided from the NASA Electronic Parts and Packaging (NEPP) Program Task “Mixed Signal Devices for Low/High Temperatures”.