THE UNIVERSITY OF CHICAGO ENGINEERING CENTER

5640 S. Ellis Avenue, Chicago, IL  60637

Phone:  (773) 834-3773

wang@oddjob.uchicago.edu

# MEMO

To:		Al Harper

From:		Shu-i Wang

Date:		4/4/01

Re:		HAWC Optical Design

The HAWC Optical Design produces diffraction-limited images in 4 selectable bands between 50µm and 250µm (53µm, 88µm, 155µm, 215µm)

.  Each passband has a different magnification to provide full spatial sampling of the PSF with 1mm square detectors.  The platescales are listed in Table 1.

Table 1.  HAWC passband platescales.

| Band | Center Wavelength ((m) | Platescale (arcsec/mm) | FWHM Image Diameter (mm) |
| --- | --- | --- | --- |
| 1 | 53 | 2.25 | 6 |
| 2 | 88 | 3.5 | 9 |
| 3 | 155 | 6 | 16 |
| 4 | 215 | 8 | 22 |

The optics uses a combination of reflective and refractive optics.  A warm, reflective fore-optics system images the telescope onto a common, cold Lyot stop for all bands.  This minimizes the size of the cryostat pressure window and facilitates alignment with the telescope optics. A Helium-cooled carousel mechanism located at the cold pupil plane allows interchanging up to seven different Lyot masks and a cold shutter.  Four sets of selectable lenses and filters mounted on a helium-cooled carousel mechanism re-image the telescope focal plane onto the detector at the appropriate image scales for the four passbands.  A schematic view of HAWC (with light entering from the telescope on the right) is shown in Figure 1.  Figure 2 shows the layout of the Band 4 optics under the normal convention of light travelling from left to right, where FFM is the folding flat mirror, and FM is the field mirror.  The four passbands, defined by metal-mesh interference filters, coincide with atmospheric windows.  In addition to the 4 bands of imaging optics, a pupil reimaging lens (PRL) can be flipped into the ray path between the detector fold mirror and the detector array to view the pupil image/illumination.  The PRL is optimized for Band 3 and forms a 12mm diameter image of the pupil at the detector.  The PRL can also be used with the other 3 bands to view an out-of-focus image of the pupil to check for uniformity of the pupil illumination.  In Band 1, the out-of-focus pupil image diameter is about 17mm.  The pupil image diameter is about 16.5mm for both Bands 2 and 4.

Figure 1.  Schematic View of HAWC Layout.

Figure 2.  Layout of HAWC Band 4 Optics.

Figures 3-6 shows the y-z plane layout of the cold optics inside the cryostat for Bands 1-4, respectively.  Figure 7 shows the raytrace of the cold optics in the pupil reimaging mode for Band 3 with the telescope secondary mirror as the object.

The window curvature is modeled with the curvature measured during the 2nd pump down on 7/11/00 (last measurement, see Window3.doc).  The optics was optimized with the criteria of best performance with all three of the last three sets of window profile measurements without refocusing the telescope.  The largest variation in rms spot diameter between the three sets of measurements was about 15 microns.

Figure 3.  HAWC Band 1 Cold Optical Layout.

Figure 4.  HAWC Band 2 Cold Optical Layout.

Figure 5.  HAWC Band 3 Cold Optical Layout.

Figure 6.  HAWC Band 4 Cold Optical Layout.

Figure 7.  HAWC Band 3 PRL Cold Optical Layout.

The re-imaging lenses in each band were optimized and evaluated in the Zemax Optical Design Software using the center wavelength, the 10% transmission short wavelength (λs10%), and the 10% transmission long wavelength (λl10%) as defined in Table 2.  (The bandwidth listed in Table 2 is the full width half maximum (FWHM) bandwidth.)  The three wavelengths were equally weighted in Zemax.  The optics is designed to accommodate a 32x32 detector array with 1mm square pixels.  All optics, field stops, and baffles are oversized by at least 5 Airy rings (in addition to the PSF central core) to minimize effects of diffraction on the uniformity of illumination of the cold pupil.  The observed field-of-view (FOV) and the 5-Airy oversized FOV used to size the optics are listed in Table 3.  The 5-Airy field was calculated for the FWHM long wavelength in each band.  In optimizing the lens designs, the field was defined by the center, 0.7(FOV), and the actual FOV for a 32x32 detector array.  (The 5-Airy FOV was given a weight of zero and did not contribute to the optics optimization.)  Simple spherical lenses sufficed to achieve the required image quality.

Table 2.  HAWC Filter Passbands.

| Bandwidth (dλ/λ) | λs10% (μm) | λc (μm) | λl10% (μm) |
| --- | --- | --- | --- |
| 0.10 | 48.1 | 53 | 58.8 |
| 0.10 | 81.3 | 88 | 95.2 |
| 0.14 | 140.8 | 155 | 172.4 |
| 0.20 | 190.5 | 215 | 246.9 |

Table 3.  HAWC Field-of-view.

|  | 32x32 Detector Array | 12x32 Detector Array |
| --- | --- | --- |
| Band | FOV (arcmin) | FOV + 5 Airy (arcmin) | FOV (arcmin) | FOV + 5 Airy (arcmin) |
| 1 | 1.2x1.2 | 2.14x2.14 | 0.45x1.2 | 1.39x2.14 |
| 2 | 1.87x1.87 | 3.42x3.42 | 0.7x1.87 | 2.25x3.42 |
| 3 | 3.2x3.2 | 5.99x5.99 | 1.2x3.2 | 3.99x5.99 |
| 4 | 4.27.4.27 | 8.25x8.25 | 1.6x4.27 | 5.59x8.25 |

Other factors in the design merit function included the positioning of the optics to correspond with the carousel mechanisms and the rms spot diameter for each field point.  A full ghost analysis was performed iteratively with the lens design to minimize and eliminate where feasible all ghosts in the system.  Lenses were shaped to reduce the ghost intensity to below noise levels.  The lenses will also be anti-reflection (AR) coated with a coating optimized for each filter band.  (The lenses will be AR coated with either thin layers of PE/PP by QMC2, or with Parylene C by Spectral Systems.)   The filters were spaced apart and tilted at 10 degrees successively in the opposite direction to eliminate ghosting.  The filters were modeled by a constant index of 2 with a thickness of 9λ/5 for hot-press filters2 and 50 μm for air-gap filters.  The filter thicknesses entered in the raytraces are listed in Table 4.  Blocking filters CD1 & CB4 were not included in the optics models.  The only significant source of ghosting is the window ghost due to the front-back window surface reflections where the maximum total ghost energy, as defined in the ghost document, is 0.2% in Band 1.

Table 4.  HAWC Filter Size Model.

| Filter | Type | 50% (µm) | Thickness (µm) | Mount Size | O.D. (mm) | C.A. (mm) |
| --- | --- | --- | --- | --- | --- | --- |
| (CB1) | Hot-press | 37 | 66.6 |  |  |  |
| CB2 | Hot-press | 40.8 | 73.44 | 7 | 76.987 | 64.999 |
| CB3 | Hot-press | 45.5 | 81.9 | 6 | 65.608 | 55.804 |
| (CB4) | Hot-press | 50.5 | 90.9 |  |  |  |
| ED_B1 | Air-gap | 55.9 | 50 | 6 | 65.608 | 55.804 |
| ED_B2 (1) | Air-gap | 92.6 | 50 | 6 | 65.608 | 55.804 |
| ED_B2 (2) | Air-gap | 84 | 50 | 6 | 65.608 | 55.804 |
| B_B2 | Hot-press | 71.4 | 128.52 | 6 | 65.608 | 55.804 |
| ED_B3 (1) | Air-gap | 166.7 | 50 | 6 | 65.608 | 55.804 |
| ED_B3 (2) | Air-gap | 144.9 | 50 | 6 | 65.608 | 55.804 |
| B_B3 | Hot-press | 111.1 | 199.98 | Custom | 89.916* | 76.2 |
| ED_B4 (1) | Air-gap | 238.1 | 50 | 7 | 76.987 | 64.999 |
| ED_B4 (2) | Air-gap | 196.1 | 50 | Custom | 89.916* | 76.2 |
| B_B4 | Hot-press | 166.7 | 300.06 | Custom | 97.409* | 82.55 |

*Actual O.D. TBD.  Used O.D.=1.18(C.A.) as a placeholder for now, where 1.18=(O.D.)/(C.A.) of the two standard size filters.

The optics for each band was modeled in Zemax with three configurations.  The first configuration models the flight configuration, with a temperature of –50C for the foreoptics and –267C for the cold optics inside the dewar.  The second configuration models cold testing in a lab, with a temperature of 20C for the foreoptics and –267C for the cold optics.  The third configuration models a warm dewar in the lab, with a temperature of 20C for both the foreoptics and the dewar optics.  The optics was optimized for the cold flight configuration.  The differences in rms spot diameters between the cold flight configuration and the cold lab configuration are less than 10 μm.

The glasses were modeled with custom glass files in Zemax for the warm and cold indices.  Index and absorption data from Loewenstein (1.5K & 300K data) were used for the silicon lenses in Band 1 and the pupil reimaging lens (PRL);  data from Brehat & Wyncke (10K & 300K data) were used for the crystal quartz lenses in the other bands.  Lens transmissions, shown in Table 5, were modeled with the TFCalc Thin Film Design Software for Windows.  In Table 5, tc is the lens center thickness, te is the lens edge thickness, tAR is the AR coating thickness, Tabs is the transmission due to material absorption, and Ttotal is the total lens transmission.  The actual transmissions could be a couple of percent less than the theoretical values shown in Table 5.  The silicon lenses were modeled with a single layer Parylene AR coating.  Parylene was modeled with a constant refractive index of 1.6 without absorption due to a lack of absorption data.  The AR coatings are quite thin and the Parylene absorption is thus expected to be less than a couple of percent.  The crystal quartz lenses were modeled with a single layer polyethylene AR coating.  The transmission data are listed in ASCII text files L1-1.txt and L2-1.txt for the Band 1 lenses, L1-2.txt and L2-2.txt for the Band 2 lenses, L1-3.txt for the Band 3 lens, L1-4.txt for the Band 4 lens, and PRL.txt for the pupil reimaging lens.  These files are available at the irma.yerkes.uchicago.edu ftp site, directory /pub/wang.  The files contain extrapolated transmission values from 0.3 – 400 μm in increments of 0.1 μm.  However, the actual silicon data only covers from about 28.57 – 33.33 μm, and extrapolation beyond this range should be cautioned.  The crystal quartz data covers from about 16.67 – 500 μm.

Table 5.  Lens AR Coating Models.

| Element | Material | tc (mm) | te (mm) | Coating Material | tAR (um) | Tabs | Ttotal |
| --- | --- | --- | --- | --- | --- | --- | --- |
| L1-1 | Silicon | 2.3 | 1.681 | Parylene C | 8.8 ± 1 µm | 0.86 | 0.79 |
| L1-2 | Crystal Quartz | 3.2 | 1.943 | PE/PP | 14.4 ± 1 µm | 0.98 | 0.96 |
| L1-3 | Crystal Quartz | 3.0 | 2.025 | PE/PP | 25.4 ± 1 µm | 0.99 | 0.98 |
| L1-4 | Crystal Quartz | 4.5 | 1.999 | PE/PP | 35.4 ± 1 µm | 0.99 | 0.98 |
| L2-1 | Silicon | 2.0 | 3.270 | Parylene C | 8.8 ± 1 µm | 0.98 | 0.97 |
| L2-2 | Crystal Quartz | 2.0 | 3.539 | PE/PP | 14.4 ± 1 µm | 0.98 | 0.98 |
| PRL-1 | Silicon | 5.3 | 1.979 | Parylene C | 25.8 ± 1 µm | 0.72 | 0.55 |
| PRL-2 |  |  |  |  |  | 0.82 | 0.47 |
| PRL-3 |  |  |  |  |  | 0.90 | 0.82 |
| PRL-4 |  |  |  |  |  | 0.92 | 0.75 |

The geometric spot diameters, diffraction encircled energy, mtf and distortion plots are listed below for documenting the design image quality of the HAWC optics.  The geometric spot diameters for the imaging optics are shown in Figures 8-11 for Bands 1-4, respectively.  Boxes in the spot diagrams are 1mm x 1mm square (single pixel).  The field positions are the corner edges of the field of view, the center of the field, and the corner edges of the 5-Airy oversized fields.  This is the geometric spot pattern only and does not take diffraction into account.  Figures 12-15 show the diffraction encircled energy vs. radius from centroid for the 4 bands.  One tick on the x-axis equals 1mm (1 pixel). The curves are for the same 9 points sampled in the geometric spot diagrams. The 80% encircled energy occurs at a radius of 2.5 pixels or less.  Figures 16-19 show the diffraction modulation transfer function (MTF) for all field points.  Distortion is characterized in Figures 20-23. The square grid is a 32x32 image at the HAWC detector of a square grid of the telescope focal plane that would just fill the HAWC field of view. The size of each square in the grid represents one pixel. The distortion is a small fraction of a pixel.

Figure 8.  Geometric Spot Diameter for HAWC Band 1 Optics.

Figure 9.  Geometric Spot Diameter for HAWC Band 2 Optics.

Figure 10.  Geometric Spot Diameter for HAWC Band 3 Optics.

Figure 11.  Geometric Spot Diameter for HAWC Band 4 Optics.

Figure 12.  Diffraction Encircled Energy for HAWC Band 1 Optics.

Figure 13.  Diffraction Encircled Energy for HAWC Band 2 Optics.

Figure 14.  Diffraction Encircled Energy for HAWC Band 3 Optics.

Figure 15.  Diffraction Encircled Energy for HAWC Band 4 Optics.

Figure 16.  MTF for HAWC Band 1 Optics.

Figure 17.  MTF for HAWC Band 2 Optics.

Figure 18.  MTF for HAWC Band 3 Optics.

Figure 19.  MTF for HAWC Band 4 Optics.

Figure 20.  Distortion Grid at Detector Array for HAWC Band 1 Optics.

Figure 21.  Distortion Grid at Detector Array for HAWC Band 2 Optics.

Figure 22.  Distortion Grid at Detector Array for HAWC Band 3 Optics.

Figure 23.  Distortion Grid at Detector Array for HAWC Band 4 Optics.

A simulated image of the pupil in Band 3 is shown in Figure 24.  The box in Figure 24 represents a 12x32 detector array.  (The cross in the center is an artifact.)  The pupil reimager was initially designed with the Lyot stop at the cold paraxial pupil image formed by the FM, and the PRL 60mm after the DFM.  The PRL was optimized for best rms spot size of the pupil image at the detector array with the telescope secondary as the object, taking into account ghosting considerations.  Although the PRL was optimized for Band 3, we intend to use the PRL with all bands.  Therefore, silicon was chosen as the PRL glass material for its small chromatic aberration.  After a tolerance analysis and truncation of some dimensions during the request for lens quotes, the pupil reimager design was reoptimized.  The Lyot stop was placed at the smallest geometric pupil image location, keeping the rest of the optics fixed.  Then, placement of the PRL was optimized for the best rms spot size of the pupil image with the telescope secondary as the object.

Figure 24.  Simulated Geometric Pupil Image.

The image quality for the pupil reimaging mode can be characterized by the rms spot diameter of the pupil image of various points on the telescope secondary mirror/pupil object.  Figure 25 shows the geometric spot diameter for the center, edge, and 0.7(edge radius) image points of the pupil.  Boxes in the spot diagrams are 1mm x 1mm square (single pixel).  Another figure of merit for the pupil reimager is the separation of the star images within the field of view, which can lead to a blurring of the effective pupil image.  The separations between the pupil image centers from various field points are listed below each pupil image in Figure 26.

A design decision I made for both the imaging optics and the pupil reimager is to keep the cold filters and lenses centered on the mechanical axis defined as the normal center to the cryostat window.  This differs slightly from the optical axis due to the tilts of the filters.  Normally in an optical design the optics are decentered to take into account the lateral shift in the optical axis due to tilted elements.  However, the tilts here are very small (very thin filters) and have minimal effect on the image quality.  It would not be worth the mechanical complexity to put the slight filter decenters into the design.

Figure 25.  Geometric Spot Diameter of the Pupil Reimager for Band 3.

Figure 26.  Geometric Pupil Images of Various Star Object Fields.

A thermal analysis was performed on the optical design using Zemax to calculate warm-cold changes in the optics.  Before we discuss the thermal analysis, it's important to understand the coordinate system conventions.

HAWC has several coordinate systems.  Below are the transformations between the SOFIA Aircraft Coordinates, (Xa, Ya, Za), the Cryostat Coordinates, (Xc, Yc, Zc), and the Optics Coordinates, (Xo, Yo, Zo).  Figure 24 shows the sign convention of the Aircraft Coordinate System.  Figure 25 shows the sign convention of the Cryostat and the Optics Coordinate Systems.  The Optics Coordinate System follows the Zemax sign convention.  It is the coordinate system used for documenting the optics.  In addition, Figure 26 shows the transformation between the Lens Wheel Coordinates and the Optics Coordinates for relating the carousel wheel to the optics.

Aircraft Coordinates:	(Xa, Ya, Za) = (-Zc, Yc, Xc)

Cryostat Coordinates:	(Xc, Yc, Zc) = (Za, Ya, -Xa)

Optics Coordinates:	(Xo, Yo, Zo)

(0, 0, 0) is defined to be the front surface center of the cryostat window.

To transform from (Xc, Yc, Zc) to (Xo, Yo, Zo):

Xo = -Xc*COS(-50) - Yc*SIN(-50)

Yo = -Xc*SIN(-50) + Yc*COS(-50)

Zo = -Zc

To transform from (Xo, Yo, Zo) to (Xc, Yc, Zc):

Xc = -Xo*COS(-50) - Yo*SIN(-50)

Yc = -Xo*SIN(-50) + Yo*COS(-50)

Zc = -Zo

Figure 24.  Aircraft Coordinate System.

Figure 25.  Cryostat and Optics Coordinate Systems.

Figure 26.  Lens Wheel Coordinates.

The 4 bands of HAWC optics and the pupil reimager were designed with a (temperature, pressure) setting of (-50C, 1 atm.) for the foreoptics, and (-267C, 0 atm.) for the cold optics inside the dewar.  To model the optics thermal behavior, two additional configurations were added to model the warm and cold configurations in a laboratory setting.  Table 6 shows the optics configurations in Zemax.

Table 6.  Optics (Temperature, Pressure) Configurations in Zemax.

| Configuration | Foreoptics | Cold Cryostat Optics |
| --- | --- | --- |
| 1 | (-50C, 1 atmosphere) | (-267C, 0 atmosphere) |
| 2 | (20C, 1 atmosphere) | (-267C, 0 atmosphere) |
| 3 | (20C, 1 atmosphere) | (20C, 1 atmosphere) |

Using the Zemax thermal pickup function, changes in the optics curvature, thickness, index (manually entered in cold and warm glass files for HAWC), diameters, and positions can be calculated for each (temperature, pressure) setting.  This is how the lens cold specifications were converted into warm manufacturing specifications.  The dividing line between the foreoptics and the cold cryostat optics is the cryostat window.  The window is modeled with the foreoptics temperature.  A global shift is applied to all the elements after the window with a dummy surface in the Zemax lens files to model the thermal positional change in the cryostat suspension system.  Tim Rennick calculated the global ((Xo, (Yo, (Zo) cold to warm shift between the warm foreoptics and the cold cryostat optics as (-0.6mm, -1.032mm, -0.029mm), in spreadsheet cte001011.xls.  (This global shift will change with the new suspension design and should be updated after the new design is completed.)  Tables 7 & 8 show the thermal coefficient of expansion (CTE) values entered into Zemax to model the camera.  The optics mounts in both the foreoptics and the cryostat were modeled with Aluminum alloy 6061.  The lenses were also manually decentered in the warm configuration to model the lens mount design.  The amount of lens decentering is shown in Figure 9.  Note, unlike lens specifications in Zemax where the center-to-center distances are used, the thermal expansion is computed along a length of material which extends from the edge of the surface to the edge of the next surface.  This means tilted surface (i.e. filters) positions may not be calculated correctly.  Because of the complexity of modeling the cryostat thermally, all optics positions should be checked separately in the Pro-E model and/or manually.

Table 7.  HAWC Material Thermal Coefficient of Expansion.

| Element | Material | TCE (m/m-K) | Temperature |
| --- | --- | --- | --- |
| Optics mounts, support structures | Aluminum alloy 6061 | 1.45E-05 | -267C |
| FFM, FM | Fused Silica | 5.50E-07 | -50C |
| Window | HDPE | 1.30E-04 | -50C |
| Filters | HPFILTER (HDPE) | 1.30E-042,11 | -267C |

Table 8.  Lens Thermal Coefficient of Expansion and Change in Diameter.

| Element | Material | avg TCE | (T | Warm Diameter | (Diameter | Decenter |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | (C-1) | (C) | (in) | (in) | (in) |
| L1-1 | Silicon | 7.53E-079 | -287 | 2.0472 | -0.0004 | -0.0002 |
| L1-2 | Crystal Quartz | 4.07E-064 | -287 | 2.0472 | -0.0024 | -0.0012 |
| L1-3 | Crystal Quartz | 4.07E-06 | -287 | 2.1260 | -0.0025 | -0.0012 |
| L1-4 | Crystal Quartz | 4.07E-06 | -287 | 3.1890 | -0.0037 | -0.0019 |
| L2-1 | Silicon | 7.53E-07 | -287 | 1.5748 | -0.0003 | -0.0002 |
| L2-2 | Crystal Quartz | 4.07E-06 | -287 | 1.7717 | -0.0021 | -0.0010 |
| PRL | Silicon | 7.53E-07 | -287 | 3.0000 | -0.0006 | -0.0003 |

Table 9.  Lens Warm-Cold Decenters.

| Element | Optics Coordinate System | Zemax Cold-Warm Decenters |
| --- | --- | --- |
|  | Xo Decenter (in) | Yo Decenter (in) | DCXo (mm) | DCYo (mm) |
| L1-1 | 0.0001 | -0.0002 | -0.0032 | 0.0046 |
| L1-2 | 0.0007 | -0.0010 | -0.0174 | 0.0249 |
| L1-3 | 0.0007 | -0.0010 | -0.0181 | 0.0258 |
| L1-4 | 0.0011 | -0.0015 | -0.0271 | 0.0388 |
| L2-1 | 0.0001 | -0.0001 | -0.0025 | 0.0035 |
| L2-2 | 0.0006 | -0.0008 | -0.0151 | 0.0215 |
| PRL | 0.0002 | -0.0003 | -0.0047 | 0.0067 |

Table 10 lists the warm and cold optics specifications from the thermal analysis.  The warm specifications were sent in request for quotes to various manufacturers.  Tables 11-14 lists the warm and cold global coordinates for each optical element in the optics coordinate system.  The warm coordinates are for a warm cryostat in the lab, and the cold coordinates are for the cooled cryostat in-flight.

Table 10.  Optics Warm and Cold Specs.

Table 11.  Warm and Cold Band 1 Optics Coordinates.

| Warm Dimensions |  |  |  |  |
| --- | --- | --- | --- | --- |
| Surf | Xow (mm) | Yow (mm) | Zow (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.131 | 5058.120 | 0 | -90 |
| Secondary | 0 | 1624.679 | 5058.120 | 0 | -90 |
| Tertiary | 0 | -139.131 | 5058.120 | 0 | -45 |
| FFM | 0 | -139.131 | 154.292 | 0 | 9 |
| TFP | 0 | -69.599 | 368.288 | 0 | 18 |
| FM | 0 | 0.000 | 582.493 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.840 | 0 | 0 |
| CB2, Center | -0.6 | -1.038 | -40.649 | 0 | 10 |
| CB3, Center | -0.6 | -1.025 | -112.582 | 0 | -10 |
| Lyot Stop | -0.6 | -1.032 | -150.840 | 0 | 0 |
| L1-1 S1 | -0.6032 | -1.027 | -165.901 | 0 | 0 |
| ED_B1, Center | -0.6 | -1.036 | -186.090 | 0 | 10 |
| L2-1 S1 | -0.6025 | -1.029 | -399.579 | 0 | 0 |
| DFM | -0.6 | -1.032 | -490.134 | 0 | -45 |
| Detector | -0.6 | -181.380 | -490.134 | 0 | -90 |
| Cold Dimensions |  |  |  |  |
| Surf | Xoc (mm) | Yoc (mm) | Zoc (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.058 | 5057.714 | 0 | -90 |
| Secondary | 0 | 1624.753 | 5057.714 | 0 | -90 |
| Tertiary | 0 | -139.058 | 5057.714 | 0 | -45 |
| FFM | 0 | -139.058 | 154.495 | 0 | 9 |
| TFP | 0 | -69.529 | 368.483 | 0 | 18 |
| FM | 0 | 0.000 | 582.470 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.832 | 0 | 0 |
| CB2, Center | 0 | -0.006 | -40.611 | 0 | 10 |
| CB3, Center | 0 | 0.007 | -112.244 | 0 | -10 |
| Lyot Stop | 0 | 0.000 | -150.342 | 0 | 0 |
| L1-1 S1 | 0 | 0.000 | -165.342 | 0 | 0 |
| ED_B1, Center | 0 | -0.004 | -185.456 | 0 | 10 |
| L2-1 S1 | 0 | 0.000 | -398.043 | 0 | 0 |
| DFM | 0 | 0.000 | -488.242 | 0 | -45 |
| Detector | 0 | -179.600 | -488.242 | 0 | -90 |

Table 12.  Warm and Cold Band 2 Optics Coordinates.

| Warm Dimensions |  |  |  |  |
| --- | --- | --- | --- | --- |
| Surf | Xow (mm) | Yow (mm) | Zow (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.131 | 5058.120 | 0 | -90 |
| Secondary | 0 | 1624.679 | 5058.120 | 0 | -90 |
| Tertiary | 0 | -139.131 | 5058.120 | 0 | -45 |
| FFM | 0 | -139.131 | 154.292 | 0 | 9 |
| TFP | 0 | -69.599 | 368.288 | 0 | 18 |
| FM | 0 | 0.000 | 582.493 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.840 | 0 | 0 |
| CB2, Center | -0.6 | -1.038 | -40.649 | 0 | 10 |
| CB3, Center | -0.6 | -1.025 | -112.582 | 0 | -10 |
| Lyot Stop | -0.6 | -1.032 | -150.840 | 0 | 0 |
| ED_B2.1, Center | -0.6 | -1.036 | -169.091 | 0 | 10 |
| L1-2 S1 | -0.6174 | -1.007 | -185.789 | 0 | 0 |
| ED_B2.2, Center | -0.6 | -1.028 | -224.158 | 0 | -10 |
| B_B2, Center | -0.6 | -1.043 | -266.703 | 0 | 10 |
| L2-2 S1 | -0.6151 | -1.011 | -402.303 | 0 | 0 |
| DFM | -0.6 | -1.033 | -490.139 | 0 | -45 |
| Detector | -0.6 | -181.380 | -490.139 | 0 | -90 |
| Cold Dimensions |  |  |  |  |
| Surf | Xoc (mm) | Yoc (mm) | Zoc (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.058 | 5057.714 | 0 | -90 |
| Secondary | 0 | 1624.753 | 5057.714 | 0 | -90 |
| Tertiary | 0 | -139.058 | 5057.714 | 0 | -45 |
| FFM | 0 | -139.058 | 154.495 | 0 | 9 |
| TFP | 0 | -69.529 | 368.483 | 0 | 18 |
| FM | 0 | 0.000 | 582.470 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.832 | 0 | 0 |
| CB2, Center | 0 | -0.006 | -40.611 | 0 | 10 |
| CB3, Center | 0 | 0.007 | -112.244 | 0 | -10 |
| Lyot Stop | 0 | 0.000 | -150.342 | 0 | 0 |
| ED_B2.1, Center | 0 | -0.004 | -168.516 | 0 | 10 |
| L1-2 S1 | 0 | 0.000 | -185.148 | 0 | 0 |
| ED_B2.2, Center | 0 | 0.004 | -223.367 | 0 | -10 |
| B_B2, Center | 0 | -0.011 | -265.732 | 0 | 10 |
| L2-2 S1 | 0 | 0.000 | -400.765 | 0 | 0 |
| DFM | 0 | 0.000 | -488.242 | 0 | -45 |
| Detector | 0 | -179.600 | -488.242 | 0 | -90 |

Table 13.  Warm and Cold Band 3 Optics Coordinates.

| Warm Dimensions |  |  |  |  |
| --- | --- | --- | --- | --- |
| Surf | Xow (mm) | Yow (mm) | Zow (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.131 | 5058.120 | 0 | -90 |
| Secondary | 0 | 1624.679 | 5058.120 | 0 | -90 |
| Tertiary | 0 | -139.131 | 5058.120 | 0 | -45 |
| FFM | 0 | -139.131 | 154.292 | 0 | 9 |
| TFP | 0 | -69.599 | 368.288 | 0 | 18 |
| FM | 0 | 0.000 | 582.493 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.840 | 0 | 0 |
| CB2, Center | -0.6 | -1.038 | -40.649 | 0 | 10 |
| CB3, Center | -0.6 | -1.025 | -112.582 | 0 | -10 |
| Lyot Stop | -0.6 | -1.032 | -150.840 | 0 | 0 |
| L1-3 S1 | -0.6181 | -1.006 | -165.901 | 0 | 0 |
| ED_B3.1, Center | -0.6 | -1.032 | -186.066 | 0 | 10 |
| ED_B3.2, Center | -0.6 | -1.032 | -241.066 | 0 | -10 |
| B_B3, Center | -0.6 | -1.032 | -403.991 | 0 | 10 |
| DFM | -0.6 | -1.033 | -490.144 | 0 | -45 |
| PRL | -0.6047 | -61.079 | -490.137 | 0 | -90 |
| Detector | -0.6 | -181.380 | -490.144 | 0 | -90 |
| Cold Dimensions |  |  |  |  |
| Surf | Xoc (mm) | Yoc (mm) | Zoc (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.058 | 5057.714 | 0 | -90 |
| Secondary | 0 | 1624.753 | 5057.714 | 0 | -90 |
| Tertiary | 0 | -139.058 | 5057.714 | 0 | -45 |
| FFM | 0 | -139.058 | 154.495 | 0 | 9 |
| TFP | 0 | -69.529 | 368.483 | 0 | 18 |
| FM | 0 | 0.000 | 582.470 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.832 | 0 | 0 |
| CB2, Center | 0 | -0.006 | -40.611 | 0 | 10 |
| CB3, Center | 0 | 0.007 | -112.244 | 0 | -10 |
| Lyot Stop | 0 | 0.000 | -150.342 | 0 | 0 |
| L1-3 S1 | 0 | 0.000 | -165.342 | 0 | 0 |
| ED_B3.1, Center | 0 | 0.000 | -185.434 | 0 | 10 |
| ED_B3.2, Center | 0 | 0.000 | -240.204 | 0 | -10 |
| B_B3, Center | 0 | 0.000 | -402.453 | 0 | 10 |
| DFM | 0 | 0.000 | -488.242 | 0 | -45 |
| PRL | 0 | -59.819 | -488.242 | 0 | -90 |
| Detector | 0 | -179.600 | -488.242 | 0 | -90 |

Table 14.  Warm and Cold Band 4 Optics Coordinates.

| Warm Dimensions (mm) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Surf | Xow (mm) | Yow (mm) | Zow (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.131 | 5058.120 | 0 | -90 |
| Secondary | 0 | 1624.679 | 5058.120 | 0 | -90 |
| Tertiary | 0 | -139.131 | 5058.120 | 0 | -45 |
| FFM | 0 | -139.131 | 154.292 | 0 | 9 |
| TFP | 0 | -69.599 | 368.288 | 0 | 18 |
| FM | 0 | 0.000 | 582.493 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.840 | 0 | 0 |
| CB2, Center | -0.6 | -1.038 | -40.649 | 0 | 10 |
| CB3, Center | -0.6 | -1.025 | -112.582 | 0 | -10 |
| Lyot Stop | -0.6 | -1.032 | -150.840 | 0 | 0 |
| ED_B4.1, Center | -0.6 | -1.036 | -186.091 | 0 | 10 |
| ED_B4.2, Center | -0.6 | -1.028 | -241.091 | 0 | -10 |
| L1-4 S1 | -0.6271 | -0.993 | -259.922 | 0 | 0 |
| B_B4, Center | -0.6 | -1.058 | -404.085 | 0 | 10 |
| DFM | -0.6 | -1.034 | -490.144 | 0 | -45 |
| Detector | -0.6 | -181.381 | -490.138 | 0 | -90 |
| Cold Dimensions |  |  |  |  |
| Surf | Xoc (mm) | Yoc (mm) | Zoc (mm) | XoTILT (deg.) | YoTILT (deg.) |
| Primary | 0 | -1129.058 | 5057.714 | 0 | -90 |
| Secondary | 0 | 1624.753 | 5057.714 | 0 | -90 |
| Tertiary | 0 | -139.058 | 5057.714 | 0 | -45 |
| FFM | 0 | -139.058 | 154.495 | 0 | 9 |
| TFP | 0 | -69.529 | 368.483 | 0 | 18 |
| FM | 0 | 0.000 | 582.470 | 0 | 9 |
| Window S1 | 0 | 0.000 | 0.000 | 0 | 0 |
| Window S2 | 0 | 0.000 | -0.832 | 0 | 0 |
| CB2, Center | 0 | -0.006 | -40.611 | 0 | 10 |
| CB3, Center | 0 | 0.007 | -112.244 | 0 | -10 |
| Lyot Stop | 0 | 0.000 | -150.342 | 0 | 0 |
| ED_B4.1, Center | 0 | -0.004 | -185.446 | 0 | 10 |
| ED_B4.2, Center | 0 | 0.004 | -240.216 | 0 | -10 |
| L1-4 S1 | 0 | 0.000 | -258.969 | 0 | 0 |
| B_B4, Center | 0 | -0.025 | -402.550 | 0 | 10 |
| DFM | 0 | 0.000 | -488.242 | 0 | -45 |
| Detector | 0 | -179.600 | -488.242 | 0 | -90 |

# Endnotes

Focus Software, Inc. Web Site:  http://www.focus-software.com.

Private Conversation with Ken Wood of QMC Instruments, Ltd (9/00).

Specialty Coating Systems.  Web Site:  http://www.scscookson.com (9/99).

"Optical constants of far infrared materials. 2: crystalline solids", E. Loewenstein, D. Smith & R. Morgan, Applied Optics, Vol. 12, No. 2, pp. 398-406 (2/73).

"Measurement of the optical constants of crystal quartz at 10K and 300K in the far infrared spectral range: 10-600 cm-1", F. Brehat & B. Wyncke, Int. J. of IR & Millimeter Waves, Vol. 18, No. 9, pp. 1663-1679 ('97).

Software Spectra, 14025 N.W. Harvest Lane, Portland, OR  97229.  Phone:  (503)690-2099 (2/98).

Private conversation with Jim Heaney.

"Optical constants of far infrared materials. 3: plastics", D. Smith & E. Loewenstein, Applied Optics, V14 #6, pp. 1335-1341 (6/75).

"Helium Cryogenics", Steven S. Van Sciver, Kluwer Academic/Plenum Publishers (11/86).

"Fabrication Methods for Precision Optics", Hank H. Karow, John Wiley & Sons, Inc. (1993).

Goodfellow Web Site:  http://www.goodfellow.com/ (7/00).