Alignment, tolerances, commissioning: PEEM-3

Hi-res.: Assume 2 um FOV, Magnification 12000 (M4=M5=M6=10) (CCD: 24x24mm)

Lo-res.: Assume 20 um FOV, Magnification 1200 (M4=M5=10, M6=0)

Aux.: Assume 17 um FOV, Magnification: 1440 (MA1+6, MA2=20) (CCD: 24x24 mm)

| Component | Description | Beam alignment (Hi-res./Lo-res.) | Physical tolerance relative to relevant magnetic axis of SM |
| --- | --- | --- | --- |
| S | Sample | - | <1 mrad (tilt), 200 nm rel. O |
| O | Objective lens (M=12) | - | 25 um |
| D1 | El. dodecapole deflector/stigmator | - | 250 um |
| D2 | El. dodecapole deflector/stigmator | - | 250 um |
| F1 | Field lens (M=1) | 2.5 um / 25 um | 25 um (or better) |
| SM | Separator magnet | 2.5 um / 25 um | - |
| F2 | Field lens (M=1) | 2.5 um / 25 um | 25 um (or better) |
| D3 | El. Dodecapole / Mag. Octupole | - | 250 um |
| D4 | El. Dodecapole / Mag. Octupole | - | 250 um |
| M | Mirror (M=1) | 2.5 um / 25 um | 25 um |
| A1 | Aux. projector lens (M=6) | 25 um | 25 um (to mirror) |
| A2 | Aux. projector lens (M=20) | 150 um | 100 um (to mirror) |
| CCD1 | Aux. CCD | - | 1000 um |
| D5 | El. Dodecapole | - | 250 um |
| I1 | Intermediate lens (M=1) | 2.5 um / 25 um | 25 um |
| AP1 | Contrast Aperture | 2.5 um / 25 um | 1 um rel. I1 |
| I2 | Intermediate lens (M=10) | 2.5 um / 25 um | 25 um |
| D6 | El. Octupole | - | 250 um |
| P1 | Projector lens (M=10) | 25 um / 250 um | 250 um |
| D7 | El. Octupole | - | 1000 um |
| P2 | Projector lens (M=10 / off) | 250 um / - | 250 um |
| CCD2 | CCD |  | 1000 um |

Beam alignment will be achieved by either using the upstream deflector and mechanical pre-alignment. The criteria for the above numbers are:

Beam alignment: 10% of local beam size at chosen magnification.

Physical tolerance: 10% of beam size at intermediate magnification. Deflector elements not used

Sample alignment: approximate number chosen from experience with PEEM-2.

Optical alignment is based on mechanical pre-alignment of PEEM. The mechanical tolerances were chosen such that a high-quality image can be seen on the detector at moderate magnification without use of deflector elements. PEEM-2 has shown that this is possible in a straight PEEM. The major unknown is the function of the separator magnet. The 90° auxiliary PEEM will help adjusting the trim coils of the 1st sector and will give hints about the alignment of the 2nd sector.

The final electron optical alignment has to be based on image quality, image distortion, stigmatism, coma, and movement while wobbling lens voltages, deflector voltages/currents and trim coil currents. Beam monitors will be useful to identify problems, which lead to a complete loss of transmission. Whether the precision and the amount of the information that they will provide are sufficient to assist with the final alignment of PEEM-3 at high magnification is questionable. The accepted electron beam size before the contrast aperture is on the order of 1 mm and its “center” would have to be determined with an accuracy of 2.5 um. 2.5 um is 10% of the accepted beam size for a FOV of 2 um and 12x magnification of the objective lens. The beam intensity is furthermore laterally modulated by the profile of the illumination. The “center point of the object” that has to be transferred through the optical center of all components of the microscope is actually not defined as the center of gravity of the electron beam but by the symmetry axis of the sample/objective lens group. In the following I therefore have to assume that the final alignment will not require any additional beam monitors in the microscope.

For the PEEM alignment a monochromatic electron source of high brightness is required. Possible sources are an electron gun in a LEEM configuration or a monochromatic UV source operating at a wavelength close to the workfunction threshold of the sample. An energy bandwidth of <0.5 eV would be ideal. Our PEEM design does not allow using an electron gun. We therefore either use a well-focused UV lamp or a UV laser. The illumination intensity should be comparable to the intensity of the x-ray source: ~100 W/cm2 (1012 ph/s at 1keV in 10um x 10um). A frequency-tripled, focused Ti:Sa at 4.5 eV (270 nm) has a CW averaged intensity of 100 W/cm2 (1mW in 30um x 30um). If necessary, tighter focusing is possible. (Will inquire about Hg arc lamps)

The microscope alignment consists of two stages: beam alignment and aberration correction. During the 1st stage the “beam” will be centered and aligned in optical components using trim coils, and the dodecapole and octupole elements. During the 2nd stage residual aberrations will be corrected using trim coils, dodecapoles and the mirror. In general, alignment will be an iterative process starting at low magnification and step-by-step optimization of all elements while ramping up the magnification until the ultimate resolution is reached. The initial alignment will be carried out using a monochromatic electron source such that only spherical aberrations will contribute. Dispersion and chromatic defocus can be studied by adding an offset to the sample voltage. Initial alignment will also be carried out without a contrast aperture to maximize the effect of misalignment and “mis-correction”. A PEEM behind the mirror arm is essential for the initial alignment and would be very useful later because it allows a) to independently test and optimize the first sector of the separator and b) to independently optimize ingoing and outgoing beam at the mirror. Samples with high-contrast structures with structure sizes of down to 30 nm (lithography) and down to 5 nm (islands, multilayer?) will be required. Below the 11 major stages of alignment are described:

| Stage | Description | Settings |
| --- | --- | --- |
| 0 | Choose default voltage, currents |  |
| 1 | Beam alignment, sample, objective, 1st sector | No aperture, UV, 90° |
| 2 | Correction 2,3 fold stigmatism, objective, 1st sector | No aperture, UV, 90° |
| 3 | Focus mirror | No aperture, UV, |
| 4 | Beam alignment 2nd sector, projector | No aperture, UV, |
| 5 | Correction 2,3 fold stigmatism, 2nd sector | No aperture, UV, |
| 6 | Correction spherical defocus, mirror | No aperture, UV, |
|  | => maximum resolution using UV illumination |  |
| 7 | Dispersion correction 1st sector, readjust D1,2 | No aperture, UV, sample bias, 90° |
| 8 | Dispersion correction 2nd sector, readjust DM3,4 | No aperture, UV, sample bias |
| 9 | Correction chromatic defocus, mirror | No aperture, UV, sample bias |
| 10 | Insert aperture, repeat 1-9 | Aperture, UV |
| 11 | Switch to x rays, repeat 1-9 | Aperture, X-rays |
|  | => maximum resolution using x ray illumination |  |

0: Choose default voltages, currents

Goal: Focused image on CCD1, CCD2. Optimize distance by focusing with z-travel at fixed O.

1: Beam alignment, sample, objective, 1st sector

Settings: No aperture, UV, 90°

Condition: Focused image on CCD1

Goal: Align the sample axis with the objective axis and transfer the point on the sample, which is thus defined through the center of the following optical elements: F1, SM(1), F2, M. The beam is steered using the D1, D2, D3, D4 electric and magnetic deflectors.

- Align sample using xy tilt:
- Maximize intensity for objective voltage well below focus (assumes PEEM alignment is perfect). 
- Wobble O and minimize drift.

- Pick sample feature in the center of image expansion while wobbling O. This is the object that needs to be transferred through the center of all following lenses. In the following it will be called SC (sample center)

- Wobble F1 and minimize drift of SC with D1.

- Wobble F2 and minimize drift of SC with SM(1) trim (xy-deflect.).

- (Wobble F2 and minimize image breathing with SM(1) trim (focus))

- Steer SC through center of mirror aperture with DE3 (low magnification)

- Wobble M and minimize drift with DE3.

2: Correction 2,3 fold stigmatism, objective, 1st sector

Settings: No aperture, UV, 90°

Condition: Focused image on CCD1

Goal: Correct 2,3-fold stigmatism

- Wobble O and use D1 dodecapole to correct 2,3-fold stigmatism

- Wobble SM(1) focus and use D3 dodecapole to correct 2,3-fold stigmatism

3: Focus mirror

Settings: No aperture, UV

Condition: Sector 1 aligned, focused image on CCD2

Goal: Mirror operates at magnification 1

- Wobble F2 and minimize image breathing using M focus

4: Beam alignment 2nd sector, projector

Settings: No aperture, UV

Condition: Sector 1 aligned, focused image on CCD2

Goal: Steer SC through F2, SM(2), I1,2, P1,2 using D3,D4,D5,D6,D7

- Wobble F2 and minimize drift of SC using D4

- Wobble I1 and minimize drift of SC using SM(2) trim

- Wobble I2 and minimize drift of SC using D5, readjust SM(2) trim

- Wobble P1 and minimize drift of SC using D6

- Wobble P2 and minimize drift of SC using D7

5: Correction 2,3 fold stigmatism, 2nd sector

Settings: No aperture, UV

Condition: Sector 1 aligned, focused image on CCD2

Goal: Correct 2,3-fold stigmatism

- Wobble SM(1) focus and use D5 dodecapole to correct 2,3-fold stigmatism

6: Correction spherical defocus, mirror

Settings: No aperture, UV

Condition: Sector 1,2 aligned, focused image on CCD2, high magnification

Goal: Optimize spatial resolution by correcting spherical aberrations with mirror. Determine curve of optimal Cs correction in Cs vs. Cc diagram in absence of significant chromatic errors.

7: Dispersion correction 1st sector, readjust D1,2

Settings: No aperture, UV, 90°

Condition: Sector 1,2 aligned, no drift of SC while wobbling lens voltages, focused image on CCD1

Goal: Minimize dispersion in 1st sector of separator.

- Wobble S and study drift. This drift should be only due to due to chromatic dispersion in the separator since all other elements have cylinder symmetry and SC passes through their optical center (step 1)

- Adjust D1, D2, and SM trim coils to find “sweet spot” of separator while keeping SC in the center of F1, F2

- Finally the image should only breath with SC as the center. This breathing is due to defocus and has to be finally corrected using M

8: Dispersion correction 2nd sector, readjust DM3,4

Settings: No aperture, UV

Condition: Sector 1,2 aligned, no drift of SC while wobbling lens voltages and no dispersion in SM(1), focused image on CCD2

Goal: Minimize dispersion in 2nd sector of separator.

- Wobble S and study drift. This drift should be only due to due to chromatic dispersion in the separator since all other elements have cylinder symmetry and SC passes through their optical center (step 1)

- Adjust D3, D4, and SM trim coils to find “sweet spot” of separator while keeping SC in the center of F2, I1

- Finally the image should only breath with SC as the center. This breathing is due to defocus and has to be finally corrected using M

9: Correction chromatic defocus, mirror

Settings: No aperture, UV

Condition: Sector 1,2 aligned, no drift of SC while wobbling lens voltages and no dispersion in SM, focused image on CCD2

Goal: Optimize spatial resolution by correcting chromatic aberrations with mirror.

- Wobble S and minimize defocus with M. Stay on curve determined in step 6 which guarantees optimum correction of spherical aberrations.

10: Insert aperture, repeat 1-9

11: Switch to x rays, repeat 1-9