11/26/2002

SURFACE METEOROLOGICAL

OBSERVATION SYSTEM

OPERATING MANUAL

Operations Safety Procedures

There are a few potential dangers, which may be encountered while servicing or maintaining the surface meteorological stations.

Lightning	Standard SGP CART safety procedures prohibit outdoor activities in the presence of cloud to ground lightning.  These procedures apply to all surface meteorological station servicing and maintenance.

Lightning protection for the station is provided by a lightning rod at the top of the tower.  Sensor inputs are also protected by spark gaps.  The station is powered by batteries, which are kept charged by an AC/DC adapter.

Exposed	The rain gauge has exposed 120 VAC inside.

120 VAC	Whenever servicing or maintenance is required for the rain gauge, disconnect the AC power to the instrument prior to removing the lid.

Corrective maintenance

The Surface Meteorological Observation System, SMOS, is designed to require a minimum amount of routine maintenance.  A preventative maintenance, system check, and sensor verification schedule is presented below.

If the system does not operate properly or conform to the routine checks, the following action steps should be taken.

1)  On-site check.  The site operator should check for obvious problems like battery voltage out of operating range (9.6 to 16 V) and loose connections.

If the system is not operating, try restarting by turning off the power and then turning it back on.  The program is automatically reloaded on power-up.  The internal clock will need to be reset by using the keypad in *5 mode or the laptop computer.

If one sensor is not operating properly, replace it using the Installation Manual as a guide.  A spare of each sensor has been included for each five stations.  When a sensor is replaced, verification will need to be performed and documented.  The site configuration form containing the serial numbers for all parts will also need to be updated.

2)  Inform the instrument mentor.  If the site operator is unable to correct the problem, inform Michael Ritsche, the instrument mentor.  Mike may be contacted at:

Argonne National Laboratory

ERD Building 203

9700 S. Cass Ave

Argonne IL 60439

(630) 252-1554			Office

(630) 252-5498			FAX

(815) 756-4300			Home

mtritsche@anl.gov       Email

3)  Contact the vendor.  If the mentor is not available, contact the vendor directly.

Campbell Scientific Inc.

815 West 1800 North

Logan UT 84321-1784

(801) 753-2342			Voice

(801) 752-3268			FAX

Depot Maintenance

Some components that have been replaced due to malfunction or being out of tolerance will be returned to the vendor and/or manufacturer for service and/or repair.  Normally this will be done after consultation with the instrument mentor and/or the vendor.  Many items can and will be repaired in the Repair Lab at the Central Facility.  All sensors will be returned to the SGP CF and given to George Sawyer.  He will either have them repaired under warranty, repaired in-house in the Repair Lab, or send them to the appropriate vendor.  If a unit or system has to be shipped to the SGP, ship to the address listed below.  An ARM SGP shipment notification form ‘must’ be filled out and sent to SGP before the shipment arrives.  A link to that form is attached below.

George Sawyer

ARM SGP CART Site

309600 EW 28

Billings, OK  74630

(580) 388-4053			        Voice

(580) 388-4052			        FAX

george@ops.sgp.arm.gov           Email

http://sdsops.arm.gov/shipping/shipnote.html   Shipment Notification

All returns will be evaluated in the lab and/or repaired.  If the unit cannot be repaired in the lab it will be sent to the manufacturer or appropriate vendor.  An RMA number must be obtained from the vendor.  A SGP PR must be filled out with all of the appropriate information including the obtained RMA number.  The PR will list the quoted repair cost and/or a not to exceed cost.  The PR must be signed by all of the appropriate managers and then faxed to Laurel Chapman at Argonne National Lab.  When notified by Argonne that a PO is in place, pack the unit appropriately and ship it to its repair destination.

Laurel Chapman

ERD  Building 203 Rm# J013

Argonne National Laboratory

9700 S. Cass Ave

Argonne IL 60439

(630) 252-2887			   Voice

(630) 252-9792			   FAX

mailto:lachapman@anl.gov    Email

Packaging, Handling, and Transportation

The original shipping boxes should be used to ship the product out if available.  The boxes for the wind monitor and the rain gauge have been saved and should be used for packing these components for shipping.  The propeller will need to be removed from the wind monitor.  Padding and/or tape should be used to prevent the rain gauge bucket from tipping during shipment.

Any other components should be placed in plastic bags, and packed in sturdy boxes with adequate shock-absorbing materials such as packing foam, Styrofoam, bubble-wrap, peanuts, or cellulose wadding.

Overnight courier is the preferred method of shipping.

Once the manufacturer, or our repair lab has repaired the item, an inventory data sheet must be filled out placing it back into our available spares.  It is also suggested that the mentor, designated inventory control person, and/or relevant technician be notified that the item has been repaired and returned and is available for use.

Logistics

The following expendables are used with the SMOS system:

1.  The Vaisala HMI31/HMP35 Temperature and Relative Humidity Meter uses 9V batteries.

2.  The RM Young 18810 anemometer drive uses AA batteries.

3.  A desiccant pack, CSI type DSC 20/4, is used in each enclosure.

The desiccant packs will require replacement approximately every two months.  The site operators should maintain an adequate inventory of this item.  Contact Laurel Chapman to obtain replacement packs when the supply levels falls below one box.

Parts List

Each SMOS consists of the following parts.  Since the parts are interchangeable, a Station Description form is kept for each station, which contains the serial numbers of all components.  The form is updated whenever a component is changed.

CAMPBELL	Item Description

Part no.

UT930	10m Tower

ENC 16/18	Fiberglass Enclosure

CR10	Measurement and Control Module with

Wiring Panel

SM716	Storage Module

6234	SM716 Bracket Kit

PS12LA	12 V Power Supply with Charger and

Battery

SC32A	Optically Isolated RS232 Interface

Special	Vaisala PTB201A Digital Barometer

Special	Mounting Bracket for PTB201A

05103-L	RM Young Wind Monitor

UT018	Mounting Arm, 7 feet in length

HMP35C-L	Vaisala Temperature and RH Probe

UT018	Mounting Arm and Bracket

UT12VA	Radiation Shield for HMP35C

UDG01-L	CSI Snow Depth Sensor

6924	Mounting Arm

6879	Mounting Bracket for UDG01-L

Special	Novalynx 260-2500E-12 Electrically Heated

Rain/Snow gauge

Special	Nova Link Alter Shield

Spare Parts

Replacement parts have been included with the system for use when a component or sensor needs to be replaced for repair or calibration.  One each of the following items is needed on site for every five stations.  As there are currently fifteen stations, three of each item needs to be maintained on site.

CAMPBELL	Item Description

Part no.

CR10 	Measurement and Control Module

CR10KD 	Keyboard/display

SM716 	Storage Module

PS12LA 	Power Supply

SC32A 	Optically Isolated RS232 Interface

SC532 	9-pin Peripheral to RS232 Interface

05103-L 	RM Young Wind Monitor

HMP35C-L 	Vaisala Temperature and RH Probe

UDG01-L 	CSI Snow Depth Sensor

Special	Vaisala PTB201A Digital Barometer

Special	Novalynx 260-2500E-12 Rain/Snow Gauge

Sensor Verification and Checkout Equipment

The following supplemental equipment is supplied for system checkout and sensor verification:

A Laptop Computer w/PC208W

CR10KD Keyboard/display

7026 9-pin Computer to 25-pin Cable Adapter

SC32A Optically Isolated RS232 Interface

Vaisala PTB-220 Digital Barometer

Vaisala HMI31/HMP35 Digital Temperature and RH Probe

RM Young 18810 Anemometer Drive

RM Young 18212 Vane Angle Fixture

500 ml Volumetric Flask with Case

System operation

The SMOS system is designed to run unattended for prolonged periods of time with a minimal amount of routine maintenance.  The system will automatically load the program from the storage module and start data acquisition and processing at power up.  The only operator intervention required is setting the clock.  This can be accomplished by using the CR10KD keyboard/display in *5 mode or by using a laptop computer and PC208W.  Refer to the CR10 or PC208W Manual for instructions.

Routine Maintenance

Every Two Weeks:

The following are required to perform the checks and preventative maintenance:

1)  A laptop computer with PC208W

2)  A CR10KD keyboard/display

3)  A soft brush (to clean rain gauge buckets)

4)  A desiccant pack

5)  One person

If equipment or sensor replacement is needed, the following additional items may be required:

1)  Replacement equipment or sensors

2)  The Installation Manual

3)  Assorted tools, e.g., wrenches and screwdrivers

4)  Verification sensors - PTB-220 digital barometer,

HMI31/HMP35 temperature and RH probe, and 18810

anemometer drive

5)  A second person

The following servicing, maintenance, and checks shall be performed.

If there has been a communication problem, download the data from the storage module to the laptop computer.  Consult the Campbell and PC208W Manual for instructions to download data.  Record the date and time of the download on the site log.

Visually inspect the system and all sensors for any signs of damage.  Observe the wind monitor for movement of the vane and propeller.  Record any problems in the site log and notify the instrument mentor.

Use the laptop computer to scan through input channels 1 through 8.  Check for reasonable values.  A table of inputs and limits is given below.  If any value is questionable, record the observation in the site log and notify the instrument mentor.

| Location | Contents | Range * |
| --- | --- | --- |
| 1 | Wind speed | 0 to 45 m/s |
| 2 | Wind direction | 0 to 360 deg |
| 3 | Temperature | -40 to +50 C |
| 4 | Relative Humidity | 0 to 103% |
| 5 | Precipitation | 0 to 1 mm |
| 6 | Snow depth | -10 to 1000 mm |
| 7 | Vapor pressure | 0 to 10 kPa |
| 8 | Barometric pressure | 80 to 110 kPa |
| 10 | Wind Speed (RPM) | 0 to ?? |

*  Location 5 is one-second total precipitation in mm.

Location 16 is total accumulation in tips.

Snow depth measured only at temperatures below 10 C.

Set flag 7 high.  Remove any debris from the rain gauge screen and/or funnel.  After disconnecting the AC power, check inside for mice nests, etc.  If the buckets are dirty, clean them with a soft brush (e.g., a toothbrush).  Manually tip the bucket 3 times.  Check input location 16 for a count of 3.  Set flag 7 low.  Record results and time period during which flag 7 was set high, in the site log.  Consult the Campbell GRAPHTERM and PC208 Manuals for instructions to set the flag and monitor the inputs.

Note!  If flag 7 is not set low on the completion of this test, some data will be lost.

Remove any obstacles or debris from the ground under the snow depth sensor.  Trim vegetation growing below the sensor as needed.  Record the action in the site log. This procedure is currently not done.

Check and replace, as needed the desiccant in the enclosure.

See Appendix A for comprehensive procedures of the Bi-weekly checks.

Every Six Months:

The following are required to perform the checks, maintenance, and sensor verification:

1)  A laptop computer with PC208W

2)  A CR10KD keyboard/display

3)  Assorted tools, e.g., wrenches and screwdrivers

4)  A soft brush (to clean rain gauge buckets)

5)  A desiccant pack

6)  Verification sensors - PTB-220 digital barometer,

HMI31/HMP35 temperature and RH probe, 18810

anemometer drive, 18212 vane angle fixture, and 500

ml volumetric flask

7)  Two people

If equipment or sensor replacement is needed, the following additional items may be required:

1)  Replacement equipment or sensors

2)  The Installation Manual

In addition to the checks and maintenance described above for two week, perform sensor verification as described in Appendix B.

If any sensor is found to be out of calibration, it should be replaced by a spare and returned to the manufacturer for recalibration.

Every 24 Months:

Once every two years the Wind Monitor should be returned to the manufacturer for replacement and or inspection of the bearings and the potentiometer.

APPENDIX A

BI – WEEKLY PROCEDURES

SMOS

PREVENTATIVE MAINTENANCE PROCEDURES

Facility:		EF

Frequency:	Bi–Weekly

Mentor: 		Michael Ritsche

The following procedures are for performing scheduled bi-weekly maintenance on the SMOS sites.  All results should be recorded on the PM checklist.

INSPECTION OF SITE GROUNDS NEAR THE INSTRUMENT:

Visually inspect the site grounds around the instrument for hazards such as rodent burrows, buried conduit trench settling, and insect nests.

Checklist response:		No Problems Noted

Problem:  See Comments

Problem:  See CM Report

VISUAL INSPECTION OF INSTRUMENT COMPONNENTS:

Tower Structure:

Check that the mast is vertical and all tower base bolts are tight.  Check for any indications of broken or cracked tower structure components.  Inspect the lightning ground wire for tightness and damage.

Checklist response:		No Problems Noted

Problem:  See Comments

Problem:  See CM Report

Conduit, Cables and Connectors:

Check that all the conduits on the bottom of the control box are secure.  Check all conduits from the control box to the sensors for damage.  Check all sensor wires inside the control box for tightness and damage.  Check all the connectors at the sensors for damage, water intrusion and tightness.

Checklist response:		No Problems Noted

Problem:  See Comments

Problem:  See CM Report

Sensor Arms and Mounts:

Check that the wind sensor arm is aligned N-S.  This is done by using a compass.  Include the magnetic declination in the reading.  Check the Snow Depth and Temperature/Humidity sensor arms and mounts for tightness and damage.

Checklist response:		No Problem Noted

Problem:  See Comments

Problem:  See CM Report

Wind Sensor:

Visually check the sensor for obvious damage and that the sensor moves freely.  DO NOT lower the tower unless there is an obvious problem.

Checklist response:		No Problem Noted

Problem:  See Comments

Problem:  See CM Report

Temperature/Humidity Sensor:

To verify the accuracy and installation of the HMP-35C sensor the following tools will be needed.

1.	HMI31/HMP35 Temperature and Relative Humidity

Meter, recently calibrated.

- Computer or CR10KD

- Portable Aspiration Unit.

Remove the sensor from the radiation shield and inspect, clean or replace the humidity cap.  Verify that the sensor is working properly using the following procedures.

The verification of the temperature sensor is done with the Vaisala HMI31/HMP35 T/RH meter.  Mount the portable aspiration unit on the SMOS tower approximately chest high.  Place the HMP-35C T/RH sensor and the HMI31/HMP35 T/RH Meter in the holders of the portable aspiration unit.  Turn on the HMI31/HMP35 meter.  Allow the sensors to stabilize such that the temperature and RH readings are changing by less that +/- 0.2 Deg C or +/- 1% RH, then wait another 5 minutes before taking readings.  The air temperature measurement is stored in the data logger in location 3 and the RH readings are stored in location 4.  You can monitor this location with either a computer or the CR10KD.  Compare the data logger measurement to the read-out of the HMI31/HMP35.  The temperature readings should match within +/- 0.5 C.  The resolution of the data logger measurement is 0.01 C.  The RH readings should match within +/- 2 % RH from 0 - 90 % RH or +/- 3 % RH from 90 - 100 % RH.  The resolution of the data logger measurement is 1.0 % RH.

If the Bi-weekly check is being done at the Central Facility the portable aspirator is not needed.  The SMOS site at the Central Facility has an aspirated shield.  The HMI31/HMP35 Meter may be held near the intake of the aspirated shield.  Take care to position the HMI31/HMP35 T/RH meter out of direct sunlight.

NOTE:  If the probe or the T/RH Meter has been exposed to a temperature significantly different from the ambient air temperature and RH, it may take an appreciable time for the sensor to stabilize.  They may require rapid waving before they will stabilize.

Test Results:

Measurement	Cal Standard	SMOS Sensor	Pass/Fail

Temperature	____________	___________	_________

Humidity		____________	___________	_________

Temperature/Humidity Cap Change Scheduled?

Checklist response:		Yes

No

Snow Depth Sensor: (not done at this time)

Inspect the cable and connector for tightness and damage.  Ensure that the vegetation under the snow depth gauge is cut as low as possible.  Make sure that taller surrounding vegetation cannot be blown under the sensor.  Check that the bottom of the sensor is parallel with the ground.

Checklist response:		No Problems Noted

Problem:  See Comments

Problem:  See CM Report

Rain Gauge:

After turning off and locking out the heater circuit breaker, remove the rain gauge funnel and ensure that both the large and small funnels are clear of debris.  Check the wiring and connector for tightness and the housing for debris and damage.  Inspect all conduits and cables.  Re-install the rain gauge funnel.

Checklist response:		No Problem Noted

Problem:  See Comments

Problem:  See CM Reports

Digital Barometer:

Inspect the wiring from the Data logger to the Barometer box.  Ensure that the intake tubes are unobstructed.

Checklist response:		No Problem Noted

Problem:  See Comments

Problem:  See CM Report

Desiccant Check:

Check the desiccant pack inside the control box and replace if necessary.

Checklist response:		No Problem Noted

Problem:  See Comments

Problem:  See CM Report

Laptop Recorded Sensor Readings:

Connect the laptop computer and record the following readings:

Wind Directions:	(degrees)		range: 0 to 359

Wind Speed:		(m/s)

Temperature:		(degrees C)

Humidity:			(%)			range: 10 to 100

Pressure:			(mbs)

Snow Depth:		(mm)			range: +/- 25mm(not done)

Also note the sky condition and present weather.

Rain Gauge Tip Test:

Set flag 7 on the SMOS CR10 to disable intermediate processing.  Log the time that flag 7 was set.  Manually tip the rain gauge bucket several times.  Check location 16, TotCounts to ensure that the bucket tips are being counted.  Reset flag 7 and log the time that the flag was reset.

Tip Test Times:

Time flag 7 set:		hhmm (GMT)

Time flag 7 reset:		hhmm (GMT)

Check list response:	No Probem Noted

Problem:  See Comments

Problem:  See CM Report

Comments:

Enter any appropriate and applicable comments for this PM activity.

APPENDIX B:

6-MONTH CALIBRATION PROCEDURES

SMOS

SIX-MONTH CALIBRATION PROCEDURES

Facility:		EF

Frequency:	Twice a year

Mentor: 		Michael Ritsche

NOTE: Prior to all testing and verification, Flag 7 must be set high to prevent false data from being recorded.  This Flag should be reset when complete.

1.	05103 RM Young Wind Monitor

A.  Wind speed

To check that the wind speed sensor is measuring properly, the following equipment is needed.

1.	9/16" wrench

2.	#18810 Anemometer drive

3.	A computer or CR10KD

4.	Two people

Tilt the tower down to access the wind monitor. Refer to Appendix C for Tower lowering procedure. Remove the propeller from the propeller shaft and install the torque nut provided.  The anemometer drive connects to the wind monitor body and the motor connects to the torque nut.  Next, select one of the following RPM levels on the anemometer drive and start the unit.  Wait for the RPM level to stabilize.

RPM	RPM Reading

500            ___________

750            ___________

1000            ___________

Compare the reading from the anemometer drive and the sensor output as read from the computer or the CR10KD.  The RPM’s measured by the CR10 (Location 10) should be within +/- 1% of that indicated in the above table.

B.  Wind Direction

Using the angle fixture and a computer or a CR10KD, check that the sensor output (Location 2) matches the direction settings determined by the angle fixture #18212.  The angle fixture is inserted between the orientation ring and the wind monitor.  Once the angle fixture is installed, make sure that the notches of the orientation ring, angle fixture and wind monitor all match up.  Use the positioning arm to hold the tail of the wind monitor in the correct position.  Rotate the wind monitor through the positions listed below and compare them to directions measured by the data logger (Location 2).

Quadrant		Measured value

0			   ___________

90			   ___________

180			   ___________

270			   ___________

355			   ___________

The accuracy of this measurement is +/- 5.0 degrees, with a resolution of 1.0 degree.

NOTE: The absolute accuracy of this measurement depends on the exactness of the orientation of the sensor with respect to true north.

2.	HMP-35C Temperature & Relative Humidity

To verify the accuracy and installation of the HMP-35C sensor the following tools will be needed.

1.	HMI31/HMP35 Temperature and Relative Humidity

Meter, recently calibrated.

- Computer or CR10KD

- Portable Aspiration Unit.

Remove the sensor from the radiation shield and inspect, clean or replace the humidity cap.  Verify that the sensor is working properly using the following procedures.

The verification of the temperature sensor is done with the Vaisala HMI31/HMP35 T/RH meter.  Mount the portable aspiration unit on the SMOS tower approximately chest high.  Place the HMP-35C T/RH sensor and the HMI31/HMP35 T/RH Meter in the holders of the portable aspiration unit.  Turn on the HMI31/HMP35 meter.  Allow the sensors to stabilize such that the temperature and RH readings are changing by less that +/- 0.2 Deg C or +/- 1% RH, then wait another 5 minutes before taking readings.  The air temperature measurement is stored in the data logger in location 3 and the RH readings are stored in location 4.  You can monitor this location with either a computer or the CR10KD.  Compare the data logger measurement to the read-out of the HMI31/HMP35.  The temperature readings should match within +/- 0.5 C.  The resolution of the data logger measurement is 0.01 C.  The RH readings should match within +/- 2 % RH from 0 - 90 % RH or +/- 3 % RH from 90 - 100 % RH.  The resolution of the data logger measurement is 1.0 % RH.

If the Bi-weekly check is being done at the Central Facility the portable aspirator is not needed.  The SMOS site at the Central Facility has an aspirated shield.  The HMI31/HMP35 Meter may be held near the intake of the aspirated shield.  Take care to position the HMI31/HMP35 T/RH meter out of direct sunlight.

NOTE:  If the probe or the T/RH Meter has been exposed to a temperature significantly different from the ambient air temperature and RH, it may take an appreciable time for the sensor to stabilize.  They may require rapid waving before they will stabilize.

Test Results:

Measurement	Cal Standard	SMOS Sensor	Pass/Fail

Temperature	____________	___________	_________

Humidity		____________	___________	_________

3.  UDG01 Snow Depth Gauge (CURRENTLY NOT DONE)

NOTE:  The grass or vegetation under the gauge should be cut short and kept manicured to obtain accurate snow depth readings.

To start this test set Flag 8 high.  This is done either using the F8 key while monitoring the input locations in PC208W or using the CR10KD.   To use the CR10KD press *6AD8 to set the flag.  The flag will automatically reset in two minutes.  Once Flag 8 is high the UDG01 snow depth gauge will start making measurements once per second for two minutes.  The snow depth measurement is found in location 17 in the data logger, this measurement should be within +/- 25 mm of zero when nothing is under the snow depth gauge.  The snow depth gauge has a resolution of 0.5 mm.  A cardboard box of a known height at least 2 feet by 2 feet square is then placed under the gauge.  The measurement with the box under the gauge should show the height of the box within +/- 25 mm.

4.	2500E Rain/Snow Gauge

Two things need to be verified on this sensor.  The first is that the sensor records each tip of the bucket, and the other is to check that the heater portion of the rain gauge functions.

A.  Rain fall

The rainfall verification is accomplished with the 500 milliliter volumetric flask and plastic cup provided with the system.

Note this test should be started just after the even hour to insure that the system does not reset the number of counts during the test (This location resets every two hours).

Fill the volumetric flask to the indicator line and place the plastic cup in the top of the rain gauge.  Pour the 500 ml of water into the cup and allow the contents to slowly drain through the rain/snow gauge.  Using the data logger to measure the counts of the rain gauge the number of tips for 500 ml of water should be 25 to 27 counts; this value is stored in Location 16.  This test requires that the amount of water used is precisely 500 ml.  Once most of the water has drained from the cup, slowly pour the last of the water through the rain gauge.

B.  Heater

To verify the operation of the heater in the rain/snow gauge, the following test should be performed.  Place an ice cube in a plastic bag and hold the ice on the thermostat for 1 minute, and continue to hold in place while checking the heater (See Rain Gauge Manual in Installation Manual).  Feel the heater element at the base of the rain gauge and ensure that it has started to get hot.  The heater should be warm to the touch within 30 to 45 seconds after the thermostat has been tripped; the thermostat is set to trip at 37 degrees F.

5.  PTB-201 Barometric Pressure

The PTB-201 barometer is located inside of the enclosure.  A computer or the CR10KD can be used to monitor the location where the barometric pressure is being displayed (LOC: 18).  Establish communications with the data logger and monitor location 18 for barometric pressure in millibars.  Use the PTB-220 Digital Barometric pressure transducer to read the current barometric pressure in millibars and compare the two values.  The two values should be with in +/- 0.5 millibars of each other; the resolution of the sensor is 0.1 millibars.  The barometric pressure that is measured and recorded in the station (Location 8) is in Kilopascals and

is 1/10 of the millibars readings used in the verification of the sensor.

SURFACE METEOROLOGICAL OBSERVATION SYSTEM

SENSOR VERIFICATION FORM

SITE ID:_____________  LOCATION:____________________________

Tested by:__________________________  Date:_____/_____/_____

Started: __________________GMT   Ended: _________________GMT

1.	05103 Wind Monitor

A.	Wind Speed

RPM      Measured RPM (LOC: 10)	PASS/FAIL

500		 _________			_________

750		 _________			_________

1000		 _________			_________

The measured RPM should be within +/- 1% of the motor drive RPM.

Acceptable: YES  /   NO

B.	Wind Direction

Quadrant	  Measured value (LOC: 2) PASS/FAIL

0		   ___________			 _________

90		   ___________			 _________

180		   ___________			 _________

270		   ___________			 _________

355		   ___________			 _________

The accuracy of this measurement is +/- 5.0 degrees, with a resolution of 1.0 degree.

Acceptable: YES  /   NO

2.	HMP-35C Temperature & Relative Humidity

Measurement	Cal Standard	SMOS Sensor	Pass/Fail

Temperature	____________	___________	_________

Humidity		____________	___________	_________

(Temperature readings are in LOC: 3, RH in LOC:4)

The temp readings once stabilized should match within +/- 0.5 C.  The resolution of the data logger measurement is 0.01 C.

The RH readings once stabilized should match within +/- 2 % RH from 0 - 90 % RH, +/- 3 % RH from 90 - 100 % RH.  	The resolution of the measurement is 1.0 % RH.

Acceptable: YES  /   NO

3.	UDG01 Snow Depth Gauge  (not done at this time)

(LOC: 17)

Object Height Measured:__________   UDG01:______________

This measurement should be within +/- 25 mm.

Acceptable: YES  /   NO

4.	2500E Rain/Snow Gauge

A.	Rain, 500 ml should produce 25 to 27 tips

Number of tips recorded:________________ (LOC: 16)

Acceptable: YES  /   NO

B.	Heater

Acceptable: YES  /   NO

5.	Barometric Pressure

PTB-220 Reading:____________

PTB-201 Reading:____________ (LOC: 8)

The readings should agree within +/- 0.05 kPa or 0.5 mb.

Acceptable: YES  /   NO

6.	Comments:

APPENDIX C:

# 10 METER TOWER OPERATION PROCEDURE

SMOS

10-METER TOWER OPERATION PROCEDURE

Facility:		EF

Frequency:	As required

Mentor: 		Michael Ritsche

NOTE:  DO NOT LOWER the 10-meter tower if the wind gusts are in excess of 12 meters per second (26.8 miles per hour).  High winds can overcome the counter balance weights and cause the tower to lower at excessive speed.

NOTE:  DO NOT LOWER the 10-meter tower if there is visible cloud to ground lightning.  Lightning can strike miles ahead of a thunderstorm and the tower is typically the tallest object in the immediate vicinity.

To lower the 10-meter tower, the following equipment is needed.

- Two ¾ " wrenches

- Hard hats and safety glasses

- Cinderblock

- Two people

- Gloves when working with rope

Using the laptop, monitor the wind speed for several minutes and ensure that it never gusts above 12 meters per second.

NOTE:  Hardhats must be in place before tower activity begins.

To Lower:

Loosen and remove the lock nut from the bolt with 3/4" wrenches.

Remove the bolt from the beam and bracket and the bracket will then be out of the way for tower lowering.

Untie and unwrap the rope and place the cinderblock where the top of tower will be when it is down.

Pull the rope until the top of the tower is in a comfortable working position and secure the rope to the cinder block. Perform necessary PM work.

To Raise:

Untie the rope from the cinderblock and use it to slowly return the tower to an upright position.

Reinstall the bracket and bolt then tighten the lock nut with 3/4" wrenches.

Rewrap and secure the rope to the tower. Place the cinderblock back in its proper position at the base of the tower.

SMOS

OKMULGEE TOWER OPERATION PROCEDURE

Facility:		EF

Frequency:	As required

Mentor: 		Michael Ritsche

NOTE:  DO NOT CLIMB the tower if the wind gusts are in excess of 12 meters per second (26.8 miles per hour).  High winds can cause the tower to move making it unsafe to walk on the top platform.

NOTE:  DO NOT CLIMB the tower if there is visible cloud to ground lightning or if you hear thunder.  Lightning can strike miles ahead of a thunderstorm and the tower is the tallest object in the immediate vicinity.

NOTE:  Contact the Central Facility Safety Officer for tower access authorization prior to climbing the Okmulgee tower.

To operate the tower instrument booms, the following equipment is needed.

- Large standard screwdriver

- Safety Harnesses and Lanyards

- At least two people

- Hard hats and safety glasses

- Disposable gloves and disposable coveralls

Using the laptop, monitor the wind speed for several minutes and ensure that it never gusts above 12 meters per second.

One person must remain on the ground and must wear a hard hat when someone is climbing or working on the tower.

Anyone who is working on the tower and is below someone else must also wear a hard hat.

Inspect the Safety Harness for excessive wear or damage. The Safety Harness should be properly adjusted to properly fit the person wearing it and the lanyard must be secured to a tower vertical leg before boom operations begin.  A Lanyard must not be secured to the horizontal members of the tower because they can come loose.

With a large screwdriver, loosen the eyebolts that secure the boom.

When the eyebolts are loose, remove the bird deterrent strip from the boom by releasing the hand clamp and allowing the strip to hang from the safety rope that is attached to the strip. Mark a line on the boom at the edge of the roller block.

Firmly grasp underneath the boom and lift upwards slightly to remove tension of the block rollers: pull inward until instruments are in a comfortable working position.  Perform the necessary PM.

Extend the boom back to the operating position by grasping underneath the boom while pushing outward until the mark on the boom is in the correct position.

Tighten eyebolts (snugly only, not extremely tight, to avoid damage) with large screwdriver and reinstall the bird deterrent strip, securing it with the hand clamp.