INTERMET, INC

7732 Castleisland Drive

Sarasota, FL 34240

USA

Phone + 941 378 1313

Business Phone/Fax + 941 379 3536

Infotiempo@comcast.net

To:		Mr. Oscar Arango B.

WMO Sub Regional Office for the Americas

San Jose, Costa Rica

Subject:	The following report describes the activities of the contractor for the period September thru November 2004.

In General:

The Contractor has accomplished all the activities proposed  for the second period of the Proposal and Terms of Reference with the exception of the missions to the Mexican Sites and San Andres Island.

Specifically:

The following summary of activities is in two parts, namely those activities performed for the GUAN sites and those for the GSN sites.

GUAN SITES (7)

On a daily basis, Monday through Friday, Intermet INC carefully reviewed the THANKS2 REPORT generated by  NOAA/NCEP Central Operations for Data Management and Quality Assessment.  See http://www.ncep.noaa.gov/dmqab/qap/realtime/thanks/

The contractor evaluated the results of the report in terms of whether the site reported a full report or only a partial report.

If the daily report showed THANKS for that particular observation then no action was taken

Daily report showing the site only reported groups TTAA, TTBB and PPBB which are the  mandatory and significant levels for Pressure/Height, Temperature and Humidity as well as the winds aloft up to 100 Millibars .  If this occurred for 3 or 4 observations in succession then the Contractor contacted the site to see why the observation was being terminated at low levels.   The Contractor would then work with the site to correct the problem and resume normal operations.

Daily reports showing the site only reported TTAA which is just the mandatory levels up to 100 millibars.  This was acted on immediately by the contractor in determining the problem and assisting in correcting the problem.

The contractor dedicated a considerable amount of time into the above activities and was very successful in supporting the GUAN sites.  This near realtime monitoring is essential in maintaining the integrity of the GUAN network.

The contractor performed non-realtime monitoring of the same GUAN sites in assessing the quality of the reports in terms of temperature, heights, and winds.  During the first week of each month the contractor carefully reviewed and analyzed the Upper Air Inventories generated by NOAA/NCEP.  The contractor investigated the report for significant errors in their observations including the parameters, Temperature, Height, and Winds aloft.   The number of observations rejected because of significant errors was noted and the site was advised of these error and corrective measure were discussed.  Large departures in the selection and reporting of significant levels in all three parameters were noted and discussed with the particular observatory.

MAINTENANCE SUPPORT

The contractor identified a major error in the software of the upper air computers at one site which may also exist at the other sites.  This is still being investigated.  The contractor working with NOAA/NCEP determined the pressure heights being reported at one observatory were not hydrostatic.  The contractor believes this holds true for three of the other GUAN sites and this is currently being investigated.  This is a very time consuming investigation and will take several experts from different organizations to act on this problem.

NOAA/NCEP identified a wind-bias error at one of the GUAN sites and had been working for several months to correct the problem.  NOAA/NCEP asked the contractor for help and the contractor immediately determined the cause of the problem and worked with the site to realign the antenna on true north and eliminate the bias.

The Contractor worked for several months to reestablish the operation and reporting of one of the GUAN upper sites.  The site became silent in January 2004 and through the efforts of the contractor it finally resumed operations in mid-October 2004.

There were several other instances where the GUAN sites needed support from the contractor to maintain the integrity of their operation.  Some requests were trivial and some were not.

CIMAT TEMP REPORTS

During the time period in question the contractor worked with the GUAN sites and the Hadley Centre in the UK to increase the number of observatories reporting the CLIMAT TEMP report and to review the quality of those reports.  Because of the efforts of the contractor all of the 7 GUAN sites are currently reporting the CLIMAT TEMP reports on a monthly basis.  The contractor dedicated a large amount of time and effort into coordinating the exchange of these reports with the Hadley Center and with Weather Records center in Asheville.  Communications links with NOAA/NCEP had to be established and all participants had to agree to this exchange.

GSN Sites  (16)

The realtime Monitoring of the exchange of synoptic reports from the 16  sites is not performed by the contractor.  This would be too difficult and time consuming for the contractor.  However, monthly evaluation of the site activities is performed  and the sites are queried if there are a significant number of missing observations.

There are several problems with the GSN network in terms of quality and satisfying the requirements of the GCOS program.  All of these problems and recommendations were reported in the Contractor’s first Report dated September 28, 2004.   Many of the solutions to these problems require funding and funding is limited.  Therefore the contractor is working with GCOS and WMO authorities to prioritize the importance of the repairs to ensure the limited funds are spent wisely.

The contractor has not performed missions to the Mexican Sites or to the GSN site on San Andres Island.  These missions have been delayed because of timing problems with funding and availability of the appropriate personal.

CLIMAT Reports

The September 28th report still holds true for the status of CLIMAT reports.

During the period in question the contractor accomplished the following with regard to the CLIMAT reports:

Identified and helped resolve communications problems which led to the resumption of CLIMAT and CLIMAT TEMP reports

Identified major communications problems which are being resolved to allow GSN site to enter their Synops and Hourlies onto the GTS

Identified and helped resolve responsibilities and communications problems between RTHs and GCOS centers that had caused loss of CLIMAT reports.

Reviewed the CLIMAT data quality control procedures at the GSN sites and made recommendations

Danny L. Foster_____________________________January 7, 2005

Intermet, Inc.