### STANDARD OPERATION PROCEDURE - 3004

#### PROPER PACKAGING

PURPOSE:  To establish guidelines for proper use of packaging materials and methods used to maintain sterility of the item sterilized until the items are used.

POLICY:  The following types of sterile packaging are the only methods approved for use:

a.  Rigid sterilization container

b.  Paper-plastic peel pouch

c.  Woven or Non-woven wrapper

PROCEDURE:

A.  Rigid Sterilization Containers

- These are metal or plastic containers used to contain the instrument set and provide

protection to maintain sterility during and after sterilization.

- When the container is returned to the decontamination area, the inner

baskets  that contain the soiled instruments must be removed and instruments cleaned according to established decontamination procedures.

3.  Release the filter holder on both the lid and the main container, then remove and discard the disposable filter.  Also remove and discard used chemical indicator and load control cards.

4.  Thoroughly clean all parts of the container.  Cleaning can be accomplished by manual or mechanical methods.

5.  Carefully inspect all parts of the container for damage or inadequate cleaning, then pass through to Preparation Room.

6.  Replace the filters and secure in place with the filter holders.

7.  Reassemble the instrument set and place into the inner mesh basket along with count sheet.  Place the inner mesh basket with instruments into the container.  If using an internal indicator, place into the basket.

8.  Place lid onto the container bottom and close the latches.

9.  Secure container by attaching tamper evident seals onto the container latches.  NOTE:  The seal can also function as an external  chemical indicator for either steam or gas sterilization.

10.  Make certain the container label matches the actual contents of the container.

B.  Paper-Plastic Peel Pouch

1.  Select a package of appropriate size to accommodate the item(s) that are being packaged.  Allow enough room to securely seal the package.

2.  Cover all sharp pointed items with tip protectors to prevent tips from puncturing package.

3.  Place item(s) in package with the “handle end” of the instrument toward the end of the package that will be opened by the user.

4.  When double peel packaging, select a larger size outer package to allow the inner package to remain unfolded.  Folding the inner package may hinder the sterilant from penetrating the packaging.

5.  Peel pouches should be heat sealed by a 5/8” seal or double sealed if sealer is less than 5/8” thick.  Use of self-sealing peel pouches are also acceptable.  All seals on peel pouches must be carefully inspected for wrinkles or “tunnels” that could allow for the contents to be contaminated.

6.  Place a piece of sterilizing indicator tape on the transparent side of the package, which has had written description of the contents and initials of the preparer on the tape.

7.  For Sterrad sterilizer, select a Tyvek peel pouch of appropriate size to accommodate the item being packaged.  Place a piece of Sterrad indicator tape on the transparent side of the package.  The description of the contents and initials of the preparer are to be written on the tape.

C.  Wrapped Items (Woven or Non-Woven)

- Muslin wrappers must be laundered, inspected for holes, and delinted between each

use.  The minimum thread count is 240 threads per square inch and must have surged edges

2.  All items wrapped in muslin, paper, spun-bond fabrics or non-woven materials will be wrapped sequentially in two layers of wrapping material.

3.  Select the appropriate sized wrappers to adequately cover the item being wrapped. Place item in wrapper and wrap the item with either the envelope or square method.  NOTE: Generally, small items are best wrapped using the envelope method and large items using the square method.

4.  Seal package using appropriate sterilization indicator tape, depending upon the method of sterilization being used, which has written description of the content, initials of the preparer on the tape, and service where the item will be used or stored (if not SPD).

Reference:  VA Handbook Directive 7176, Supply, Processing and Distribution

Operation Requirements