Choice of High Voltage Bias Polarity for the SNS BLM Ion Chambers

R. L. Witkover and D. Gassner

March 5, 2001

# Background

After the FNAL style ion chambers were selected for the RHIC BLM detectors, acceptance tests were made on the first units received. The units were tested with the high voltage bias polarity preferred by their designer, Bob Shafer. He recommended that the positive voltage be applied to the center electrode of the cylindrical geometry chamber to minimize the effects of space charge and stray fields.

Figure 1 shows the RHIC version of the FNAL ion chamber. Changes included a G-10 plate to insulate the signal connector and break ground loops, and multiple connectors for both signal and HV to allow easy daisy-chaining. Another difference was that the signal was taken from the center electrode and the HV was applied to the outer electrode, opposite to FNAL. This was done to shield the signal (center) electrode for improved low signal characteristics, and to provide a printed circuit board on which to mount resistors to isolate the HV. A board already existed in the FNAL design but it was only used  to provide the mount for the chamber via the pins to the outer electrode. Making the outer cylinder the HV electrode allowed a single PC board to serve both needs.

Figure 1. RHIC version of the FNAL Ion Chamber

In the FNAL application, positive voltage was applied to the center electrode, allowing the electrons to move towards the region of increasing electric field. With negative high voltage on the outer electrode, the same electric field orientation could be achieved in the BNL design.

Tests by Mike Plum at the REX pulsed X-ray source Facility at LANL, which provided a dose of up to 2 Rad in a 50 nsec pulse, indicated that + HV on the center electrode reduced saturation effects in high dose rate situations. His data is reproduced in Figure 2.

Figure 2. Response of LANL ion chamber to high dose rate radiation. (M. Plum, et al.)

While the LANL chamber was filled with nitrogen and the RHIC/FNAL chamber is filled with argon, the conclusion could be expected to be the same. That is, the use of “positive polarity” will minimize saturation effects under high dose rates.

# Tests of the First RHIC Ion Chambers

The RHIC ion chambers were purchased using the same specifications and from the same vendor as FNAL and packaged as shown in Figure 1. Acceptance testing of the first batch indicated that the signal versus bias voltage was very different from the FNAL data and varied considerably from unit to unit. Typical test data supplied by FNAL showed a signal which leveled off at about 100 V bias and stayed flat until 2500-2700 V, after which it rose by 10-20% at 3000V. The tests of the RHIC detectors showed an extended, rounded “knee” which didn’t level off until 500-700 V and began rising again at 1500-2000V depending on individual detectors. Some increased by 50% at 3000V. Clearly this was not the response of the FNAL chambers and was unacceptable.

Discussions with Bob Shafer (now at LANL) and Alan Hahn (FNAL) indicated that the FNAL chambers were “baptized” before being tested. No clear statement of this “holy rite” could be obtained but it was essentially an HV “cleaning” of the detector. To simulate this at BNL a Tesla coil was used on the detector for 30 minutes, followed by pulse cleaning with a “Spot Knocker”. The “Spot Knocker” consisted of a HV power supply charging a 0.5 (F capacitor through a string of  resistors totaling 100 k(. The HV was raised until breakdown occurred, forming a relaxation oscillator. The results of this cleaning are shown in Figure 3. “Neg HV” is with the FNAL field direction but  with negative HV on the outer electrode rather than positive on the inner electrode. “Pos HV” is with the opposite field direction but the curve has been inverted for ease of plotting.

Figure 3. Effects of High Voltage Cleaning on a RHIC Ion Chamber

This was a particularly poor chamber, as can be seen from the upper most curve. There is virtually no flat region in which to work. The curve immediately below this is the same chamber after cleaning with the Tesla coil. The third curve down is the same chamber after cleaning with the Spot Knocker. A definite plateau is now observed.

It was clear that with nearly 500 chambers to be installed in RHIC, it would not be practical to spend more than 1 hour cleaning each one. The measurements with the opposite HV polarity showed that chambers tested as poor with –HV on the outer electrode, all showed acceptable flat operating regions in the opposite polarity even before the cleaning. This can be seen from Figure 3. Note that even this polarity showed some improvement after cleaning. At this point it was decided to use the ion chambers biased so that the electrons would be collected on the outer, rather than the inner, electrode. Referring to Figure 2, it was expected that the chambers would stay within the linear range for either polarity for the RHIC application. No further study of this unresolved question was made at that time.

# The SNS BLM  System Beam Loss Range

The SNS BLM System must cover a wide range of beam loss. The low end is set by the condition that the operators must be notified when the loss rate exceeds 1 W/m, or about 0.01% beam loss. The high end is not clearly defined. A limit of 1% beam loss at a single location has been taken. This represents a range of approximately 2.5 x 104. Now, the loss of 1 W/m roughly corresponds to 100 mRem/hr activation. A “rule of thumb” indicates that there is a factor of 500-1000 between the beam-off activation and the beam-on dose rate. Thus 100 mRem/hr implies 100 Rem/hr beam-on dose rate for 0.01% loss. Scaling this by the 2.5x104 range gives 2.5 x 106 Rem/hr (average) or 194 Rem/pulse (peak). Since the electrons are collected in a time the order of a microsecond,  pulse irradiation tests on ion chambers normally consider this the time of interest. For a SNS pulse length of 1 msec, this amounts to 0.194 Rem/microsecond, which is within the linear range of Figure 2. However, there has been a lot of hand waving to arrive at this number, and, since the figure of 1 % loss at a single point is completely arbitrary, it was considered important to again consider the question of HV bias polarity.

# Further Tests of the FNAL and RHIC Ion Chambers

The only FNAL ion chambers available at BNL had an unknown history and showed evidence of heavy irradiation so recently purchased glass detectors bottles were placed in an FNAL housing. This unit was tested with both positive and negative voltage on the inner conductor. Similarly, a typical RHIC unit was tested with both positive and negative voltage on the outer conductor. The four configurations tested are shown in Figure 4.  Test results are shown in Figure 5. The curves marked FNAL(+) and BNL(+) are with positive HV on the inner electrode (the preferred polarity), and negative on the outer electrode, respectively. The curves marked FNAL(-) and BNL(-) are with the opposite electric field. Again, all curves have been plotted positive regardless of sign.

Figure 4. BLM Test Configurations

Figure 5. Test results for FNAL and BNL BLM electrode configurations.

The normal configurations are FNAL(+) for the Tevatron and BNL(-) for RHIC, which, ignoring measurement noise excursions, show a plateau across nearly the range from several hundred volts to 3 kV.  The other two configurations show much smaller plateau regions. In the case of the BNL(+) data, the curve starts rising at about 2kV. This rise marks the start of the proportional region for the chamber in which the some electron gain enough energy between collisions to cause further ionization. The BNL(+) configuration is the one which will be less susceptible to saturation under high dose rate, however, since it cannot be used above 2 kV this advantage compromised.

# Conclusions

Clearly the FNAL(+) configuration offers the best response as well as being less susceptible to saturation under high radiation. While the BNL(+) configuration should have the same good high-dose rate saturation resistance, since the E-field is in the same direction, it clearly does not have as good plateau characteristics as either the FNAL(+) or the BNL(-) configuration. The cause of the poor plateau is not understood but might be due to local high gradient points resulting from mounting the outer electrode.

While the BNL(-) configuration BLMs should work linearly up to about a 1% local beam loss, modification of the design to use the FNAL(+) configuration is being studied. This would entail the use of 2 printed circuit boards rather than the present single one. It should be noted that the FNAL design only uses one PC board, to support the tube at the outer electrode end, with the HV isolation resistor string free-standing in the air. This is not good practice and a second PC board must be used to mount the resistors if this configuration is to be used.

# Acknowledgements

Paul Ziminski of BNL did all the BLM testing and made many helpful suggestions. Discussions with Bob Shafer and Mike Plum of LANL and I-Hung Chang of BNL were of great value both during the RHIC development and in understanding the present polarity problem.

# Endnotes

Shafer, Robert, “Doubler Beam Loss Monitor Detector”, 8/19/82. Private Communication

Plum, M., et al, “Fail-Safe Ion Chamber Errant Beam Detector Tailored for Personnel Protection”,  PAC 1989, p1556