Valeri Makeev

August 14, 2001

NUMI  BPM  detectors

# The goal.

- Continuous monitoring of the beam centroid position.

- Beam intensity measurement.

Specifications

- 12 detectors @ 0.2 mm resolution – beam transport

- 4 detectors @ 0.05mm resolution – beam targeting

##### The requirements for the BPM’s functional accuracy

- BPM’s  along  the transport line    +/- 0.2mm.

- BPM’s along the target line-up    +/- 0.05mm.

3.   Beam intensity measurement        +/- 3%.

What type BPM will be used?

Two different sizes, but the same in style of BPM’s will bused. For the first part transport line will be used standard BPM’s  (4 inches in diameter) as in current

MI transfer lines. And for the target line-up will use the BPM’s 2 .125 inches in diameter. For the last one no experimental data is available.

Both BPM ’s have diagonally cut circular cross-section pickup electrode. That electrode shape was chosen to have a linear response to beam displacement.

To define the electronic specifications (which must be reliable and affordable) a simple analytical calculation for was done output signal, noise budget and position sensitivities for that type BPM’s.

To be convinced of calculations the calculated values are compared to the measurements of  standard BPM.

Signal analysis and the main BPM’s parameters

We have made the approximation that the beam acts as a perfect current source in generating image current on the inside surface of the electrodes.

The analysis was done for Gaussian bunch shape:

Ib (t) = eNp /Exp(- t 2/22)

where  e  is the electron charge in Coulombs, Np  is the number of  the protons, is the RMS bunch temporal length (as bunch length we take 2, that correspond to 95% protons in the bunch).Further analysis will be done in the frequency domain.

Assuming that the bunch is symmetric in time, centered at time t =0

and is in a pulse train with the bunch spacing T, the Gaussian bunch shape can be expended in a cosine series with angular frequency0 = 2/T.

Ib (t) = Ib

where Am  =Exp(-m2022  /2) is the harmonic amplitude factor and

Ib= eN/T is the average (dc) bunch current.

The image current flowing over the characteristic impedance

is equal to the time derivative of the induced charge Q( t ) on the electrode

is (t) = dQ (t)/dt = d{Ib (t) /bc  Z(x, m0)}/dt

or

is (t) = 2Ibm0 /bcExp(-m2022  /2) Z(x, m0) sin(m0 t)

where q(t)= Ib (t) /bc is the beam charge density, Z(x, m0) is the “transfer function” from beam current to BPM current and  x is the offset center beam in horizontal plane.  That is true in some approximation when the inter -electrode coupling capacitance is ignored .

The peak output BPM current for harmonic m of the bunch angular

frequency 0  occurs at  t = +/- T/4m and is

i peak = 2Ib Z(x, m0) m0 /bc Exp(-m2022  /2)

What harmonic number is better to use?

To decrease the influence of the possible bunch length variation (38)ns on output signal it is preferable to use the main harmonic 52.8MHz  (m=1). In that case the signal variation will be 1.7dB. For m=2  the signal variation will be 6.6dB.

Remark, if we consider Microwave power sensors as the including many harmonics, the peak output current which occurs at the time t = +/- , in time domain representation, varies inversely as square of bunch length RMS

ipeak = Z(x)  eNp /bc2Exp(-1/2).

As a result the output signal would have a large variation (up to 17dB) at the same condition as before.

The transfer function Z(x,0).

To determine the diametrical part of the transfer function we used the expression/ 1/

where r is the displacement from the BPM center, b is the inside

electrode radius and  is the usual polar coordinate.  is  the angle from the median electrode plane to the plane defined by the beam and the BPM center.

The transfer function was calculated analytically for the real pickup electrode design. Also taken into account was the electrodes axes misalignment and inter-electrode coupling capacitance. The estimation was done for the zero beam displacement on y =rsincoordinate (=0) from the center of the BPM.

For the diagonal type electrode L()=L0 [1+cos()]/2.

The transfer functions for hypothetical Right and Left electrodes are:

ZR(x,0) = L0[ 1+KgKc(x +/2) /b ]  / 2Kg

ZL(x,0) = L0[ 1 - KgKc(x -/2) /b ]  / 2Kg

Where  Kg =(1 + 2 f /L0 ) –1   and

Kc = [R0 (2 + 1/ R00 Ce  ) 0 Ce]–1

(Kt = KgKc )

are the geometrical and the inter-electrodes capacitance factors  accordingly.

The   f  and L0  are the lengths of the flat part and the diagonally inside part

of the BPM electrodes, accordingly. The  is the electrodes axes misalignment.

Ce is the inter-electrodes capacitance and R0 is the load impedance.

The output signal voltage

(All calculations were done for beam intensity 1e10ppb)

Table 1. Some BPM and Beam parameters

| Parameter | Symbol | Standard BPM | Targeting BPM |
| --- | --- | --- | --- |
| Inside BPM radius | b | 49.149mm | 25.349mm |
| Diagonal electrode part | L0 | 98.298mm | 103.947mm |
| Distance between electrode | d | 4.49mm | 5 .57mm |
| Diagonally angle |  | 450 | 260 |
| Flat electrode part | f | 14.351mm | 6.968mm |
| Offset electrodes |  | 0.4mm | 0.2mm |
| Required resolution | x | 0.2mm | 0.05mm |
| Bunch frequency | f0 | 52.8MHz | 52.8MHz |
| RMS bunch length |  | 2ns | 2ns |
| Measurable bandwidth | B | 1MHz | 1MHz |
| Beam velocity |  | 1 | 1 |
| Inter-electrodes capacitance | Ce | 2.97Pf | 2.09pF |
| Geometrical factor | Kg | 0.774 | 0.882 |
| Inter-electrodes capacitance factor | Kc | 0.907 | 0.935 |
| Total correction factor | Kt | 0.702 | 0.825 |
| Load impedance | R0 | 50 | 50 |

If we use the previous formulas and parameters from Table 1 the peak signal in voltage  is given by

UR(L) =9.36210 -3Exp(-0.0552)  L0[1 +(-) Kt]/2 Kg

The output signals amplitude UR(L)  for Standard and Targeting BPM’s are presented in Table 2.

The displacement sensitivity and Beam position computing

There are two formulae for displacement sensitivity:

(UR - UL )/( UR + UL ) = Kt (1+ Kt  /2b )-1 x/b

and

20 Log10 ( UR / UL ) = x Sx .

The displacement sensitivity (Sx) is measured in dB/mm.

These parameters for two size BPM’s are presented in Table 2.

The wire measurements done by Jim Crisp and the calculated values  Sx  for standard BPM’s are Sx = 0.243 (0.277) dB/mm and  Sx = 0.249, respectively.

The position resolution

The position resolution can be derived as

(x) = ( b /Kt  ) (U)/U

Where (U)/U  is the relative accuracy for amplitude voltage measurement.

Assuming a required resolution x and the parameters in Table 1 the value (U)/U was specified. The results for both size BPM’s are presented in Table  2.

The thermal noise and loss budget

The thermal peak voltage noise was calculated as

U noise = = 0.64V    ( t=300 0 K )

Translating the thermal voltage noise into a position resolution (x)

(x)noise = ( b /Kt  )  U noise/ UR(L)min

For standard BPM’s (x)noise = 0.16 m and for targeting BPM’s

(x)noise = 0.06 m for lowest beam intensity(6e9 ppb).

Assuming a required resolution x (Table 1), we have a noise and loss budget

62dB and 58dB accordingly.

Dynamic range in beam position and dynamic range

available for beam intensity.

x [ 2N -1x (1+ Kt/2b) - b / Kt - /2 ] / (+2N -1 x Kt/b)

where N is bit digitization number,  is the beam intensity ratio(dynamic range).

Other values are specified in Table 1.

For   = 3.2 (Beam intensity range  9.5e10ppb to 3e10ppb) and N=12, the dynamic range in beam position is +/- 6.0mm.

Table2.

| BPM type | Standard BPM | Targeting BPM |
| --- | --- | --- |
| Output signal amplitude UR(L) (Volts) | 0.47 (x = 0) (x =+20mm) 0.34 (x = -20mm) | 0.444 (x = 0) 0.530 (x =+6mm) 0.358 (x = -6mm) |
| Sensitivity conversion factor Sx (dB/mm) | 0.249 | 0.58 |
| Measurement accuracy (U)/U (%) | 0.4 | 0.23 |

REFERENCES

1.   R.Shafer, “Beam Position Monitoring”, in API  Conf. Proc. 212, 26-58(1990)