dtpredict

SBC01-????-00

May 10, 1995

Prepared by:		J.W. Pflugrath

Molecular Structure Corporation

3200 Research Forest Drive

The Woodlands, Texas 77381

(713) 363-1033

(713) 364-3628 FAX

jwp@msc.com

Prepared for:		Contract No. 943072401

Mary Westbrook

Argonne National Laboratory

9700 South Cass Avenue

Argonne, Illinois 60439

(708) 252-8914

(708) 252-4021 FAX

westbrook@anlel.el.anl.gov

Reviewed by:

Approved by:

Copyright ( 1995, 1994 Molecular Structure Corporation

RESTRICTED RIGHT NOTICE SHORT FORM (JUNE 1987)

Use, reproduction, or disclosure is subject to restrictions set forth in Contract No. W-31

109-ENG-38 and Contract No. 943072401 with the University of Chicago, Operator of

Argonne National Laboratory

TABLE OF CONTENTS

1  Scope and Purpose

1.1. Definitions and Abbreviations

2  Background

3  Input images

3.1. Source keywords

3.2. Crystal keywords

3.3. Crystal goniometer keywords

3.4. Detector keywords

3.5. Detector goniometer keywords

3.6. Rotation keywords

4  Running dtpredict

4.1. Command line options

4.2. Examples

4.3. Results

4.4. Overcoming difficulties

5  Accessory programs

6  File formats

6.1. Reflection file format

1 Scope and Purpose

This document describes how to use the dtpredict module of d*TREK. dtpredict produces a list of reflections which will appear on a 2-dimensional position sensitive detector for a given rotation range.  The reflection list contains the Miller indices (hkl), the detector pixel coordinates, the detector millimeter coordinates, the starting, central and ending rotation angle values, the polarization correction factor, the Lorentz correction factor, the oblique incidence correction factor and the resolution  in Ångstroms.

In order to predict the reflection list, dtpredict requires knowledge of the source, the crystal properties, the crystal goniometer, the crystal rotation axis, the crystal rotation range, the detector and the detector goniometer.  dtpredict gleans all this information from the header of an image file, so that additional input is not required to predict reflections which will fall on an image.

1.1 Definitions and Abbreviations

Image	An image is the output of a position-sensitive detector. It has a header that describes the image and how it was made. dtpredict reads image headers to get information about images, so that users do not need to describe the image to the program.

Rotation	A rotation describes the goniometer movement while an image is exposed. It has the following properties:

Rotation axis

The 3D vectorial representation of the rotation axis in world coordinates when the crystal goniometer is at the datum position. The axis goes from the crystal position (0,0,0) in the axis direction such that a right-hand rotation is the positive direction. For example, if the omega axis points along X, then Rotation axis is (1, 0, 0). General  rotation axes may be easily specified this way. In addition, the crystal goniometer datum position specifies a rotation angle of 0 around this axis.

Start angle	The angle of the rotation axis at the start of exposure

End angle	The angle of the rotation axis at the end of exposure

Increment	The difference between the Start angle and End angle

Time		The exposure time in seconds of the image

Oscillations	The number of times the axis oscillated from the Start angle 			to the End angle back to the Start angle. A simple rotation 			will have Oscillations equal to 0.

Dark		The number of dark images collected before the start of the 			rotation and used for non-uniformity correction.

DarkUpdate	The dark image update interval. After this number of 			images, a new dark image should be made.

DarkLimit	The dark image change limit. If the dark current changes by 			more than this fraction, then a new dark image should be 			collected.

DCoffset	The number of DC offset images collected before the start 			of the rotation and used for non-uniformity correction.

DC Update	The DC offset image update interval. After this number of 			images, a new DC offset image should be made. This 			should probably always be 0.

Scan	A scan is a set of contiguous (in rotation angle) images with similar properties. The images are generated by a single detector so they have the same dimensions, exposure time, angular increment and non-uniformity of response. A scan has the same global properties as described for Rotation, but the Scan End Angle is the end angle of the last image in the scan.  Other Scan properties which are not pertinent to dtpredict are described in the dtfind documentation.

2 Background

In order to integrate the intensities of Bragg reflections in a single crystal diffraction experiment, their positions in images must be predicted accurately.  This can be done provided information about the source, the crystal, the crystal goniometer, the detector, the detector goniometer, and the rotation is known.  Internally, dtpredict uses the vectorially description of these objects in order to predict the positions of Bragg reflections.  This vectorial description has been described in the Proceedings of the EEC Cooperative Workshop on Position-Sensitive Detetor Software (Phases I & II) , ibid. (Phase III), and in Computational Aspects of Protein Crystal Data Analysis DL/SCI/R25.

In the following discussion, the origin of the world coordinate system is at the ideal crystal position.

Source Vector

The source vector goes from the origin along the primary beam in the same direction as the photons or neutrons.  If the world Z axis goes from the crystal towards the beam, then the design source vector sd is (0, 0, -1);

Source Orientation

The source may be mis-oriented slightly away from the source vector.  This can be parameterized by rotations around axes perpendicular to the source vector u1 and u2. which in turn give rotation matrices U1, U2.  The actual source vector used  by dtpredict s0 = U2U1Sd.

Crystal

The coordinates r of a reciprocal lattice point with Miller indices h for a crystal in a standard setting can be derived from the following equation:

r = Bh								(1)

where B is a matrix.  The reciprocal lattice of the crystal in a standard orientation at crystal goniometer angles all 0 as described by Busing and Levy (1966) is used.  In this orientation a* is along X, b* is in the XY plane.  The so-called B matrix looks like:

a*	b*cosg*	c*cosb*

B	=	0	b*sing*	-c*sinb*cosa			(2)

0	0		c*sinb*sina

Crystal Orientation

In the real world, crystals are never aligned as described by the B matrix.  They are generally rotated away from the standard setting.  This rotation is described by the crystal rotation matrix C which can be decomposed into 3 matrices C1, C2, C3 which are themselves rotations around 3 principle vectors c1, c2, c3.  In d*TREK and dtpredict, the vectors c1, c2, c3 are specified.  The orientation matrix encountered often in the literature is the C matrix times the B matrix.

Crystal Goniometer

The crystal goniometer is composed of at most 3 axes which are described vectorially as g1, g2, g3, where g1 is the axis closest to the baseplate (e.g. w) and g3 is the axis closest to the crystal (e.g. f)  The crystal goniometer datum position describes the values of rotations around these axes when the crystal rotation axis has a value of 0.  If the rotations around the three goniometer axis, the overall goniometer rotation matrix G is G3G2G1.

Crystal Rotation

In a rotation experiment, the crystal is rotated around an axis r1 to continually bring new Bragg reflections into the diffracting condition.  This rotation can be described by the matrix R.

With all the rotations applied, the equation to convert a Bragg reflection to a reciprocal lattice point in the diffracting position is:

rd = RGCBh					(3)

With regards to predicting reflections, the matrices G, C, and B are known, while the R matrix must be determined and the rotation angle derived from it.

Detector

A vectorial way to relate detector pixel coordinates (in the detector focal plane) has been described.  Let d1 and d2 be the basis vectors to convert the two pixel directions to the two focal plane millimeter directions.  Let do be the vector from the world origin to the origin of the coordinates on the detector (e.g. the 0.0, 0.0 millimeter coordinate).  The coordinates on the ideal flat focal plane of the detector that correspond to the point rd (3) satisfy (4):

Xd1 + Yd2 + do = a t

X

[ d1 |  d2 | d0]   Y  = at				(4)

1

where t is a vector coincident with the scattered beam wavevector s = s0 + rd

Let

[ d1 |  d2 | do] = d						(5)

D = [d]-1 = [D1 | D2 | Do]-1				(6)

(u (

v    = 	( v (    Dt					(7)

.			 (w(

X = u /w					(8)

Y = v / w

Now the rotation axis angle coordinate is determined.  The diffraction condition is

t ( t = s ( s						(9)

r ( r - 2 r ( s = 0

Detector Goniometer

The vectors d1, d2 and do in the preceding paragraph describe an unrotated detector.  Since the detector is often rotated by a goniometer, especially around the 2q axis, these vectors must be rotated, too.  Let E be the matrix that describes the detector goniometer matrix.  It is specified in the same way as the crystal goniometer matrix.

Thus, numerous vectors and rotation angles around these vectors are needed by dtpredict in order to predict a list of reflections.  It gets this information from an image header.

3 Input image

The dtpredict requires one or more images as input. It reads the headers of input images to get the experimental conditions of the source, crystal, crystal goniometer, rotation, detector and detector goniometer. dtpredict does not read the binary pixel data in the image which may be even be missing (i.e. the image is dimensioned 0 by 0).  If the detector spatial distortion and non-uniformity of response information in a header requires additional images to be read in, then dtpredict does so automatically.  A typical image header is shown below followed by descriptions of the keywords required by dtpredict.

{

HEADER_BYTES= 2048;

TYPE=mad;

SIZE1=512;

SIZE2=512;

COMMENT=Header edited by dtheaderedit;

SCAN_TITLE=start end inc time nOsc nDark nDup nDlim nDC nDCup;

SCAN_ROTATION=0.0 12.0 0.2 4 0 1 0 100 1 0;

SCAN_ROTATION_AXIS_NAME=omega;

SCAN_ROTATION_VECTOR=1.0 0.0 0.0;

SCAN_TEMPLATE=proteinX##.img;

SCAN_SEQ_INFO=1 1 1;

ROTATION=0.0 0.2 0.2 4 0 1 0 100 1 0;

ROTATION_VECTOR=1.0 0.0 0.0;

ROTATION_AXIS_NAME=Omega;

CRYSTAL_UNIT_CELL=82.34 88.29 103.65 90.00 90.00 90.00;

CRYSTAL_MOSAICSPREAD=0.145;

CRYSTAL_DESCRIPTION=Test of dtpredict;

CRYSTAL_SPACEGROUP=19;

CRYSTAL_REFINE_FLAGS=0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0;

CRYSTAL_ORIENT_ANGLES=10.04 -26.37 31.81;

CRYSTAL_ORIENT_VECTORS=1 0 0 0 1 0 0 0 1;

SOURCE_VECTORS=0.0 0.0 -1.0 1 0 0 0 1 0;

SOURCE_POLARZ=0.5 1.0 0.0 0.0;

SOURCE_SPECTRAL_DISPERSION=0.0002 0.0002;

SOURCE_SIZE=0.0 0.0 0.0 0.0;

SOURCE_CROSSFIRE=0.0 0.0 0.0 0.0;

SOURCE_WAVELENGTH=1 1.54178;

SOURCE_REFINE_FLAGS=0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0;

DETECTOR_NUMBER=1;

DETECTOR_NAMES=D0_;

D0_NONUNF_TYPE=Dark_only;

D0_NONUNF_INFO=nonunf.img dark.img;

D0_SPATIAL_DISTORTION_TYPE=Interp_spatial;

D0_SPATIAL_DISTORTION_INFO=distor;

D0_DETECTOR_DIMENSIONS=512 512;

D0_DETECTOR_SIZE=50.0 50.0;

D0_DETECTOR_VECTORS=1 0 0 0 1 0;

D0_DETECTOR_DESCRIPTION=ANL-SBC gold detector single module;

D0_DETECTOR_REFINE_FLAGS=0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0;

D0_GONIO_NUM_VALUES=6;

D0_GONIO_NAMES=RotX RotY RotZ TransX TransY TransZ;

D0_GONIO_UNITS=deg deg deg mm mm mm;

D0_GONIO_VECTORS=1 0 0 0 1 0 0 0 1 1 0 0 0 1 0 0 0 -1;

D0_GONIO_VALUES=0.0 0.0 0.0 0.0 0.0 102.3;

D0_GONIO_DESCRIPTION=A simple detector goniostat;

CRYSTAL_GONIO_NUM_VALUES=3;

CRYSTAL_GONIO_NAMES=Omega Chi Phi;

CRYSTAL_GONIO_UNITS=deg deg deg;

CRYSTAL_GONIO_VECTORS=1 0 0 0 1 0 1 0 0;

CRYSTAL_GONIO_VALUES=0.0 0.0 0.0;

CRYSTAL_GONIO_DESCRIPTION=A 3-circle eulerian crystal goniostat;

FILENAME=proteinX01.img;

BYTE_ORDER=big_endian;

Data_type=short int;

COMPRESSION=None;

}

3.1 Source keywords

The following keywords related to source properties are required in the header of the input image specified on the command line with the -scan or -image option:

SOURCE_VECTORS s0x s0y s0z s1x s1y s1z s2x s2y s2z

Specifies the direction s0 of the incident beam plus two vectors s1 and s2 perpendicular to s0 around which the beam rotation is refined.   For now, always have s0 = (0, 0, -1), s1 = (1, 0, 0) and s2 = (0, 1, 0).  This keyword need not be present as the vectors will default to these values.

SOURCE_ORIENT_ANGLES s1 s2

The small angles around s1 and s2 that the source vector s0 makes.  In practice, these are usually 0, 0.

SOURCE_POLARZ fp pnx pny pnz

The degree of polarization of the source and a vector normal to the plane of polarization as described by Kabsch (1988).

SOURCE_CROSSFIRE c11 c12 c12 c22

The source crossfire specified in degrees.  This is not used yet by dtpredict but may be used in the future.

SOURCE_WAVELENGTH w

The source wavelength in Ångstroms.

3.2 Crystal keywords

The following keywords related to crystal properties are required in the header of the input image specified on the command line with the -scan or -image option:

CRYSTAL_UNIT_CELL a b c a b g

This specifies the crystal unit cell in Ångstroms and degrees.

CRYSTAL_ORIENT_ANGLES c1 c2 c3

Three angles which specify the rotation of the crystal away from the standard orientation when the crystal goniometer angles are all 0.

CRYSTAL_ORIENT_VECTORS c1x c1y c1z c2x c2y c2z c3x c3y c3z

3D vectors which correspond to the directions of rotation for the CRYSTAL_ORIENT_ANGLES.  For now, these should be defined at (1, 0, 0), (0, 1, 0) and (0, 0, 1).  This keyword need not be present as the vectors will default to these values.

CRYSTAL_SPACEGROUP  n

The spacegroup of the crystal from the International Tables.  This is used to test for systematic absences which will not be predicted.

CRYSTAL_MOSAICSPREAD mosaicity

The effective mosaic spread of the crystal in degrees as described by Greenhough and Helliwell (1982).

3.3 Crystal goniometer keywords

The following keywords related to the crystal goniometer are required in the header of the input image specified on the command line with the -scan or -image option:

CRYSTAL_GONIO_NUM_VALUES n

The number of crystal goniometer axes.

CRYSTAL_GONIO_NAMES   name1 name2 name3 ... namen

The names of the axes.  They must appear in the order from the goniometer baseplate to the crystal.  The rotations of the axes will be applied in the reverse order.

CRYSTAL_GONIO_VECTORS g1x g1y g1z g2x g2y g2z g3x g3y g3z ... gnx gny gnz

The vectors describing goniometer rotation axes.

CRYSTAL_GONIO_VALUES  g1 g2 g3 ... gn

The goniometer axes angles at the datum position.  That is, the goniometer angles when the crystal is at a rotation angle of 0.

CRYSTAL_GONIO_UNITS   units1 units2 units3 ... unitsn

The units of the goniometer axes.  This MUST be ‘degs’.  Other possible units are ‘mm ‘which is reserved for translations.

3.4 Detector keywords

The following keywords related to the detector goniometer are required in the header of the input image specified on the command line with the -scan or -image option:

DETECTOR_NUMBER n

The number of detectors.  dtpredict will predict reflections for each detector in turn.

DETECTOR_NAMES   name1 name2 name3 ... namen

Specifies the names of each of the detectors.  The names are important as each name will be prepended to a keyword in order to get the detector properties from the image header.  For readability of the image header, each name should end with underscore character.  For the following keywords, the names found here are substituted into the dname_ characters.

dname_DETECTOR_DIMENSIONS dim1 dim2

The pixel dimensions of the detector image with dim1 being the dimension of the fastest varying direction.  In the header listing above, there is one detector whose name is D0_, so the keyword is D0_DETECTOR_DIMENSIONS.

dname_DETECTOR_SIZE mm1 mm2

The nominal size of the detector in millimeters.

dname_DETECTOR_VECTORS  p1x p1y p1z p2x p2y p2z

The vectors which relate pixel directions 1 & 2 to laboratory coordinates. For example, if pixel direction 1 is along Y and pixel direction 2 is antiparallel to X, the values would be 0 1 0 -1 0 0.

dname_NONUNF_TYPE  nonunf_type

Specifies the type of non-uniformity correction to apply to images.  dtpredict will use the NONUNF_TYPE and NONUNF_INFO values to flag reflections which are predicted to be in a bad non-uniformity regions.  The possible values for nonunf_type are:

None

Simple_scale

Dark_only

Dark_dcoffset

dname_NONUNF_INFO  dependent on nonunf_type

The values here depend on the option specified by dname_NONUNF_TYPE keyword, but generally they contain the non-uniformity files in the order: non-uniformity image, dark current image, and Dcoffset image.  In the example above, the type is Dark_only, so only the non-uniformity image (nonunf.img) and the dark current image (dark.img) are specified there.

dname_SPATIAL_DISTORTION_TYPE spatial_distortion_type

Specifies the type of spatial-distortion correction to apply to when converting from millimeters to pixels and the reverse.  The possible values for spatial_distortion_type are:

Simple_spatial

Interp_spatial

Other possibilities may be added in the future.  With type Simple_scale a simple linear scaling of pixels is performed to convert to millimeters. With type Interp_table the spatial distortion tables of the dtcalibrate program are used.  The extra information needed to complete the spatial distortion description is found in the next keyword.

dname_SPATIAL_DISTORTION_INFO dependent on spatial_distortion_type

If the spatial_distortion_type is Simple_spatial, then 4 values are read in which specify the primary beam center in pixels and the pixel sizes in millimeters for the fast and slow varying pixel directions in the image.  If the type is Interp_spatial as in the example above (line 35), then keyword value specifies the interpolation table fileset basename (line 36).

3.5 Detector goniometer keywords

The following keywords related to the detector goniometer are required in the header of the input image specified on the command line with the -scan or -image option:

dname_GONIO_NUM_VALUES n

The number of crystal goniometer axes.

dname_GONIO_NAMES   name1 name2 name3 ... namen

The names of the axes.  Rotations must appear in the sequence first and must be in the order they will be applied.  Translations must appear after all rotations.  Translations can be in any order as their values will be summed to arrive at the detector translation vector.

dname_GONIO_VECTORS g1x g1y g1z g2x g2y g2z g3x g3y g3z ... gnx gny gnz

The vectors describing goniometer rotation and translation axes.

dname_GONIO_VALUES  g1 g2 g3 ... gn

The goniometer values for the angles and translations at the datum position.  That is, the position of the detector when the image was exposed.

dname_GONIO_UNITS   units1 units2 units3 ... unitsn

The units of the goniometer values.  Rotation axes MUST be have units ‘degs’, while translations must be ‘mm‘.

3.6 Rotation keywords

The following keywords related to the crystal rotation are required in the header of the input image specified on the command line with the -image option.  If the -scan option is used, then the keywords below are prepended with SCAN_.  For example,  instead of the ROTATION, the keyword would be SCAN_ROTATION.  As a reminder of this in the text below, this keywords have the optional SCAN_ prepended to them.

SCAN_ROTATION Start End Incr time nosc ndark ndarkup dlim ndc ndcup

The 10 values specify the Start and End angles in degrees relative to the CRYSTAL_GONIO_VALUES, the angular increment per image, the exposure time in seconds, the number of oscillations, the number of dark images taken at the beginning of the rotation, the dark image update interval, the dark image change limit, the number of DC offset images taken at the beginning of the rotation and the DC offset image update interval.  If this keyword applies to a scan (SCAN_ROTATION), then the Start and End specify the overall start and end of the scan.  If this keyword applies to a single image (ROTATION), the Start and End specify the rotation angles within a scan for the image.

SCAN_ROTATION_VECTOR  rx ry rz

Specifies the rotation axis vector in world coordinates at the crystal goniometer datum position.

SCAN_ROTATION_AXIS_NAME  name

Specifies the axis name.  Usually the crystal is rotated around a crystal goniometer hardware axis and this name helps reminds one which one was used.   The name is not used by dtpredict.

4 Running dtpredict

After dtpredict has been installed and placed in your PATH, just enter dtpredict along with command line options to run it. It parses the command line arguments, reads any required input images, spatial distortion files, non-uniformity files and writes a reflection file if reflections are predicted. Messages are written to stdout and stderr as required. The syntax for running dtpredict is:

dtpredict [-scan scanfile | -image imagefile] [ -reso rmin rmax |

-rot start end | -list]

4.1 Command line options

Although dtpredict will run without any command line options if scan information is found in the header of a file called scan.img in the current working directory, most situations will benefit from the use of one or more command line options described below.

Command line options: Description and default

-scan sName	The header of image sName (no default) is read for the keywords described in section 3 to set up the experiment properties required for predicting reflections.  The keywords SCAN_ROTATION, SCAN_ROTATION_VECTOR are used to specify the rotation properties.  If neither a -scan option nor an -image option is specified, then dtpredict uses '-scan scan.img' as the default.

-image sName	The header of image sName (no default) is read for the keywords described in section 3 to set up the experiment properties required for predicting reflections.  The keywords ROTATION, ROTATION_VECTOR are used to specify the rotation properties.  If neither a -scan option nor a -image option is specified, then dtpredict uses '-scan scan.img' as the default.  Up to 20 -image options may be used on a single command line.

-list 	The option cause the experimental conditions to be echoed to stdout before a prediction is made.  This is useful to confirm that any image header has been interpreted properly.  The default is not to list this information.

-ref sReflectionFileName

Specifies an alternate output reflection file instead of the default dtpredict.ref. The file format is a standard d*TREK reflection file format described in section 7.

-reso fReso1 fReso2

A resolution range in Ångstroms which overrides the default resolution range determined from the source wavelength and the detector(s) position.  In order not miss any reflection on the image, do not use this option.

-rot fStart fEnd

A rotation range in degrees which overrides the default rotation range determined from the SCAN_ROTATION keyword.  fStart must be less than fEnd.  The keyword CRYSTAL_GONIO_VALUES in the image header specifies the position of the crystal at an fStart of 0.

4.2 Examples

dtpredict -scan ribozyme.0003 -rot 0 5

The above command line reads image ribozyme.0003 for scan information and overrides the rotation range value found in the image with a new range of 0 to 5 degrees.  Any non-uniformity correction information is found in the header of image ribozyme.0003 which in turn points to the non-uniformity files.  Results are written to the file dtpredict.ref.

dtpredict -scan hivrt.0001 -reso 2 50

The above command line reads image hivrt.0001 for scan information and overrides  the default resolution value of the entire detector surface with a resolution range of 2 to 50 Ångstroms.

dtpredict -image hivrt.0002 -image hivrt.0003 -list -ref hivrt.ref;

The above command line reads predicts reflections that are found on images hivrt.0002 and hivrt.0003.  It lists the experimental setup for each image.  All reflections are written to the file hivrt.ref instead of dtpredict.ref.

4.3 Results

The command:

dtpredict -scan proteinX01.img -list

produced the following output which is explained below.

File proteinX01.img successfully opened.

Header of file proteinX01.img successfully read.

Image is 512 by 512 pixels.

Data_type in header is short int.

Compression_type is None.

Byte_order is big_endian.

File distor.calpar successfully opened.

Warning in Cimage_header::nRead, header does not have proper ending!

Info in Cimage_header::nRead, able to repair ending!

File distor.x_int successfully opened.

Warning in Cimage_header::nRead, header does not have proper ending!

Info in Cimage_header::nRead, able to repair ending!

Header of file distor.x_int successfully read.

Image is 256 by 256 pixels.

Byte_order is big_endian.

File distor.y_int successfully opened.

Warning in Cimage_header::nRead, header does not have proper ending!

Info in Cimage_header::nRead, able to repair ending!

Header of file distor.y_int successfully read.

Image is 256 by 256 pixels.

Byte_order is big_endian.

File distor.inv_x_int successfully opened.

Warning in Cimage_header::nRead, header does not have proper ending!

Info in Cimage_header::nRead, able to repair ending!

Header of file distor.inv_x_int successfully read.

Image is 256 by 256 pixels.

Byte_order is big_endian.

File distor.inv_y_int successfully opened.

Warning in Cimage_header::nRead, header does not have proper ending!

Info in Cimage_header::nRead, able to repair ending!

Header of file distor.inv_y_int successfully read.

Image is 256 by 256 pixels.

Byte_order is big_endian.

D0_NONUNF_TYPE: >>Dark_only<<

Cnonunf::nInitValues called

File nonunf.img successfully opened.

Header of file nonunf.img successfully read.

Image is 512 by 512 pixels.

Data_type in header is short int.

Compression_type is None.

Byte_order is big_endian.

File dark.img successfully opened.

Header of file dark.img successfully read.

Image is 512 by 512 pixels.

Data_type in header is short int.

Compression_type is None.

Byte_order is big_endian.

There were 262144 pixels, with 0 bad pixels in the nonunf file.

Predict listing:

Crystal listing:

Unit cell lengths: 82.34, 88.29, 103.65

Unit cell  angles: 90, 90, 90

Unit cell  volume: 753515

Spacegroup number: 19

Orientation angles: 10.04, -26.37, 31.81

Mosaic Spread: 0.145

Description: Test of dtpredict

CRYSTAL_Goniometer listing:

Description: A 3-circle eulerian crystal goniostat

Number of values: 3

Value 0 name: Omega

units: deg

Vector: 1, 0, 0

Datum value: 0

Current value: Unknown

Value 1 name: Chi

units: deg

Vector: 0, 1, 0

Datum value: 0

Current value: Unknown

Value 2 name: Phi

units: deg

Vector: 1, 0, 0

Datum value: 0

Current value: Unknown

Source listing:

Single wavelength: 1.54178

Direction vector: 0, 0, -1

Spectral Dispersion: 0.0002, 0.0002

Crossfire: 0, 0, 0, 0

Polarization: 0.5, 1, 0, 0

Size: 0, 0, 0, 0

Intensity: 0

Refinement flags: Not yet available

D0_Detector listing:

Pixel dimensions: 512, 512

Nominal size in mm: 50, 50

Description: ANL-SBC gold detector single module

Rotation: 0, 0, 0

Translation: 0, 0, -102.3

Fast pixel vector: 1, 0, 0

Slow pixel vector: 0, 1, 0

D0_Goniometer listing:

Description: A simple detector goniostat

Number of values: 6

Value 0 name: RotX

units: deg

Vector: 1, 0, 0

Datum value: 0

Current value: Unknown

Value 1 name: RotY

units: deg

Vector: 0, 1, 0

Datum value: 0

Current value: Unknown

Value 2 name: RotZ

units: deg

Vector: 0, 0, 1

Datum value: 0

Current value: Unknown

Value 3 name: TransX

units: mm

Vector: 1, 0, 0

Datum value: 0

Current value: Unknown

Value 4 name: TransY

units: mm

Vector: 0, 1, 0

Datum value: 0

Current value: Unknown

Value 5 name: TransZ

units: mm

Vector: 0, 0, -1

Datum value: 102.3

Current value: Unknown

DetResolution min: 154178

DetResolution max: 5.23813

Rotation list:

Start: 0

End: 12

Increment: 0.2

Time: 4

Oscillations: 0

Axis vector: 1, 0, 0

Axis name: Omega

Dark start: 1

Dark update: 0

Dark change limit: 100

DCoffset  start: 1

DCoffset update: 0

Predict resol min: 999999

Predict resol max: 0.5

dtpredict: There were 828 reflections predicted.

dtpredict: Reflections written to dtpredict.ref

Explanation

Lines 1-9	The file proteinX01.img was opened and read in.

Lines 10-35	The detector spatial distortion correction requires that an interpolation table fileset be read in.  These lines show these files have been successfully read.

Lines 36-48	The detector non-uniformity correction requires that non-uniformity image and a dark image be read in.  These lines show these files have been successfully read.

Lines 49-142	These lines show the properties of the crystal, crystal goniometer, source, detector, and rotation which is the result of the -list command line option.  Most of this information is found in the header of the input image proteinX01.img although some is calculated from values in the header.

Lines 64, 69, 74, 99, 104, 109, 114, 119, 124

These shows the values of the crystal and detector goniometers when the rotation angle is 0.

Lines 65, 70, 5, 100, 105, 110, 115, 120, 125

The current values of the crystal and detector goniometers are unknown since dtpredict is not physically connected in any way to these hardware devices.  These values are not used by dtpredict in any case.

Lines 129-130	Reflections wll be predicted which diffract between a rotation angle of 0( and 12(.

Lines 126-127	The calculated minimum and maximum resolution of pixels on the detector.  These will ultimately limit the resolution range of the predicted reflections although the range may be restricted further by a -reso command line option.  In no case will dtpredict predict reflections that do not fall on a detector.

Lines 141-142	The dtpredict resolution range which may be restricted further by the detector size and position (see Lines 126-127).

Line 143	There were 828 reflections predicted.

Line 144	The reflections were written to a file called dtpredict.ref which is the default filename.

4.4 Overcoming difficulties

Since dtpredict gets almost all of its input information from an image header, difficulties may arise when the header is incomplete.  Be sure that the detector spatial distortion information and non-uniformity information is correct and the files exist in the current directory or the filenames in the header include the path or directory.

The output of dtpredict clearly shows when any files have not been read properly.  Use the -list command line option to display the source, crystal, goniometers, detector and rotation properties.  If these are incorrect, then the image header is incorrect and should be modified.

In these cases, add or correct the information (use dtheaderedit, see section 5) and run dtpredict again.

5 Accessory programs

Since dtpredict gets all image information from the image header, it may be useful to create and edit image headers from time to time.  Since headers are ASCII characters, use a text editor such as GNU emacs to do this.  The program dtheaderedit may also be used.  It keeps track of the length of the header and prevents the modification of keywords that describe the binary data in the image file.  Input to dtheaderedit comes from stdin and consists of  the name of the input image file, the name of the output image file, then keyword value pairs until an end-of-file is found on the stdin.  Separate keyword and value by a whitespace (not an equal sign).  Any value that contains whitespace must be enclosed in single quotes.

If a keyword exists in the header, the value is replaced.  If the keyword does not exist, the keyword=value; pair is added.  To delete a keyword=value; pair enter the keyword followed by 2 single quotes:  keyword ‘’.

An example of a dtheaderedit session can be found in the dtfind documentation.

One potential problem is that dtheaderedit requires an image file.  If one is not handy, one may be created with the dtcreate program:

% dtcreate

dtcreate: Copyright (c) 1995 Molecular Structure Corporation

dtcreate: Enter output image name,

its dimensions and constant value: predict.img 0 0 100

File predict.img successfully opened.

dtcreate: Done

As shown above it is perfectly OK for an image file to have no pixels.  The file contains only a header which must be modified with dtheaderedit to serve as input to dtpredict.

6 File formats

Image file formats are discussed in the calibrate documentation.  An example of an image header is given in Section 3.

6.1 Reflection file format

Reflection files have the following format.  First they are all ASCII characters.  The first line consists of 3 numbers designating the number of integer, float and string fields that each reflection has.  After this are the field labels as character strings, one per line for the integer fields first, then the float fields, then the string fields.  The first 3 fields of the reflection file are H, K, L.  The first 2 float fields are Intensity and SigmaI.  No other fields are required.  Field labels must contain no whitespace.  After the field labels, comes the reflection data, one reflection per line in free format.  This format is a prelude to using a true CIF reflection file format.

dtpredict adds the following fields to a reflection list and calculates their values for each reflection

Detector_number

The detector number from the DETECTOR_NAMES field.  The first detector is numbered 0.

Nonunf_flag

A non-uniformity of response flag.  If not 0, the predicted pixel position is flagged as bad in the non-uniformity file.

Calc_pixel1

Calc_pixel2

These are the calculated pixel coordinates of the reflection centroid on the detector and in the image.  The first value corresponds to the fast varying pixel direction and the second value to the slowest varying pixel direction.  These directions may be along a row, a column, horizontal, vertical or none of these depending on the detector orientation, position and spatial distortion.

Calc_1mm

Calc_2mm

The calculated millimeter coordinates for the reflection centroid on an ideal focal plane detector.  The directions of the first and second mm axes are system dependent, but can be generally considered to lie along the Xmm and Ymm laboratory axes.

Calc_rot_start

The calculated rotation angle in degrees when the reflection first begins to diffract.

Calc_rot_end

The calculated rotation angle in degrees when the reflection stops diffracting.

Calc_rot_mid

The calculated rotation angle when the reflection lies exactly on the Ewald sphere.

Calc_rot_width

The calculated rotation angle width in degrees.

Calc_polarz

The calculated polarization correction factor P, such that Itrue = P Iobs.

Calc_lorentz

The calculated Lorentz correction factor L, such that Itrue = L Iobs.

Calc_oblique

The calculated oblique incidence correction factor O, such that Itrue = O Iobs. Be careful as it may not be necessary to correct for this.

Calc_partial

A field for the partiality of a reflection which is not calculated by dtpredict.

Resolution

The resolution of the reflection in Ångstroms which is dependent on the source wavelength (of course).

Calc_recip1

Calc_recip2

Calc_recip3

The reciprocal lattice coordinates of the reflection on the Ewald sphere in dimensionless units (i.e. the Ewald sphere has a radius of 1 Å-1).

Below is a truncated example of a reflection file created by dtpredict.  Lines 24-34 are too long to fit across a page so they are wrapped below.

H   5   18   0

K

L

Detector_number

Nonunf_flag

Intensity

SigmaI

Calc_pixel0

Calc_pixel1

Calc_0mm

Calc_1mm

Calc_rot_start

Calc_rot_end

Calc_rot_mid

Calc_rot_width

Calc_polarz

Calc_lorentz

Calc_oblique

Calc_partial

Resolution

Calc_recip1

Calc_recip2

Calc_recip3

2  -12    0    0 -999 -999 5.539 2.48132 -21.647 -21.7509 2.84386 3.053 2.94843 0.209144 0.979688 4.96143 -1.05489 -999 5.33334 0.200591 0.201555 -0.0520377

1  -11    0    0 -999 -999 1.5939 29.8988 -22.061 -19.4289 0.843981 1.06717 0.955578 0.223193 0.98344 5.49488 -1.04359 -999 5.53591 0.206642 0.181987 -0.0417703

2  -11    0    0 -999 -999 31.5765 23.4825 -19.5596 -20.087 3.60743 3.81433 3.71088 0.206898 0.982499 5.34514 -1.04954 -999 5.81014 0.182173 0.187086 -0.0472011

2  -11    0    0 -999 -999 14.9786 13.5911 -20.9124 -20.8533 -0.146523 0.0620606 -0.0422311 0.208583 0.980817 5.10531 -1.04069 -999 5.50354 0.19643 0.195875 -0.0390982

3  -11    0    0 -999 -999 44.9524 7.10276 -18.4115 -21.517 2.3765 2.57117 2.47383 0.194676 0.979806 4.97597 -1.04661 -999 5.74824 0.17196 0.200966 -0.0445335

-1  -10    0    0 -999 -999 0.385201 82.9582 -22.2876 -14.9281 3.94615 4.21199 4.07907 0.265837 0.990235 7.1555 -1.04417 -999 6.05412 0.20865 0.139753 -0.0422987

0  -10    0    0 -999 -999 13.6456 66.6067 -21.1375 -16.365 2.64951 2.89076 2.77013 0.241251 0.988198 6.50901 -1.04125 -999 6.06841 0.198437 0.153633 -0.0396151

1  -10    0    0 -999 -999 26.9955 50.292 -19.9939 -17.7943 1.45457 1.67635 1.56546 0.221777 0.985969 5.96953 -1.03835 -999 6.05412 0.188224 0.167517 -0.0369372

2  -10    0    0 -999 -999 40.2075 34.1359 -18.8567 -19.2161 0.34025 0.546486 0.443368 0.206236 0.983546 5.51256 -1.03548 -999 6.01187 0.178012 0.181404 -0.034265

3  -10    0    0 -999 -999 69.6918 27.5982 -16.3569 -19.8671 2.93466 3.12618 3.03042 0.191514 0.98261 5.36213 -1.04135 -999 6.29788 0.153542 0.186493 -0.0397066

4  -10    0    0 -999 -999 82.7698 11.3977 -15.2267 -21.287 1.65424 1.83562 1.74493 0.181387 0.979925 4.99062 -1.03847 -999 6.18864 0.14333 0.200376 -0.0370447

...

The Intensity, SigmaI and Calc_partial values have not been set by dtpredict, so they have default values of -999.