C1:  Preparation of Panel Testing Slides with Known Contents

1.	Introduction

This procedure is a self-explanatory laboratory method for producing multiple test slides from AFB positive and negative samples.  Your laboratory staff should read and understand both the procedure and the testing protocols before developing test slides.  This procedure has been reproduced/validated in state and national laboratories.   If your laboratory has difficulty in producing slides that meet the requirements for consistency you should either: 1) review the procedure with special attention to the steps of heating and   re-suspension; or 2) select patient specimens with less mucus.   The sample development procedure  requires materials that are routinely available in a national or regional reference laboratory in a low-income country.  If your laboratory has continued difficulties with clumping of AFB that prevents slide to slide consistency, the use of N-acetyl-L cysteine (NALC) may improve the quality of the slides.  Your laboratory should demonstrate proficiency in producing samples with a minimum of 25-30 slides that are consistent for negative and low numbers of AFB before proceeding to developing test slide sets.

NaOH method

(ref Dr. Nguyen Ngoc Lan, Pham Ngoc Thach Hospital, Ho Chi Minh City, Vietnam and Dr. Alex Sloutsky, Massachusetts Dept. Health)

2.	Materials Required

Note: Processing should be performed in a Biological Safety Cabinet.

50 ml plastic screw cap tubes

40% Formaldehyde

4% NaOH

Vortex

Water bath at 55-60°C

Distilled water

Centrifuge

Slides

Positive specimen (fresh specimens, no more than 2 days old, are preferred)

Amount: 	3 ml or more;

AFB load: 	>2+ AFB by Ziehl-Neelsen direct smear;

Color:	White to light green; blood stained specimens should be avoided;

Thickness:	Watery (less mucous) specimens are preferred to increase consistency.

Negative specimen (fresh specimens, no more than 2 days old, are preferred)

Amount:	5 ml or more;

Color:	white to green;

Thickness:	Watery (less mucous) specimens are preferred to increase consistency

Note:	An AFB negative specimen with 20 or more white blood cells per field is preferred.

3.  Preparation of AFB Positive Stock

a.	Place 3 ml of AFB positive specimen into a 50 ml screw cap plastic tube. If volume of the specimen is more than 3 ml, aliquot it into separate tubes.

b.	Add 1 drop (approx. 50 µl) of 40% Formaldehyde per 1 ml of sputum, vortex well.

c.	Incubate for 1 hour at room temperature (25- 30oC).

d.	Add 1 ml of 4% NaOH (if the sputum is too thick, add up to 2 ml of NaOH solution so that the final concentration of NaOH is always 1-2%).

e.	Vortex thoroughly for 4-5 min.

f.	Add up to 20 ml of distilled water, mix  well.

g.	Incubate in a water bath for 30 min. at 55-60oC, mix occasionally by inverting the tube during incubation. If there is no water bath available, boil a beaker of water, cool to 90-95oC and place the tube in the beaker for 20-25 min.  It is important to maintain the incubation temperature in the 55-90oC range.

h.	Add distilled water to a total volume of  40 ml, mix by inversion.

i.	Centrifuge @ 3,000 x g for 20 min. at room  temperature (25-30oC).

j.	Decant supernatant carefully, add 0.5-1 ml of distilled water to resuspend pellets. If initial sputum was aliquoted into portions, pellets from the same specimen are combined, prior to resuspending.

Note:	It is advisable to avoid specimens containing impurities (food remains etc.)  However if the impurities are still found in the sediment after it is dissolved in distilled water, filter the specimen through the gauze and recentrifuge it.

4.  Preparation of AFB Negative Stock

a.	Distribute 3-4 ml aliquots of AFB-negative sputum into 50 ml screw cap tubes.

b.	Note: Several good quality negative sputa can be pooled together and then split into 3 ml aliquots.   Sputa should be checked for AFB prior to pooling.

c.	Add 1 drop ( approx. 50 µl) of 40% Formaldehyde per 1 ml of sputum, vortex well.

d.	Incubate for 1 hour at room temperature (25-30oC).

e.	Add 1 ml of 4% NaOH (if the sputum is too thick, add up to 2 ml of NaOH solution so that the final concentration of NaOH is always 1-2%).

f.	Vortex for 2-3 min.

g.	Add up to 20 ml of distilled water, mix well.

h.	Incubate in a water bath for 10 min. at 55-60oC (Note: the negative specimen should be heated for a shorter period than the positive specimen to preserve white blood cells). If there is no water bath available, boil a beaker of water, cool to 90-95oC and place the tube in the beaker for 5-10 min.

This preparation is used as a diluent in the Dilution Procedure (step 7).

5.	Evaluation of Positive Stock Preparations

a.	If foam has formed on top of the stock solution, pipette the contents from beneath the foam into a fresh tube.

b.	Using a standard microbiological loop make 2-3 test smears (approx. 1x2 cm in size) from the suspension for evaluation of the stock preparations.

c.	Use a well leveled surface for drying the smears.

Positive stock:  It is optimal to have concentration 50-60 AFB per microscope field.

6.	Dilution Procedure

a.	Using negative preparation as a diluent make dilutions according to WHO Guidelines for AFB quantification:

0 AFB/100 fields:	negative

1-9 AFB/100 fields:	exact # of AFB required

10-99 AFB/100 fields:	1+

1-10 AFB/field:	2+

>10 AFB/field :	3+

b.	Choose suitable AFB concentration on a case-to-case basis within suggested range.  For better results, however, it may be recommended using 20 AFB/field for 3+ smears, 5 AFB/field for 2+ smear, 50 AFB/100 fields for 1+ smears, and 5 AFB/100 fields for “exact” smears.

c.	Make 3-4 ml of each suspension in order to be able to generate sufficient amount of smears.

d.	For easy calculations both AFB-positive and AFB-negative aliquots are measured in drops. Calibrate one typical disposable Pasteur pipette by measuring the number of drops in 1 ml of sputum suspension.  Note: do not use water for calibration since the amount of drops may be different from sputum due to the lack of viscosity.

e.	For calculation of the dilution factor use the following formula :

N = ( DC / AC ) * A

where :

N   - is amount of drops of positive sputum to be added.

DC - is desired AFB concentration.

AC - is actual AFB concentration.

A   - is the amount of drops in a given volume that was estimated during

calibration.

Example:	AFB concentration in the stock suspension (AC) is 65 AFB/field and we have to  prepare  4 ml (A = 60 drops) of  2+ suspension (DC=5 AFB/field ).

In this case  N = (5 AFB /65 AFB )*60 drops

N = 4.6 drops (approx 5 drops). So, 5 drops of the positive prep is mixed with 55   (60 - 5 = 55 ) drops of the negative prep.

Procedural notes:

1.	It is important for reading and interpretation of results that appearance of the smears is more or less consistent, and that is why it would be beneficial to keep the amount of leucocytes as stable as possible in various dilutions.  In order to achieve this, it is suggested to dilute negative sputum with distilled water (prior to adding NaOH) when the amount of leukocytes is relatively high and avoid dilution if the amount of leukocytes is low.

2.	It would be also useful when making 1+  suspension to consider making two different concentrations: 50 AFB/100 fields for 1+ smear preparation and 15 AFB/100 fields for further dilution to “exact” count smear.

7.	Prepare and Validate Batches of Slides

a.	Using diluted stock preparations, prepare slide batches (50-100 slides per batch is recommended).  Note: If laboratories are proficient in developing consistent slides, then developing many slides from fewer samples will help to save time.   Heat fixed slides should last for months if stored in a cool/dry location.

b.	The consistency of each batch of slides must be validated by selecting  a sample of = 6 slides from each batch to be stained and read by different technicians to document consistency.  Some samples that are produced and tested will not be of sufficient consistency and should be discarded.

Apppendix C.1 Form 1: Validation Log for AFB Panel testing slide batches can be used to record results for the test slides and determine if consistency standard is acceptable.

Number of Slides made The laboratory should record how many slides were made from each sample to determine how many slides are available for test slide sets.   We recommend that laboratories prepare 50-100 slides so that sufficient slides are available to put duplicate samples (one stained and one unstained) in test slide sets.

Date slides made This is the date that the test slides were produced.   The length of time that slides can be stored without affecting performance has not been determined, but we estimate that 4-6 months is practical with proper storage.

Slide test results (columns 1-6) Each column represents the number AFB/100 fields for 6 separate slides selected for the sample and preferably read by 2-6 different technicians.  For high positives (2+ or 3+) the technicians may estimate the number AFB/100 fields by selecting a sufficient number of representative fields.   For low positives (exact count AFB/100 fields and 1+) and AFB negatives slides the technicians should read a minimum of 300 fields per slide and record the average number AFB/100 fields.

Average/Mean average is computed from slide test results 1-6 (see example)

Standard deviation The standard deviation is computed from slide test results 1-6 (see example).

Insert formula

Consistency The consistency column result is computed using the following formula:

Mean [M] minus 2 standard deviations [SD]

If M - 2 SD is > 0 then consistency is true (sufficient)

If M - 2 SD is < 0 then consistency is false (insufficient)

If the consistency is false—then there is too much variation in the number of AFB per slide and this sample is not of sufficient consistency to use in a PT test for a reliable evaluation of performance.  This formula provides an objective evaluation of consistency, but the laboratory should still review and determine what is acceptable variation within a sample of slides.

Report Result This is the slide test result for all the test slides.  This test result should be representative of the 6 slides tested and the sample should meet the consistency criteria.

8.  Prepare Panel Testing Sets

Sets of slides with identical composition of positives and negatives can be made from the prepared batches of slides.

Appendix C.2: Logbook of Test Slide Sets can be used to select slide sets and record the original batch numbers and expected results for a 10 slide panel testing exercise.  This form can also be used to record and evaluate the results from one or more peripheral laboratories that perform the PT test.   Form 2 serves as the official record of results for multiple slide sets that are sent to different laboratories.

## Alternate Procedure using NALC (ref Dr. Sang Jae Kim, South Korea)

a.	Collection of sputum specimens: sputum specimens with numerous AFB should be collected from the patients and be stored for not more than 2 days after collection in order to prevent destruction of sputum cells. Fresh AFB negative sputa also must be selected from the routine specimens.

b.	Preparation of mucolytic solution: 2% of N-acetyl-L-cysteine is mixed with an equal amount of 2.9% sodium citrate.2H2O right before use.

c.	Liquefaction of sputum specimen: AFB positive and negative sputum samples are mixed with an equal amount of mucolytic solution separately and shake gently to liquefy specimens.

d.	Dilution of AFB positive sputum homogenate: the liquefied AFB-positive sputum is diluted with varying proportions of AFB-negative specimen.

e.	AFB counts of sputum dilutions: one drop of each sputum dilution is spread on a slide with a smear size of 2 cm2  and dry and sterilize in a hot oven for one hour without scorching. 10 smears are prepared with each sputum dilution and stained with Ziehl-Neelsen staining method and count AFB per 1, 10 or 100 microscopic fields. Sputum dilutions whose average AFB counts fall into “1-9/100 fields”, “10-99/100 fields (1+)”, “1-10 per microscopic field (2+)” or “more than 10 per microscopic field (3+)” are selected and used to prepare as many smear slides as possible. Negative slides must also be prepared. AFB counts should be rechecked with randomly selected 10 to 15 slides again after completion of smear preparation in order to confirm AFB counts of every batch of slides.

C.2  Validation Log for AFB Panel Testing Slide Batches

| Validation Log for AFB Panel Testing Slide Batches |  |
| --- | --- |
|  | Slide Preparation | Slide evaluation |  |
| Batch No. | Number of slides made | Date Slides Made | Slide test results | Standard | Consistency | Report |
|  |  |  | (AFB per 100 fields) | deviation | (average minus 2 | result |
|  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | Average |  | stand deviations) |  |
|  |  |  | 10 | 10 | 50 | 15 | 10 | 11 | 17.7 | 16.0 | FALSE | 1+ |

C3: Logbook of Test Slide Sets

(Record of a set of 10 slides selected from Form 1)

Central Laboratory administering test :_________________________________________

Slide Set No. ______

Date slide set sent to peripheral laboratories: _____/____/_______

| Slide no. | Stained or unstained | Batch no. | Expected result | Peripheral Laboratory Results | Comments |
| --- | --- | --- | --- | --- | --- |
|  |  |  | (from Form 1.) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |  |
| Slide 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 7 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 9 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Slide 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

C.4: Panel Testing Recording and Feedback Form

Peripheral Laboratory:________________________________________________

Date PT received by your laboratory_________________________ (DD/MM/YY)

Date PT results returned to Central Laboratory _________________( DD/MM/YY)

Name of technician reading test smears: ___________________________________________

Note:  If more than one technician performs AFB microscopy in the laboratory, each technician should read all 10 smears and record their results on a separate form.  Technicians should not discuss results or share forms until all results have been sent back to the central laboratory.  Forms for all technicians should be sent to the central laboratory for evaluation.

|  |  | Reference Laboratory Only |
| --- | --- | --- |
| Slide Number | Result | Expected Result | Error Type | Points |

| Feedback |
| --- |
| Total Points: | Pass/Fail: |
| HFP | HFN | LFP | LFN | QE |
| Recommended Action: |

C.5:  Panel Testing Report of Multiple Laboratories for District Supervisor & NTP

District :

District Supervisor:

Test Slide Set:

Supervising Laboratory:

| Peripheral Lab | Annual Volume | SPR | PT Score | HFN | HFP | LFN | LFP | QE | Total Errors |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| District Averages |  |  |  |  |  |  |  |  |  |

#### SPR :slide positivity rate    PT: Panel Testing 	HFN: High False Negatives	HFP: High False Positives	LFN: Now False Negatives	LFP: Low False Positives	QE: Quantitation Errors

# Text boxes

Central Laboratory Use Only:

Test Slide set # _____________________________

Date Sent:_________________________________

Date results received:_________________________