## venting calculations
Revised March 5, 2008 by D. Olis with corrections provided by M. Cooper.
Another revision March 28, 2008, by Tom Peterson (rev's in red)
Method: calculate the pressure drop step-by-step through the vent line
adjust density for new pressure at each step
iterate to find vessel pressure which results in approx. 1 bar at vent exit
Adjusted mass flow rate helium | 0.8462 | kg/sec | Divided predicted flow by 0.9
| 6717.1343 | lbm/hr
Helium temperature | 8 | K | (assume constant through entire vent line)
Helium pressure in 1.7-inch pipe | 2 | bar
absolute viscosity | 0.002 | cP
d_1-684 | 1.684 | inch
Re_1-684 | 12584654.7705 | (for 1.684-inch diameter pipe)
f_1-684 | 0.02 | (for 1.684-inch diameter pipe)
d_1-76 | 1.76 | inch
Re_1-76 | 12041226.4963 | (for 1.76-inch diameter pipe)
f_1-76 | 0.02 | (for 1.76-inch diameter pipe)
Helium pressure in 2.16-inch pipe | 1.66 | bar
absolute viscosity | 0.002 | cP
d_2-16 | 2.16 | inch
Re_2-16 | 9811369.7377 | (for 2.16-inch diameter pipe)
f_2-16 | 0.019 | (for 2.16-inch diameter pipe)
d_2 | 2 | inch
Re_2 | 10596279.3167 | (for 2-inch diameter pipe)
f_2 | 0.019 | (for 2-inch diameter pipe)
Helium pressure in 3-inch pipe | 1.3402 | bar
absolute viscosity | 0.002 | cP
d_3-26 | 3.26 | inch
Re_3-26 | 6500784.8569 | (for 3-inch diameter pipe)
f_3-26 | 0.0175 | (for 3-inch diameter pipe)
Helium pressure in 4-inch pipe | 1.3172 | bar
absolute viscosity | 0.002 | cP
d_4-26 | 4 | inch
Re_4-26 | 5298139.6584 | (for 4-inch diameter pipe)
f_4-26 | 0.0161 | (for 4-inch diameter pipe)
Helium pressure in 6-inch pipe | 1.2985 | bar
absolute viscosity | 0.002 | cP
d_6-357 | 6 | inch
Re_6-357 | 3532093.1056 | (for 6-inch diameter pipe)
f_6-357 | 0.015 | (for 6-inch diameter pipe)
Beta is smaller dia / larger dia for a fitting
[9 empty] | Ref is reference dia (1.68) / actual or downstream dia |  |  |  | Assume T=8K at each section
| K | K x Ref^4 | L | d | fT | r | r/d | Beta | Ref | K_R | Total K for section | He pressure | He density | He spec vol | Press drop | Press drop
|  |  | (inch) | (inch) | [7 empty] | (bar) | (kg/m^3) | (ft^3/lbm) | (psi) | (bar)
Piping within HTC | [10 empty] | Vent line on vessel - entrance | 2 | 13.45 | 1.1909 | 1.568 | 0.10811
Entrance into vent pipe | 1 | 0.83814 | [6 empty] | 0.95682 |  | 0.83814
Vessel supply pipe | 0.034628 | 0.029023 | 3.0472 | 1.76 |  |  |  |  | 0.95682
Sudden contraction | 0.042249 | 0.042249 | [5 empty] | 0.95682 | 1
SR 90 deg pipe bend | 0.4 | 0.4 |  | 1.684 | 0.02 | 1.5 | 0.89074 |  | 1 |  | Vent line on vessel - elbows, run & contract | 1.8919 | 12.6524 | 1.266 | 3.3061 | 0.22795
Straight run | 0.041568 | 0.041568 | 3.5 | 1.684 |  |  |  |  | 1 |  | 1.6624
LR 90 deg bend | 0.28 | 0.28 |  | 1.684 | 0.02 | 2.25 | 1.3361 |  | 1 |  | Vent line on vessel - run to dewar | 1.6639 | 10.986 | 1.458 | 0.056571 | 0.0039004
Gradual enlargement | 0.084924 | 0.031375 | [5 empty] | 0.77963 | 0.77963
Vent line to LL dewar | 0.066852 | 0.024698 | 7.6 | 2.16 |  |  |  |  | 0.77963 |  | 0.024698
Thru LL dewar | [10 empty] | Vent line through LL dewar 1 | 1.66 | 10.9577 | 1.4618 | 0.34824 | 0.02401
straight run | 0.040463 | 0.014949 | 4.6 | 2.16 |  |  |  |  | 0.77963 |  | 0.15164
tee thru | 0.37 | 0.1367 |  | 2.16 | 0.0185 |  |  |  | 0.77963 |  | Vent line through LL dewar 2
2-90 deg bends | 0.57047 | 0.21076 |  | 2.16 | 0.0185 | 2.3 | 1.0648 |  | 0.77963 |  | 0.21856 | 1.636 | 10.7835 | 1.4854 | 0.51 | 0.035164
straight run | 0.021111 | 0.0077995 | 2.4 | 2.16 |  |  |  |  | 0.77963
Sudden contraction | 0.071331 | 0.035853 | [5 empty] | 0.92593 | 0.842 |  | Flexhose
Flexhose assy | 0.9766 | 0.49087 | 25.7 | 2 |  |  |  |  | 0.842 |  | 0.53695 | 1.6009 | 10.5288 | 1.5213 | 1.2833 | 0.088479
Sudden expansion | 0.027689 | 0.01023 | [5 empty] | 0.92593 | 0.77963
Pumping line weldment | [10 empty] | Components of pumping line weldment
helicoflex flange | 0.010556 | 0.0038997 | 1.2 | 2.16 |  |  |  |  | 0.77963 |  | 0.16107 | 1.5124 | 9.8905 | 1.6195 | 0.40979 | 0.028254
90 deg bend | 0.37 | 0.1367 |  | 2.16 | 0.0185 | 2 | 0.92593 |  | 0.77963 |  | 0.41009 | 1.4841 | 9.6874 | 1.6535 | 1.0652 | 0.073444
straight run | 0.055417 | 0.020474 | 6.3 | 2.16 |  |  |  |  | 0.77963
tee branch | 1.11 | 0.41009 |  | 2.16 | 0.0185 |  |  |  | 0.77963
Helium vent line branch | [10 empty] | Components of helium vent line branch
straight run | 0.11523 | 0.042572 | 13.1 | 2.16 |  |  |  |  | 0.77963 |  | 0.25158 | 1.4107 | 9.1614 | 1.7484 | 0.69099 | 0.047642
90 deg bend | 0.37 | 0.1367 |  | 2.16 | 0.0185 | 2 | 0.92593 |  | 0.77963 |  | 0.11627 | 1.363 | 8.8217 | 1.8158 | 0.33165 | 0.022867
straight run | 0.19572 | 0.072307 | 22.25 | 2.16 |  |  |  |  | 0.77963
Expansion 2 to 3-inch | 1.6329 | 0.11627 | [5 empty] | 0.66258 | 0.51656
3-inch diameter piping | [10 empty] | Cmpts of 3-in diameter piping
BS&B LPS burst disk | 0.79 | 0.078435 |  | 3 |  |  |  |  | 0.56133 | 0.79 | 0.078435 | 1.3402 | 8.659 | 1.8499 | 0.22793 | 0.015716
Straight run | 0.3865 | 0.02752 | 72 | 3.26 |  |  |  |  | 0.51656 |  | 0.02752 | 1.3245 | 8.5474 | 1.874 | 0.081018 | 0.005586
Expansion 3 to 4-inch | 0.25554 | 0.0080278 | [5 empty] | 0.815 | 0.421 |  | 0.0080278 | 1.3189 | 8.5077 | 1.8828 | 0.023744 | 0.0016371
4-inch diameter piping | [10 empty] | Discharge line to atmosphere
Straight runs total length | 0.8694 | 0.027312 | 216 | 4 |  |  |  |  | 0.421 |  | 0.027312 | 1.3172 | 8.4961 | 1.8853 | 0.080889 | 0.0055772
90 deg elbow | 0.483 | 0.015173 |  | 4 |  |  |  |  | 0.421 |  | 0.015173 | 1.3117 | 8.4566 | 1.8941 | 0.045149 | 0.0031129
90 deg elbow | 0.483 | 0.015173 |  | 4 |  |  |  |  | 0.421 |  | 0.015173 | 1.3085 | 8.4345 | 1.8991 | 0.045267 | 0.0031211
90 deg elbow | 0.483 | 0.015173 |  | 4 |  |  |  |  | 0.421 |  | 0.015173 | 1.3054 | 8.4124 | 1.9041 | 0.045386 | 0.0031293
90 deg elbow | 0.483 | 0.015173 |  | 4 |  |  |  |  | 0.421 |  | 0.015173 | 1.3023 | 8.3902 | 1.9091 | 0.045506 | 0.0031376
Expansion 4 to 6-inch | 1.5625 | 0.0029972 | [5 empty] | 0.66667 | 0.28067 |  | 0.0029972 | 1.2992 | 8.3679 | 1.9142 | 0.0090127 | 0.00062141
6-inch diameter piping
Straight run | 0.24 | 0.0014893 | 96 | 6 |  |  |  |  | 0.28067 |  | 0.0014893 | 1.2985 | 8.3635 | 1.9152 | 0.0044807 | 0.00030894
Exit | 1 | 0.0062053 |  | 6 |  |  |  |  | 0.28067 |  | 0.0062053 | 1.2982 | 8.3613 | 1.9157 | 0.018675 | 0.0012876
Total K |  | 3.7459 | 473.6972 | [7 empty] | Final calculated P at discharge (bar) | 1.2969 |  | Total dP | 10.1969 | 0.70306
|  |  | 1203.191 | [10 empty] | Subtraction arithmetic check |  | 0.70306
Notes (Tom P.)
In cell B52 fixed one more inverted use of Beta
Multiplied K for flex hose by 4 to account for bellows convolutions.
The following all look over-estimated, but I left them as is in the spreadsheet.
3-inch piping above burst disk looks like 14 inches long total.
4-inch stainless and PVC vent line looks like 184 inches long with three (not four) 90 degree elbows.
The last elbow is actually in the 6-inch PVC, which does total about 96 inches long.

## L` table
Determine the range of L' for relief pressure 12-psig
Based on _Technology of Liquid Helium_ Monograph 111
3.9-GHz helium vessel surface area |  |  |  | 785 | in^2 |  | 3.9-GHz helium vessel surface area |  |  |  | 148 | in^2
|  |  |  | 5064.506 | cm^2 | [5 empty] | 954.8368 | cm^2
Heat flux for loss of insul vacuum |  |  |  | 2 | W/cm^2 |  | Heat flux for loss of insul vacuum |  |  |  | 4 | W/cm^2
|  | T1 |  | v1 | h1 | m1 | Tx |  | vx | hx | mx | 0.5*(m1+mx) | (m1-mx) | (hx-h1) | L` | L' | Mass flow rate | Mass flow rate
Relief pressure |  | TEMP | DENSITY | Sp volume | ENTHALPY | Mass | TEMP | DENSITY | Sp volume | ENTHALPY | [6 empty] | for helium vessel | for RF cavity
psia | bar | [R] | [Lb/ft3] | [ft^3/lbm] | [BTU/Lb] | [lbm] | [R] | [Lb/ft3] | [ft^3/lbm] | [BTU/Lb] | [lbm] | [lbm] | [lbm] | [BTU/Lb] | [BTU/Lb] | [J/g] | [g/sec] | [g/sec]
26.7 | 1.8414 | 6 | 8.878 | 0.11264 | -3.775 | 1 | 7 | 8.443 | 0.11844 | -2.986 | 0.951 | 0.9755 | 0.048998 | 0.789 | 15.7084 | 36.5282 | 277.2933 | 104.559
|  | 7 | 8.443 | 0.11844 | -2.986 | 1 | 8 | 7.782 | 0.1285 | -1.924 | 0.92171 | 0.96086 | 0.07829 | 1.062 | 13.034 | 30.3092 | 334.1891 | 126.0127
|  | 8 | 7.782 | 0.1285 | -1.924 | 1 | 9 | 1.971 | 0.50736 | 6.204 | 0.25328 | 0.62664 | 0.74672 | 8.128 | 6.8209 | 15.8613 | 638.6001 | 240.797
|  | 9 | 1.971 | 0.50736 | 6.204 | 1 | 10 | 1.391 | 0.71891 | 8.756 | 0.70573 | 0.85287 | 0.29427 | 2.552 | 7.3964 | 17.1996 | 588.9109 | 222.0607
|  | 10 | 1.391 | 0.71891 | 8.756 | 1 | 11 | 1.145 | 0.87336 | 10.6 | 0.82315 | 0.91157 | 0.17685 | 1.844 | 9.5048 | 22.1025 | 458.2737 | 172.8013
|  | 11 | 1.145 | 0.87336 | 10.6 | 1 | 12 | 0.9903 | 1.0098 | 12.23 | 0.86489 | 0.93245 | 0.13511 | 1.63 | 11.2493 | 26.1591 | 387.2075 | 146.0044
|  | 12 | 0.9903 | 1.0098 | 12.23 | 1 | 13 | 0.8793 | 1.1373 | 13.76 | 0.88791 | 0.94396 | 0.11209 | 1.53 | 12.8851 | 29.9629 | 338.0515 | 127.4691
|  | 13 | 0.8793 | 1.1373 | 13.76 | 1 | 14 | 0.7942 | 1.2591 | 15.23 | 0.90322 | 0.95161 | 0.096782 | 1.47 | 14.4538 | 33.6109 | 301.3606 | 113.6341
|  | 14 | 0.7942 | 1.2591 | 15.23 | 1 | 15 | 0.7261 | 1.3772 | 16.65 | 0.91425 | 0.95713 | 0.085747 | 1.42 | 15.8504 | 36.8585 | 274.8081 | 103.6219
|  | 15 | 0.7261 | 1.3772 | 16.65 | 1
50 |  | 6 | 9.112 | 0.10975 | -3.4012 | 1 | 8 | 8.264 | 0.12101 | -1.7582 | 0.90694 | 0.95347 | 0.093064 | 1.643 | 16.833
|  | 8 | 8.264 | 0.12101 | -1.7582 | 1 | 10 | 6.108 | 0.16372 | 1.563 | 0.73911 | 0.86955 | 0.26089 | 3.3212 | 11.0696
|  | 10 | 6.108 | 0.16372 | 1.563 | 1 | 12 | 2.354 | 0.42481 | 9.33 | 0.3854 | 0.6927 | 0.6146 | 7.767 | 8.7539
|  | 12 | 2.354 | 0.42481 | 9.33 | 1 | 14 | 1.672 | 0.59809 | 13.3 | 0.71028 | 0.85514 | 0.28972 | 3.97 | 11.7179
|  | 14 | 1.672 | 0.59809 | 13.3 | 1 | 16 | 1.346 | 0.74294 | 16.56 | 0.80502 | 0.90251 | 0.19498 | 3.26 | 15.09
|  | 16 | 1.346 | 0.74294 | 16.56 | 1 | 18 | 1.142 | 0.87566 | 19.56 | 0.84844 | 0.92422 | 0.15156 | 3 | 18.2941
|  | 18 | 1.142 | 0.87566 | 19.56 | 1
Effective
|  |  | latent heat | Heat input | Mass flow
|  |  | (J/g) | (W) | (g/sec)
| 2.2 |  | 13 | 10129.012 | 779.1548
| 2 |  | 13.3 | 10129.012 | 761.5798 |  | Note (TomP.): I leave the above table and associated chart for reference but use this mass flow for 2.0 bar
| 1.8 |  | 13.7 | 10129.012 | 739.3439