# STANDARD OPERATING PROCEDURE—6013

# PROCEDURES FOR THE HANDLING AND REPROCESSING OF ITEMS USED ON A PATIENT WITH CREUTZFELDT-JAKOB DISEASE

PURPOSE: To outline procedures for the cleaning and sterilization of reusable equipment, instrumentation, and supplies contaminated with infective material from a patient who is known or suspected to have Creutzfeldt-Jakob Disease (CJD).

POLICY:  In as much as possible, disposable instruments, drapes, and supplies must be used.  For those items that can be sterilized and reused, the following procedures will be implemented.

DEFINITIONS:

- Medium or low-risk infectious material: cerebrospinal fluid (CSP), kidney, liver,

spleen, lung, and lymph node tissues of patients with known or suspected diagnosis

of CJD.

- High-risk infectious material: brain, spinal cord, and eye tissues of patients with known or suspected diagnosis of CJD.

PROCEDURES:

- Initial decontamination requires steam sterilization at 134 C (272 F) for 18 minutes exposure in a prevacuum sterilizer or 121 C (250 F) for one-hour exposure time in a gravity displacement sterilizer/cycle.

- Items are then to be put through routine decontamination and sterilization processes.

- Items that cannot be reprocessed are to be bagged and disposed of in biohazard trash that is to be incinerated.

- Counter tops and sinks that are contaminated with high-risk infectious material should be cleaned with 1:10 dilution of sodium hypochlorite (bleach).

- Regulated medical waste (i.e., bulk blood, pathological waste, microbiological waste, and sharps) will be disposed of in accordance with local, state, and federal regulations.

- Laundry will be managed as required by the Occupational Safety and Health Administration (OSHA) rule on blood borne pathogens.  No additional precautions are required.

REFERENCES:

American Society of Hospital Central Service Professionals. (2002).  Guidance document for the care and handling of surgical instruments and environmental recommendations for known or suspected cases of Creutzfeldt-Jacob Disease.  Retrieved January 9, 2003 from http://www.ashcep.org.

Association of Preoperative Nurses. (2003).  Standards, recommended practices, and guidelines.  Denver: AORN.

Centers for Disease Control and Prevention. (2001).  Questions and answers regarding Creutzfeldt-Jacob Disease infection-control practices.  Retrieved from cdc.gov/ncidod/diseases/cjd/cjd_inf_ctrl_qa.htm

Roos, R. (2001).  Infection control and CJD: Sterilization measurers for a tough, lethal agent.

SPD Advisory Group, Best Practices

10/25/03