# Magnetic Measurements during Production of VLHC Low Field Combined Function Magnets

For this study we focus discussion on magnetic field measurements of the combined function magnets during production. No doubt measurements of other magnetic elements of the ring are necessary. One would also expect to perform an extensive series of measurements, possibly using several measurement techniques, during R&D.  Here, however, we have chosen to present one concept for measurements of production magnets which utilizes a rotating probe of specific design, try to define the most important issues related to this concept, and examine the consequences. Thereafter, a brief discussion of other approaches.

Main parameters of the combined function magnets are 2 T field,  (10% horizontal gradient, 65 m length, and  20 mm gap. Measurement of the integrated dipole field is needed with an accuracy of a few parts in 104. Desired measurement accuracy for the higher order harmonics is 0.1 units. This implies the horizontal gradient is known relative to the main field to a part in 105; other harmonics (of order 1 unit), to 10%. We begin by examining the implications of requirements for field strength measurements. Based on this discussion we present a concept for a measurement system. Given this system we then address the issue of measurement of field harmonics.

With respect to measurement of the integrated field strength, there are two issues. Since the accuracy of the strength is given by

where R is the radius of the probe. This requires that the probe radius (actually the radius of the coil winding) be known with high accuracy. Since manufacturing and mechanical measurement accuracy are fixed, (R is fixed. As the radius of the probe becomes smaller, these tolerances become an issue. The following table illustrates the problem.

Typical manufacturing and measurement accuracy achieved for magnetic measurements is 1 mil (0.001in, 0.025 m. For a probe of 10 mm maximum radius, this is inadequate. Calibration of the probe with at least an order of magnitude greater precision (3e-4 strength error) is thus required. This level of precision can only be achieved using a calibration magnet in which the strength is known at least as well as one desires to know it in the subject magnets. The second issue is unique to measurement of combined function magnets. Due to the horizontal field gradient, an error in the horizontal position of the probe translates into an error in measured field strength. In a magnet with 10% gradient, measurement of the strength to 1 part in 104 requires positioning the probe with 1 mm accuracy (see table).

Although it may be difficult, this can be achieved using mechanical means. Vertical placement of the probe is not as critical as long as the good field region is significantly larger than the probe.

The main conclusions from the preceding discussion are that (1) a rotating coil of the size we are considering can be calibrated with sufficient precision, and (2) that the horizontal position of the probe can be set with mechanical fixtures. For reasons we list below, we choose to insert the rotating probe into the aperture from the side rather than from the end. The probe is supported on a frame attached to the magnet as shown in Fig. 1-3. Principle elements of the design are the following.

- The probe body is clamped to a precisely machined strong-back on an edge which describes the magnet sagitta.

- The strong-back is slides horizontally in a series of jaws attached to the vertical member of the frame. The strong-back is spring loaded so that it and the probe are forced against the referenced surface on the inner edge of the magnet aperture.

- The vertical position of the probe is defined by the precision with which the frame is attached to the magnet and the precision of the jaws relative to the vertical member of the frame and the horizontal surfaces of the strong-back.

- The motor for rotation and encoder for probe angle readout are mounted on the frame and connected to the probe by a mechanism such as a toothed belt or right angle drive. Other drive components would be similarly mounted (e.g. gearboxes). A tilt sensor is mounted on the strong back to allow correction of the probe angle.

- Development of slip rings small enough to fit the probe bore is assumed although one might choose to digitize and transmit the signals from inside the bore. A pre-amplifier circuit sized to fit the bore is also assumed allowing a boost in signal size of up to 1000.

Figure 1:  Measurement Fixture

Figure 2:  Probe and Strong-Back

Figure 3:  Probe Assembly

A consequence of insertion from the side is that one loses a few extra mm of probe radius due to the reduction in gap at the magnet edge of the focusing quad. The following allocation of real estate is made in arriving at a more or less realistic probe radius.

| horizontal aperture at beam line | 20 mm |
| --- | --- |
| horizontal aperture at gap edge | 18 mm |
| clearance | 1 mm |
| lost to clamp clamping probe to strong-back | 2 mm |
| probe diameter | 15 mm |

Given a probe radius, we need to examine the accuracy of the measurement of field harmonics. From [1] we see that probe calibration matters little in measuring harmonics. Accuracy of 10% or so is adequate. (I.e. If the harmonic is 1 unit; you don’t care if it is 0.9 or 1.1.) This is probably true even of the gradient although it needs to be considered. What is of greater interest is the evaluation of the interplay of signal size with harmonic measurement. This has implications for the required level of excitation current during test. Field measurements of model magnets made during the R&D stage of the LHC IR quadrupole programs indicate that that for harmonics above n=13 or 14, uncertainty from measurements dominate the random errors [2]. The same measurements also show good agreement between the measured and design values of the allowed harmonic b14 (-0.01 unit).  From these observations one can infer a specific signal size which is adequate but below which measurements are difficult. We use this level (12 uV) to calculate a required excitation level for measurement of the transmission line magnets. We use the probe parameters in the table below.

| probe radius | 7.5 mm |
| --- | --- |
| length | 6.5 m |
| number of turns | 1 |
| angular rotation period | 3 s |
| tangential winding opening angle | 13( |
| transfer function | 2T/100kA |

Sensitivity as a function of harmonic number is shown in Fig. 4. The sensitivity versus harmonic number is a function of the opening angle and the ratio of probe to reference radius. In particular, if Rref > Rcoil, the sensitivity falls with increasing harmonic number. For Rref=18 mm, the sensitivity is down by a factor of 10 through n=11. One may assume pre-amplifiers with a reasonable gain may be inserted in the probe bore. With a gain of 100, measurement accuracy of 0.1 unit through n=11 requires 200 A excitation current. This should be possible for room temperature measurements.

Figure 4 Sensitivity

A probe length of  6.5 m was assumed above.  Ten measurements are then required to construct the integral field. The time required for measurement of a production magnet is estimated in the following table. Two crews of 3 technicians are assumed; each works one aperture. Four sets of measurement apparatus are assumed (2/side) so 1 can be set up while the other measures.

| measurement setup | 15 min. |
| --- | --- |
| data acquisition | 100 rotations x 3 s./rotation=5 min. < 15 min. for setup |
| time/location | 15 min. + 5 at the last station |
| locations/aperture | 10 |
| #probes/aperture | 2 |
| total/aperture | 155 min./aperture |
| total | 155 (2 hr. 35 min.) |

This certainly allows for the required production throughput. Apparatus and procedures would be optimized to minimize this time during production.

We have chosen a probe inserted in the magnet bore from the side and supported by an external frame. An alternative approach would be a "mole", such as have been developed for measurements of SSC [3] and LHC magnets [4,5] which are probes with drive and angular alignment systems built in. The LHC apparatus includes both warm and cold moles for a magnet bore of 70 mm. There are some advantages to this approach. There is no external frame. The probe is inserted at one end and pulled through the magnet. The probe can presumably be larger as it doesn't have to be inserted through the narrow gap at the outer edge of the bore. As was discussed above, sensitivity increases with probe radius (depending on the reference radius). The effects of imperfections in manufacturing or motion are also damped by some factor scaling with radius. Development of the necessary apparatus of the size necessary for a mole that can operate in a strong magnetic field is not trivial. The LHC apparatus would have to be scaled down by a factor of 3.5 or more. We have no doubt that it could be done. However, we have chosen to remove the active elements of the measurement system from the bore. The only pieces in the bore are slip rings and pre-amplifiers. This seems like a less daunting task.

# References

[1] P. Schlabach, "Effect of Probe Manufacturing Errors on Field Measurements", TD-98-060, 1998.

[2] P. Schlabach, " Field Errors of HGG Model Magnets," TD-00-052, 2000.

[3] G. Ganetis., et. al., "Field Measuring Probe for SSC Magnets", Particle Accelerator Conference, 1987.

[4] L. Bottura, et. al., "A Mole for Warm Magnetic and Optical Measurements of LHC Dipoles", Eleventh International Magnet Measurement Workshop, 1999.

[5] L. Walkiers, "Measurements for the Acceptance Tests of the LHC Superconducting Magnets', Eleventh International Magnet Measurement Workshop, 1999.

# Footnotes

For this exercise, we specify a probe has windings of dipole and quadrupole symmetry sensitive to n=1, 3, 5, ( and n=2, 6, 10, 14( respectively as well as a tangential winding with sensitivity to all harmonics. The opening angle of the tangential winding is 13(. This also allows for bucking the harmonic signal.

One unit is 104 ( the harmonic normalized to the main field.

Additional details can be found in [1]. We also remark that what is said here about probe radius is also true of probe length. Length is typically much greater than radius so a fixed manufacturing or measurement error matters little.

Note that this problem could be separated into two separate requirements:  (1) absolute accuracy of a single probe and (2) relative accuracy between a set of probes that might be used in production measurements. The former is an error relative to overall design energy; the latter, a more or less random error magnet to magnet or, more likely, magnet section to magnet section. We choose not to differentiate the two.

The goal of this exercise is only to demonstrate that the required time is reasonable.