### 5.1.7	Magnetic Measurements

For this study we focus discussion on magnetic field measurements of the combined function magnets during production. No doubt measurements of other magnetic elements of the ring are necessary. One would also expect to perform an extensive series of measurements, possibly using several measurement techniques, during R&D.  Here, however, we have chosen to present one concept for measurements of production magnets which utilizes a rotating probe of specific design, try to define the most important issues related to this concept, and examine the consequences. Thereafter, a brief discussion of other approaches.

Main parameters of the arc dipole magnets are ~2 T field,  (12% horizontal gradient, 65 m length, and  20 mm gap. Measurement of the integrated dipole field is needed with an accuracy of a few parts in 104. Desired measurement accuracy for the higher order harmonics is 0.1 units. This implies the horizontal gradient is known relative to the main field to a part in 105; other harmonics (of order 1 unit), to 10%. We begin by examining the implications of requirements for field strength measurements. Based on this discussion we present a concept for a measurement system. Given this system we then address the issue of measurement of field harmonics.

With respect to measurement of the integrated field strength, there are two issues. The accuracy of the measured field strength is given by  where R is the radius of the probe [1]. This requires that the probe radius (actually the radius of the coil winding) be known with high accuracy. Since manufacturing and mechanical measurement accuracy are fixed, (R is fixed. Typical manufacturing and measurement uncertainty for magnetic measurement apparatus is ~0.001 inch (25 m) which translates into a field error of 3e-3 for a probe of 10 mm maximum radius. Calibration of the probe with at least an order of magnitude greater precision (3e-4 strength error) is thus required. This level of precision can only be achieved using a calibration magnet in which the strength is known at least as well as one desires to know it in the subject magnets. The second issue is unique to measurement of combined function magnets. Due to the horizontal field gradient, an error in the horizontal position of the probe translates into an error in measured field strength. In a magnet with 10% gradient, measurement of the strength to 1 part in 104 requires positioning the probe with 0.1 mm accuracy. Although it may be difficult, this can be achieved using mechanical means. Vertical placement of the probe is not as critical as long as the good field region is significantly larger than the probe.

The main conclusions from the preceding discussion are that (1) a rotating coil of the size we are considering can be calibrated with sufficient precision, and (2) that the horizontal position of the probe can be set with mechanical fixtures. For reasons we list below, we choose to insert the rotating probe into the aperture from the side rather than from the end. The probe is supported on a frame attached to the magnet as shown in Fig. 1. Principle elements of the design are the following:

The probe body is clamped to a precisely machined strong-back on an edge which describes the magnet sagitta.

The strong-back slides horizontally in a series of jaws attached to the vertical member of the frame. The strong-back is spring loaded so that it and the probe are forced against the referenced surface at the edge of the magnet gap.

The vertical position of the probe is defined by the precision with which the frame is attached to the magnet and the precision of the jaws relative to the vertical member of the frame and the horizontal surfaces of the strong-back. A tilt sensor is mounted on the strong-back to monitor its alignment.

The motor for rotation and encoder for probe angle readout are mounted on the frame and connected to the probe by a mechanism such as a toothed belt or right angle drive. Other drive components would be similarly mounted (e.g. gearboxes).

Development of slip rings small enough to fit the probe bore is assumed although one might choose to digitize and transmit the signals from inside the bore. A pre-amplifier circuit sized to fit the bore is also assumed allowing a boost in signal size.

Figure 1:  Measurement Fixture

A consequence of insertion from the side is that one loses a few extra mm of probe radius due to the reduction in gap at the magnet edge of the focusing magnet. One also needs to allow space for the clamps holding the probe to the strong-back and for clearance during insertion.

Given a probe radius, we need to examine the accuracy of the measurement of field harmonics. From [1] we see that probe calibration matters little in measuring harmonics. Accuracy of 10% or so is adequate. (i.e., If the harmonic is 1 unit; you don’t care if it is 0.9 or 1.1.) This requires knowing the probe radius to 10e-3. This is probably true even of the gradient although it needs to be considered. What is of greater interest is the evaluation of the interplay of signal size with harmonic measurement. This has implications for the required level of excitation current during test. For a reasonable set of probe parameters, pre-amplifier gain of 100, measurement accuracy of 0.1 unit through n=9 requires at least 200 A excitation current [3]. This should be possible for room temperature measurements.

A probe length of  6.5 m was assumed above.  Ten measurements are then required to construct the integral field. The time required for measurement is estimated as 2.5 hours with 2 crews of 3 technicians using 4 sets of measurement apparatus [2]. This certainly allows for the required production throughput. Automation and optimization of procedures could reduce the time required. We also note that the apparatus described could be used for either warm or cold measurements. The only special consideration for cold measurements would be to construct the strong-back of a nonmetallic material to prevent the formation of eddy currents during ramping.

We have chosen a probe inserted in the magnet bore from the side and supported by an external frame. An alternative approach would be a "mole", such as have been developed for measurements of SSC [3] and LHC magnets [4,5] which are probes with drive and angular alignment systems built in. The LHC apparatus includes both warm and cold moles for a magnet bore of 70 mm. There are advantages to this approach. There is no external frame. The probe can be larger by as much as 25% as it is inserted in the magnet from the end instead of from the side. As was obvious from the discussion above, a larger radius is better although an increase of 25% doesn't help all that much. However, development of apparatus of the size necessary for a mole that can operate in a strong magnetic field is not trivial. The LHC apparatus would have to be scaled down by a factor of 3.5 or more. While we have no doubt that this could be done, we have chosen an approach that removes elements of the measurement system other than slip rings and pre-amplifiers from the bore. This seems like a less daunting task.

References

[1] P. Schlabach, "Effect of Probe Manufacturing Errors on Field Measurements", TD-98-060, 1998.

[2] P. Schlabach, "Magnetic Measurements during Production of VLHC Low Field Combined Function Magnets" TD-01-007, 2001.

[3] G. Ganetis, et al., "Field Measuring Probe for SSC Magnets", Particle Accelerator Conference, 1987.

[4] L. Bottura, et. al., "A Mole for Warm Magnetic and Optical Measurements of LHC Dipoles", Eleventh International Magnet Measurement Workshop, 1999.

[5] L. Walkiers, "Measurements for the Acceptance Tests of the LHC Superconducting Magnets', Eleventh International Magnet Measurement Workshop, 1999.

# Footnotes

current file: Q:\VLHC\TDDoc\DesignStudyReport\05_Stage1_components\Draft2.1\5.1.7_Magnetic Measurements_ed010328.doc

For this exercise, we specify a probe has windings of dipole and quadrupole symmetry sensitive to n=1, 3, 5, ( and n=2, 6, 10,( respectively as well as a tangential winding with sensitivity to all harmonics. This allows for bucking the harmonic signal.

One unit is 104 ( the harmonic normalized to the main field.