|  | Pressure Vessel Engineering Note (short) For the 3.9-GHz Helium Vessel, Cavity #8 | Doc. No. Rev. No. 0 Date: 13 Oct 2008 Page 1 of 28 |
| --- | --- | --- |

FERMILAB

Technical Division

Pressure Vessel Engineering Note

For the

3.9-GHz Helium Vessel, Cavity #8

(short version referencing cavity #5)

| Authors: M. Wong, G. Galasso, Harry Carter, Tom Peterson, Phil Pfund, Dan Olis, Tug Arkan | Date: 13 Oct 2008 |
| --- | --- |

Table of Contents

Section											Page Number

General introduction – exceptional vessel						Pg 3

Pressure vessel engineering note form (FESHM 5031) 				Pg 4

Appendix A – Extended Engineering Note

Introduction 									Pg 8

Exceptional vessel discussion and note summary				Pg 9

Vessel description								Pg 10

Analysis									Pg 14

Welding information								Pg 16

Appendix B – 2.3 bar Pressure Test Results						Pg 21

Appendix C – Cavity Processing Information					Pg 24

General introduction – exceptional vessel

This engineering note for dressed cavity #8 references the note for cavity #5 where the information is almost identical.  For the same reasons as Cavity #5, Cavity #8 is an exceptional vessel.  Starting with the Extended Engineering Note on page 8, only the differences from cavity #5 are detailed here.  Specifically, the highlighted test details the differences.  The non-highlighted next is identical to its counterpart in the Cavity #5 note.

Amendment No.:                 Reviewed by:              Date:

_________________              _____________             _______________

_________________              ______________            ________________

Lab Property Number(s):

Lab Location Code:  Meson Detector Building	(obtain from safety officer)

Purpose of Vessel(s):  LHe containment for nine-cell 3.9-GHz superconducting

Radio frequency cavity

Vessel Capacity/Size:  3.4-L  Diameter:  5.5”(140mm) Length:  19.9”(505.92mm)

Normal Operating Pressure (OP)  0.02-bar (0.25-psia)

MAWP-OP =  28.75 PSID

List the numbers of all pertinent drawings and the location of the originals.

Drawing #	Location of Original

5525-ME-440598 Rev B 		 IDEAS FERMI TDM

5520-ME-426450 Rev F		 IDEAS FERMI TDM

5520-ME-426321 Rev B		 IDEAS FERMI TDM

5520-MB-426252 Rev C		 IDEAS FERMI TDM

(Above are the top-level drawings, see Appendix B, “drawing tree”, in the design note)

2.	Design Verification

Is this vessel designed and built to meet the Code or “In-House Built” requirements?

Yes     No  X

# If “No” state the standard that was used ASME Code, Section VIII, Div 1 and Div 2 with finite element analysis and reference to published mechanical data for niobium and NbTi (where code data were not available)

Demonstrate that design calculations of that standard have been made and that other requirements of that standard have been satisfied.

Skip to part 3 “system venting verification.”

Does the vessel(s) have a U stamp?  Yes_____ No  X .  If "Yes", complete section 2A; if "No", complete section 2B.

A.	Staple photo of U stamp plate below.

Copy "U" label details to the side

Copy data here:

Figure 1.  ASME Code:  Applicable Sections

2B.

Summary of ASME Code

CALCULATION RESULT

(Required thickness or stress

Reference ASME	level vs. actual thickness

Item	Code Section	calculated stress level)

(see Design Note)

vs

vs

vs

vs

3.	System Venting Verification  (See Figure A4 for vent system schematic

- drawing 4906.320-ME-440302)

Does the venting system follow the Code UG-125 through UG-137?

Yes  X  No___

Does the venting system also follow the Compressed Gas Association Standards S-1.1 and S-1.3?

Yes  X 	No_____

A “no” response to both of the two proceeding questions requires a justification and statement regarding what standards were applied to verify system venting is adequate.

List of reliefs and settings:

Manufacturer	Model #  Set Pressure	 Flow Rate	Size

BS&B		 LPS		 12-psig	 	 2188-SCFM air	 3-inch

Hylok 700 	CV5-F12N-25				Cv=5.2		3/4"

4.	Operating Procedure

Is an operating procedure necessary for the safe operation of this vessel?

Yes_____  No  X  (If "Yes", it must be appended)

5.	Welding Information

Has the vessel been fabricated in a non-code shop?  Yes  X  No

If "Yes", append a copy of the welding shop statement of welder qualification (Procedure Qualification Record, PQR) which references the Welding Procedure Specification (WPS) used to weld this vessel.

See “Welding Information” Section in Appendix A

6.	Existing, Used and Unmanned Area Vessels

Is this vessel or any part thereof in the above categories?

Yes_____ No  X

If "Yes", follow the requirements for an Extended Engineering Note for Existing, Used and Unmanned Area Vessels.

7.	Exceptional Vessels

Is this vessel or any part thereof in the above category?

Yes  X  No

If "Yes", follow the requirements for an Extended Engineering Note for Exceptional Vessels.

Appendix A – Extended Engineering Note

Introduction

The 3.9 GHz “dressed cavity”, cavity #8, is a niobium superconducting radio frequency (RF) cavity surrounded by a helium vessel which will be filled with liquid helium (so liquid helium surrounds the niobium RF cavity) at temperatures as low as 1.8 K.  The cavity will be performance tested in the Horizontal Test Stand (HTS) at the Meson Test Area at Fermi National Accelerator Laboratory (Fermilab) before being installed in a cryostat with three other 3.9 GHz dressed cavities.  The cryostat with four dressed cavities will be shipped to DESY, in Hamburg, Germany, to be installed in the FLASH-linac (former TTF2-linac).  This document summarizes how the helium vessel for cavity #8 follows the requirements of the FESHM 5031 [ref 1].  This document and supporting documents for the 3.9 GHz helium vessel engineering note and also for the 3.9 GHz cryomodule may be found at:

http://ilc-dms.fnal.gov/Workgroups/CryomoduleDocumentation/3rd-Harmonic-Cryomodule-for-DESY/

Cavity #8 is the third dressed cavity that will be tested in HTS.  The designs of the dressed cavity are almost identical to Cavity #5.  The only difference is in one part of the helium vessel design.  The difference is detailed in the “Vessel Description” of Appendix A of this engineering note.  The fabrication of Cavity #8 follows almost the exact the same procedure as the fabrication for Cavity #5.  Several steps in the welding procedure differ in an effort to optimize the fabrication process.  How the fabrication for Cavity #8 differs from Cavity #5 is detailed in the section titled “Welding Information” of Appendix A of this engineering note.

…

Exceptional Vessel Discussion and Extended Engineering Note Summary

Reason for Exception

The 3.9-GHz dressed cavity as a helium vessel has materials and complex geometry that are not conducive to complete design and fabrication following the ASME Code [ref 2].  Parts of the vessel are made of niobium and an alloy that is 45% niobium/55% titanium.  Neither of these materials is approved in the Code.

However, we show that the vessel is safe at a level that is similar to the Code, as recommended by the DOE’s Federal Register for contractors on Worker Safety and Health Program [ref 3] and in accordance with FESHM 5031. Since the vessel design and fabrication cannot exactly follow the guidelines given by the Code, the vessel requires a Director’s Exception.  (The page numbers in Table A1 refer to the pages in the Cavity #5 engineering note.)

Table A1 – Specific areas of exception to the ASME code

| Location or procedure | Reference | Explanation for exception |
| --- | --- | --- |
| Niobium material | pg 17 | Used for its superconducting properties, not explicitly allowed by the code |
| Calculated stress in Nb cavity at the iris. | pg 41 | Worst case tuning extension (0.5 mm) with internal pressure at room temperature (82.4 MPa vs 60 MPa allowable). This stress is self-limited by the tuner and results at worst in an offset of the neutral tuning position. |
| Niobium-Titanium material | pg 19 | Used for compatibility in welding with niobium and titanium, a transition between the two metals |
| Partial penetration welds in Ti and between NbTi and Ti | pp 79 - 85 | Care not to burn through the material |
| Weld procedure specification (WPS), procedure qualification record (PQR), and welder performance qualification (WPQ) were either not performed, or initiated after completion of the work | pp 74 | The 0.8 factor taken on allowable material stress per FESHM 5031 is intended exactly for this issue, to mitigate incomplete code-specified information. |
| Incomplete testing of bulk and weld affected zone material | pp 74 | See note above regarding 0.8 factor. |
| No X-ray or ultrasonic inspections of welds (Div 2 requires 100% radiography, plus 10% ultrasonic testing on welds.) | pg 74 | Access difficulty. Samples were welded, cut, polished, and photographed instead in order to establish degree of weld penetration. |

…

Fabrication

Detailed welding information is included in this engineering note.  A fabrication procedure, associated material specifications, and welders’ certifications are compiled in a second document with the file name “Cavity-8-Matl_Weld_Certs.doc”.

...

Pressure Test

The dressed cavity was pressure tested to 2.3 bar, 1.15 times the MAWP of 2.0 bar.

...

Vessel Description

…

A complete drawing tree is shown in Figure A1.  The top level and second level assembly drawings are copied in the following figures for reference here.  A complete set of drawings is at:

http://ilc-dms.fnal.gov/Workgroups/CryomoduleDocumentation/3rd-Harmonic-Cryomodule-for-DESY/DressedCavityDrawings/  and is also archived at:

http://tdserver1/project/TDarchivedDrawings/drawings/

The design of Cavity #8 is almost the same as for Cavity #5.  The drawings shown in Figure A1 have been updated to reflect the changes.  The only design change for Cavity #8 is the thickness of the bellows cuff design (drawing MB-457115) at the bellows-to-shell joint in the helium vessel.  The thickness of the cuff is 2-mm.  The cuff design allows for a full penetration TIG weld of the bellows assembly to the helium vessel shell.  The joint detail is shown in Figure A29 (weld #7, #8).

…

Figure A1.  Dressed cavity #8 drawing tree

Figure A2 – Helium Vessel for the 3.9-GHz Cavity, assembled in the welding fixture (Drawing 5520.000-ME-440598)

…

Figure A4 – Helium Vessel Weldment for 3.9-GHz Dressed Cavity (Drawing 5520.000-ME-426450)

…

Welds Simulation

This section is identical to that for cavity #5 but is included here since the welds # 7 and 8 are different for cavity #8.  (See “Welding Information” starting on pg. 16 in this engineering note.)  The simulation is identical since minimal weld penetration and a conservatively small joint area are assumed.  For cavity #8, the actual fused areas for these welds increased, so the analysis is just more conservative and was not modified.

…

Analysis

…

2nd Load Case

Primary + Secondary stress evaluation at room temperature (worst case)

All identical to cavity #5 engineering note, pp. 38 – 49.  Table A12, which shows the cavity iris calculated stress, is copied here since this includes the one exceptional stress.

…

Table A12 -2nd Load Case: Niobium Cavity. Linearized Stress Intensity summary

| SCL | Component and Evaluation Region | Material | Stress Category and Value Linearized Stress Intensity [MPa] | Acceptance Criteria [MPa] |
| --- | --- | --- | --- | --- |
| J | Cavity: Cavity/Nb-Ti Joint Weld Region | Nb RRR | PM+Pb+Q | 11.22 | ≤ 3∙S=60 |
| Y | Cavity: Cavity/Nb-Ti Joint Weld Region | Nb RRR | PM+Pb+Q | 11.31 | ≤ 3∙S=60 |
| Z | Cavity: End Cells Region | Nb RRR | PM+Pb+Q | 20.81 | ≤ 3∙S=60 |
| D | Cavity: Mid Region- Iris | Nb RRR | PM+Pb+Q | 82.43 | ≤ 3∙S=60 |
| F | Cavity: Mid Region- Equator | Nb RRR | PM+Pb+Q | 25.59 | ≤ 3∙S=60 |

...

Welding Information

All welds were designed, welded, and qualified as per Cavity #5 except those highlighted in Tables A24-1 and A24-2.  …

Table A24-1 - 3rd harmonic dressed cavity #8 weld summary

| Weld No. | Weld description | Drawing or reference | Type of weld | How qualified and other notes |
| --- | --- | --- | --- | --- |
| 1 | Nb cavity end cell to Nb tube end flange, input coupler end | MD-442202 | e-beam | Parameters developed based on microscopic or metallographic analysis of cut, etched and polished weld samples. Welded at Sciaky. |
| 2 | Nb cavity end cell to Nb tube end flange, non-input coupler end | MD-442201 | e-beam | Same as above. Welded at Sciaky. |
| 3 | NbTi large conical end cap A to Nb tube end flange, input coupler end | MD-442202 | e-beam | Same as above. Welded at Sciaky. |
| 4 | NbTi small conical end cap B to Nb tube end flange, non-input coupler end | MD-442201 | e-beam | Same as above. Welded at Sciaky. |
| 5 | Nb end tube to Nb conical flange A, coupler end | MD-442202 | e-beam | Same as above. Welded at Sciaky. |
| 6 | Nb end tube to Nb conical flange B, non-coupler end | MD-442201 | e-beam | Same as above. Welded at Sciaky. |
| 7 | Ti bellows to Ti tube | MC-457114 | TIG | Same as above. Welded at Hi-Tech. |
| 8 | Ti bellows to Ti tube | MC-457114 | TIG | Same as above. Welded at Hi-Tech. |
| 9 | NbTi conical end piece to Ti cylinder, coupler end | ME-440598 detail C and “He_vessel to cavity weld.ppt” | e-beam | Same as weld numbers 1 - 6 above. Welded at Sciaky. |
| 10 | NbTi conical end piece to Ti spacer ring, non-coupler end | ME-440598 detail D and “He_vessel to cavity weld.ppt” | e-beam | Same as weld numbers 1 – 6 above. Welded at Sciaky. |
| 11 | Ti spacer ring to Ti cylinder, non-coupler end | ME-440598 detail D and “He_vessel to cavity weld.ppt” | TIG | Welded by Dan Watkins, Fermilab, in the A0 glove box. WPQ and PQR from Alloy Weld Inspection Company. |

The following information explains the welding procedure of the helium vessel for cavity #8 for weld numbers 7 – 11 (in Table A24-1):

All the parts for this specific helium vessel were machined and TIG welded at Hi-Tech.

There are three welds connecting the helium vessel to the cavity:  an EB weld joining the bottom of the vessel to the large conical end cap (weld #9), an EB weld joining the titanium slip ring to the small conical end cap (weld #10), and a TIG fillet weld joining the vessel to the slip ring (weld #11).  We are not allowed to do full penetration EB welds because the RF people want to avoid any NbTi or Ti vapor deposition on the half-cells.  They believe it will negatively affect heat transfer to the liquid helium in the vessel.

The two EB welds were done at Sciaky.  In the Power Point presentation you will see pictures of the cut, etched and polished weld samples performed using the same weld parameters.  The TIG weld was done at Fermilab by Dan Watkins.  You will also see pictures of the preliminary weld samples Dan produced in the Power Point presentation.  His WPQ is provided.

Weld #11, where there is a final fillet weld joining the titanium spacer ring to the titanium cylinder, is a complete circumferential weld.  The procedure for Cavity #8 to arrive at the weld is slightly different than for the same weld in Cavity #5.  The weld was built up in increments, then allowed to cool so that the frequency spectrum can be monitored.  The additional steps of cooling the weld and monitoring the frequency spectrum were taken for Cavity #8 and were not done for Cavity No. 5.  The final resulting weld is the same for both cavities.

Table A24-2 below lists the welds that are different between Cavities 8 and 5.

Table A24-2 – Summary of Differences in the Welds of Cavities #5 and #8

| Weld No. | Description | Cavity No. 5 | Cavity No. 8 |
| --- | --- | --- | --- |
| 7 | Ti bellows to Ti tube | E-beam weld at Sciaky – 75% penetration | TIG weld at Hi-Tech – full penetration |
| 8 | Ti bellows to Ti tube | E-beam weld at Sciaky – 75% penetration | TIG weld at Hi-Tech – full penetration |
| 9 | NbTi conical end piece to Ti cylinder, coupler end | Welds later certified by radiographic analysis, tensile and bend tests performed on weld samples via Packer Engineering (pending). | No radiographic analysis. Weld procedure is the same as for Cavity #5. |
| 11 | Ti spacer ring to Ti cylinder, non-coupler end | For final fillet weld, weld in increments around circumference, working towards a complete weld | For final fillet weld, weld in increments around circumference, allow part to cool, monitor frequency spectrum, repeat to work towards a complete weld |

Weld documents are available at:

http://ilc-dms.fnal.gov/Workgroups/CryomoduleDocumentation/3rd-Harmonic-Cryomodule-for-DESY/helium-vessel-engineering-note/

These include:

“Helium Vessel to Cavity No. 8.”  This PowerPoint file by Mike Foley describes the helium vessel weld setup and process.  Weld samples are also shown.

“Fermilab Welder Performance Qualification (WPQ).”  The final closure titanium TIG weld was performed by Dan Watkins.

“Procedure Qualification Record (PQR) for GTAW / Manual Process for Fermilab”.  The final closure titanium TIG weld was created at Fermilab.

“Procedure Qualification Record (PQR) for Electron Beam Welding process at Sciaky.”  The niobium cavity was welded at Sciaky.  Also, the NbTi / Ti welds were created at Sciaky.

“Procedure for TIG Welding Helium Vessels to 3.9 GHz 3rd Harmonic Cavities.”  The procedure details the steps in making the final three welds to attach the cavity to the titanium shell.

Figure A29 - Welds forming the Cavity #8 Helium Vessel Pressure Boundary

…

Weld Details:  Weld numbers 7, 8 from Figure A29

| Weld Number | Weld Description | Type of weld | Full Penetration or Partial Penetration | Location of Weld Stress Calculation in this Eng. Note |
| --- | --- | --- | --- | --- |
| 7 | Titanium bellows to titanium tube | TIG | Full | - |
| 8 | Titanium bellows to titanium tube | TIG | Full | - |

…

Appendix B – 2.3 bar Pressure Test

Pressure test results for Cavity 8 TBD.

Test Procedure

See attached sketch for test flow schematic. Following paragraph 8.3 of Section VIII, Div 2 of the ASME pressure vessel code: The pressure in the vessel shall be gradually increased to not more than half the test pressure (which means not more than 17 psig). Then increase pressure in steps of 2 to 3 psi until the final test pressure of 33.5 psig is reached. Then back the pressure off to the MAWP of 29 psig and stop. Verify that no leaks occur by watching the system pressure for 3 minutes. Reduce pressure to half of the test pressure (to 17 psig), then enter the cave for a visual inspection of the vessel. Note 1: since this vessel and test are a joint TD and AD effort, safety officers from both TD and AD will witness the test. Note 2: the pressure test will also include an RF frequency measurement before, during, and after pressurization to check for any permanent deformation of the cavity.

Report on the pressure test results.

People present were:  Elvin Harmes, Andy Hocker, Timergali Khabiboulline (RF test), Raymond Lewis, Tom Peterson, Allan Rowe, Rich Ruthe, Bill Soyars.  Roger Milholland and Steve Cozzens were the AD cryo technicians who performed the test.

Pressure tests and pressure hold at the MAWP follow the general guidelines presented in ASME pressure vessel code, Section VIII, Div 2, Part 8-3.  The table below shows the actual pressure levels for each pause and what was done at that pressure.  This table or something similar may be incorporated into procedures for future dressed cavity pressure tests.  Total time for the test, not including setup and tear-down time, was about 20 minutes.

Appendix C – Cavity processing information

The bare niobium cavity #8 was processed and baked as described in the tables below, excerpted from Allan Rowe’s spreadsheet, “SCRF 3rd Harm Cavity Processing Chart.xls”.

# Text boxes

FERMILAB

Technical

Division

Weld #7 & #8

Seamless Titanium Tube

Titanium Bellows Assembly

#1, #3, #5

#9

#7, #8

#2, #4, #6

#11

#10