##### QADATA Summary – March 2003

The Inorganic Blind Sample Project (IBSP) reviews 95 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 84 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-assurance samples.  The QADATA Summary is a report that includes a short summary of only those analytical lines for which the quality-assurance charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 09/15/02 through 03/03/03; however, longer-range charts are shown when they add clarity to problem.  The attached quality-assurance charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality control-review as the data released to the districts.  Thus, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-assurance charts.  Samples for which the results are outside of control limits (+/-2 f-pseudosigma) are reanalyzed but the values are not updated unless it is determined that the sample had deteriorated or had been contaminated before submittal, or a problem at IBSP had occurred.

In this report, the “Problem” is determined and stated by IBSP.  The “Problem” is sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report when those comments help to provide a clearer perspective of the overall concern.

Charts containing all data received as of 03/25/03 may be accessed at http://bqs.usgs.gov/bsp/WY02_03charts.htm and

http://bqs.usgs.gov/bsp/WY03charts.htm.

Lab personnel and customers can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp/mainpage.html (select QADATA Application button).  Charts pertaining specifically to this report can be accessed by following the blue, underlined hyperlinks embedded within the report and associated with each analysis discussed in the report.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes (in blue) are hyperlinked to their respective positions within the report.  After following a link, use the back-arrow (if shown) to return the cursor to its previous position.

NWQL

Analyte	Page

Orthophosphate, Filtered, Labcode 1974,  Parmcode 0671H . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Phosphorus, Filtered, Low-level, Labcode 2331, Parmcode 0666G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Phosphorus, Whole-Water Recoverable, Low-level, Labcode 2333, Parmcode 0665G . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	4

Antimony, Filtered, ICP-MS, Labcode 1785, Parmcode 1095G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	5

Barium, Whole-Water Recoverable, ICP, Labcode 2352, Parmcode 1007C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Beryllium, Filtered, ICP-MS, Labcode 1787, Parmcode 1010G  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .   7

Manganese, Filtered, ICP-MS, Labcode 1793, Parmcode 1056G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	7

Strontium, Filtered, ICP-MS, Labcode 2507, Parmcode 1080I. . . . . . . .  . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .   7

Chloride, Filtered, IC, Labcode 1571, Parmcode 0940J. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  8

Potassium, Filtered, Flame AA, Labcode 0833, Parmcode 0935A . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	8

OWQRL

Analyte	Page

Ammonia+Organic-N, Filtered, Labcode 1985, Parmcode 0623D. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	9

Ammonia+Organic-N, Whole-Water Recoverable, Labcode 1995, Parmcode 0625E. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .   	9

Arsenic, Whole Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	 9

Calcium, Filtered, Labcode 3016, Parmcode 0915(1). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	10

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Orthophosphate, Filtered Colorimetric labcode 1974 parmcode 0671H | The high variability, starting October 2002, has still not improved. Since October 2002, almost one-third of the data are outside of control limits. The OWQRL has no points outside of +/-2 SD for this parameter and is analyzing the same reference materials. | The Nutrients Unit reran N76 and NS10 samples on three different instruments. The results were consistently low for NS10 and consistently high for N76. Samples prepared from certified reference standards at concentrations close to NS10 and N76, and reruns of N75 gave values close to their MPVs. On-line QC charts for ERA certified reference samples with concentrations close to NS10 and N76 show little change in variability for the period Jan 2002 to present. The photometer, pump, tubing and fittings for the instrument have been checked and are functioning properly. Reagent chemistry was also checked with no problems noted. | The NWQL has ordered certified reference samples from NIST and will prepare samples at concentrations close to NS10, N75 and N76. These samples, the certified references used by the Nutrients Unit and reruns of the three IBSP samples will be run together on several runs and the results will be reported to IBSP. The OWQRL was contacted to compare methods and instrumentation. The method will continue to be examined to see if any problems can be found. |
| Previous comments for orthophosphate, filtered, labcode 1974, parmcode 0671H |
| From QADATA Summary January 2003 | High variability, starting October 2002, still demonstrated. | The Nutrients Unit has been monitoring the results for the IBSP samples using the IBSP control charts. N76 and the on-line QC sample prepared from a certified reference material continue to run higher than the MPV. NS10 is still running lower than the MPV and has variability that is not exhibited on the other instruments. Examination of the ortho-phosphate line analytical system was performed. | The Nutrients Unit, with the assistance of the MRD Program, corrected deficiencies found on the orthophosphate line. On-line QC samples, including commercial certified reference samples, will be monitored for improvement in the variability. The BQS charts will also be looked at weekly for improvement. |
| From QADATA Summary November 2002 | Increased variability overall starting 10/01/02, and high bias demonstrated for N76. Variability problems not as evident for OWQRL, however, less data are available for assessment. | The IBSP control chart for LC 1974 shows a high bias for N76, a low bias for 3 of 4 data points for NS10 and no bias for N75. Analysis of N76 samples obtained from IBSP on the 4-channel, and low-level 4-channel (at x5 dilution) instruments give results similar to those on the IBSP chart. N75 samples give results very close to the MPV on both instruments. | The Nutrients Unit will continue to run N76 and N75 samples obtained from BQS on-line. In addition, a sample has been prepared at the same concentration level as N76 from a commercially prepared certified reference material that will also be run on-line. Results will be reported to IBSP. |
| Phosphorus, Filtered Colorimetric Low-level labcode 2331 parmcode 0666G | Corrective actions taken by the NWQL seem to have corrected the long-running (since April 2002) positive bias. Analysis should still be closely monitored until the next QADATA Summary is distributed. | New stock standards were obtained from a different source and seem to have corrected the bias. The analysis will continue to be monitored. | None required at present. |
| Phosphorus, Whole-Water Recoverable Colorimetric Low-level labcode 2333 parmcode 0665G | Corrective actions taken by the NWQL seem to have corrected the long-running (since April 2002) positive bias. Analysis should still be closely monitored until the next QADATA Summary is distributed. | New stock standards were obtained from a different source and seem to have corrected the bias. The analysis will continue to be monitored. | None required at present. |

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Antimony, Filtered ICP-MS labcode 1785 parmcode 1095G | The positive bias is still present; however, only two values have been received since implementation of the new calibration standards. Since March 2002, 63 out of 65 points are above the MPV. | We have identified the bias to be related to the vendor-supplied standards we have been using. Using calibration standards prepared from standard material used by the Method Research Development Program gave unbiased SRWS results after one check. | We will complete verifying this new calibration source, and continue to monitor for bias. |
| Previous comments for antimony, filtered, ICP-MS, labcode 1785, parmcode 1095G |
| From QADATA Summary January 2003 | Positive bias persists. Fifty-three out of 55 points are above the MPV since March 2002. | Bias has continued on in-house control charts. | New standards have been verified as acceptable. The NWQL started using the new standards the last week of January 2003 and will keep monitoring QC for bias. |
| From QADATA Summary November 2002 | A positive bias (which appeared near the beginning of March 2002) is still present. New calibration standards made in August seemed to have no effect on the bias. | Bias is apparent on in-house control charts also. | Newly purchased standards are prepared and currently undergoing verification against the newest SRS and other third party standards. The lab will monitor QC for bias. |
| From QADATA Summary September 2002 | A positive bias, which appeared near the beginning of March 2002, is still present. | Lab QC charts also confirm bias. | Although calibration standards were made in February (2002) and tested within the +/- 2% acceptance, new calibration standards were recently prepared. Antimony will be closely monitored. |
| From QADATA Summary July 2002 | This analytical line is exhibiting a high bias. The bias started near the beginning of March. | We have been looking at QC charts concerning the high bias noted for dissolved antimony - ICPMS. It appears to have been a result of the standards that have been in use since about that time. | New standards have been prepared and we will continue to monitor closely. |
| Barium, Whole-Water Recoverable ICP labcode 2352 parmcode 1007C | Although there are two points above the MPV, this line is still exhibiting a negative bias. No new values have been received since the implementation of new standards (last week of February 2003). Analysis should be closely monitored to determine the effects of the new standards. | New standards were prepared and verified by ICPMS. New standards were accepted. | The NWQL will continue to monitor. More data from the IBSP since these actions were implemented will be examined in order to better assess if the bias is improving. |
| Previous comments for barium, whole-water recoverable, ICP, labcode 2352, parmcode 1007C |
| From QADATA Summary January 2003 | The historical negative bias persists. | NWQL charts still show a slight negative bias. The standards approved in August, 2002 were recently re-checked and found to be slightly outside acceptance limits on a new ICP instrument in development status. | New standards will be prepared and checked the first week in February, 2003. They will be compared with the old standards on the ICP as well as by ICPMS (alternate method). SRS performance will be evaluated using the new standards. The OWQRL lab will be called to compare procedures as well. |
| From QADATA Summary September 2002 | Magnitude of historical negative bias has decreased although a negative bias is still demonstrated. Analysis should still be closely monitored to ensure magnitude of bias does not increase. | Adjustments to past descriptions of bias from IBSP have been to digest the calibration standards and QC to match the digestion process of the samples. | Calibration standards were made and accepted on August 20, 2002. Standards will be cross-checked by a different method. Using lab QC data and IBSP, we will continue to monitor and, additionally, obtain more current SRWS from BQS for internal lab QC. |
| Beryllium, Filtered ICP-MS labcode 1787 parmcode 1010G | A positive bias is still evident, however, the most recently received data looks favorable. Analysis should be closely monitored until the next QADATA Summary is distributed. | Third party QC checks still look unbiased. | The NWQL will continue to closely monitor beryllium for bias. |
| Previous comments for beryllium, filtered, ICP-MS, labcode 1787, parmcode 1010G |
| From QADATA Summary January 2003 | The positive bias that developed around the beginning of July 2002 has somewhat decreased, however, a positive bias is still evident. | Third party QC checks in lab remain unbiased. | New standards are prepared and verified against the newest SRSs and other third party standards. The new calibration standards were implemented in January 2003. The NWQL will monitor QC for improvement in bias. |
| From QADATA Summary November 2002 | A high bias has developed that started the beginning of July 2002. This bias may be aggravated by the new mix introduced in October. | New mix in October shows high bias. Third party QC checks in lab do not show a bias. | Newly purchased standards are prepared and currently undergoing verification against the newest SRS and other third party standards. The lab will monitor QC for bias. The lab will monitor beryllium closely for improvement. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | Positive bias still evident although the two most recently received values are close to MPV. Cadmium by other methods at the NWQL and the OWQRL do not show the significant bias that this line does. | Third party QC charts show improvement since new standards started in January 2003. | The lab will continue to closely monitor third party and other quality assurance samples. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary January 2003 | The slight positive bias that developed around April 2002 appears to have increased even more around mid-September 2002. | Third party QC charts show a 5% bias at 2 ug/L and 3% bias at 20 ug/L since October 2002. All in-house QC data have been acceptable. New SRSs were started December 2002. Not enough data for the newest SRSs to evaluate bias. | New standards are acceptable and have been in use since January 2003. The NWQL will monitor cadmium closely for improvement in bias. |
| From QADATA Summary November 2002 | A slight positive bias that developed around mid-April 2002 is increasing. | In-house charts show bias starting at about the same time. All in-house QC data has been acceptable. Bias most noticeable in the two newest mixes. | Newly purchased standards are prepared and currently undergoing verification against the newest SRS and other third party standards. The lab will monitor cadmium closely for improvement. |
| Manganese, Filtered ICP-MS labcode 1793 parmcode 1056G | A upward shift in the data occurs mid-December 2002. Analysis should still be closely monitored to ensure magnitude of the bias does not increase. | Third party QC checks do not show a shift. | The NWQL will closely monitor for bias. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | Increased variability starting mid-October 2002. Since October, six out of 19 points (32%) are outside of +/-2 SD. | NWQL monitors performance at 250(g/L and 2500(g/L. QC Charts for these concentration levels are in control. The majority of environmental samples received are near 250ug/L concentration. IBSP monitors much lower concentrations. | The NWQL will monitor at the lower concentrations that IBSP monitors and base run acceptance on these performance checks. We will continue to monitor for variability. We are continuing to investigate issues related to the increased variability including instrumentation, procedures, and blank QC. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Chloride, Filtered IC labcode 1571 parmcode 0940J | A negative bias may be developing. Analysis should be closely monitored until the next QADATA Summary is distributed. | We will continue to monitor and review the line QC to see if a similar trend is noted. | No corrective action is determined as yet. |
| Potassium, Filtered Flame-AA low-level labcode 0833 parmcode 0935A | A negative bias of approximately 10% is noted. | Control charts (QC and blanks) for the past year (3/2002 – 3/2003) do not show an overall trend toward a negative bias. The most recent chart (from the IBSP website, 03/26/03) shows the negative bias with IBSP samples improving. | The NWQL will continue to monitor for bias. |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Ammonia+Organic-N, Filtered Colorimetric labcode 1985 parmcode 0623D | Variability has increased as of October 2002. This pattern is not demonstrated by the NWQL. | Our internal control charts do not show the increase in variability the IBSP charts show. The variability is also not shown by labcode 1986, which is run on the same system at the same time. A possible problem was noted with the mercuric chloride reagent for this determination. | A new bottle of mercuric chloride was ordered and is now being tested. We will continue to monitor this closely. |
| Ammonia+Organic-N, Whole-Water Recoverable EPA labcode 1995 parmcode 0625E | The slight positive bias may be increasing. Analysis should be closely monitored until the next QADATA Summary is distributed. | This is also run on the same system as the above analysis. There are only 2 samples currently being used for QA by the IBSP on this line, one of which is below the MDL. | Some of the bias here may be due to rounding error. This parameter is only reported to 1 significant figure in the range of the samples that are used for QA. We will continue to monitor this closely. |

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole Water Recoverable GFAA labcode 3119 parmcode 1002(1) | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 SD. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode1002(1) |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |

| Major Ion Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered, ICP labcode 3016 parmcode 0915(1) | A positive bias, starting in November 2002, is demonstrated. The magnitude of the bias increased in December. | Positive bias may be due to increased sensitivity that decreased our linear range. | Calibration range has been adjusted. The high standard has been lowered from 100 to 80 mg/L. We will continue to monitor this. |