##### QADATA Summary – September 2003

The Inorganic Blind Sample Project (IBSP) reviews 95 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 84 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-assurance samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-assurance charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 02/01/03 through 08/20/03; however, longer-range charts are shown when they add clarity to the problem.  The attached quality-assurance charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality control-review as the data released to the districts.  Thus, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-assurance charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have deteriorated or are known to have been affected by a problem at IBSP are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report when they help provide a clearer perspective of the overall concern.

Charts containing all data received as of 09/05/03 may be accessed at http://bqs.usgs.gov/bsp/WY03charts.htm.  Lab personnel and customers can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp/mainpage.html (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the blue, underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Orthophosphate, Filtered, Colorimetric, Labcode 1974,  Parmcode 0671H . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Barium, Whole-Water Recoverable, ICP, Labcode 2352, Parmcode 1007C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	5

Lead, Filtered, GFAA, Labcode 1560, Parmcode 1049F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	6

Lead, Whole-Water Recoverable, GFAA, Labcode 1561, Parmcode 1051F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	6

Alkalinity, Filtered, Titrimetric, Labcode 2109, Parmcode 29801A . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	7

Alkalinity, Whole-Water Recoverable , Titrimetric, Labcode 0070, Parmcode 90410A . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	7

OWQRL

Analyte

Phosphorus, Filtered, Colorimetric, Labcode 1983,  Parmcode 0666D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	8

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	9

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Orthophosphate, Filtered Colorimetric labcode 1974 parmcode 0671H | Results for NS10 are still demonstrating a negative bias with about one-half of the NS10 data outside of -2 f-pseudosigma. The OWQRL is not demonstrating a negative bias of this magnitude and has no data outside of -2 f-pseudosigma. | A new vial of the NIST certified reference sample purchased by the NWQL was opened and a sample at a concentration similar to NS10 was prepared and analyzed. Values were slightly higher than the calculated mean value. Samples made at three concentrations from a new commercially prepared certified reference material were also analyzed. The values for these samples were close to but slightly lower than the calculated mean. | The Nutrients Unit will continue to monitor IBSP control charts and will analyze the new QC samples as well as on-line NIST QC samples. We will continue to look for other causes for the low bias for NS10. |
| Previous comments for orthophosphate, filtered, labcode 1974, parmcode 0671H |
| From QADATA Summary July 2003 | Over the last two months (April and May 2003) results for NS10 show improvement relative to earlier data. Analysis should be closely monitored to ensure values stay within +/- 2 f-pseudosigma. | Values for N77 and N78 are distributed closely around the median. QC samples prepared from the NIST certified PO4 reference material purchased by the NWQL also run close to their certified value. | The Nutrients Unit will continue to run NIST samples and monitor the IBSP charts. |
| From QADATA Summary May 2003 | N76 has been replaced by a new matrix (N78) which now indicates that this line has improved variability. However, for the past two months, about 50% of the data for NS10 are outside of -2 f-pseudosigma, whereas for the OWQRL no data for NS10 are outside of 2 f-pseudosigma. | The NIST certified reference sample purchased by the NWQL was prepared at concentrations close to the PO4 concentration in N75, N76 and NS10. The NIST sample near the NS10 level ran slightly high on average and the NIST sample near the N75 level gave results that averaged close to the certified value. Results for the NIST sample near the N-76 concentration indicate a higher than expected concentration. All NIST samples were cross-checked on the low-level ortho-P line with results that were close to the regular line values. | The Nutrients Unit will continue to analyze NIST samples in addition to the on-line QC samples. The results from these samples and the IBSP results will continue to be monitored. |

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Barium, Whole-Water Recoverable ICP labcode 2352 parmcode 1007C | The negative bias (-5.2% for WY03) has still not improved. | New ICP instrumentation went online beginning in August. There is no IBSP data provided from this time to date. Online QC data for these new instruments indicate no appreciable bias at these concentrations. A Rapi-Note and an NWQL Technical Memorandum will be issued in the near future detailing the use of the new ICPs for Whole-Water Recoverable analyses. | The NWQL will continue to monitor WWR barium QC on the new instruments to determine if a bias continues. |
| Previous comments for barium, whole-water recoverable, ICP, labcode 2352, parmcode 1007C |
| From QADATA Summary July 2003 | The historical negative bias has not improved and may be getting worse. It is again recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | The NWQL agrees with IBSP that the barium bias is getting worse. New ICP instrumentation has come online for the dissolved line. | The NWQL will continue to monitor the barium WWR line on the different instrument to see if there is an improvement. The ICP WWR method has been moved to a different instrument that has had recent maintenance. We will reanalyze IBSP samples on this instrumentation. We plan to have the new Perkin Elmer ICP instrumentation online this summer. |
| From QADATA Summary May 2003 | The historical negative bias persists. It is recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | Online QC charts show that the negative bias increased significantly during mid-March. Several troubleshooting measures were checked on the instrument including: plasma power, pump efficiency, cleaning of nebulizer and spray chamber, dark current test, all of which showed no indication of a problem. The only test that showed a significant problem was that the Purged Optical Path, which needed adjusting. | The NWQL will reanalyze samples on the new ICPs to compare the data, once we have assessed the WWR prove-out data for the new instruments. We will continue to monitor online QC charts to see if the negative bias has improved as a result of the optical path adjustment. |
| From QADATA Summary March 2003 | Although there are two points above the median, this line is still exhibiting a negative bias. No new values have been received since the implementation of new standards (last week of February 2003). Analysis should be closely monitored to determine the effects of the new standards. | New standards were prepared and verified by ICPMS. New standards were accepted. | The NWQL will continue to monitor. More data from the IBSP since these actions were implemented will be examined in order to better assess if the bias is improving. |
| From QADATA Summary January 2003 | The historical negative bias persists. | NWQL charts still show a slight negative bias. The standards approved in August, 2002 were recently re-checked and found to be slightly outside acceptance limits on a new ICP instrument in development status. | New standards will be prepared and checked the first week in February, 2003. They will be compared with the old standards on the ICP as well as by ICPMS (alternate method). SRS performance will be evaluated using the new standards. The OWQRL lab will be called to compare procedures as well. |
| From QADATA Summary September 2002 | Magnitude of historical negative bias has decreased although a negative bias is still demonstrated. Analysis should still be closely monitored to ensure magnitude of bias does not increase. | Adjustments to past descriptions of bias from IBSP have been to digest the calibration standards and QC to match the digestion process of the samples. | Calibration standards were made and accepted on August 20, 2002. Standards will be cross-checked by a different method. Using lab QC data and IBSP, we will continue to monitor and, additionally, obtain more current SRWS from BQS for internal lab QC. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | The results for February 2003 through May 2003 demonstrate a positive bias (7.8%). The results for the same reference materials analyzed at the OWQRL do not exhibit a bias. | Began using new standards July 9, 2003. On line QC charts prior to July 9 show some bias, but bias has decreased since new standards were made. | Improvement is noted since standards were made. NWQL will continue to monitor data quality. |
| Previous comments for cadmium, whole-water recoverable, GFAA, labcode 1555, parmcode 1027F |
| From QADATA Summary July 2003 | As of January 2003, a high bias has developed and variability increased. | NWQL QC charts show variable but acceptable data. | Instrument has recently been optimized, and the NWQL will continue to monitor. |
| Lead, Filtered GFAA labcode 1560 parmcode 1049F | The four most recent values suggest that the variability has improved. Analysis should be closely monitored until the next QADATA Summary is distributed. | NWQL QC charts also show improvement in variability. | NWQL will continue to monitor line. |
| Previous comments for lead, filtered, GFAA, labcode 1560, parmcode 1049F |
| From QADATA Summary July 2003 | The most recently received data suggest that the variability is not improving and may be increasing. | Variable, but acceptable data quality is also noted on NWQL charts. | Instrument has recently been optimized and the NWQL will monitor for improvement. |
| From QADATA Summary May 2003 | Variability has increased somewhat since January 2003. Variability should be closely monitored until next QADATA Summary is distributed. | Change in variability not noticeable on NWQL charts. | NWQL will continue to monitor. |
| Lead, Whole-Water Recoverable GFAA labcode 1561 parmcode 1051F | The three most recent values suggest that the variability has improved. Analysis should be closely monitored until the next QADATA Summary is distributed. | NWQL QC charts also show improvement. | NWQL will continue to monitor line. |
| Previous comments for lead, whole-water recoverable, GFAA, labcode 1561, parmcode 1051F |
| From QADATA Summary July 2003 | The data indicate increased variability and positive bias starting March 2003. | Variable, but acceptable data quality is also noted on NWQL charts. | Instrument has recently been optimized and the NWQL will monitor for improvement. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Filtered Titrimetric labcode 2109 parmcode 29801A | Actions taken in March appear to have corrected the problem reflected in the February and March data; however, irregular spikes in the data are still occurring. | Irregular spikes are due to instrument stirring variability. The titration dispenser arm position varies slightly and can cause variable stirring, which causes the samples to over titrate, resulting in spikes in the data. A slight bump to the arm can cause the stirring variability. | Analysts have adjusted the titration dispenser arm to allow for the proper stirring. The arm adjustments are checked more frequently and the analysts are watching for slight over titrations to catch the arm variations sooner. |
| Previous comments for alkalinity, filtered, titrimetric, labcode 2109, parmcode 29801A |
| From QADATA Summary May 2003 | The data indicate a severe problem starting late in February 2003. Most recently received data (early March) suggest problem has been corrected. | In late February and early March we determined that there was a problem with the alkalinity line. We were getting high values for some samples; however, this was not immediately shown in the on-line QC. Upon investigation, we determined that the titrant was not remaining stable due to an autoclave steam vent that had been moved next to the titrant storage area. | We have remade and moved the titrant away from the steam vent and the problem has been corrected. We are continuing to closely monitor the on-line QC and the BQS charts. |
| Alkalinity, Whole-Water Recoverable Titrimetric labcode 0070 parmcode 90410A | See above (lab code 2109) | Irregular spikes are due to instrument stirring variability. The titration dispenser arm position varies slightly and can cause variable stirring, which causes the samples to over titrate, resulting in spikes in the data. A slight bump to the arm can cause the stirring variability. | Analysts have adjusted the titration dispenser arm to allow for the proper stirring. The arm adjustments are checked more frequently and the analysts are watching for slight over titrations to catch the arm variations sooner. |
| Previous comments for alkalinity, whole-water recoverable, titrimetric, labcode 0070, parmcode 24101A |
| From QADATA Summary May 2003 | The data indicate a severe problem starting late in February 2003. Most recently received data (early March) suggest problem has been corrected. | In late February and early March we determined that there was a problem with the alkalinity line. We were getting high values for some samples; however, this was not immediately shown in the on-line QC. Upon investigation, we determined that the titrant was not remaining stable due to an autoclave steam vent that had been moved next to the titrant storage area. | We have remade and moved the titrant away from the steam vent and the problem has been corrected. We are continuing to closely monitor the on-line QC and the BQS charts. |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Phosphorus, Filtered Colorimetric labcode 1983 parmcode 0666D | All data received since May 1st, 2003, indicate a negative bias. The NWQL is not showing this bias. | Most of the data points exhibiting the negative bias are a commercially obtained sample that has also exhibited a negative bias when run at the NWQL. Other blind samples are also showing a negative bias, though not as much as the commercial one. Whole-water recoverable phosphorus, which is analyzed on the same instrument at the same time, is not showing this same bias. The same commercially obtained sample is not submitted for Total Phosphorus. | The cause of the bias is currently being investigated. This line is being monitored closely. |

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 SD with 3 of 19 (16%) outside of +/-3 SD. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode 1002(1) |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 SD with 3 of 13 (23%) outside of +/-3 SD. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 SD. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |