SUMMARY OF SAFETY AND EFFECTIVENESS

- General Information

#### Device Generic Name:           Total temporomandibular joint implant

## Device Trade Name:		Total Temporomandibular Joint (TMJ)

## Replacement System

Applicant’s Name:		Biomet , Inc.

56 East Bell Drive

P.O. Box 587

Warsaw, Indiana  46582  U.S.A.

Premarket Approval Application (PMA) Number:

Date of Panel Recommendation:

Date of Notice of Approval to the Applicant:

Indications for Use

The Total TMJ Replacement System is indicated for use in cases of:

- 1.   arthritic conditions: e.g. osteoarthritis, rheumatoid arthritis, or traumatic

- arthritis,

- malignancy (e.g. post-tumor excision),

- benign neoplasms,

- functional deformity,

- revision procedures where other treatments (e.g. alloplastic reconstruction, autogenous grafts) have failed,

- avascular necrosis,

- ankylosis including but not limited to recurrent ankylosis with excessive heterotopic bone formation,

- fracture,

- multiple operated joints, and

- developmental abnormality

- Device Description

# The Total Temporomandibular Joint (TMJ) Replacement System is implanted in

# the jaw to functionally reconstruct a diseased and/or damaged temporomandibular

# joint. The Total TMJ Replacement System is a two component system comprised

# of mandibular condyle and glenoid fossa components. Both components are

# available in multiple sizes as right and left side specific designs and are attached

# to bone by screws. The individual components are not for use in partial joint

# reconstruction.

# Materials:

# Mandibular Component – Cobalt-Chromium-Molydenum (Co-Cr-Mo) alloy per

# ASTM F 1537 with titanium alloy (Ti-6Al-4V per

# ASTM F 136) plasma spray porous coating

# Fossa Component – ArCom® ultra-high-molecular-weight (UHMWPE) per

# ASTM F 648

# Screws – Titanium alloy (Ti-6Al-4V per ASTM F 136)

### Contraindications

1.   Active or chronic infection.

- 2.   Patient conditions where there is insufficient quantity or quality of bone to

- support the components.

- 3.   Systemic disease with increased susceptibility to infection.

- 4.   Patients with perforations in the mandibular fossa and/or bony deficiencies in

- the articular eminence or zygomatic arch comprising support for the artificial

- fossa component.

- 5.   Partial TMJ joint reconstruction.

- 6.   Known allergic reaction to any materials used in the components.

NOTE:  Patients with known or suspected nickel sensitivity should not have

Co-Cr-Mo devices implanted since this material contains nickel.

- 7.   Patients with mental or neurological conditions who are unwilling or unable to

- follow postoperative care instructions.

- 8.   Skeletally immature patients.

- 9.   Patients with severe hyper-functional habits (e.g. clenching, grinding etc.)

- Warnings

The following risks are associated with the use of a total TMJ system.

1.  Implant loosening or displacement can occur.

- 2.  The screws used to anchor the implant may loosen causing changes in bite,

- difficulty in chewing, limited joint function and/or unpredictable wear on

- implant components.

- 3.  Implant breakdown may result in erosion or resorption of the glenoid fossa,

- which may result in intense pain.

- 4.  A foreign body reaction may occur resulting in implant deterioration and

- migration of materials.

5.  If the implant is not properly sterilized, infection may result.

- 6.  If the implant materials are unable to withstand the forces or pressure placed

- on the implant, the implant can be torn, worn, perforated, fragmented,

- fatigued, or fractured resulting in failure of the device to function properly.

- 7.  Degenerative changes within the joint surfaces and components of the TMJ

- due to implant breakdown may result in chronic pain.

- 8.  Degenerative changes in the joint cartilage and/or bone from disease or

- previous implants may lead to failure of this device.

- 9.  If the implant materials are subject to the production of particles or corrosion,

- toxic elements may migrate to various parts of the body.

- 10.  Placement of the implant in one joint only may result in harmful effects to the

- joint on the opposite side.

- 11.  Placement of the implant may produce an improper relationship between the

- teeth surfaces that should contact during biting.

- 12.  Implant breakdown may cause bony erosion, heterotopic bone formation, or

- reactive bone within the joint.

- 13.  Use of implants may result in tinnitus or other ear problems.

- 14.  Limited range of motion and chronic pain may continue after total TMJ

- surgery.

- 15.  Infection can occur which may result in implant removal.

- 16.  Damage to the facial and/or trigeminal nerve with temporary or permanent

- paralysis of the facial muscle and/or loss of feeling in the chin, teeth, tongue,

- or lower jaw may occur.

The surgeon must be thoroughly knowledgeable with the components,

instruments and surgical procedure. In all cases sound medical practice is to be

followed and the surgeon must select the type of device appropriate for treatment.

Existing mandibular and/or zygomatic arch screw holes may compromise

fixation. The Total TMJ Replacement System is designed for total joint

reconstruction and components are to be used as a system. Do not use the

individual components for partial joint reconstruction.

The patient is to be warned that the system does not replace normal healthy bone

in their TMJ and they may continue to have chronic pain and limited range of

motion. The system can break or loosen as a result of stress, activity, or trauma.

Patients with severe hyper-functional habits may have an undesirable outcome.

The patient is to be made aware of surgical risks and possible adverse effects prior

to surgery and warned that failure to follow postoperative care instructions can

cause failure of the implant and the treatment.

Precautions

- DO NOT USE if there is loss of sterility of the devices.

- Discard and DO NOT USE opened or damaged implants and only use

implants that are packaged in unopened or damaged containers.

- DO NOT USE the individual components of this total system (e.g. mandibular

components,  fossa components, or screws) for partial joint reconstruction.

- Infection can lead to failure and subsequent removal of the devices.

- Damage to the implant can occur as a result of traumatic injury or excessive activity.

- Neurovascular injuries can occur due to surgical trauma.

- Metal sensitivity or foreign body reaction can occur due to the device materials or materials from previously implanted devices.

- Implant breakdown and/or degenerative changes in the TMJ may cause pain, which may lead to re-operation.

9.   Use of the system with filler material:

The fossa component may be used with a filler material when it is desired to

fill voids between the fossa prosthesis and the glenoid fossa bone. The filler

should never be used for fixation of the device or in any load bearing

application. If a filler is used in the fossa region, screws are placed after

polymerization of the filling material, if applicable.  Use of any legally

marketed craniomaxillofacial filler material is recommended.

- Alternative Practices and Procedures

Alternative practices and procedures include autogenenous or allogeneic bone

grafting, soft tissue grafts or implantation of other devices for partial or total TMJ

reconstruction.

VIII.   Potential Adverse Effects

Adverse events can occur following placement of TMJ implants and specific risks

are associated with this type of surgery. See the package insert for adverse events

reported with the use of this system during an approved Food and Drug

Administration (FDA) clinical study. The occurrence of a complication may be

related to or influenced by prior medical conditions or treatment and may require

further treatment. These adverse events include but are not limited to the

following:

(   Infection

(   Post-operative pain, swelling, bruising, jaw muscle spasm, or hematoma

formation

(   Chronic or recurring pain

(   Heterotopic bone formation, neuroma formation, adhesions, resorption or bony

erosion, and/or ankylosis

(   Facial nerve dysfunction

(   Dislocation

(   Removal of components(s) and/or revision

(   Harmful effects to the contralateral joint in unilateral cases

(   Implant wear, loosening, damage, migration or displacement

- (   Changes in bite, difficulty in chewing, limited joint function

- (   A foreign body or allergic  reaction to implant materials

(   Ear problems

- (   Degenerative changes in the joint cartilage and/or bone

### IX.      Marketing History

#### Approval for marketing has been granted by Europe (EC-Certificate issued  November 23, 2000) and marketing efforts are just beginning. The system has been marketed in South Africa since January 2000.

- Summary of Preclinical Studies

The following biomechanical tests were conducted on the Total TMJ Replacement System. Test results were all determined to be sufficient for the intended use of the construct/component.

- Fatique Testing of Fossa and Mandibular Component Construct

- Static Testing of the Mandibular Component

- Fatigue Testing of Bone Cement

- Fossa Screw Head Pull-Through Test

- Compression Strength of Fossa Component Flange

- 2.7mm Self-Tapping Screw Pull-Out Strength

- Fatique Testing of Fossa and Mandibular Component Construct

Initial fatigue testing was performed on five joints with mandibular components minus the titanium plasma spray coating. No failures were seen after 10 million cycles at a maximum load of 145 lbs. at frequencies between 10 and 30 hertz. The same testing was repeated on four joints with mandibular components coated with titanium plasma spray. There were no failures after 10 million cycles and no porous coating delamination was observed as expected from previous testing on orthopedic devices.

B. Static Strength Testing of the Mandibular Component

A mandibular component was fixed to porcine bone using four 2.7mm diameter screws and a load was applied to ultimate failure. At 575.9 lbs. the component’s neck portion bent with neither fracture/pull-out of the bone screws or fracture of the component.

- Fatigue Testing of Bone Cement

A total of five joints were tested with the fossa components secured to the test fixture using bone cement. The testing was conducted under cyclic compression with a maximum load of 145 lbs. at frequencies between 10 and 30 hertz for 10 million cycles. Following the 10 million cycles fatigue test, no cement failure (e.g. cracking or chipping) was noted.

- Fossa Screw (2.0mm) Head Pull -Through Test

Twelve specimens were tested to determine the force required to pull the fossa screw head through the UHMWPE zygomatic arch flange of the fossa component.  A standard static tensile test was performed using a cross-head travel rate of 0.05(/minute and the ultimate tensile loads were recorded. The mean tensile strength was 79.8 ± 2.5 lbs. It was noted that clinical failure is highly unlikely because of the large shearing load and displacements required.

F.  Compression Strength of Fossa Component Flange

A fossa component was tested to establish the load required to collapse an

unsupported fossa body and assure that failure in this fashion does not cause

tearing or cracking of the UHMWPE junction between the body and flange of

the fossa component. The fossa body collapsed against the flange at 83 lbs.

without material failure at the body/flange junction.

G.  2.7mm Self-Tapping Screw Pull-Out Strength

Five 2.7mm screws used to fixate mandibular components were tested for

pull-out strength in fresh frozen bovine cortical bone. This substrate was

chosen to mimic the clinical application. The mean pull-out strength was 373.2

± 68.8 lbs.

Summary of the Clinical Studies

A.  Objective

The study was designed to obtain clinical data to support the safety and effectiveness of this device.

B.  Inclusion / Exclusion Criteria

Inclusion Criteria:

- Patients requiring total joint reconstruction due to:

arthritis (osteo-, rheumatoid, traumatic)		malignancy

ankylosis						functional deformity

avascular necrosis					revisions

benign neoplasms					fracture

multiple operated joints

2.   Patients who are skeletally mature.

- Patients must have at least one of the following criteria for surgical TMJ treatment.

- presence of considerable pain and/or limited function in the joint area.

- clinical and imaging evidence consistent with anatomic joint

pathology.

- previous failure of non-surgical treatment/therapy or a failed implant.

- high probability of patient improvement by surgical treatment.

- Patients must be able to return for follow-up examinations.

- Patients without serious compromising general medical conditions.

Exclusion Criteria:

- Patients with active infection.

- Patient conditions where there is insufficient quantity or quality of bone to support the device.

- Patients with perforations in the mandibular fossa and/or bony deficiencies in the articular eminence compromising support for the artificial fossa component.

- Patients with mandibular and/or zygomatic arch screw holes compromising component fixation.

- Patients requiring partial joint reconstruction or other TMJ procedures not listed as an indication.

- Patients who are NOT skeletally mature.

- Patients who are incapable or unwilling to follow postoperative care instructions.

- Patients who are unable to return for follow-up examinations.

- Patients with severe hyper-functional habits.

- Patients on chronic steroid therapy.

C.  Patient Population and Demographics

A total of 180 cases (268 joints) with a mean patient age of 40 years (range 12-82 years) were enrolled into the study. There were 161 females (89%) and 19 males (11%) comprised of 88 (49%) bilateral cases and 92 (51%) unilateral cases. Of the 92 unilateral cases, 46 (50%) are the right side and 46 (50%) are        left sides only. Demographic data are summarized in Table 1. Most cases had multiple diagnoses with osteoarthritis and ankylosis being the most common. See Table 2 for a complete listing of diagnoses.

The mean duration of symptoms prior to implantation with this device was 11 years (range 0.1- 34 years) with the mean number of 5.2 (range 0-29) prior surgeries.

Patients were categorized according to the Wilkes Classification. There were 2 (1%) cases in Class I, none in Class II, 4 (2%) in Class III, 68 (38%) cases in Class IV, and 106 (59%) cases in Class V.

D.   Evaluation Schedule

Patients were evaluated preoperatively and postoperatively at 1 month, 3

months, 6 months, 1 year, 1.5 years, and 3 years.  All data collected past the 3

years follow-up are included. The assessments carried out at each visit

labeled as Visit 1-Visit 11 are summarized in Table 3.

E.   Study Design

The study was a prospective, multi-center, non-randomized controlled study. It was designed to compare baseline clinical and radiographic assessments to assessments made postoperatively.

F.   Patient Accountability

Table 4 shows the number and percentage of cases with follow-up data at each of the visits.  Compliance ranged from 95.0 % at the 1 month follow-up  visit to 82.5 % at 3 years follow-up.

G.   Efficacy and Safety Parameters

- Primary efficacy endpoints include:

·  Jaw pain intensity as measured on a 10 cm visual analogue scale (VAS)

from preoperative assessment to assessment 3 years postoperative,

adjusted for baseline at preoperative assessment,

·  Interference with eating as measured on a 10 cm VAS from preoperative

assessment to assessment 3 years postoperative, adjusted for baseline at

preoperative assessment,

·  Maximal incisal opening (MIO) measurement (in mm) from

preoperative assessment to assessment 3 years postoperative, adjusted

for baseline at preoperative assessment

Analysis was performed on cases with 3 years follow-up postoperatively.

These cases were defined as two groups.  One is the cohort unimputed

group comprised of 45 cases and the second group, cohort imputed, is

comprised of  59 cases.  The cohort imputed group used data points obtained

at the follow-up visit closest to but not after the 3 years visit for analysis for

the 14 cases missing data at the 3 years visit. The primary endpoints are

summarized on the following chart.

| Primary Efficacy Endpoints | Cohort Imputed Cases n=59 | Cohort Unimputed Cases n=45 |
| --- | --- | --- |
|  | Difference between Vs 1 & Vs 8 ( SD | Difference between Vs 1 & Vs 8 ( SD |
| Jaw pain | 5.43 ( 2.73 cm | 5.70 ( 2.40 cm |
| Interference with eating | 5.59 ( 2.95 cm | 5.80 ( 2.27 cm |
| MIO | 10.61 ( 8.44 mm | 10.27 ( 8.33 mm |

These primary efficacy endpoints showed a significant improvement from

baseline to 3 years postoperative.  Multiple analyses (t-test and repeated

measures) demonstrate that significant improvement is evidenced after

implantation of the Total TMJ Replacement System, same patterned effect for

the cohort imputed and unimputed groups.

Further t-test analysis shows that in both the total group (n = 180) and the

cohort imputed group (n = 59), there was a statistical difference (p<.0001) in

all three primary endpoints between baseline (Vs 1) and assessments at all

time points from 1 month follow-up to 3 years follow-up.

Figures 1, 2, and 3 graphically display the three primary endpoints for the

total study group and the two cohort groups from baseline to the 3 years visit.

2.  Secondary efficacy endpoints include (at visits Vs 1, Vs 3 – 8):

·  Jaw pain intensity, interference with eating, and maximal incisal

opening

·  Patient satisfaction, with a focus on the comparison from postoperative

baseline (Vs 5) to 3 years follow-up (Vs 8):

- in hindsight, whether the patient would choose to have this

surgery;

- degree of satisfaction with surgery across time

The secondary efficacy endpoints demonstrated a gradual improvement

over time in terms of jaw pain, interference with eating, and MIO. Table

5 lists the means and standard deviation for the three endpoints at visits

Vs 1 - Vs 11.

Most patients were satisfied with their outcome as demonstrated with

over 90 % of cases reporting at least satisfied or better at every follow-up

visit. Furthermore, over 90 % of the cases in hindsight would choose to

have this surgery at all time points. More specifically for Vs 3 – Vs 8,

between 94 –99 % of the cases said yes to the question “ In hindsight

would you choose to have this surgery?”

- 3.  Safety

- Radiographic assessment (position of components, heterotopic bone formation, osseous erosion, fossa resorption)

The position of mandibular and fossa components and the mandibular

and fossa screws were assessed by investigators in comparison to

immediate postoperative radiographs.  There were three mandibular

components reported as having a change in position: two at Vs 4 and

one at Vs 7. One of these cases noted at Vs 4 also had a change of

position of the mandibular screws and the joint was removed at 6

months postoperative. No change of position was reported for fossa

screws.

Heterotopic bone formation was found in 13 joints, 7 rights and 6 left

joints.  There are no reports of osseous erosion or fossa resorption.

b.   Adverse events

Adverse Events (AEs) were documented for all cases throughout the

duration of the study. There have been no unanticipated device related

adverse events reported. Overall, 90 AEs were reported in 55 cases

(30.6 %) of the 180 cases. Four cases (2.2 %) terminated the study due

to their permanent total joint removal AEs. Table 6 summarizes AEs

requiring device removal. Table 7 summarizes AEs not requiring

device removal.

- Patient and Study Success

- Patient Success

A patient was determined to be a success if:

1.  patient has not had a permanent total joint removal, and

- 2.  patient meets two of the following three criteria:

(  reduction of pain by 1 cm (VAS) from baseline to 3 years follow-up

(  reduction of interference with eating by 1 cm (VAS) from baseline

to 3 years follow-up

(  increase in MIO of 10% from baseline to 3 years follow-up

- Study Success

The study was deemed to be a success with 60% or more of the patients receiving the device having met the above Patient Success at 3 years follow-up.

In the cohort unimputed group, 44 of 45 (97.8%) cases are patient successes. In the cohort imputed group, 57 of 59 (96.6%) cases are patient successes. These patient success rates surpass the criteria for study success.

H.   Safety Analysis

1.	Deaths

There have been three deaths reported in the study.

- Case # 7 died from complications from back surgery at five years TMJ postoperative.

- Case # 24 died approximately 3 years postoperative. Death was

secondary to a recurrent brain tumor.

3.   Case # 46 died 11 days postoperative due to hepatic coma.

2.  Revisions/Removals

- Total joint removed

- Case # 20

Bilateral patient first had the right fossa component removed (10 months postoperative) due to infection. A year later the right mandibular component was removed also due to infection. The right side (case #103) was re-implanted 7 months later. Case # 20 is now a left side only.

- Case # 61

Unilateral patient first had her fossa removed 10 months postoperative and subsequent mandibular removal 6 months later due to infection. This case is lost to follow-up.

- Case # 84 and # 134

Bilateral patient had only the left prosthesis removed at 4 months postoperative due to infection. Case # 84 is now a right side only. A second prosthesis (case # 134) was implanted 7 months after the removal. An early infection occurred again and the device was removed approximately 9 months postoperative.  Case #134 is lost to follow-up due to the permanent joint removal.

- Case # 100

Unilateral (right side) patient had removal of prosthesis at 6 months postoperative due to chronic swelling and pain. This case is lost to follow-up.

- Case # 145

This bilateral patient had difficulties with chronic left TMJ

dislocation immediately after implantation of the devices. At his 6

month visit he had a fistula in his right ear canal. At 10 months

postop the fistula (late infection) was still present along with

significant ankylosis and adhesions and both total joints were

removed. This case is now lost to follow-up.

- Fossa component only revised/removed

- Case # 1

Bilateral patient had removal of left fossa component due to aseptic necrosis at almost 2 years postoperative and 3.5 years after the removal had it replaced.

- Case # 13

Fossa component was removed secondary to infection in the ear canal at 2.5 years postoperative and was replaced a month later.

- Case # 19

Fossa component removed 4 years postoperative due to a late infection of the ear.

- Case # 117

At 11 months postoperative the fossa was removed to see if this would decrease swelling. There were no signs of infection but heavy encapsulation was noted.

c.   Mandibular component only revised

- Case # 183

This bilateral case was treated for an anterior dislocation by removing the right 50mm mandibular component and replacing it with a 45mm component.

3.  	Additional Safety Measurement

- Surgical Site (wound healing)

Most surgical wounds healed by 3 months postoperative with 99 %

(right side) and 98 % (left side) healed.  Redness and drainage

accompanied with infection are documented as adverse events.

XII.	Conclusions Drawn from Studies

The pre-clinical and clinical data provide reasonable assurance of the safety and effectiveness of the Total TMJ Replacement System for the stated indications.