Load Lock Gun Meeting

December 14, 2004

Attending: Phil Adderley, Joe Grames, Danny Machie, Matt Poelker

- Danny reported that the preparation chamber has been received from Kurt Lesker. Phil wants to clean with hot water, purge with N2 and then leak check the chamber.  [After meeting, Phil reported pumpdown to 300 mA ion pump current in 1 hour following the hot water cleaning.]

- Phil reported that Jim cleaned the flanges of the 2nd high voltage chamber yesterday. Phil reported that Ralph cleaned chamber in chemical lab. Once complete it will be electropolished, then vacuum fired. The flanges will also be vacuum fired.

- Phil noted that we should use vacuum melt stainless steel for the electrodes so that polishing does not reveal hidden defects.  He said that we have 2 proud anodes and will need a 2nd “tee”-electrode assembly for the 2nd gun.

- Phil reported that we have on 4.5” all metal valve and need one 2.75” all metal valve. Phil subsequently corresponded,

“Hello Danny, we ordered the wrong compliment of Vat valves. I have in my possession an All metal 2.75" Flanged / 1.5" bore valve for the new Load lock prep to H.V.  We need 2@  2.5" bore / 4.5" flanged manual valves for the new load lock Gun. One for the suitcase and one for the Prep chamber. Can we get procurement to edit the Req / and P.O. ?? If not we will need to order another 2.5" bore All metal valve for the suitcase. Thanks. Phil.”

- Danny reported that the manipulators are due on Dec. 23rd.

- Danny reported that the springs for the stalk heater are not done. The original material we delivered for fabrication was not sufficient, so we sent different material.

- Danny suggested that we need a new gun table.  Following discussion we decide that we can modify that table in the EEL lab that presently supports the 1st high voltage chamber.

- Danny reported that the 2nd activation chamber is due Jan. 7th.

- Danny reported that the bellows for the manipulators are due next week. All except one passed leak checking and the remaining bellows is being remade.

- We discussed the spring on the rotational manipulator. The spring was damaged in BTLLPEG because it conflicted with the copper fingers. After Phil repaired the spring in BTLLPEG, and since we are careful not to press the copper fingers against the spring, the spring has worked.  We all decided it is important to test the originally chosen spring method further before pursuing a different method.

- We discussed the rotation of the puck in the high voltage electrode. Danny will think about a puck with a flat or notched surface. He suggested we can test ideas with an stainless puck, rather than with a costly moly puck.

- We discussed the suitcase design. We considered the following:

  - A baked docking chamber will be attached to the preparation chamber by a 4.5” all metal valve.

  - The baked docking chamber will have a bellows on the chamber side of the valve for alignment.

  - The baked docking chamber will have an ion pump with a turbo pump-out port.

  - When not in use the opposing side of the docking chamber will hae a 4.5” blank.

  - The suitcase chamber will have a 4.5” all metal valve which would connect to the docking chamber at the blank location.

  - The suitcase chamber will have a manipulator which can suitable reach the center of the preparation chamber.

  - The suitcase chamber will have a means protecting the manipulator such that it cannot be bend.

  - The suitcase chamber will have handles so that two people may easily move it.

  - The suitcase manipulator will have a puck carrier suitable for four pucks, such that the pucks do not readily fall out of the carrier.  Danny will consider a suitable carrier.  Joe proposed the idea of a “deep” puck carrier.

  - A baked load chamber will exist permanently in the EEL lab.

  - The load chamber will have a window on one end and the opposing will have a 4.5” blank.

  - The suitcase 4.5” all-metal valve will mate to the load chamber at the location of the blank.

  - The load chamber will have an RGA and ion pump with a turbo pump-out port.

  - The load chamber will have a support to transfer 4 pucks to the suitcase and 4 pucks from the suitcase without venting the chambers.

  - The support will ideally provide heating to four pucks simultaneously.

    - We discussed how to locate the puck, specifically the face of the photocathode, into the cathode electrode. No solution was reached.