Date: 23 June, 2003

FROM:	Malaysian Rubber Glove Manufacturers Association (MARGMA)

A2-20 PJ Industrial Park,

Jalan Kemajuan, Sec.13,

46200 Petaling Jaya

## Selangor, Malaysia

Tel: 603-7957 8362

E-mail: margma@tm.net.my

TO: 	Dockets Management Branch (HFA-305),

Food and Drug Administration,

5630 Fishers Lane, Rm. 1061,

Rockville, MD 20852.

Reference : Docket:  No. 00D-1384.

Subject: Import Alert IA 80-04

The Malaysian Rubber Glove Manufacturers Association (MARGMA) is an association of 52 manufacturers of medical gloves. See Attachment 1 for a list of current glove manufacturers who are members of MARGMA.

MARGMA manufacturers shipped about 1.2 billion pairs surgeons’ gloves and 8.7 billion pairs examination gloves to the Unites States in year 2002.

The following comments are submitted by MARGMA, representing 52 manufacturers, and were developed with the collaboration of our US Regulatory Affairs Consultant, Mr. Andrew Lowery.

MARGMA notes that the present or proposed 800.20 is implemented with 21 CFR Part 820 GMP/QS and Import Alert #80-04. Therefore, MARGMA respectfully requests that FDA review the current Import Alert #80-04 guidance, “Surveillance and Detention without Physical Examination” [DWPE] together with the new proposed rule identified immediately below:

800.20 Medical Devices; Patient Examination and Surgeons’ Gloves;  Test Procedures and Acceptance Criteria.

Reference: Docket No. 03N-0056.

Federal Register: March 31, 2003 (Volume 68, Number 61) (Proposed Rules) (Page 15404-15417) 21 CFR Part 800

Sincerely Yours,

President

Malaysian Rubber Glove Manufacturers' Association

COMMENTS ON IA 80-04

MARGMA has submitted two proposed changes/approaches for Level 3:

one, delete Level 3 and enhance Level 2; and,

two, modify Level 3.

If the first one is not accepted, then MARGMA respectfully requests that FDA consider the second request.

GENERAL COMMENTS

FDA Review IA 80-04 with 800.20

In view of the lowering of AQL levels for both surgeon and examination glove, MARGMA wishes to propose that FDA review the current Import Alert #80-04 guidance, "Surveillance and Detention without Physical Examination" together with the new proposed rule 800.20.

Rationale: 21 CFR 800.20 is not a standalone regulation. It is implemented in conjunction with IA 80-04. Therefore, in order to be synergistic, avoid conflicts, and not be overly burdensome, the two documents should be reviewed together.

FDA Inform Manufacturers of Sampling

MARGMA requests that FDA make appropriate changes in IA 80-04 and associated import and analyst forms so that manufacturer contact information is collected. MARGMA manufacturers request that FDA, in a timely manner, inform them by letter, Fax, or email when one of their shipments is detained and inform them of the results of the sampling and testing.

Rationale: Manufacturers wish to be informed/notified directly by the FDA whenever one of their containers is held for sampling and examination; and, we wish to be directly informed of the outcome of the visual inspection and water leak testing. Currently, often after significant delays, MARGMA manufacturers learn about the detention of their gloves from their customers or shippers or sometimes in the OASIS webpage. This delay, in turn, delays initiation of our CAPA activities.

IA 80-04 Accept Enhanced Test Passes

MARGMA requests that reconditioned gloves that do not undergo physical reworking and that pass the testing under the tightened sampling plan be allowed to be distributed as passed; and, the passed lot be classified as a passed lot under IA 80-04.

Rationale: The gloves have passed enhanced sampling and testing and have enhanced assurance of meeting quality requirements. Also, this change would make IA 80-04 consistent with the Food, Drug, and Cosmetic Act.

Classification Versus 510(k)

When a manufacturer has one type of examination glove detained, all of their various examination gloves are detained because detention is presently performed based on one classification regulation for all examination gloves. The same is true for surgeons’ gloves.  MARGMA requests that detention be applied per glove 510(K) instead of per glove classification regulation.

Rationale: Detention should be applied per glove 510(K) instead of per classification regulation. Gloves are not manufactured per a classification regulation. Gloves are manufactured according to the specifications and information in the 510(k) submission.

For example, under the classification approach, a detention of Nitrile powdered gloves results in NRL powder-free gloves being detained. Nitrile and NRL gloves are made with:

- different materials,

- different formulations,

- different process parameters, and

- usually with different equipment.

A failure with one glove type usually has little or no relationship to the quality of another glove type with a different 510(k) clearance.

Detention by classification regulation instead of by glove 510(k) results in extremely unfair punishment to glove manufacturers. After a failure for one type of glove, this classification approach can force a manufacturer into bankruptcy and/or out of medical glove production when all of their other various types of medical gloves meet all regulatory and quality requirements. MARGMA is not aware of FDA routinely applying the same “per classification treatment” to manufacturers of other medical devices.

IA 80-04 Incompatible with QS/GMP

IA 80-04 Time Element

IA 80-04 is based on quality-related events and time. The GMP is based only on quality-related events. MARGMA recommends that IA 80-04 be modified to be more related to quality events rather than time. The addition of the reconditioning enhanced sampling and testing program is a step in the right direction because it is a quality event-related program. Likewise the 5 and 10 good shipments criteria are quality events. MARGMA also requests that the detention probation period be reduced from 24 to 12 months.

Rationale: IA 80-04 contains time elements whereas the GMP is based on events. The time element makes IA 80-04 a very complex policy and very punitive for medical glove manufacturers. Changing IA 80-04 to be based more on quality events and less related to time would make it more compatible with the GMP and less punitive.

IA 80-04 Has Late CAPA; GMP Requires Early CAPA

IA 80-04 needs to emphasize CAPA at Level 1 rather than at Levels 2 and 3.

Rationale: The emphasis on corrective action in IA 80-04 occurs at Level 2 and 3.

CAPA in the GMP regulation emphasizes corrective action when the manufacturer becomes aware of a problem which would be Level 1.

COMMENT BY TOPIC

The REASON FOR REISSUANCE states that “A new procedure was developed for detention without physical examination of gloves manufactured by firms who repeatedly import non-compliant products.”

In contrast, the NOTE under Guidance states that, “Because the presence of defects in medical gloves represents a potential hazard to health, only one (1) violative sample is needed to submit a recommendation for detention without physical examination to the Division of Import Operations & Policy (DIOP).

Recommendation: The REASON FOR REISSUANCE and the NOTE under Guidance need to be rewritten so that they become compatible statements.

Rationale: “Repeatedly import non-compliant products”  and “only one (1) violative sample is needed to submit a recommendation for detention”  are incompatible statements. That is, one violation is not repeated violations.

Level 1 Detention(INDICATED BY * IN ATTACHMENT A)

We recommend that the above heading be changed to:

‘Level 1 DWPE and Probation’(INDICATED BY * IN ATTACHMENT A).

[Where DWPE is detention without physical examination]

MARGMA recommends that this section be re-written as identified by the following underlined text:

When a district encounters a shipment of medical gloves which is found violative for defects by FDA analysis, the shipment should be detained and a recommendation for detention without physical examination (DWPE) should be forwarded to DIOP, HFC-170.

If the recommendation meets current guidance, the manufacturer/shipper will be placed on Level 1 DWPE for the specific type of glove and 510(k), attachment "A" with a single asterisk(*), and placed on 12 months probation. The 12 months of probation starts on the date the firm was placed on DWPE. Subsequent medical glove shipments from that manufacturer/shipper may be detained without physical examination under this import alert.  The manufacturer/ shipper may obtain admission of subsequently detained shipments of medical gloves by presenting evidence that the individual shipments are not adulterated, such as sample analyses performed by an independent testing laboratory, following the sampling plan and test method contained in FDA current guidance or current regulation.

In order for the responsible manufacturer's/shipper's name to be removed from Level 1 DWPE and attachment "A," documentation should be provided with sufficient evidence that their specific type of medical glove as identified by a 510(k) number is not adulterated (for example, five consecutive non-violative shipments analyzed as described in the preceding paragraph may be considered adequate evidence for manufacturers/shippers to be removed from Level 1 DWPE for the specific type [specific 510(k)] of glove. The manufacturer/shipper remains on Level 1 probation for 12 months).

Rationale:

IA 80-04 needs to make it clear that the end of Level 1 or 2 DWPE is not the same as the end of Level 1 or 2 “probation.” IA 80-04 is not clear with respect to the end of Level 1 DWPE after 5 consecutive good shipments and the end of Level 1 “probation” after 24 months with no entry failures. For example, IA 80-04 states:

In order for the responsible manufacturer's/shipper's name to be removed from attachment "A," documentation should be provided with sufficient evidence that their medical gloves are not adulterated (for example, five consecutive non-violative shipments analyzed as described in the preceding paragraph may be considered adequate evidence for manufacturers/shippers to be removed from Level 1 detention).

This statement appears to state that 5 good shipments removes the firm from Level 1 detention when the firm is actually on ‘Level 1 detention-probation’ for 24 months.

IA 80-04 and the GMP/QS Regulation

We further recommend that the following be added to the Level 1 section:

Without regard to the status of the 5 consecutive good shipments, during the 12 months probation FDA may ask the manufacturer to submit to CDRH an outline and schedule of the CAPA activities proposed or underway per 21 CFR 820.100 that are intended to identify and correct the cause of the deficiency. The cover letter or document to FDA shall state that the manager with executive responsibility is aware of the deficiency or deficiencies and CAPA activities. FDA may or may not reply to the manufacturers documents.

Rationale: IA 80-04 is not consistent with 21 CFR §820.100 of the GMP/QS regulation. The corrective actions in the present IA 80-04 are too late – corrective action in §820.100 CAPA are intended to start as soon as the manufacturer becomes aware of a quality or system problem. The change recommended above is intended to help make IA 80-04 more consistent with the GMP regulation.

Level 2 Detention

We recommend that the heading be changed to: Level 2 DWPE and Probation

We recommend that this section be re-written for the same reasons (rationale) given above for Level 1. The revised text is as follows:

Level 2 DWPE and Probation --  (Recidivist firms which are currently on Level 1 DWPE, or which have been removed from Level 1 DWPE) (INDICATED BY **  IN ATTACHMENT A)

If a manufacturer/shipper, while on Level 1 DWPE, has a violative sample analyzed by a private laboratory, the district should notify DIOP and submit supporting documentation. DIOP will place the firm on Level 2 DWPE and 12 months probation for the one or more specific violative glove types with specific 510(k) numbers and identify the manufacturer/shipper on attachment "A" with two asterisks (** ). The manufacturer/shipper goes to Level 2 DWPE if the glove type is the same or different from the glove type that triggered the Level 1 status. Level 2 is based on the sum of failures of all glove types; for example, 1failure of textured pink powdered NRL, and 1 failure of purple powder-free Nitrile results in Level 2 status. However, Level 2 DWPE does not apply to gloves types that have not failed [are not violative].

Similarly, if a manufacturer/shipper that has been removed from Level 1 DWPE has a violative sample analyzed by FDA within 12 months from the date they were removed from Level 1 DWPE, the district should notify DIOP and submit supporting documentation. DIOP will verify the manufacturer/shipper was listed under Level 1 detention probation during the past 12 months and if confirmed will place the manufacturer/shipper on Level 2 DWPE and 12 months probation and identify the manufacturer/shipper on attachment "A" with two asterisks(** ).

DIOP will bring this action to the attention of CDRH.

Delete the paragraph below because CAPA is started under Level 1 and continues under Level 2.  CAPA is more detailed and, therefore, is more effective than this following paragraph.

Based on this second listing within 24 months, CDRH has agreed to notify the foreign firm in writing of FDA's concerns about potential deficiencies in the manufacturing practices and process controls which may be affecting the quality of the medical gloves shipped to the U.S. A copy of the Good Manufacturing Practices (GMP) Regulations will be provided with the letter for their information.

After a manufacturer/shipper is placed on Level 2 DWPE and placed on 12 months probation, FDA may need greater assurance that the medical gloves are not adulterated before removing the manufacturer/shipper from DWPE status. For example, 10 consecutive non-violative shipments for each type [each 510(k)] of glove involved, analyzed by an independent testing laboratory and CAPA information as stated below, may be considered adequate evidence that the manufacturer/ shipper is shipping medical gloves to the U.S. which are not adulterated. The firm remains on detention probation for the remainder of the 12 months.

FDA will ask the manufacturer to submit an outline and schedule of the CAPA activities proposed or underway per 21 CFR 820.100 that are intended to identify and correct the cause of the deficiency or deficiencies; and, FDA will ask at the same time that follow-up CAPA reports and summaries of the data for each lot of the subject glove types be submitted within 6 months. Each cover letter or document to FDA shall state that the manager with executive responsibility is aware of the deficiency or deficiencies and CAPA activities. FDA may or may not reply to the manufacturers document submissions.

Additional and subsequent CAPA data and other types of evidence to remove the appearance of violations may be requested FDA and will be evaluated by CDRH on a case-by-case basis.

Manufacturers may submit additional and subsequent CAPA data and QS improvement information to FDA and request removal from Level 2 at the end of 12 months on a case-by-case basis for low volume specialty gloves or the specific glove type where it would take more than 12 months to accumulate 10 shipments.

MARGMA requests that Level 3 be removed from IA 80-04 and our rationale follows. MARGMA proposes that the manufacturers/importers in default remain under detention Level 2, regardless of additional failures, until the manufacturers/ importers corrects their glove quality  problems. All shipments of the violative types undergo enhanced sampling and testing by private independent labs or by FDA.   [However, if FDA elects to retain Level 3, MARGMA requests that Level 3 be changed as noted in the next section with underlined text.]

Rationale for removal of Level 3 from IA 80-04:  See comments to docket No. 00D-1384.  MARGMA had observed and commented on the pragmatism of current recidivist policy dated 20 Oct 2000, docket No. 00D-1384. The policy supports the current QS regulation, specifically the §820.100 CAPA requirements. When FDA has placed manufacturers/importers on detention, it is the manufacturer/importers responsibility to provide greater assurance that they meet the required AQL. MARGMA proposes that the manufacturers/importers in default remain under detention Level 2 DWPE, regardless of additional failures, until the manufacturers/ importers correct their glove quality problems. Then the manufacturer/shipper requests FDA to remove them from DWPE. The request must be supported by evidence as set forth in the revised IA 80-04 or in FDA guidance. The recidivist policy should not be punitive in nature but corrective in practice, allowing the manufacturers/importers the opportunity to rectify their problems immediately, especially at the early stages of violation. In all instances American consumers are fully protected during the detention period because 100% of the shipments of the violative types are evaluated by enhanced sampling and testing by private independent laboratories before release/distribution of the gloves into the US market.

In addition, medical gloves are now commodity products manufactured by a mature industry; therefore, profit margins are small. Commodity product manufacturers cannot sustain the cost of continuous testing by private labs and the disrupted flow of product to their clients. Therefore, they must immediately solve their quality problems and get off DWPE or face dire consequences.

If FDA elects to retain Level 3, MARGMA requests that Level 3 be changed as noted below with underlined text:

Level  3 No Entry for Failed Types -- (Recidivist firms which are currently on Level 2 DWPE and probation, or which have been removed from Level 2 DWPE but still on probation) (** * IN ATTACHMENT A DENOTES A WARNING LETTER IS BEING CONSIDERED BY CDRH). Level 3 is based on the sum of failures of all glove types; for example, 1failure of blue powdered NRL, 1 failure of buff textured powder-free polymer-coated NRL and 1 failure of green powder-free Nitrile results in Level 3 status. However, Level 3 entry denial does not apply to glove types with different 510(k)s where such glove types have not failed.

If a manufacturer/shipper, while on Level 2 DWPE or Probation, has another violative sample analyzed by an independent testing laboratory, the district should notify DIOP and submit supporting documentation.

Similarly, if a manufacturer/shipper removed from Level 2 DWPE has another violative sample analyzed by FDA within 12 months from the date they were removed from Level 1 DWPE, or the date they were placed on Level 2 detention, DIOP will bring these instances to the attention of CDRH.

Based on the failure of the manufacturer's/shipper's medical gloves to pass FDA analysis [Section 501(c)] after being listed on Level 2 detention, or having failed an independent testing laboratory examination while under Level 2 detention, CDRH may then issue a Warning Letter to the firm.

Additionally, after reviewing the manufacturer/shipper's export and inspectional history, CDRH may elect to charge the firm for failure to conform with GMP’s [Section 501(h)], based on the recurring GMP deficiencies as revealed by gross failures (detailed below) or the repeated failures of analyses.

If CDRH determines a warning letter should issue to the foreign manufacturer/shipper, DIOP will place the manufacturer/shipper on Level 3 detention (attachment B) for the specific glove types cleared under their specific 510(k)s. FDA will deny entry for the specific violative glove types cleared under their specific 510(k)s.

Non-violative analyses may not be sufficient to remove the appearance of a violation for manufacturers/shippers on attachment B for the specific glove types cleared under the specific 510(k)s.  Manufacturers/shippers of these specific glove types will remain on Level 3 detention until evidence is provided to CDRH which demonstrates that the specific medical glove types are being manufactured in accordance with Good Manufacturing Practices (for example, an acceptable FDA inspection or a written certification from the manufacturer/shipper together with the results of an independent audit performed by a qualified third party).

After a foreign manufacturer/shipper satisfactorily demonstrates that the apparent GMP deviations have been corrected, the manufacturer/shipper will be removed from attachment "B" and placed on attachment "A" under the Level 1 detention for individual shipment analysis and document submissions to confirm their specific medical glove types, cleared under specific 510(k)s, do not contain defects until they provide adequate evidence to be removed from Level 1 (which see above) DWPE for each type of glove. The manufacturer/shipper remains on detention probation for 12 months from the date they were placed on DWPE

Grossly Defective Gloves

MARGMA recommends that a manufacturer, without regard to previous detention status, be placed on Level 3 [or, Level 2 if there is no level 3] if a shipment of a specific type of glove is grossly violative by visual inspection or water leak testing. However, Level 3 entry denial does not apply to glove types with different 510(k)s where such glove types have not failed.

Grossly defective gloves usually can be visually identified. Examples of grossly violative glove conditions are:

- Large numbers or package of fused/blocked of gloves;

- Blocked package or large numbers of gloves that cannot be donned;

- Gloves that have disintegrated into flakes or chips; and

- Large number of gloves that tear when correctly donned.

[A large number can be a number greater than the largest failure number for the lot size as established by FDA in the Preamble to 800.20 or in the IA 80-04 Policy.]

Rationale:  Such manufacturers need to assure by performing basic life testing that each type of glove that they produce meet user needs when delivered. Timely corrective action [CAPA] may be delayed by waiting for the manufacturer to regress through Level 2 and on to Level 3. FDA needs to be able under IA 80-04, along with other remedies, to immediately see that CAPA is implemented and that incoming shipments of the specific glove type are evaluated by enhanced sampling and testing by a private lab [or not allowed entry for the specific type of glove depending on the requirements in the next version IA 80-04].

Copy of IA 80-04 for reference.

IA #80-04

REVISED 9/11/97, IMPORT ALERT #80-04, "SURVEILLANCE AND DETENTION WITHOUT PHYSICAL EXAMINATION OF SURGEON'S AND OR PATIENT EXAMINATION GLOVES

ATTACHMENT A - 3/15/01,  ATTACHMENT B - 8/1/00

(NOTE:  This import alert represents the Agency's current guidance to FDA field personnel regarding the manufacturer(s) and/or product(s) at issue.  It does not create or confer any rights for or on any person, and does not operate to bind FDA or the public.

REASON FOR REISSUANCE: A new procedure was developed for detention without physical examination of gloves manufactured by firms who repeatedly import non-compliant products.

TYPE OF ALERT:  Surveillance (Physical Examination), and Detention Without Physical Examination

PRODUCT:  Surgeon's and Patient Examination Gloves

PRODUCT CODE:  80LZA, 80FMC, 80LYY, 80LZC, 80LYZ, 79KGO

PAC:  82Z003

PROBLEM:  Defective per 21 CFR 800.20 - Leaks, tears, mold, embedded foreign objects

COUNTRY:  Surveillance - All Countries - (ZZ) Detention Without Physical Examination - See Attachments (A & B)

MANUFACTURER/ SHIPPER:  See Attachments (A & B)

MANUFACTURER/ SHIPPER ID#:  N/A

IMPORTER'S ID #:  N/A

CHARGE:  "The article is subject to refusal of admission pursuant to section 801(a)(3) in that it appears to contain defects/holes [Adulteration, section 501(c).]"

When FDA has documented repeated violations, and the manufacturer/shipper has been issued a Warning Letter and listed in Attachment B of the Recidivist Policy described below, future shipments may be detained charging the following:

"The article is subject to refusal of admission pursuant to section 801(a)(1) in that it appears that the methods used in, or the facilities and controls used for, the manufacture, packaging, storage, or installation of the device do not conform with Good Manufacturing Practices [Adulteration, section 501(h), Good Manufacturing Practices, Section 520(f).]"

RECOMMENDING OFFICE:  Center for Devices and Radiological Health (CDRH)

REASON FOR ALERT:  CDRH has determined that many foreign manufacturers and shippers of medical gloves have failed to consistently provide medical gloves of adequate quality for distribution in the U. S.  Surgeon's and examination gloves are widely used by health care professionals and others as a barrier to blood-borne diseases and pathogens.  Defective medical gloves present a potential hazard to health for users as well as patients.  Therefore, continuous monitoring of these medical gloves is needed.

GUIDANCE: Districts may detain without physical examination, all shipments of surgeon's and patient examination gloves from the manufacturers/shippers listed on the attachments to this alert.

Surveillance sampling of other manufacturers/shipper's surgeon's and patient examination gloves per the most recent guidance is indicated.

NOTE: Because the presence of defects in medical gloves represents a potential hazard to health, only one (1) violative sample is needed to submit a recommendation for detention without physical examination to the Division of Import Operations & Policy (DIOP).  DIOP has direct reference authority for detention without physical examination of these medical gloves.  Violative medical gloves are subject to RPM Chapter 9, Detention Without Physical Examination, subchapter "Recommendations Based on One Sample."

The following strategy has been established to address those manufacturers/shippers who are found to repeatedly ship violative medical gloves to the U.S. (recidivist firms.)

Level 1  Detention (INDICATED BY * IN ATTACHMENT A)

When a district encounters a shipment of medical gloves which is found violative for defects by FDA analysis, the shipment should be detained and a recommendation for detention without physical examination should be forwarded to DIOP, HFC-170.

If the recommendation meets current guidance, the manufacturer/shipper will be placed on attachment "A" with a single asterisk(*), and subsequent medical glove shipments from that manufacturer/shipper may be detained without physical examination under this import alert.  This is referred to as "Level 1 detention."  The manufacturer/ shipper may obtain admission of subsequently detained shipments of medical gloves by presenting evidence that the individual shipments are not adulterated, such as sample analyses performed by an independent testing laboratory, following the sampling plan and test method contained in FDA guidance.

In order for the responsible manufacturer's/shipper's name to be removed from attachment "A," documentation should be provided with sufficient evidence that their medical gloves are not adulterated (for example, five consecutive non-violative shipments analyzed as described in the preceding paragraph may be considered adequate evidence for manufacturers/shippers to be removed from Level 1 detention).

Level  2  Detention  --  (Recidivist firms which are currently on Level 1 detention, or which have been removed from Level 1 detention) (INDICATED BY **  IN ATTACHMENT A)

If a manufacturer/shipper, while on Level 1 detention, has a violative sample analyzed by a private laboratory, the district should notify DIOP and submit supporting documentation. DIOP will place the firm on Level 2 detention and identify the manufacturer/shipper on attachment "A" with two asterisks (** ).

Similarly, if a manufacturer/shipper that has been removed from Level 1 detention has a violative sample analyzed by FDA within 24 months from the date they were removed from Level 1 detention, the district should notify DIOP and submit supporting documentation.  DIOP will verify the manufacturer/shipper was listed under Level 1 detention during the past 24 months and if confirmed will place the manufacturer/shipper on Level 2 detention and identify the manufacturer/shipper on attachment "A" with two asterisks(** ).

DIOP will bring this action to the attention of CDRH.

Based on this second listing within 24 months, CDRH has agreed to notify the foreign firm in writing of FDA's concerns about potential deficiencies in the manufacturing practices and process controls which may be affecting the quality of the medical gloves shipped to the U.S. A copy of the Good Manufacturing Practices (GMP) Regulations will be provided with the letter for their information.

Once a manufacturer/shipper is placed on Level 2 detention, FDA may need greater assurance that the medical gloves are not adulterated before removing the manufacturer/shipper from detention without physical examination status.  For example, 10 consecutive non-violative shipments, analyzed by an independent testing laboratory, may be considered adequate evidence that the manufacturer/shipper is shipping medical gloves to the U.S. which are not adulterated.  Other types of evidence to remove the appearance of a violation will be evaluated by CDRH on a case-by-case basis.

Level  3  Detention  -- (Recidivist firms which are currently on Level 2 detention, or which have been removed from Level 2 detention) (** * IN ATTACHMENT A DENOTES A WARNING LETTER IS BEING CONSIDERED BY CDRH)

If a manufacturer/shipper, while on Level 2 detention, has another violative sample analyzed by an independent testing laboratory, the district should notify DIOP and submit supporting documentation.

Similarly, if a manufacturer/shipper removed from Level 2 detention has another violative sample analyzed by FDA within 24 months from the date they were removed from Level

1 detention, or the date they were placed on Level 2 detention, DIOP will bring these instances to the attention of CDRH.

Based on the failure of the manufacturer's/shipper's medical gloves to pass FDA analysis [Section 501(c)] after being listed on Level 2 detention, or having failed an independent testing laboratory examination while under Level 2 detention, CDRH may then issue a Warning Letter to the firm.

Additionally, after reviewing the manufacturer/shipper's export and inspectional history, CDRH may elect to charge the firm for failure to conform with GMP’s [Section 501(h)], based on the recurring GMP deficiencies as revealed by the repeated failures of analyses.

If CDRH determines a warning letter should issue to the foreign manufacturer/shipper, DIOP will place the manufacturer/shipper on Level 3 detention (attachment B).

Non-violative analyses may not be sufficient to remove the appearance of a violation for manufacturers/shippers on attachment B.  Manufacturers/shippers will remain on Level 3 detention until evidence is provided to CDRH which demonstrates that the medical gloves are being manufactured in accordance with Good Manufacturing Practices (for example, an acceptable FDA inspection or a written certification from the manufacturer/shipper together with the results of an independent audit performed by a qualified third party).

Once a foreign manufacturer/shipper satisfactorily demonstrates that the apparent GMP deviations have been corrected, the manufacturer/shipper will be removed from attachment "B" and placed on attachment "A" under the Level 1 detention for individual shipment analysis to confirm their medical gloves do not contain defects until they provide adequate evidence to be removed from Level 1 detention (see above).

For questions or issues concerning science, science policy, sample collection, analysis, preparation, or analytical methodology, contact the Division of Field Science at (301)

443-3320 or 3007.

For questions concerning 21 CFR 800.20, CPG 7124.31, medical glove labeling, or other compliance issues, contact CDRH, Office of Compliance, Division of Enforcement II at (301)

594-4618.

SAMPLING:  See 21 CFR 800.20(c) for sample plans.

All sizes should be represented as closely as feasible to the proportion which they exist in the shipment, however, exact statistical representation is not necessary.  If the sample is found violative, all sizes should be detained.

When an entry consists of only one size, attempt to collect as many lot numbers as possible.  For example, if during a random sample collection three lot numbers are observed, represent them all as subs within one sample. If the sample is found violative, all lots should be detained.

If a shipment is detained based on the analysis of one lot and subsequently it is learned that there were 10 lot numbers present, the detention stands.

ATTACHMENT 1:

LIST OF MEMBER COMPANIES

[ORDINARY (Manufacturing) MEMBERS]

MALAYSIAN RUBBER GLOVE

MANUFACTURERS’ ASSOCIATION

(MARGMA)

| - Alliance Rubber Products Sdn Bhd - Ansell (Kedah) Sdn Bhd - APL Healthcare Sdn Bhd - Apollo Rubber Sdn Bhd - Brightway Holdings Sdn Bhd - Comfort Rubber Gloves Industries Sdn Bhd - Concept Rubber Products Sdn Bhd - Contract Latex Dippers Sdn Bhd - Cranberry (M) Sdn Bhd - FELDA Rubber Products Sdn Bhd - Flexitech Sdn Bhd - GB Industries Sdn Bhd - Glovco (M) Sdn Bhd - Guthrie Medicare Products (NS) Sdn Bhd - Handsafe Products Sdn Bhd - Hartalega Sdn Bhd - IGA Overseas Sdn Bhd - JB Star Sdn Bhd - KL-Kepong Rubber Products Sdn Bhd - Koon Seng Sdn Bhd - Kossan Latex Industries (M) Sdn Bhd - Latexx Manufacturing Sdn Bhd - Longcane Industries Sdn Bhd - MAPA (gloves) Sdn Bhd - Marcon Rubber Industry Sdn Bhd - N.S. Uni-Gloves Sdn Bhd | - Oon Corp Resources (M) Sdn Bhd - Pan Century Rubber Products Sdn Bhd - Perusahaan Getah Asas Sdn Bhd - Perusahaan Pelindung Getah (M) Sdn Bhd - Protection Gloves (M) Sdn Bhd - Purnabina Sdn Bhd - Quality Latex Products Malaysia Sdn Bhd - Riverstone Resources Sdn Bhd - Sanchem Corporation Sdn Bhd - Seal Polymer Industries Sdn Bhd - Seltom Pacific Sdn Bhd - Smart Glove Corporation Sdn Bhd - Sri Johani Sdn Bhd - SSN Gloves (M) Sdn Bhd - Super Latex Sdn Bhd - Supermax Corporation Berhad - Tekmedic (M) Sdn Bhd - Terang Nusa Sdn Bhd - TG Medical (M) Sdn Bhd - Top Glove Corporation Berhad - Ultrawin Sdn Bhd - Wear Safe (Malaysia) Sdn Bhd - WRP Asia Pacifc Sdn Bhd - WRP Senetimed Sdn Bhd - YTY Industry Sdn Bhd - Bonric Sdn Bhd |
| --- | --- |