Module 11: Use and Care of Equipment at the HIV Rapid Testing Site

| Purpose | To help participants understand exactly what is expected of them concerning the care and use of equipment at the HIV rapid test site and to equip them with the necessary skills |
| --- | --- |
| Pre-requisite Modules | Module 5: Assuring the Quality of HIV Rapid Testing |
| Module Time | 55 minutes |
| Learning Objectives | At the end of this module, participants will be able to: Specify their responsibilities related to equipment at the HIV rapid test site Routinely monitor the temperatures of refrigerators or freezers Confirm auto pipettes deliver specified volumes Properly use and maintain centrifuges |

Module Overview

| Step | Time | Activity/ Method | Content | Resources Needed |
| --- | --- | --- | --- | --- |
| 1 | 5 min | Presentation | Module introduction | Slides 1-6 Prepared flipchart – content outline |
| 2 | 5 min | Presentation Discussion | Responsibilities concerning equipment | Slides 7-8 |
| 3 | 25 min | Presentation Activity Demonstration | Refrigerator & freezer Pipettes Centrifuge | Slides 9-23 Pipettes Centrifuge |
| 4 | 15 min | Activity | Create a Maintenance Activity List | Slide 24 Flipcharts |
| 5 | 5 min | Q&A | Summary | Slide 25 |

Material/Equipment Checklists:

- PowerPoint slides or transparencies

- Overhead projector or computer w/LCD projector

- Flipchart for content outline and small-group activity

- Handouts: Daily Temperature Check Chart, Generic Maintenance Form

- Precision and disposable pipettes – have several sets to pass around

- Centrifuge

Teaching Guide

| Slide Number | Teaching Points |
| --- | --- |
| Customization Notes | Modify this module based on the equipment your participants will be using at the test sites. For example, consider: What type of pipettes will they be using Whether a centrifuge will be available on site |
| Customization Notes | Also modify this module based on participants’ jobs and responsibilities. For lab management or supervisors, put emphasis on the oversight responsibilities. For lab technicians or testers, put emphasis on execution and following procedures. |
| 1 | Module 11: Use and Care of Equipment DISPLAY this slide before you begin the module. Make sure participants are aware of the transition into a new module. |
| 2 | The Lab Quality System REMIND participants that equipment is a component of the Lab Quality System. This component addresses: Equipment selection Equipment acquisition Equipment installation and initial calibration/validation Maintenance service and repair Troubleshooting Retiring equipment and disposition |
| 3 | Learning Objectives STATE the objectives on the slide. |
| 4 | Content Overview EXPLAIN the topics that will be covered in this module. |
| Flipchart | WRITE the content outline on a flipchart prior to training. REFER to it frequently to orient participants to where they are in the module. |
| 5 | Functioning Equipment is Vital to Quality Service EXPLAIN the rationale for properly maintaining equipment. Reliable results aid clinical diagnosis. Unreliable results may result in incorrect diagnosis and treatment of the patient. If equipment is properly maintained, it is less likely to breakdown before its next service and is less likely to perform inadequately due to lack of maintenance. |
| 6 | Equipment at HIV Rapid Testing Site STATE HIV Rapid Testing can be performed with minimal equipment. At a minimum, there are 4 pieces of equipment that may be used at the rapid testing site. EXPLAIN the use of each type of equipment. Refrigerators store reagents, kits, and quality control materials requiring refrigerator storage (such as Capillus). Freezers store specimens collected for EQA purposes prior to transport to a reference laboratory. Quality control specimens may also be stored in a freezer. Pipettes collect or transfer specimen to test device. Centrifuges separate cells from serum/plasma EMPHASIZE all equipment used at the testing site must be properly maintained. Using equipment that has not been properly maintained may compromise the quality of test results. |
| 7 | Management Responsibilities: Ensure Test Site Readiness EXPLAIN lab management is responsible for making sure the test site is ready to receive and install a new piece of equipment. This includes: Assigning responsibilities – this includes oversight of all lab equipment and individual responsibilities. Establishing inventory record – Each piece of equipment must have an inventory record. This record contains pertinent information such as make and model, maintenance and service record, and manufacturer contact information. Training the operators – Everyone using any piece of laboratory equipment must be properly trained. Training must include troubleshooting. |
| 8 | Your Responsibilities: Execute at Test Site EXPLAIN the responsibilities of the test site staff. DISCUSS the reason against using malfunctioning equipment. |
| 9 | Function Checks Verify that Equipment is Working Properly EXPLAIN function checks – they are activities performed periodically to ensure that: Equipment is working properly before use Equipment is properly maintained for peek performance PROVIDE examples of function checks at the HIV rapid test site. |
| 10 | Refrigerator and Freezer: Use and Care EXPLAIN the points on the slide. |
| 11 | Refrigerator & Freezer: Temperature Checks EXPLAIN this photo illustrates routine monitoring of temperatures of this refrigerator. It is good practice to attach the form for recording temps directly on the front of the refrigerator for easy access. Inserting it into a protector page will guard against tearing of paper. |
| 12 | Refrigerator & Freezer: Temperature Log EXPLAIN this form is an example of what should be used to record and monitor refrigerator/freezer temps. REFER participants to a sample temperature log in the participant manual. |
| 13 | Types of Pipettes EXPLAIN the difference between the two types of pipettes. EMPHASIZE never re-use disposable items. Doing so will cause cross contamination. |
| Activity 3 minutes | PASS around a few sets of precision and disposable pipettes for everyone to see and touch. |
| 14 | Pipette: Use and Care EXPLAIN the points on the slide. EXPAND the first bullet: Select a pipette with delivery range close to the volume required. For example, if 50 micro liters (ul) of specimen is required, use a 100 ul pipette. EMPHASIZE never lay the pipette on its side when liquid is in the tip – doing so will cause the specimen to flow into to the pipette shaft and damage the pipette. |
| 15 | Pipette: Use and Care (Cont’d) EXPLAIN the points on the slide. EXPAND the third bullet: Air bubbles in the tip can greatly reduce pipetting accuracy. If an air bubble is trapped within the tip during intake, do the following: Dispense the sample into the original vessel Check the tip immersion depth Pipette more slowly If an air bubble appears a second time, discard the tip and use a new one. EMPHASIZE the following safety practices. (REMIND participants of the safety module) Care should be taken to discard pipette tips in the appropriate container. Used pipette tips should not be found on the floor, as this poses a safety hazard. Never re-use a pipette tip, which causes cross contamination and will compromise patient results. A fresh tip should be used for each sample. |
| Customization Notes 16-17 | If precision pipettes are used at test sites, provide hands-on practice. |
| 16 | Precision Pipettes Require Performance Checks The analytical balance should have a scale of 0.1 to 0.0001 mg |
| 17 | Pipette: Steps for Checking Reproducibility EMPHASIZE performance checks include reproducibility and calibration. EXPLAIN the procedures for checking reproducibility. |
| 18 | Pipette: Troubleshooting EXPLAIN the points on the slide. |
| Customization Notes 19 | For participants without a lab background, or who have never seen a centrifuge, add the following steps before showing the next few slides: SHOW a centrifuge. DEMONSTRATE how it works. NAME all parts that will be mentioned in the following slides. |
| 19 | Centrifuge: Use and Care EXPLAIN the bullets on the slide. Always operate with the lids closed – Operating a centrifuge without the lid closed poses an unnecessary safety hazard. Balance contents before turning on. – For example, if there is only one sample to be centrifuged, a tube identical in size and volume must be placed in the rotor opposite the tube. Note: The rotor is the part of the centrifuge that holds the tubes and rotates during operation. Check for vibration – There may be several reasons why a centrifuge vibrates. When vibration occurs, you’ll need to: Stop operation of the centrifuge. Determine the cause of the noise or vibration. Correct immediately to prevent severe damage to the centrifuge or injury to the worker. Refer to the owner’s manual for possible causes aside from improper balancing Do not open until the lid until the rotor has come to a complete stop Keep lids on tubes when spinning – Do not take the tops off the tubes before spinning. Doing so will cause splashing and creating of aerosols from potentially infectious material |
| 20 | Centrifuges: Function Checks EXPLAIN separation activity is a function of both centrifugal force and timing. Proper balance, lubrication and rotor function are essential for proper centrifugation to occur. |
| 21 | Centrifuge: Routine Maintenance EXPLAIN how to clean the centrifuge: Clean interior daily with soap and water, wipe with a disinfectant Wipe spills using 10% bleach solution After cleaning, run the centrifuge at varying RPMs to check the braking mechanism and ensure a smooth gradual stop EXPLAIN what to do when noticing unusual noises: Stop operation of the centrifuge Follow manufacturer’s recommendation on activation and release of brakes Correct immediately to prevent severe damage to the centrifuge or injury to the worker MENTION brushes need to be inspected every 3-6 months and replaced according to manufacturer specifications. |
| 22 | Centrifuge Safety STATE the bullets on the slide. STATE simply turning the power off does not remove power to the centrifuge. |
| 23 | Keep a Log for All Maintenance Activities MENTION the need for documenting problems, corrective action, preventive maintenance, cleaning, and inspections. REFER participants to a sample, generic maintenance log in the participant manual. |
| 24 Activity 15 minutes | Exercise: Create a Maintenance Activity Log FOLLOW the procedure below when conducting the activity: INFORM participants of the activity. READ the instructions on the slide. PROVIDE a flipchart for each work group. ALLOW 10 minutes for the small group to work. DEBRIEF by asking volunteers to present their lists and solicit comments from other participants. |
| Customization Notes 24 | Alternative way to conduct this activity: If all the participants come from the same test site, or if you have time constraint, you may modify this activity by facilitating a large group discussion (instead of having small work groups) to create the activity list. |
| 25 | Summary ASK participants to answer the questions on the slide. ANSWER any questions participants may have. |