Guy J. Graening, P.E.

Brown and Caldwell

10540 White Rock Road, Suite 180

Rancho Cordova, CA 95670

Mr. Graening,

Here is the information you requested in your email of May 10, 2007 regarding the first batch of Yerington Teflon filters.

# Teflon Filter Handling and Filter Pack Preparation

Filter packs are assembled and disassembled in accordance with DRI SOP 2-112.2, Filter Pack Assembly.

All samples received at DRI for analysis are logged on the day that they are received following the procedure described in the DRI Shipping and Receiving SOP, 2-113.2.

Any unusual deposits or filter conditions should be noted on the sample list.  These conditions include, but are not limited to:

•	scratches or smudges

•	holes

•	wet spots

•	foreign particles (e.g., insects, metal shavings, hair, etc.)

Damage to the filters after receipt at DRI must be noted as such.  The conditions of the samples must be documented so that an accurate assessment of analytical quality can be made.

# Filter Analysis

### Gravimetric Analysis

Gravimetric analysis is performed in accordance with DRI SOP 2-114.2, as summarized below:

Unexposed and exposed Teflon-membrane filters are equilibrated at a temperature of 21.5 ± 1.5 °C and a relative humidity of 35 ± 5% for a minimum of 24 hours prior to weighing.  Weighing is performed on a Mettler MT-5 electro microbalance with ±0.001 mg sensitivity. The charge on each filter is neutralized by exposure to a polonium source for 30 seconds before the filter is placed on the balance pan. The balance is calibrated with a 200 mg Class S weight and the tare is set prior to weighing each batch of filters. After every 10 filters are weighed, the calibration and tare are re-checked. If the results of these performance tests deviate from specifications by more than ±5 mg, the balance is re-calibrated.

All initial filter weights are checked by an independent technician. Samples are re-weighed if these check-weights do not agree with the original weights within ±0.010 mg. At least 30% of the  exposed filter weights are checked by an independent technician. Samples are re-weighed if these check-weights do not agree with the original weights within ±0.015 mg. Pre- and post-weights, check weights, and re-weights (if required) are recorded on data sheets and are directly entered into a data base via an RS232 connection. All weights are entered by filter number into the DRI aerosol data base.

### Elements by XRF

Elemental analysis by energy dispersive x-ray fluorescence is performed in accordance with DRI SOP 2-209.2 as summarized below:

After gravimetric analysis, samples collected on the Teflon-membrane filters were analyzed by energy dispersive X-ray fluorescence (ED-XRF, PanAlytical Epsilon 5) for the following 53 elements: sodium (Na), magnesium (Mg), aluminum (Al), silicon (Si), phosphorus (P), sulfur (S), chlorine (Cl), potassium (K), calcium (Ca), scandium (Sc) titanium (Ti), vanadium (V), chromium (Cr), manganese (Mn), iron (Fe), cobalt (Co), nickel (Ni), copper (Cu), zinc (Zn), gallium (Ga), arsenic (As), selenium (Se), bromine (Br), rubidium (Rb), strontium (Sr), yttrium (Y), zirconium (Zr), niobium (Nb), molybdenum (Mo), palladium (Pd), silver (Ag), cadmium (Cd), indium (In), tin (Sn), antimony (Sb), cesium (Cs), barium (Ba), lanthanum (La), Cerium (Ce), samarium (Sm), europium (Eu), terbium (Tb), hafnium (Hf), tantalum (Ta), wolfram (W), iridium (Ir), gold (Au), mercury (Hg), thallium (Tl), lead (Pb), radium (Ra), thorium (Th), and uranium (U).

Calibration is performed using thin film standards from Micromatter Inc.  A multielement thin film standard is analyzed with each run to monitor for calibration drift and is used as the indicator for routine calibrations.

### Inorganic Ion  Analyses

Water soluble ions are obtained by extraction and analysis following DRI SOPs 2-109.5 and 2-203.5 as summarized below:

Water-soluble sulfate, is obtained by wetting the filter with 200 µl of ethanol and then extracting the filter in 10 ml of distilled-deionized water (DDW).  The extraction vials are capped and sonicated for 60 minutes, shaken for 60 minutes, then aged overnight to assure complete extraction of the deposited material in the solvent.  The ultrasonic bath water is monitored to prevent temperature increases from the dissipation of ultrasonic energy in the water.  After extraction, these solutions are stored under refrigeration prior to analysis.

### Ion Chromatographic Analysis for Chloride, Nitrate, and Sulfate

Water-soluble sulfate (SO4=) is measured with the Dionex 2020i (Sunnyvale, CA) ion chromatograph (IC). In IC, an ion-exchange column separates the sample ions in time for individual quantification by a conductivity detector.  Prior to detection, the column effluent enters a suppressor column where the chemical composition of the component is altered, resulting in a matrix of low conductivity.  The ions are identified by their elution/retention times and are quantified by the conductivity peak area. Approximately 2 ml of the filter extract are injected into the ion chromatograph. The resulting peaks are integrated and the peak integrals are converted to concentrations using calibration curves derived from solution standards.  The Dionex system for the analysis of Cl-, NO3-, and SO4= contains a guard column (AG4a column, Cat. No. #37042), an anion separator column (AS4a column, Cat. No. #37041) with a strong basic anion exchange resin, and an anion micro membrane suppressor column (250 ´ 6 mm ID) with a strong acid ion exchange resin. The anion eluent consists of sodium carbonate (Na2CO3) and sodium bicarbonate (NaHCO3) prepared in DDW.  The DDW is verified to have a conductivity of less than 1.8 ´ 10-5 ohm/cm prior to preparation of the eluent.  For quantitative determinations, the ion chromatograph is operated at a flow rate of 2.0 ml/min.

The primary standard solution containing NaCl, NaNO3, and (Na)2SO4 is prepared with reagent grade salts dried in an oven at 105 °C for one hour and then brought to room temperature in a desiccator.  These anhydrous salts are weighed to the nearest 0.10 mg on a routinely calibrated analytical balance under controlled temperature (~20 °C) and relative humidity (±30%) conditions.  These salts are diluted in precise volumes of DDW.  Calibration standards are prepared at least once each month by diluting the primary standard solution to concentrations covering the range expected in the filter extracts.  These standards are then stored in a refrigerator.  The calibration concentrations prepared are at 0.1, 0.2, 0.5, 1.0, and 2.0 mg/ml for each of the analysis species.

Calibration curves are performed weekly.  Chemical compounds are identified by matching the retention time of each peak in the unknown sample with the retention times of  peaks in the chromatograms of the standards. A DDW blank is analyzed after every 20 samples and a calibration standard is analyzed after every 10 samples.  These quality control checks verify the baseline and the calibration, respectively.  Environmental Research Associates (ERA, Arvada, CO) standards are used daily as an independent quality assurance (QA) check.  These standards (ERA Wastewater Nutrient and ERA Mineral WW) are traceable to NIST simulated rainwater standards.  If the values obtained for these standards do not coincide within a pre-specified uncertainty level (typically three standard deviations of the baseline level or ±5%), the samples between that standard and the previous calibration standards are re-analyzed.

After analysis, the printout for each sample in the batch is reviewed for the following:   1) proper operational settings, 2) correct peak shapes and integration windows, 3) peak overlaps, 4) correct background subtraction, and 5) quality control sample comparisons. When values for replicates differ by more than ±10% or values for standards differ by more than ±5%, samples before and after these quality control checks are designated for re-analysis in a subsequent batch.  Individual samples with unusual peak shapes, background subtractions, or deviations from standard operating parameters are also designated for re-analysis.

# Batch 01 Specific Results

Three Teflon filters (TEOM001,002 and 003) were received on 4/10/2007, these filters had DRI IDs, BCTT004, 010 and 011 respectively.  After final gravimetric and XRF analysis of these three filters it was determined that filters 004 and 010 were actually filters 001 and 002.  The initial weights for filters 001 and 002 were used in the determination of the mass concentrations for these two exposed filters.  Microscopic examination of the filters under an optical microscope confirmed the apparent switching of these filters.  All three filters were very lightly loaded, with filters 010 and 011 (TEOM 002 and 003) having loadings typical of a field blank.  The heaviest filter 004 (TEOM001) had only 68 micrograms of deposit, and during the login procedure this filter was flagged for having part of the Teflon membrane separated from the support ring.

XRF analysis of the three filters confirmed the gravimetric results with filters 011 and 010 showing essentially no elements above detection limits.  Filter 004 did have a few of the typical crustal elements above detection limits.  Wet chemical  analysis also showed filter 004 with sulfate above detection limits.

Please let me know if I may be of further assistance.

Regards,

Steven D. Kohl

Associate Research Scientist

Desert Research Institute