| No. of R76-1 | Country | Comments | Response of TC9/SC1 Secretariat |
| --- | --- | --- | --- |
|  |  |  | The abbreviations used in the following mean: P+ = Proposal accepted and considered in the 2CD P– = Proposal not accepted |
| General | AU | We have recently become very concerned regarding aspects of instrument design and functioning which have been approved by other authorities, but which we feel either clearly do not meet current requirements, or do not meet even more basic criteria of being suitable for use for trade (that is in some cases it may be difficult to provide an objective criteria in R76 which is not met – unfortunately some of the most important aspects in R76 are rather subjective). We feel that the emphasis of type approval evaluation has become too fixated on ‘performance’ aspects and that a return to careful consideration and enforcement of usage and suitability aspects is essential. | We assume that your concerns are expressed in your specific comments and proposals below. These have – to a great extent - been accepted and considered in the 2CD. |
| General | CECIP | Although CECIP has no vote we would like to inform you that we support the voting “yes” in respect to the 1 CD but ask for some changes as listed below. | OK |
| General | CA | Please Note: We have included comments that were not addressed in the CD-1, but were ultimately accepted by the Secretariat. This is simply to ensure they get carried forward to the next revision. Comments missed in the devlopment of CD-1 appear in ORANGE Canada’s New comments appear in GREEN And our continuing concerns appear in RED Canada is concerned that there is potential for confusion and inconsistency of application by having Annex B duplicate requirements from OIML D-11. We are not sure why this recommendation does not simply reference D-11? This problem will be exasperated when either R76 or D-11 is modified in the future. If a specific requirement is changed in one, what happens to the corresponding requirement in the other? Canada is concerned that there is potential for confusion and inconsistency of application by having Annex F duplicate requirements from OIML D-SW (correct reference?). We are not sure why this recommendation does not simply reference D-SW? This problem will be exasperated when either R76 or D-SW is modified in the future. If a specific requirement is changed in one, what happens to the corresponding requirement in the other? See T.2.8. | There is a special situation with the Canadian comments, because we didn’t receive the respective Email in time. The secretariat has therefore responded separately (March 2005) and has now considered both the old comments (on the WD) and the new comments (on the 1CD). Chapter 1 (Introduction) and Chapter 2 (Scope and field of application) of OIML document D11 say that D11 has to be taken into account when drafting OIML recommendations. This should be done by implementing the D11 requirements into the relevant Recommendation (see Note (1) in Chapter 2, D11). The same will apply to the new OIML document on software (SW) once this will have been published, i.e. the software requirements proposed in an OIML D SW document would have to be implemented at the next but one revision of R76. We think this is not a problem, because if a respective requirement is changed in a general OIML document, this can and certainly will be considered at the next revision of the respective (specific) recommendation. |
| General | FR | The word “key” (T.2.2.3, T.2.2.4, …) should be changed into a more general one (unless it is confirmed by the english speaking colleagues that it covers all kinds of means). | P+ Has been considered under T.2.2.3 |
| General | NL | Also refer to our “not accepted” general remark on the Working draft 15 December 2003: Taking into account the primary purpose of OIML Recommendations: to be a recommendation for legal requirements for measuring instruments, it is our strong opinion that there should be a strict distinction between: - the requirements (legislation), and - the tests / evaluation to investigate whether the requirements are met, and - the Test Report Format as a tool for the testing laboratories. In this draft, there is still a problem that in several clauses with requirements for the instrument, aspects of the tests are included. For instance: Clause 3 “Metrological requirements” contains in 3.7 information about the test standards to be used for type examination and verification, in 3.9.1.1 about the tilting-test, and in 3.10 information about “type evaluation tests and examinations”. These should be moved to clause 8. Clause 4 ”Technical requirements ....”: 4.18.1 contains some information on tests and “OIML Test Report”. The information about the tests should be moved to 8. The last sentence (about the OIML Test report) should be moved to Annex A. Clause 5 “Requirements for electronic instruments”: in 5.3.3 “shall be subjected to the span stability test ...” and 5.4 “Performance and span stability tests” and the information about the tests as well as the documentation (5.5.2.1, Table 11) should be moved to clause 8. Clause 6 “Technical requirements for a non-self-indicating instrument”: the last sentence of 6.1 “The sensitivity tests shall be ....” should be moved to clause 8 This opinion has also been included in the comments per (sub)clause below. | We principally agree with you and we welcome a clear structure of OIML recommendations. But if the structure were changed this should be done completely, and not by just moving some single requirements to other parts. The major question is, whether we would really like to change the structure of the R76 recommendation at this point of time. The consequence would be a further delay of the revision process by certainly more than 1 year. However, we were ready to do the job if a majority of the TC9/SC1 members would support that proposal. |
| General | NL | In this draft, decimal signs are sometimes given as a comma, and sometimes as a dot. Please refer to the OIML “Directives for the Technical Work”, Clause 6.6.1: In English text the decimal sign shall be a dot and in French a comma. | P+ Commas have been turned into dots in this English version |
| General | NO | I think before we introduce the “family” we have to define a good definition for type, then family | P+ See 2CD, T.3.4. We have taken the definition from OIML P1 (2003), No 2.6. |
| General | UK | Section 3.3 uses lower case “mpe”, but Section T.5.5 uses upper case “MPE”. (there may be other references as well). This should be consistent throughout the Recommendation. Need to check for consistency in the use of the words “mass” and weight”. | P+ “mpe” is now used, see 2CD. P+ Has been checked and, if necessary, corrected throughout the entire document in accordance with the terminology defined in the new R111 and D28, see new Note under T.1.1 in the 2CD. |
| General | US | Note: 1. All changes in 1 CD against R76-1 (1992) are marked red 2. All changes in 1 CD against Working Draft Revision (2003) are marked with blue background 3. Language and amendments proposed by the U.S. are marked with yellow background. | OK |
| Question B.1 |  | Do we need a separate Annex for the digital modules “digital data processing device”, “terminal” and “digital display” (see Figure 1 under T.2.2) in addition to the existing Annex C? | Result of the vote on B.1: 10 answers: 7 Yes, 3 No (see below) Therefore a new Annex D for purely digital devices has been added, and the former Annexes D to F have been renumbered to E to G. |
| Answer to B.1 | CH | We do not think that a separate Annex for the digital modules “digital data processing device”, “terminal” and “remote display” is necessary in addition to the existing Annex C |  |
| Answer to B.1 | DE | Yes, but one Annex for all digital devices mentioned (digital data processing device, terminal, digital display); |  |
| Answer to B.1 | FR | Yes, but one Annex for all digital devices mentioned (digital data processing device, terminal, digital display); |  |
| Answer to B.1 | JP | If specific descriptions for each test (descriptions of necessary tests and unnecessary tests) are provided, there would be no need to have separate Annexes. |  |
| Answer to B.1 | KR | I think we need a separate Annex for them. |  |
| Answer to B.1 | NL | Yes, we are in favor of a separate annex for (purely) digital modules. In this annex it would be (more) clear which tests need to be performed (eg not temperature and damp heat). |  |
| Answer to B.1 | RO | yes |  |
| Answer to B.1 | ZA | Annex C is sufficient as it is. |  |
| Answer to B.1 | SI | We support the idea to separately define applicable requirements, principles of testing, tests and guidance’s for OIML certificate for digital data processing devices, terminals and remote displays. |  |
| Answer to B.1 | US | The U.S. agrees that criteria is needed for digital data processing devices either as a separate Annex or combined with Annex C. A combined Annex C could be titled as shown below to clarify that applicable requirements in Annex C, clause C1 are equivalent for both digital and analog data processing devices. This would include interfaces, connectable devices, and descriptive markings and control marks: ANNEX C TESTING AND CERTIFICATION OF INDICATORS, DIGITAL DATA PROCESSING DEVICES, AND ANALOG PROCESSING DEVICES AS MODULES OF NON-AUTOMATIC WEIGHING INSTRUMENTS |  |
| Question B.2 |  | Should the introductory note proposed under B.3.3 “Surge” be kept? | Result of the vote on B.2: 10 answers: 7 Yes, 3 No (see below) Therefore the introductory note will be kept, see 2CD, B.3.3 |
| Answer to B.2 | CECIP | Industrial customers force the manufacturers of instruments to meet not only the product related standards like R76 but also generic IEC standards for compliance with electromagnetic effects. Therefore OIML should avoid starting to copy any EMC standard existing in the world. That is not possible and useful. As Dutch NMI pointed out a surge testing is not essential for normal industrial applications. In extreme cases the general clause 4.1.1.1 or 5.1.1 comes into force. Please delete B.3.3 completely. | P- see our response under B.3.3 CECIP |
| Answer to B.2 | CH | We would agree to keep the introductory note in B.3.3 Surge. |  |
| Answer to B.2 | DE | Yes |  |
| Answer to B.2 | FR | Yes |  |
| Answer to B.2 | JP | There is no need to delete the note under B.3.3 “Surge”. However, as the module type indicators could be used indoor and outdoor, it seems that there would be necessary to add a description such as “Tested for surge” in the descriptive markings under section 7.1 |  |
| Answer to B.2 | NL | Yes, we want to keep the introductory note. |  |
| Answer to B.2 | RO | yes |  |
| Answer to B.2 | ZA | We think this test (Surge) needs to be compulsory for all instruments and therefore do not agree with the inclusion of the introductory note. The risk of significant influence of surges or when a DC powered instrument is powered from a DC network cannot be predicted (expected) because it is not always known beforehand where an instrument will be installed. |  |
| Answer to B.2 | SI | We agree to keep the introductory note. |  |
| Answer to B.2 | US | The U.S. believes that the introductory note proposed under B.3.3 “Surge” not be kept unless additional language is developed to distinguish between instruments that have and have not been tested for “surge.” Similar language has been included in the humidity marking requirements in OIML R 60. Additionally, attempting to specify conditions where surge testing is needed is still subject to different interpretations. The U.S. suggests that conditions where surge testing is not needed, such as battery operated device and indoor installations where signal lines are shorter than 30 meters. The language from D11 sub clause 8.4.5 is similar to other examples of disturbances in D 11 clause 8.4. Severity levels. | P- We think that a special marking on the instrument is not adequate for the surge test, and that the information provided in the respective test report can be considered as sufficient. However, we have added under 8.2.1.2 No 7.3 the length of signal lines as a necessary information to be provided by the manufacturer. |
| Explanary Note on page 5 | UK | “Repubic of Korea” should be “Republic of Korea”. | P+ Has been corrected. |
| Terminology | AU | Add a definition of ‘metrologically relevant’: Metrologically relevant: Any aspect of an instrument which is referred to within this document is considered to be metrologically relevant. Similarly all functions, properties, qualities, parameters and characteristics referred to in this document are considered to be ‘metrological’ unless they are excluded either by specific mention, or by the context of use of the term. Explanation: Throughout this recommendation there are many places in which there is reference to ‘metrologically relevant’ aspects (other terms such as ‘metrological functions’, parameters etc are also used). However nowhere is this defined – one interpretation (in the absence of anything specific this is our interpretation) is that all aspects in this document are relevant to ‘legal metrology’ and hence anything covered by this document is ‘metrologically relevant’. Clause 3.10.4.5 includes a listing of metrologically relevant characteristics – however this is only applicable to testing of a family of instruments or modules, and it clearly does not include aspects which are surely by any definition metrological (e.g. operation of zero and tare devices). Or perhaps metrologically relevant is intended to only apply to aspects mentioned in T.4 to T.4.6 (we believe this would clearly be inadequate). | P+ but we suggest a definition which cannot be confused with “legally relevant” (see 2CD, No T.2.9): ”Any device, module, part, component, function or software of a weighing instrument that influences the weighing result or any other primary indication is considered as metrologically relevant.” |
| T.1.2 | AT | We think that the sentence “The weighing process allows the operator to take an action which influences the weighing result in the case where the weighing result is not acceptable.” is critical, because “an action which influences the weighing result” could be “regulating the belt speed or adjusting the feeding device. We think the note is clear enough without this sentence. | P- But we think that the wording has been improved by the US proposal below. |
| T.1.2 | CA | Definition of Non-Automatic Weighing instrument (device) is also causing us some concern within Canada. Concern is that both R76 (and our own current definition) are based upon device usage as opposed to device design. This means that a given device could change from NAWI to AWI in use. This is unacceptable as evaluation is different for each instrument (device) type. Propose that definition be based upon device design. Unsure of best wording? | P- We understand your point; it is indeed problematic to define a NAWI in terms of an “operator”. The only solution that would solve the problem would be to distinguish between “static” and “dynamic” weighing; with all positive consequences (including reduction of size of AWI recommendations and test reports) and negative consequences (contradiction to existing AWI recommendations and many national laws). This proposal was, however, proposed and discussed earlier and it was refused by a clear majority vote of the P-members of TC9/SC1, see “Summary of votes” (29 April 2003), No 2.1. |
| T.1.2 | FR | Last line of the note : Delete “non-automatic” | Settled by new wording (Note 2), see T.1.2 US |
| T.1.2 | JP | The definition of non-automatic weighing instrument is the instrument that requires the intervention of an operator during weighing process to decide that the result is acceptable. However, there is weighing instrument to measure mass by stationary measurement without intervention of an operator to decide that the result is acceptable. Is this type of weighing instrument supposed not to be non-automatic weighing instrument? If so, this definition is against Japanese policy of classification of non-automatic and automatic weighing instruments. | See our response to CA |
| T.1.2 | ZA | 1st Par: Add: …weighing process to decide that the weighing result it is acceptable. | P+ see T.1.2 US |
| T.1.2 | UK | Is the word “dynamic” in the last new (red) sentence in this section appropriate? Dynamic weighing instruments tend to be considered as automatic, e.g. a dynamic catchweigher. “...weighed load, it also includes to make...” should be “...weighed load. It also includes making...” The wording of the “Note” to this section is at best ambiguous. It implies that the operator must have the capability of adjusting the weight of the weighed load whilst that load is on the load receptor. This may not always be possible either for hygiene or health and safety reasons. If left as it is the wording would mean that any “non-automatic” instrument where it is not possible for the operator to adjust the weight of the load whilst it is on the load receptor would fall outside the definition of nonautomatic given in this Recommendation. What is important is that the operator has to accept the weighing or take some action after the weighing is complete to adjust or correct it. Reinstate the definition as given in R76-1, Edition 1992 | P+ “dynamic” deleted P- see US proposal to T.1.2 P- Because of improved wording, see T.1.2 US |
| T.1.2 | US | The U.S. appreciates the earlier clarifications provided by the Secretariat on the definition of the additional “Note” of the term non-automatic weighing instrument. We believe that additional clarification is needed regarding the phrase “operator intervention during the weighing process.” U.S. manufacturers and regulators have reported that they have difficulty in consistently determining the “automatic” or “non-automatic” status of weighing instruments that require operators to place objects on the load receptor of a weighing instrument. After the weight has been determined, the instrument automatically removes the object from the load receptor to be automatically wrapped and labelled. The technical experts in the U.S. are divided as how to classify these types of instruments (automatic or non-automatic). The U.S. offers the following to clarify that the instrument in the above example is not an automatic weighing instrument: T.1.2 Non-automatic weighing instrument Instrument that requires the intervention of an operator during the weighing process to decide that it the weighing result is acceptable. Note: Deciding that the weighing result is acceptable includes any intelligent action of the operator that affects the result, such as taking an action when an indication is stable or adjusting the weight of the weighed load, it also includes and to make a decision regarding the acceptance of each weighing result on observing the indication or releasing a print out. The A non-automatic weighing process allows the operator to take an action (i.e. adjust the load, adjust the unit price, determine that the load is acceptable, etc.) which influences the weighing result in the case where the weighing result is not acceptable. The necessity to give an instruction to start the weighing process, or release a load, or place the item to be weighed on the load receptor, or to change the function of the instrument (static, dynamic, automatic loading, automatic start of weighing, etc.) are not relevant in deciding the category of a non-automatic weighing instrument. | P+ See also ZA proposal P+ See Note 1 in 2CD P+ See Note 1 in 2CD P- Instead of this text, which would probably lead to misinterpretations, we propose a new Note 2 relating to the definition of AWIs (see 2CD, No T.1.2) |
| T.1.2.11 | CA | We still find this section somewhat ambiguous. Last part of first paragraph is rather restrictive in that the vehicle must be equipped with wheels and somewhat unclear what is meant by ‘can be moved without other tools’. Is a platform scale, equipped with wheels a mobile instrument? Suggest that the 1st paragraph of this section be reworded: Non Automatic weighing instrument mounted on or incorporated into a vehicle or similar utensil that enables the instrument to be moved on wheels without other tools or other mobile equipment. Suggest dropping examples as it is not readily apparent what each of these examples actually represents. | No, such a platform scale would not be a mobile instrument, because it were not mounted on or incorporated in a vehicle. P+ But we suggest not to mention “other mobile equipment” to avoid confusion. We also suggest to improve the “Note” as follows: “Vehicle mounted instrument is a complete weighing instrument that is firmly mounted on a vehicle, and that is designed for that special purpose.” P- We think that just the examples help for a better understanding of the definition |
| T.1.2.11 | FR | 1st example of the note : after “Postal scale” add “price computing instrument” | P- But we suggest to add “(mobile post office)” |
| T.1.2.11 | UK | Not sure that all of the examples of a vehicle incorporated or a vehicle mounted instrument “enables the instrument to be moved on wheels”. In some applications, e.g. garbage weighers, the instrument is not incorporated in the vehicle to be “moved on wheels”; this is just a consequence of being on the vehicle. However, this definition is true for some of the examples given. The whole definition of mobile instrument should be reviewed accordingly. | P+ We hope that your concerns are met by the improvements (see CA). |
| T.1.2.12 | CA | Portable instrument definition is limited to cover portable vehicle scale. Canada still believes that the definition should be expanded to cover both portable vehicle scales and other portable scales. However, in either case, the definition in R76 remains somewhat misleading as even a stationary weighbridge may be moved with suitable equipment. Suggested wording: Portable Vehicle Weighing instrument – Self contained, Non-automatic weighing instrument having a load receptor – in one or several parts – that determines the mass of road vehicles, and that is designed to be moved to other locations. Rationale: The differentiation should not be can the device be moved, but rather was it designed to be moved. As previously stated, the note in this section regarding multi draft weighing should be a separate definition as it applies to any weighing activity and not simply those weighments involving portable instruments. | P- The restriction to portable vehicle scales was intended by TC9/SC1 members (as far as we understood the respective proposal); this makes sense because there is a demand for special, additional requirements only for axle or wheel load weighers. Other instruments that are portable must meet R76 in general. P+ |
| T.1.2.12 | RO | By deleting the text from the previous version of the note “This excludes the certification of portable instruments that calculate the total mass by summing up the partial axle (or wheel) loads determined one after the other“ and by using in parallel the terms “parts of load receptor”, “load weighers” and “load platforms”, the impression is created that the latter are different objects and that the axles may be determined one after the other provided that all the wheels should be simultaneously supported by appropriate platforms (even if these are not parts of the load receptor). We suggest to re-phrase as follows: “Note: This Recommendation covers only weighbridges and groups of associated non-automatic axle (or wheel) load weighers that determine simultaneously the mass of a road vehicle with all axles (or wheels) being simultaneously supported by appropriate parts of load receptor.“ | P+ |
| T.1.2.12 | UK | The note implies that this Recommendation only applies to weighbridges and groups of associated axle load weighers. The word “platforms” at the end of the sentence does not match up with the Terminology. Add “In regard to instruments for weighing road vehicles” at the beginning of the sentence in the Note. Replace “platforms” with “receptors”. | P+ We hope your concerns are covered by our responses to RO and CA |
| T.1.2.13 | CA | In Canada, the term ‘Grading Instrument’ is used for instrument (device) which determines ‘quality’ or ‘dockage’. (e.g. grain dockage scale, egg grading scale, etc.) It is not used for determination of a tariff or toll. These devices should be called ‘weight classifiers’ or something similar. Prefer to see wording changed as the term ‘grading’ is clearly not accurate in this case. See also comment 3.2 below. | P- Term is in line with R51 (DR), No 2.2.2. |
| T.1.3 | UK | The note is ambiguous. It presumably means that the word “indication” etc. can mean displaying or printing or both, but it could alternatively be understood to mean that that word “indication” etc. must mean both displaying and printing. Also, change “displaying device” to “indicating device”. | P+ is changed to read “…includes both displaying and/or printing.” P- But we have clarified the note by deleting the last part “; see also...” |
| T.2.2 | CECIP | A digital data processing device does not necessarily include No. 4. In our understanding 5 is defined as such as well if it is separate from 4 and has to be taken into account. | P+ Has been changed to read: (4) + 5 + (6) in Figure 1 |
| T.2.2 | FR | In the table, we propose to introduce a line for 3 (ADC) alone. | P- 1. We have no experience at all with testing of pure ADCs as modules, 2. But w suggest to change the heading of Figure 1 to “Definition of typical modules...(further combinations are possible).”; this gives us more flexibility as regards future developments; see also response to US comment on T.2.2. |
| T.2.2 | JP | Weighing sensors such as string sensor, capacitance sensor other than analogue load cell and digital load cell are not included. Doesn’t this definition likely impair the development of weighing sensors with higher accuracy and lower power consumption than load cells based on technical progress in future | P- “Analogue load cells” as defined in R60 are not restricted to strain gauge load cells; string sensors or capacitance sensors are considered as analogue or digital load cells, too, and are included. Please refer also to the response to FR. |
| T.2.2 | SI | In the table under Figure 1 there are some wrong references to terminology. Replace T.2.2.5 with T.2.2.6, T.2.2.6 with T.2.2.7 and T.2.2.7 with T.2.2.5. | P+ Have been corrected and renumbered |
| T.2.2 | UK | Page 9: The sentence “For the present...” should be left-just or fully justified. | P+ |
| T.2.2 | US | The U.S. supports the proposed amendments to T.2.2 and proposes a new term that consists of the combination of a load cell and the mechanical and electrical connecting elements be defined include the new term in T.2.2. Figure 1. Note that the combined effects of load cells and connecting elements are already established in Table 7. T.2.2.8. Weighing element (is there a better term?) The part of a weighing instrument that comprises all mechanical devices and load cells (i.e. load receptor, load-transmitting device, and analog or digital load cells). Figure 1 Definition of modules according to Terminology T.2.2 and 3.10.2 analogue load cell (T.2.2.1) 2 digital load cell (T.2.2.1) 2 + 3 + (4)*) Indicator (T.2.2.2) (3) + 4 + (5) + (6) + 7 analogue data processing device (T.2.2.3) 3 + 4 + (5) + (6) digital data processing device (T.2.2.4) 4 + (5) + (6) Terminal (T.2.2.5) (5) + 6 + 7 remote display (T.2.2.6) 7 weighing module (T.2.2.7) 1 + 2 + 3 + 4 + (5) + (6) weighing element (?) (T.2.2.8) 1 + 2 + (3)*) *) Numbers in brackets indicate options | We principally agree that this is a possible combination, but we rather suggest to introduce a more general term in the heading of Figure 1 (see response to FR proposal). |
| T.2.2.1 | CN | We suggest the note in T.2.2.1 is explained more clearly according to 2.1.3 of OIML R60, namely, “Load cells equipped with electronics including amplifier , A/D converter and data processor (optionally) etc., are called digital load cells (see figure 1).” | P+ We suggest, however, to use data processing device instead of “data processor”, and to leave out “etc”. |
| T.2.2.3 | AT | “… It may optionally have one or more keys to operate the instrument …” | P+ |
| T.2.2.3 | FR | The word “key” (T.2.2.3, T.2.2.4, …) should be changed into a more general one (unless it is confirmed by the english speaking colleagues that it covers all kinds of means). | P+ Clarified by adding examples to “keys” (or mouse , touch-screen, etc.) |
| T.2.2.3 | UK | “It may optionally have one more keys to operate the instrument” should be “It may optionally have one or more keys to operate the instrument”, to match Section T.2.2.4. | P+ see T.2.2.3 AT |
| T.2.2.4 | FR | The word “key” (T.2.2.3, T.2.2.4, …) should be changed into a more general one (unless it is confirmed by the english speaking colleagues that it covers all kinds of means). | P- We think that it is sufficient to explain “key” one times in T.2.2.3 |
| T.2.2.4 | UK | Need a definition of “data” in the sentence “… that further processes the data, …” | P- Cannot be misunderstood in this context |
| T.2.2.5 | CN | In T.2.2.5, We propose to add the word - “digital data processing device” to “…( ie. load receptor, load-transmitting device, load cell, analogue data processing device and digital data processing device)…”. | P+ Considered in new T.2.2.7 (re-numbered) |
| T.2.2.6 | CA | Terminal definition is too restrictive. Using a keyboard as the differentiating piece of hardware is going to cause problems as touch screens, mouses (mice?) and other alternate human interfaces are developed and exploited. Touch screens are already relatively common. Canada agrees with clarification as proposed by Secretariat. | P+ Has been clarified in T.2.2.3 (see response to FR) |
| T.2.2.7 | CA | Should remote displays be considered as a means of primary indication? Canada currently requires an approved and inspected primary indicator but allows non-approved remote displays as secondary indicators only. It is not clear why remote displays would be considered as the primary indication means. Prefer term “duplication” rather than “replication”. Canada agrees with clarification as proposed by Secretariat. | P+ Figure 1 has been changed for further clarification, please refer to our new proposal in Figure 1 and T.2.2.6 (re-numbered) To avoid confusion “remote display” has been exchanged by either “primary display”, or “secondary display”, or “digital display” in the entire document (see changed definition under re-numbered T.2.2.6). |
| T.2.3.6 | UK | Needs to be slightly reworded. | We are not aware of any necessary changes; we expect editorial and linguistic improvements be done by the BIML. |
| T.2.4 | UK | “displaying” be “indicating”? “Displaying” is only used in the list in 2.4 and nowhere else in the document. In the sentence “Component displaying”, change to “Component that displays …” or reinstate the definition as given in R76-1 Edition 1992. | P- In T.2.4 “displaying device” is used correctly; the term cannot be deleted because it is required in several contexts. In the entire document “indicating device” has been replaced by “displaying device” if the meaning is just “displaying” (and not “printing”). P+ see below |
| T.2.4.1 | UK | “displaying” be “indicating”? “Displaying” is only used in the list in 2.4 and nowhere else in the document. In the sentence “Component displaying”, change to “Component that displays …” or reinstate the definition as given in R76-1 Edition 1992. | See T.2.4 UK above P+ |
| T.2.4.2 | CA | Minor editorial change, “A line or other mark on an displaying component…” | P+ |
| T.2.5 | CA | What is the purpose for this section? It may cause issue as Canada currently uses very different terminology. | The purpose is to explain the terms used e.g. in 3.4. |
| T.2.8 | CA | Canada has many concerns with the entire metrological software section, including terminology and application. As software requirements are applicable to many instrument types, not only NAWI’s, we wonder why software is included in the recommendation and not kept as a separate document. D-SW? Suggest replacing phrase ‘Legally Relevant Software’ with ‘Metrological Software’ which is more descriptive. | 1. We refer to our response to your proposal under “General” (D11 and D-SW). 2. We’d like to remind you of the clear vote of TC9/SC1 members concerning software requirements in R76 (“Summary of the votes” dated 29 April 2003, No 4.3). The proposed requirements and terminology are in line with the pre-draft document OIML D-SW on software (TC5/SC2). 3. Even if a separate D-SW document were adopted, there would be a need to transfer its software requirements into the respective instrument-specific OIML recommendation. P- Both phrases are also used in the pre-draft document OIML D-SW on software (TC5/SC2). |
| T.2.8.1/T.2.8.3 | CN | The end of the sentences missed out period (“.”) in T.2.8.1 and T.2.8.3. | P+ |
| T.2.8.6 | CN | In T.2.8.6, what’s the meaning of “checksum”? | Simply speaking, a checksum is the result of a calculation, addition or other logical operation of all bits or bytes of the legally relevant software. |
| T.2.8.6 | UK | delete “, and”. | P+ |
| T.2.8.7 | UK | In the final sentence after the comma, change to “, the whole software is to be considered as legally relevant.” | P+ |
| T.3.2.3 | AT | It could lead to confusion to have the first regulations of the verification scale interval under “Auxiliary indication devices”. Maybe we should have a supplementary sentence in 3.4.2 saying “e = 1 ´ 10k, 2 ´ 10k, or 5 ´ 10k, the index k being a positive or negative whole number or zero.” or something like that | see our response under 3.4 AT |
| T.3.4 | CH | A parenthesis is not closed in T.3.4. | P+ |
| T.3.4 | JP | In the 3rd line of the sentence, “…device-but…” should be “…device) but”… | See above |
| T.3.4 | UK | Replace “...device-but...” with “...device), but...” last sentence: Replace “...possibility to list...” with “...possibility of listing...”. | See above P+ |
| T.4.6 | UK | Could be reworded as follows: “The weighing value that is achieved when the instrument is completely at rest and balanced, with no disturbances affecting the indication”. | P+ |
| T.5.5.1 | KR | Proposed change: The indication of an instrument minus the (conventional) true value of the corresponding mass | P+ |
| T.5.5.2 | KR | Proposed change: The error of an instrument, determined under reference conditions | P+ |
| T.8 | CH | Consider moving the index of terms defined T.8 at the end of the recommendation in order to reduce the size of the chapter Terminology which became really long. | P- This chapter has been moved from the end to this position on an earlier demand; it should therefore not be moved back again. |
| T.8 | CN | Redaction erratum in T.8: Remote display……T2.2.6 shall be revised into Remote display……T.2.2.7 Terminal .………….T2.2.5 shall be revised into Terminal .………….T.2.2.6 Weighing module….T2.2.4 shall be revised into Weighing module ….T.2.2.5 | P+ has been corrected and re-numbered |
| T.8 | JP | The section number for “Remote display” T.2.2.6 should be T.2.2.7. | P+ see above |
| 2.1 | UK | Uses the spelling “ton”, but it should be “tonne”. In the second sentence, modify the wording such that “… may be used as the unit of …” | P+ P+ |
| 2.2 | FR | The sentence “Loads are regarded as gross or net load” seems to be misleading. We propose to clarify it. | P+ See 2CD |
| 2.3 | AU | 2.3 Principles of the technical requirements Include a note: Note: ‘Technical progress’ implies improvement, not all changes should be considered to be technical progress. The intention to allow additional functions and not impede technical progress shall always be subservient to the need to comply with these requirements, ensure suitability for use and satisfactory performance, and the need to maintain appropriate legal metrological control. Explanation:We feel that this clause is being used to justify the approval of equipment which does not meet basic requirements of suitability of use, and is underpinning a reluctance to fail equipment which has implemented undesirable technical changes either without consideration of such basic principles, or on the basis of other considerations (cost, time) which are not related to ‘technical progress’, but which instead represent technical regression . | P+ However we suggest a slightly shortened version added to the second para in 2CD, No 2.3: “…, and if suitability for use and appropriate metrological control is ensured.” |
| 2.3 | UK | In the first sentence replace “completed” with “supplemented”. | P+ |
| 3.2 | CA | Classification of Instruments Canada still has concern with the terminology used in this section. The minimum capacity is reduced to 5e for instruments in class II and III used strictly for determining a transport tariff and class III and IIII for garbage weighers weighing waste material. New concern with ‘grading instrument’ see T.1.2.13 above. Further suggest that the wording on the note be amended as shown. This would allow for all types of waste, whether residential garbage or commercial waste. Does this include recyclables? Canadian requirements include recyclables, except differentiate between ferrous and non-ferrous metals. Non-ferrous have higher minimum capacity (weighment) than general waste due to their higher value. General Comment – the minimum capacities expressed in R76 are much smaller than Canadian requirements (Canadian are stricter). We like the R76 approach, but question how to ensure a correct class of device is used for a given application? See also T.1.2.13 above | P+ See new wording in 2CD See our response under T.1.2.13 P+ We understand that your concern has been addressed by changing the wording to “weighing waste material”; in addition we’d like to remind you that each member country has the freedom to define for special applications special national regulations deviating from R76. |
| 3.2 | JP | What is the reason for the specification that the minimum capacity could be reduced to 5e for grading instruments? There could be some corresponding cases with other types of instruments. Is it allow to decide flexibly the capacity in accordance with specific applications of instruments? | There must be some good metrological reasons for reducing Min for grading instruments, because this exception has also been put into the new (revised) OIML R51. The wording in 3.2 is in line with the latest draft revision R51-1 (No 2.2.2). |
| 3.2 | UK | Is the sentence beginning with “The minimum capacity is reduced to 5 e …” applicable to all accuracy classes? | Yes, it does. |
| 3.3.3 | CECIP | The direction of the signs of table 4 are wrong and have to be modified. (the smaller sign has to be a larger sign) The new example should also define any mpe for the transition regions. Now there seem to be not any mpe for load between 2000g to 2002g, 5000g to 5010g and so on. | P+ New example: P+ for m = 0 g to 500 g mpe = ( 0,5 e1 = ( 0,5 g for m > 500 g to 2000 g mpe = ( 1 e1 = ( 1 g for m > 2000 g to 4000 g mpe = ( 1 e2 = ( 2 g for m > 4000 g to 5000 g mpe = ( 1,5 e2 = ( 3 g for m > 5000 g to 15000 g mpe = ( 1 e3 = ( 10 g |
| 3.3.3 | CH | In order to shorten the recommendation, consider moving examples (like in 3.3.3, 4.6.12) or remarks to the selection (like in 3.10.4.5) into an Appendix or into a separated guide. They are neither metrological nor technical requirements. | P- For two reasons we think that this will not find the support of a majority of TC9/SC1 members: 1. The recommendation would be shortened only “virtually”, 2. the relation of the examples to their context would get. |
| 3.3.3 | CN | Issue on English abbreviations: mpe shall be revised into MPE. There are 5 places on Page 27, 2 places on Page 113, 3 places on Page 126 and 3 places on Page 128. Table 4 Class Maxi/ei+1 ( 50 000 ( 5 000 ( 500 ( 50 Change into Table 4 Class Maxi/ei+1 ≥50 000 ≥ 5 000 ≥ 500 ≥50 | P+ It is made consistent throughout R76, but in OIML recommendations “mpe” in lower case letters is the preferred version. P+ |
| 3.3.3 | FR | In table 4, the marks “(” shall be changed into “(”. | P+ |
| 3.3.3 | SI | (example for a multi-interval instrument) Change the sentence “The maximum permissible errors on initial verification (mpe) (see 3.5.1) for some examples of load are:” to “The maximum permissible errors on initial verification (mpe) (see 3.5.1) are:”, because the maximum permissible errors for a complete weighing range are presented in the example. | P+ |
| 3.4 | AT | It could lead to confusion to have the first regulations of the verification scale interval under “Auxiliary indication devices”. Maybe we should have a supplementary sentence in 3.4.2 saying “e = 1 ´ 10k, 2 ´ 10k, or 5 ´ 10k, the index k being a positive or negative whole number or zero.” or something like that. | (see also T.3.2.3 AT) P+ But we propose a slightly different addition (see reference added to 4.2.2.1 in 3.4.2). |
| 3.4.2 | US | The U.S. requests that language be added that clarifies which part of the requirement does not apply to a Class I instrument with d < 1 mg, where e = 1 mg. The following amendment to 3.4.2 represents the U.S. position about the exception language. 3.4.2 Verification scale interval The verification scale interval e is determined by the expression: d < e ≤ 10 d (see table 5) e = 10k kg k being a positive or negative whole number, or zero. Table 5: The values of e, calculated following this rule, are, for example: d = 0.1 g 0.2 g 0.5 g e = 1 g 1 g 1 g e = 10 d 5 d 2 d This requirement does not apply to an instrument of class with d < 1 mg, where e = 1 mg, for example: d = 0.01 mg 0.02 mg 0.05 mg e = 1 mg 1 mg 1 mg e = 100 d 50 d 20 d | P+ P+ But reference to “tables 5a and 5b” (see below) P+ P+ In addition we suggest to add another column to table 5b: d = < 0.01 mg / e = 1 mg / e > 100 d |
| 3.5.1 | UK | Note below table: In this note (and elsewhere in the new additions), the comma has been used as the decimal separator, whereas in the existing text the full-stop has been used. This should be consistent. | P+ |
| 3.5.3.2/3.5.3.3 | CA | Unclear as to what this means? | It is a basic rule in metrology that the digital rounding error of 0,5e has to be taken into account; only if the digital rounding error is smaller or equal 0,2e it can be neglected.. |
| 3.6.2.2 | CA | It is not readily apparent that this test is not suitable for vehicle scales. It is suggested that the paragraph be prefaced with the phrase ‘subject to 3.6.2.4’ to clearly indicate that vehicle scales are subject to the heavier loading pattern. P- comment: with the wording as in the CD1, it would be expected that both 3.6.2.2 AND 3.6.2.4 would be performed on a vehicle scale. This seems to be excessive? By changing the wording as suggested, it is clear that only 3.6.2.4 need be performed. | P- We can still not agree, because of the note to 3.6.2: It says that if the instrument can be used in different ways, it may be appropriate to apply more than one test. Only for railway scales one test acc. to 3.6.2.4 is sufficient. |
| 3.6.2.3 | CA | Shows a loading example using the words “one-tenth”. Similar sections show fractions using numeric form “1/10”. This should be standardized and “1/10” is preferred. Consideration should be made to using decimal form whenever possible “0.1”. Canada agrees with Secretariats response. | OK |
| 3.6.2.4 | FR | 2nd line : read : “… a rolling test load corresponding …”. | P+ |
| 3.6.2.4 | CA | We must restate our concern with this section. This load pattern could be too extreme. Canada uses a loading pattern for eccentricity testing on vehicle scales that places the load across a ‘section’ (pair of load cells on each side of the scale). The maximum loading for each section is 75% Max for a 2 section scale (4 load cells) and 50% Max for a scale with 3 or more sections. Consider a 6 section vehicle scale rated at 100 tonnes. Would it be realistic to assume that each section should be rated for 80 tonnes, even though the loading in use would never approach 80 tonnes for a given section? In Canada we would require each section be rated at 50 tonnes minimum. 80 tonne loading on each section of this scale would be considered extreme. In addition, there is little detail as to what ‘applied at different points on the load receptor’ actually means. Does this mean you could concentrate 80 tonnes on an end? Prefer to have the load applied over ‘sections – each group of 2 load cells)’. (see also A.4.7.4) Canadian manufacturers have also indicated concern with this section. 80% loading on a vehicle scale, with no constraints regarding loading patterns advising that it could cause immediate and or premature device damage. | We are sure that there is a misunderstanding: 3.6.2.4 clearly supports your viewpoint, because it states that “a test load corresponding to the usual rolling load” shall be applied! Just in order to avoid excessive testing it also states that this test load should not exceed 80% of Max, i.e. it is usually less than 80% of Max, if appropriate. |
| 3.6.3 | UK | At the end of the sentence, use of the word “or” seems to imply that indicating and printing devices could be different. | P+ |
| 3.7 | NL | These requirements for the tests standards are not a “Metrological requirement” for the weighing instrument. So this should be moved to clause 8. | See our response under “General” |
| 3.7.1 | CECIP | For better understanding please change the wording: “The standard weights or standard masses used for the type examination or verification of an instrument shall principally meet the metrological requirements of OIML R111. They shall not have an error tolerance greater than 1/3 of the maximum permissible error of the instrument for the applied load. If class E2 equivalent weights or better are used as standard weights it is allowed that their uncertainty rather than their error tolerance is not greater than 1/3 of the maximum permissible error of the instrument for the applied load, provided that the error actual weight value and the estimated long-term stability is taken into account.” | “Tolerance” would not bring further clarification, because “error” is even used in R111. The relationship between “error” and “uncertainty” will certainly be considered at the next but one revision of R76, after publication of the new respective OIML document elaborated by TC3/SC5. |
| 3.7.3 | AT | So it is possible to test any kind of instrument using this procedure? Each accuracy class with every capacity? Even if fraudulent use is quite unlikely, we would prefer not to have special and high class instruments tested not using standard weights | Yes, you are right, and we see no problem with this. Why shouldn’t it be possible to test also low capacity instruments in this way, when the requirements are fulfilled? But normally nobody will do this, because it takes too much time. |
| 3.7.3 | CA | Substitution of standard weights at verification. We are still confused by the phrase ‘… the place of use’ . What is place of use’? Does this section apply to instruments tested within a factory or dealers premises? Why differentiate as to where it is tested? If it only applies to testing at the traders premise, then what happens when testing at the factory? It is not readily apparent if this section applies to initial verification, subsequent verification or both? The new wording has eliminated some previous concerns, but unfortunately has created some new ones. There is no longer a stipulation that 100% standards are required for devices up to a given size. Using the new wording, even a small device could be tested with 50% standards. Is this what was desired? We suggest that 100% standards be required for all devices up to 1 000 kg. Canada still has an issue with the potentially large number of standards that could be required to test large scale installations. Scales of 100 000 kg and larger are common throughout North America. Requiring even 20 000 kg (20% Max) is a substantial load and may be considered unreasonable. A conventional test truck in North America will carry 10 000 kg of known standards. | P+ We hope that this is clarified by adding the word “(application)” after “place of use”. It applies to both types of verification, therefore the heading is correct. P- When drafting the 1CD we had changed that point according to a large number of comments opposing our previous proposal in the WD that was going in your direction. We are convinced that most TC9/SC1 members can now live with 3.7.3 as it stands, and that we should accept living with some remaining uncertainties. |
| 3.7.3 | FR | 2nd paragraph : read : “… reduced to 33 % of Max.” According to our experience there is no significant variation of repeatability at different loads. So we suggest a simplification : repeatability is to be tested at only one value between 1/3 Max and 2/3 Max. Concerning the quantity of standard mass we prefer 1/3 Max instead of 35 %. | P+ P- concerning the second part of the proposal (see our response to 3.7.3 CA) P+ |
| 3.7.3 | PL | We would like to obtain answers for the following questions: Why only ’at verification’ ? What about type evaluation ?. Does the same concern loads below 1 tonne ? Such radical change concerning the previous version of this draft should be explained. | At verification it is sometimes not possible and/or very expensive to have 100 t or more standard weights, therefore this substitution method is permitted. If additional small deviations occur, it will be no problem, because after the verification the error limits are twice to those on initial verification. On type evaluation no additional small deviations can be accepted, therefore this method is not permitted. For clarification we have added a crossreference to 3.7.1 in 8.2.2, 3rd para, and have added a reference to 3.7 in 8.3.3 after para “The authority… the tests.” Yes, it is possible to test also low capacity instruments in this way, when the requirements are fulfilled. But nobody will do this, because it takes too much time. |
| 3.7.3 | RO | The quotation marks should be deleted in the first paragraph. | P+ |
| 3.7.3 | SI | Is there any special reason to limit the use of the substitution load only to the place of use of the instrument? | At verification it is sometimes not possible and/or very expensive to have 100 t or more standard weights, therefore this substitution method is permitted. If additional small deviations occur, it will be no problem, because after the verification the error limits are twice to those on initial verification. |
| 3.7.3 | UK | Delete the inverted commas at the end of the first sentence Re-word the last sentence to improve the meaning and to clarify that standard weights do not need to be used, i.e. it can be any load, as long as the same load is used for each of the three weighings. | P+ P+ We propose: The repeatability error has to be determined with a load (weights or any other load) of about the value where the substitution is made, by placing it 3 times on the load receptor. |
| 3.8 | FR | Adding values 1 mg and 5 mg seems to us unjustified. Why NAWIs with a high resolution should not comply with the general requirements ? Manufacturers who claim they can obtain such a resolution shall provide instruments in accordance with same quality of others. We propose to delete “but not less than 1 mg” in paragraphs 3.8.1 and 3.8.2.1 and the last sentence “This applies only to instruments with d ( 5 mg”. | P- The reason for this exception is that very unusual or very small weights (i.e. 1,4 mg, 0,01 mg) would be required which is not practical. We have to accept the fact that such test weights are neither available nor usable, be this unjustified or not. See also response to 3.8.2.2 JP |
| 3.8.1 | CA | Non self-indicating instruments Suggest that wording be changed to “….shall produce a visible movement displacement of the indicating element.” Canada agrees with the Secretariat’s response. See also 4.3.2 | P+ “displacement” is OK; consequently it is changed in 4.3.2, too |
| 3.8.2.2 | JP | For the proposed instruments with 5 mg ≥ d, 0.5 mg weight would be necessary to obtain the changing point of indications. In the present situation that R111 does not recommend weight less than 1 mg, 10 mg ≥ d would be a technically consistent inequality. | We are sure that there is a misunderstanding concerning 3.8.2.2 where it is written:”… d ≥ 5 mg …”. The reason is that the additional load must be 1.4 times of the actual scale interval. At d = 5 mg the additional load is 1.4 x 5 mg = 7 mg. This can be realized with 1 mg weights. But if d is smaller than 5 mg, the additional load is 2.8 mg and cannot be realized with 1 mg weights (1.4 x 2 mg = 2.8 mg). |
| 3.8.2.2 | ZA | The result of the requirement shall not only “change the indication” but we believe the indication shall respond in a specific way therefore we suggest:-Insert: …at equilibrium shall cause an unambiguous change in the indication to the next higher or lower verification scale interval. the indication. | P+ But because also a change of 2d is possible, we suggest to change to “… at equilibrium shall change the indication unambiguously.” |
| 3.9.1.1 | CA | Instruments liable to be tilted. The last sentence in section (b) reads in part: “(b)…if the limiting value of tilting has been reached or exceeded.” The words reached or should be struck out and only exceeded left in. If a manufacturer specifies a limiting value of tilting, this would seem to indicate that the device is capable of working at this point (reached), but not beyond (exceeded). Canada agrees with Secretariat’s response. Similar change also required in 4.18.1 | P+ Has been changed, see also 4.18.1 |
| 3.9.1.1 | FR | Paragraph a), note :. We propose to modify the whole note as follows : “If technical reasons forbid the level indicator to be fixed in a visible place, it can be accepted to install it below the load receptor in a place representative for the tilt sensitive part”. Paragraph b) : The compensation by the automatic tilt sensor shall be possible only if actually related to the tilting. | P+ P- We think the present wording is sufficient. |
| 3.9.1.1 | NL | 1st paragraph: “the influence of tilting shall be determined” (type test and/or verification?) is not a requirement, so this should be moved to clause 8. 3.9.1.1 a): Delete the Note: “If technical … “ Our previous comment is still valid: This is not necessary and is likely to be misused. In our experience of type-approval testing we have always found it is technically possible to make the level indicator visible. And this requirement should be taken into account by the manufacturer in the design stage. | P- We cannot understand why this should not be a requirement. It must therefore remain in Chapter 3 “Metrological Requirements”. Also your proposal to add “(type test and/or verification)” cannot be accepted because the tilt test has never been included in the list of tests under 8.3.3. P- This item has been introduced on demand of many TC9/SC1 members. |
| 3.9.1.1 | SI | a): It is not clear for us, when the limiting value of the level indicator is reached: when the edge of the bubble touches the ring, when the centre of the bubble reaches the ring or when the bubble is more or less outside the ring. An additional clarification in the text can be helpful. | P+ Proposal to a): “… the limiting value of tilting is defined by a marking (e.g. a ring) on the level indicator which shows that the maximum permissible tilt has been reached or exceeded when the bubble is displaced from a central position and the edge touches the marking. The limiting value of the level indicator shall be obvious, so that tilting is easily noticed.” |
| 3.9.1.1 | UK | In the first sentence change the phrase “a. through d.” to “Sections a. to d.” d) The phrase “In case of...” appears twice in the last sentence. It might be better as “In the case of...” or just “If...” | P+ P+ |
| 3.9.1.2 | CA | Suggested editorial change in the first item – Class I instruments must be fitted with a leveling device and a level indicator, but these need not be tested. , because these instruments require special environmental and installation conditions and skilled operating staff anyway. The additional information is speculative and not required. Add the word ‘these’ to the sentence to ensure it is clear that only the levelling equipment need not be tested. | P+ P+ |
| 3.9.1.2 | US | Delete “anyway” as follows: Class instruments must be fitted with a levelling device and a level indicator but need not be tested, because these instruments require special environmental and installation conditions and skilled operating staff anyway. | P+ |
| 3.9.3 | NL | This is not in line with D11 for battery operated instruments For the normal (disposable) batteries, we do not see any justification for requiring 1,2 U for battery supply. | See also A.5.4.2 NL P+ see 2CD, new differentiation between “disposable batteries” and “rechargeable batteries” according to D11 |
| 3.9.3 | ZA | 3.9.3 refers to “between the lower and upper limit of the nominal voltage marked on the instrument (U)…..” In clause 7 there is not a requirement for the nominal voltage to be marked. Does this mean you don’t have to do the test if it is not marked? Suggestion: a) ….and upper limit of the nominal voltage marked on of the instrument (U)…. or b) In clause 7 make it compulsory if applicable for the nominal voltage (U) to be marked on the instrument. | P+ You are right, “no marking” does not automatically mean “no testing”, therefore we agree with your suggestion a). |
| 3.9.3 | UK | Modify this Section in line with Section A.5.4.1 to clarify what happens when a voltage range is specified. | P+ see 2CD |
| 3.9.4.2 | KR | Put a full stop after 0.5 ei : …the deviation on returning to zero from Maxi shall not exceed 0.5 ei. Furthermore, … | P+ P+ |
| 3.9.5 | CA | Other influence quantities and restraints. The example at the end of this section references clauses 3 and 4 although these clauses do not appear to exist. Is there an issue with numbering in this section? Canada agrees with Secretariat’s response – ‘Chapter 3 and 4’ is better than ‘Clause 3 and 4’. Although we agree with adding ‘In general’ to the paragraph, we would prefer to see reference to 3000d removed from final version of recommendation. Let device performance be the determining factor. Specifying 3000d is contradictory to the Device Class Table 3. | P+ Editorial: “Clause“ is replaced by the more appropriate term “Chapter” in the entire document. (Here Chapters 3 and 4 are meant) P+ We think the note is very important. 10000e and e=10 kg are extremely critical for unprotected outdoor weighbridges, because external influences such as wind and rain may have a big influence on the weighing result. Therefore we suggest to change the wording (refer to the 2CD). |
| 3.9.5 | UK | Note at end: Clarify that “clauses 3 and 4” are “Section 3 Metrological requirements” and “Section 4 Technical requirements …” | P+ Has been clarified, see our response to 3.9.5 CA |
| 3.9.5 | US | 2nd paragraph “Note.” The U.S. agrees with the NL comments that the paragraph could be confused as a strict requirement. Since all instruments shall comply with clauses 3 and 4 installed in customary and suitable locations, the U.S. recommends deleting the examples from the note as follows: Note: Instruments installed outdoors without suitable protection against atmospheric conditions may normally not comply with the requirements of clauses 3 and 4 if the number of verification scale intervals n is relatively great. (In general, a value of n = 3 000 should not be exceeded. Furthermore for road or rail weighbridges the verification scale interval should not be less than 10 kg). These requirements limits should also apply to each weighing range of combinations of instruments or of multiple range instruments or to each partial weighing range of multi-interval instruments. | P+ See our response to 3.9.5 CA |
| 3.10 | NL | These “Type evaluation tests and examinations” are no “Metrological requirements” for the weighing instrument. So this should be moved to clause 8 | Please refer to our response under “General” NL |
| 3.10.1 | SI | While C.4.2 defines the content of the test report for indicators and D.4.2 the content of the report for weighing modules, there are no requirements defined for the content of OIML certificate for complete instrument. Required relevant information can be defined in 3.10.1. | P- We think that this is a misunderstanding: Complete instruments are state of the art since 1993 and normally the entire R76-2 applies to complete instruments. The reason for defining the content for modules is that R76 should provide clear rules which requirements should be selected and which need not be considered. |
| 3.10.1 | UK | “Upon type evaluation ...” would be better as “For type evaluation...” | P+ |
| 3.10.2.1 | CN | Issue on creep influence (Table 7) Load receptor is one of three main parts same as load cell and indicator. Load receptor is mainly made of metal material. Any structural member made of metal material has creep problem. Creep not only influences load cell, but also influences metal structure. In most instances, influence degree on metal structure is larger than that on load cell. And 2.3 3.10.2.1 mentions that: for mechanical structures such as weighbridges, load transmitting devices and mechanical or electrical connecting elements evidently designed and manufactured according to sound engineering practice, an overall fraction pi = 0.5 may be applied without any test. It is recommended that when considering the distribution of creep factor pi of weighing instrument with typical modules, pi = 0.8 for load cell pi = 0.5 for connecting part. For electronic weighing instruments with several load cells (more than one), it is required that load cells have basically equal linear error. In the type approval test, the test about linearity of each load cell should be added for them. In electronic weighing instruments with several load cells, the consistence of linearity of each load cell can make the error of weighing instruments be linearly corrected or calibrated by the display of weighing instruments, or else it is hard to do it. For the producer of load cells, it is required that each load cell have the data of their linear error to select correctly them in the process of fit of weighing instruments. | P- Considering the serious consequences of such a change we very much recommend to keep the regulation as it is. This is supported by the fact that no serious problems are known with the present regulation which is state of the art since many years. P- We fully understand your point but we think that at type approval it is not necessary to introduce such a special test, because at initial verification each complete individual instrument is tested - among others - for weighing performance and eccentric loading (see No 8.3.3). We should not forbid additional linearization of the complete instrument by the software of the indicator under the modular concept. |
| 3.10.2.4 | NL | “OIML certificates for modules must be clearly distinguishable from OIML certificates for complete instruments.” Please indicate how (with an example or definition of titles) And as this is not a requirement for the instrument (module), this should be moved to 8 or to a new annex with format(s) for the OIML Certificate(s). | We fully agree, and we hope to receive support by the BIML (see 1CD, page 4, question 3). |
| 3.10.2.4 | UK | Add comma after third bullet point (first and second bullet points have commas, and fourth does not need one.) | P+ |
| 3.10.3 | UK | second paragraph: “They need neither be tested...” should be “They need not be tested...”, to be the same as in Section 3.10.2.2. | P+ |
| 3.10.4.1 | KR | Proposed change: ….see example in acceptable solution of 3.10.2.1 | P+ yes, but “of 3.10.4.5” |
| 3.10.4.1 | UK | Use of the term “highest metrological characteristics” can be mis-leading. May be better to use the term “most sensitive”? | P+ |
| 3.10.4.2 | NO | The ratio between the tested capacities shall not exceed 10. Why do you choose 10? Could you comment it? Their capacities (Max) are not more than 5 times above the largest capacity tested. Where did you get 5 from and why? Could you comment it? | This proposal is oriented at OIML R60. The reasoning is: If the ratio is ≤10 it is possible to interpolate to the variants in between. If the ratio is >10 the interpolation is becoming critical because the mechanical dimensions would change too much. Extrapolation is considered more critical, therefore only a factor up to 5 is acceptable, but not more. |
| 3.10.4.2 | ZA | Last sentence: “If the metrological characteristics are lower this factor may be higher” We have problems to understand or interpret this sentence. What does this mean? | It means, for instance, that if an EUT has been tested with n = 6000 and the variants have only n = 3000, the Max values of these variants may exceed the factor 5, i.e. the factor may then be 10. |
| 3.10.4.2 | UK | The final sentence of the final bullet point needs to clarified, i.e. what is lower and what is higher. At present it is too ambiguous. | P+ Please refer to the 2CD. |
| 3.10.4.4 | NL | This sub-clauses describes requirements for tests (type tests or verification?); not for the instruments. So this should be moved to clause 8 | See our response under “General” NL |
| 3.10.4.4 | US | The paragraphs should be formatted to list the examples of other metrologically relevant and non-relevant features similar to 3.10.4.5 Summary of Metrological Characteristics. | P+ see 2CD |
| 3.10.4.5 | CH | In order to shorten the recommendation, consider moving examples (like in 3.3.3, 4.6.12) or remarks to the selection (like in 3.10.4.5) into an Appendix or into a separated guide. They are neither metrological nor technical requirements. | P- See our response under 3.3.3 and 4.6.12 CH |
| 3.10.4.5 | CN | Issue on class symbol: II and III shall be revised into and . There are 1 place in Table 8 on Page 37, 1 place in Table 9 on Page 38, 2 places on Page 48, 3 places on Page 49, 1 place on Page 50, 2 places on Page 126, 2 places on Page 127, 2 places on Page 128 and 2 places on Page 129. | P+ changed in the whole 2CD |
| 3.10.4.5 | CN | Column “Family 1”and“Family 2” in table 8 of 3.10.4.5 can be changed as follows: Family 1 Variant Max e d n EUT Accuracy class II Temperature range: 10° / 30° C 1.1 200 g 0,01 g 0,001 g 20000 1.2 400 g 0,01 g 0,001 g 40000 X 1.3 2000 g 0,05 g 0,05 g 40000 Family 2 2.1 1,5 kg 0,5 g 0,5 g 3000 X Accuracy class III Temperature range: -10° / 40°C 2.2 3 kg 1 g 1 g 3000 2.3 5 kg 2 g 2 g 2500 2.4 15 kg 5 g 5 g 3000 X 2.5 60 kg 20 g 20 g 3000 | P+ |
| 3.10.4.5 | FR | In Table 9, the Max values 1 g (family 1) and 50 g (family 2) are not specified in the initial application (table 8) : These values shall be changed respectively into 200 g and 1,5 kg. | This seems to be a misunderstanding. We hope it is resolved with our additional explanation in the note under Table 9. |
| 3.10.4.5 | KR | Change table 8 as proposed: Family 1 Accuracy class II Temperature range: 10°/ 30℃ Variant 1.1 1.2 1.3 Family 2 Accuracy class III Temperature range: -10°/ 40℃ 2.1 2.2 2.3 2.4 2.5 | P+ See 3.10.4.5 CN |
| 3.10.4.5 | UK | The third bullet point says “for strain gauge load cells only”, but is it a load cell or the indicator that is actually being referred to? | P+ - lowest input signal (V/e (only for indicators when connecting strain gauge load cells) |
| 4 | AU | 4. Technical requirements for a self- or semi-self-indicating instrument Change this to say: The following requirements relate to the design and the construction of instruments, and are intended to ensure that instruments give correct and unambiguous weighing results (weighing results including all primary indications), under normal conditions of use. It may be expected that instrument users are trained in the proper handling of the instrument (with the exception of self-service instruments), however the equipment shall not incorporate features which facilitate deliberate fraud, or which introduce substantial risk of inadvertent provision of incorrect or ambiguous results. Note: The situation of instruments which are used or intended to be used for trading direct with the public are a particularly critical category, in which the requirement for unambiguous results is to be judged using the criteria of “would a typical member of the public encountering a transaction using the instrument, consider that the primary indications presented are clear, unambiguous and of a size suitable for reading at a typical distance”. It is not intended to prescribe solutions, but to define appropriate functioning and design criteria for the instrument. Certain solutions that have been tried over a long period have become accepted; these solutions are marked “acceptable solution”; while it is not necessary to adopt them, they are considered to comply with the requirements of the applicable provision. Explanation: It is very evident that the requirement for ”unambiguous weighing results under normal conditions of use and proper handling by unskilled users” is being given insufficient attention by approval authorities. We do feel that the requirement implies that equipment is designed to be able to be used properly by an unskilled user and feel that this is unrealistic for many types of instrument – hence we have altered this aspect. In addition we have included specific mention of the situation of instruments used for trading direct with the public where we are most concerned at the inadequacy (ambiguity and lack of clarity) of the instruments recently being presented for approval. The possibilities presented by technology and allowed by R76 should ensure that the readability and presentation of primary indications is enhanced – in reality the opposite appears to be happening in a substantial number of case. | P+ as regards the first part of the first sentence of clause 4; we suggest, however, a slightly changed wording: "The following requirements relate to the design and the construction of instruments, and are intended to ensure that instruments give correct and unambiguous weighing results and other primary indicatons, under normal conditions of use and proper handling by unskilled users." P- as regards the assumption of "unskilled users" and other proposed changes: In the “old” (1992) version it was certainly not unintentionally written “proper handling by unskilled users”. We cannot expect instrument users to be trained in proper handling, especially in case of class III and IIII instruments. In addition, clause 4 emphasizes the intention not to prescribe special solutions (except for examples of "acceptable solutions") but to provide requirements for an appropriate functioning of NAWIs. We doubt whether a majority of TC9/SC1 members would agree to a change of clause 4 towards the direction proposed. Concerning instruments for direct sales to the public we fully agree that these are a critical category of instruments, but this needs not be specially addressed in clause 4, because it is obvious by the additional chapters 4.13 and 4.14 especially devoted to direct sales instruments. |
| 4 | CN | We consider that “Hanging (Crane) Instrument “ is one of the non-automatic weighing instruments. It would be specially added in the item of “Hanging ( Crane ) Instrument “ in this recommendation, because of its testing methods (e.g. EMC etc.) differ from general non-automatic weighing instruments. | P- We do not agree, because 1. there are already special precautions for freely suspended instruments (see No 3.9.1.2 last hyphen), 2. except for some tests (eccentricity and tilting) the same tests are required for them. |
| 4.1.1.1 | AU | 4.1.1.1. Suitability for application Add a note to this, to say: An instrument shall be designed to suit its intended purpose of use. Note: All elements of ‘intended purpose’ should be considered, this should include aspects such as the nature and needs of the user (and customer), environment (e.g. an instrument with a temperature range of 10°C to 40°C would not be suitable if its intended purpose was meat packaging). Where the intended purpose needs to be restricted, a marking stating such restriction may be required according to national regulation. Explanation: It is evident that consideration of suitability of use is being given insufficient attention by approval authorities. Emphasis needs to be given to this requirement. | P+ We propose, however, a slightly shortened “Note” as follows: ‘Intended purpose’ includes aspects such as the nature and needs of the application and environment. Where the intended purpose needs to be restricted, a marking stating such restriction may be required according to national regulation. |
| 4.1.1.3 | US | There is a missing period in the next to the last paragraph. | P+ |
| 4.1.2.4 | CA | “National legislation may specify the sealing securing that is required.” The word sealing should now be replaced by securing to remain consistent with the other sections. | P+ |
| 4.1.2.4 | NO | Acceptable technical solution: The actual counter number can be displayed for comparison with the reference number by a procedure described in the manual and in the OIML certificate with test report. Please change “can be displayed” with “should be displayed”. | P- Could lead to the misunderstanding that the actual counter number needs not be indicated at all. |
| 4.1.2.4 | SI | In the second paragraph replace “devices to adjust sensitivity” with ”devices to perform external span adjustment”. a): The term “Test Report” is used (as well as in Contents, 4.18.1, 7.2, 8.2.1.2, A.4.13, C.3, C.4.2, D.3, D.4.2, F4), while in some other clauses (2.3, A.3.1, A.3.1, A.3.3., A.4.1.6, A.4.1.7, A.4.7, C.2.5, D.2.2) the term “Evaluation Report” is used. It is necessary to use uniform terminology. | P+ add ”(or span)” P+ Changed to “Evaluation report” in the entire document. However BIML will be asked to take a final decision, see 2CD, page 4, question 4. |
| 4.1.2.4 | UK | “resp.” should be worded in full. c) Note: The note states that the instrument “shall not longer be used under legal control” from this note, this is a matter for national legislation. if the displayed and fixed reference numbers do not correspond. Firstly this statement appears to be badly worded. Our interpretation is that the authors are trying to say that the instrument should not be used for a legally controlled purpose. Secondly, this goes against current UK legislation for both non-automatic and automatic weighing instruments which allows adjustment, alteration etc without the need for mandatory re-verification if the local Chief inspector of Weights and Measures has been advised in writing. Similar provisions exist in other OIML Member States. Delete the second sentence c): In the last paragraph, the phrase “...firmly mounted to the instrument...” could be better worded. In addition, in the Note under c), the final words should read “the instrument shall no longer be used …” | P+ P+ We refer to our proposal in the 2CD Further improvements by native English speakers are, of course, welcome. P+ |
| 4.1.2.4 | US | “National legislation may specify the sealing securing that is required.” The word sealing should now be replaced by securing to remain consistent with the other sections. | P+ yes, see 4.1.2.4 CA above |
| 4.2.1 | AU | 4.2.1. Quality of reading Change this to the following: Reading of the results primary indications (including weighing results) shall be reliable, easy and unambiguous under conditions of normal use: - the overall inaccuracy of reading of an analogue indicating device shall not exceed 0.2 e, - the figures, units and designations forming the results primary indications shall be of a size, shape and clarity for reading to be easy. The scales, numbering and printing shall permit the figures which form the results to be read by simple juxtaposition. Explanation: It is evident that consideration of the quality of reading is being given insufficient attention by approval authorities. Examples are occurring of instruments presenting results without including units for primary indications such as tare, others where the size of units given is very substantially smaller than the numbers comprising the indication, yet others where ‘pictograms’ are used to indicate various elements of an indication creating substantial ambiguity (such pictograms are un-standardised and largely incomprehensible to untrained users or customers). There are also examples where when taken individually elements of the display may be acceptable, but when combined the spatial location and proximity results in ambiguity. Regrettably we feel that tighter requirements now need to be provided, as the existing requirements are apparently not sufficient to advise manufacturers of acceptable practices, not sufficient to ensure that approval authorities do not approve clearly unacceptable designs.. | P+ We propose, however, a slightly modified text: “Reading of the primary indications (see T.1.3.1),...” P+ |
| 4.2.2 | AU | 4.2.2. Form of the indication Change this to the following: 4.2.2 Form of the indication 4.2.2.1 Each primary indication (including weighing results) shall contain the names or symbols of the units in which they are expressed. This is necessary to ensure an adequate form for prices and unit prices as well as gross, net and tare values. The size of the names or symbols of units comprising primary indications shall not be less than 75% of the size of the figures. This is necessary to ensure units are not provided in a size which is inadequate, hence resulting in ambiguity and potentially in facilitation of fraud. The display elements which are legally or metrologically relevant shall whenever presented, be presented in the same locations in the display. This is necessary to prevent situations in which primary indications and weighing results change size and location according to particular operational modes of the instrument – hence resulting in ambiguity. Where a number of display elements are of a similar nature (e.g. actually or apparently with the same units), provision shall be made to remove any potential for confusion or ambiguity by including appropriate designations for the display elements. Such provision shall be in a size, prominence and location suitable to remove the potential confusion or ambiguity. This is necessary to prevent potential confusion such as between two kg values, one representing the net weight and one representing the tare. Similarly we have encountered a situation where two price values expressed as $ (one qualified by a small and remote arrow indicating ‘/kg’) are easily confused. | P+ We suggest, however, a slightly changed wording, because “primary indications” according to T.1.3.1 comprise much more symbols, indications, etc. (see 2CD) P- We think that this problem has been sufficiently addressed by your own proposal under 4.2.1 (new wording: “...the figures, units and designations forming the primary indications shall be of a size, shape and clarity for reading to be easy). P- We think that this requirement would be too strict (if you think, e.g., of PC monitors); in addition we think that obvious ambiguity can be rejected by applying 4.2.1 (revised wording). P- We understand your point, but again, we think that the proposed wording is too strict, because it may exclude even reasonable technical solutions. We see, e.g., no problem to reject a confusing technical solution (as you explained) by applying 4.2.1 (revised wording). |
| 4.2.2.1 | CN | Change“…1 x 10k, 2 x 10k or 5 x 10k…”in 4.2.2.1 into “…1 x 10k, 2 x 10k or 5 x 10k…”. | P+ |
| 4.2.2.2 | NL | Non-significant zero: When dealing with multi-interval/-range instruments the current proposal will lead to a misunderstanding in the reading the weight from the display when switching to a greater (partial) range. The current proposal mentioned in the draft is as follows: Only for the lowest (partial) range the non-significant zero requirement applies. For the greater range this requirement does not apply. For the user/operator it is not immediately clear (despite several markers, indicating that the instrument has changed to another (partial) range) The following is done/interpreted by NMi: Example Range 1: e1 = 50 g and Max1 = 150 kg Range 2: e2 = 100 g and Max2 = 300 kg Indication of weight in display: xxx,xxx kg When operating in range 1, there is only one non-significant zero present. The display shows xxx,050. When switching to a greater range there are now 2 non-significant zeros. The display shows xxx,100. This is not allowed according Article 4.2.2.2. This has been solved as follows: The decimal marker has a fixed position. When switching to another (partial) range the decimal stays in the position when the instrument was turned on. When switching to a greater (partial) range the last two digits are turned off and not used until switching to the lower (partial) range. The display now shows xxx,1. We propose to maintain the procedure mentioned in the above mentioned example, because: For a user/operator it is (extra) clear that the instrument has changed to another (partial) range. The indication on the display change from xxx,050 to xxx,1. For the greater (partial) ranges the requirement of the non-significant zero does not apply, since there are no non-significant zeros present. | We do not agree, because the 1CD in 4.2.2.2, last sentence, says that the requirements of the 5th para apply only to the smallest (partial) weighing range, not to the greater ranges. Therefore, xxx,100 is permitted for multi-interval instruments and multiple range instruments with automatic change over. We suggest, however, to add two examples in 4.2.2.2 in order to avoid misunderstandings (see 2CD). P- We hope that the explanation and the examples solve your concerns |
| 4.2.2.2 | UK | Do we need to specify that the decimal sign should be on one line with the bottom of the figures, especially if a decimal point is used instead of a comma? | P- This was included on request of a TC9/SC1 member |
| 4.2.3 | JP | Does the last sentence “Indications below zero in the order of - Min are not permitted unless a tare device is in operation” mean the values less than zero and Min at minus side? | Yes. Please refer also to 4.2.3 ZA and 4.2.3 DK |
| 4.2.3 | DK | Paragraph 4.2.3 states that on a multiple range instrument there shall be no indication using ei above Maxi = ni x ei for the lower partial weighing range(s) i, for automatic change over. We find this inconsistent. Indication up to Maxi + 9ei for each weighing range is generally allowed, and it must be concluded that the authors of the standard have found this to be beneficial for the owner/operator. Anong other things, it allows the owner of the instrument to determine the error of the instrument at Maxi, without having acces to a number of small weights. This could be an issue for the owners operating a quality system. Automatic range change over occurs only once during a weighing operation. The change over operation, automatic or not, should not be compared with that of a multi-interval instrument. It is obviously confusing if a multi-interval instrument is changing resolution back and forth at Maxi + 9ei. There is no such confusion when operating a multiple range instrument. There is obviously no metrological reason, either, for having automatic change over limited to Maxi, since operation up to Maxi +9ei is otherwise generally allowed. Paragraph 4.2.3 also states that indications below zero in the order of –Min are not permitted unless a tare device is in operation. Is there any metrological reason for this new requirement? A requirement like this should not be left open for interpretation. “In the order of –Min” is not precise. If such a limit on indication below zero is required, then it should be unambiguously definded. Otherwise inspectors may fault an instrument based on a whim. | This has been discussed intensively, and the new wording as it stands in the 1CD has obviously been accepted by a majority of TC9/SC1 members. see also 4.2.3 UK and ZA The reason is that one TC9/SC1 member country requested a requirement for indications below zero for instruments with a digital display. Normally (and unless a tare function is in operation) indications below zero cannot be allowed because there is no error regime defined. But for practical reasons an instrument must be allowed to make use of a small part of the negative weighing range in order to find stable equilibrium. It is now arbitrary to define a proper criterion. We thought that “in the order of -Min” would help, but because of some comments we now suggest “-20d” and the new wording in the 2CD. |
| 4.2.3 | UK | Says that for a multiple range instrument with automatic changeover it is not allowed for a range to change to the next highest at Maxi+9ei. WELMEC 2 guide Section 3.1.36 accepts that it may be done, although prefers it not to be done. What does the sentence “Indications below zero in the order of -Min are not permitted unless a tare device is in operation.” mean? It seems to imply that indication below zero, not “in the order of -Min”, is allowed. | Yes, you are right. We hope that this has been clarified by the new wording in the 2CD (see also 4.2.3 DK and ZA) |
| 4.2.3 | ZA | 3rd Par: Delete “not”: Indications below zero in the order of –Min are not permitted unless a tare device is in operation. | P+ But we suggest a slightly different wording: “Indications below zero are not permitted unless a tare device is in operation.”, and to add the “Note: Nevertheless, a temporary indication of negative numbers down to –20d is accepted.” |
| 4.2.4 | CECIP | Please include definition of “approximate indication device” in terminology. | P- This is a supplementary indication device providing only secondary indications, and thus not being subject to requirements of this recommendation, except No 4.2.4. Not to be mixed up with bar graphs. |
| 4.3.2 | CA | “- on an instrument of Class I or II: …..in this case i0 is the relative movement displacement between…” The word ‘movement’ should be replaced by ‘displacement’. It is not movement that is desired but rather displacement of the indicating component. | P+, see also 3.8.1 |
| 4.4.1 | CA | Change of Indication After a change in load, the previous indication shall not persist for longer than 1 second. Suggested wording change After a change in load, of a magnitude sufficient to overcome the threshold of sensitivity of the instrument, the previous indication shall not persist for longer than 1 second. Rationale Placing a load, too small to be detected, on the instrument would cause an issue with this section in its current wording. Canada would prefer to see clarification. | P- It is obvious that changes > 1,4d (discrimination test) are meant. For these changes the present wording is correct and should not be supplemented. |
| 4.4.2 | UK | The first sentence of this section is imprecise; what is meant by “sufficiently close to”? The final sentence is also imprecise, does this mean that an instrument must store or print a weighing value within 1e of the final weighing value even while there is continuous or temporary disturbance? Our view is that this is not what is meant. Delete the first sentence. Replace the word “under” with “after”. If the change) above is not accepted: For both bullet points, would it be better to put “In the case of...” rather than “In case of...”, as the latter has a rather different meaning. In the final sentence, if there is continuous disturbance then the two conditions can not be met. Under continuous disturbance the instrument should not print/store data, nor should it be able to set zero or tare. | P- The first sentence is a kind of introduction. The following sentence and bullet points clearly define what is meant. P+ In order to avoid possible misunderstandings we suggest to change the wording of the final sentence (see 2CD). |
| 4.4.3 | CA | Extended Indicating Device The last sentence in this section should be amended as follows: In any case printing shall not be possible while the extended indicating device is in operation. Rationale: The current wording is too broad. Canada agrees with Secretariat’s response. | P+ |
| 4.4.4 | AU | 4.4.4 Multiple use of indicating devices Change this to say: Indications other than primary indications may be displayed in the same indicating device, provided that: any additional indications do not lead to any ambiguity in regard to the primary indications. quantities other than weight values are identified by the appropriate unit of measurement, or symbol thereof, or a special sign or designation, weight values that are not weighing results (T.5.2.1 through T.5.2.3) shall also carry the appropriate unit of measurement, or symbol thereof and shall be clearly identified, or they may be displayed only temporarily on manual command and shall not be printed. If the instrument is set into a mode in which weighing is clearly and unambiguously not taking place (in the case of instruments intended for trading direct with the public this shall also be clear and unambiguous to the member of the public), the above restrictions do not apply. Explanation: There is potential for this requirement (in its previous form) to be used to justify display arrangements and indications which are ambiguous or potentially fraudulent. It is was not clear what constitutes a ‘special command’ and in what way this differs from any other command. | P+ P+ P- this could lead to problems; “clearly identified” is sufficient. P+ However, we suggest a new wording that is closer to the previous wording (see 2CD, last sentence under 4.4.4): “No restrictions apply if the weighing mode is made inoperative and this is clear and unambiguous (including the customers in case of direct sales).” |
| 4.5.1 | UK | The Note, about Class IIII, has been put into the body of the text, whereas it used to be a footnote. That footnote made it clear that it only applied to the sentence “The overall effect...” as it referred to “This provision...”. This Note now seems to apply to both sentences above the Note. Also, does the Note now apply to the sentence below it? | P+ The note applies only to the 2nd para. This has been reflected in the 2CD. |
| 4.5.2 | ZA | Add: …shall be not more than 0.25e, however on an instrument with auxiliary indicating devices this effect shall be not more than 0.5d. | P- 1. There was a respective demand to define a requirement in terms of e only; 2. There is really no need to have a more stringent requirement for instruments with digital displays having e>d. |
| 4.6.3 | ZA | Add: ±0,5 d for mechanical instruments with digital indication and instruments with auxiliary indicating devices. | P- see above |
| 4.6.5 | AU | 4.6.5. Visibility of operation (of a tare device) Add the following: Where the tare value (or preset tare value) is displayed other than temporarily, the value shall include the appropriate units (or symbol), and shall be designated TARE (an alternative word according to the appropriate national language may be used). The location and arrangement of the value, units and designation shall be such as to remove any possibility of confusion with the (NET) weighing result. Where the instrument is not to be used for trading direct with the public, the symbols T or PT (in the case of tare and preset tare respectively) may be used instead of TARE. Explanation: We have recently seen an instrument intended for trading direct with the public in which the tare display (designated PT although not only preset tares were displayed) was presented in exactly the same manner as the net weight value clearly causing confusion and ambiguity. Symbols such as PT or T are not known to the general public, and therefore should not be used on instruments intended for trading direct with the public. | P- This would be a rather important change. As there was no such vote at the beginning of the revision, the acceptance by a majority of TC9/SC1 members is unlikely. |
| 4.6.11 | CECIP | It would be helpful to give a hint to the regulation concerning preset tare in 4.7 already in 4.6.11. When not having read 4.7 already the example 4.6.12.5 is a bit confusing. | P+ The heading in 4.6.12.5 has been supplemented: “… with a preset tare device (4.7)” |
| 4.6.12 | CH | In order to shorten the recommendation, consider moving examples (like in 3.3.3, 4.6.12) or remarks to the selection (like in 3.10.4.5) into an Appendix or into a separated guide. They are neither metrological nor technical requirements. | P- See our response under 3.3.3 and 3.10.4.5 CH |
| 4.6.12 | UK | Footnote 3: “...in in...”. | P+ |
| 4.6.12 | US | The U.S. recommends that an example should be included for a multiple range instrument with tare weighing or preset tare where e1would be rounded to zero in e2 or e3 to comply with 4.7.1. and proposed footnote 3 below. Lack of clarification as to the proper rounding of tare has resulted in conflicting opinions regarding rounding tare to zero. Are these instruments required to round up to the nearest e to avoid sales by gross weight? Possible example: 4.6.12.X Multiple range instrument with a tare-weighing device Specifications of the instrument: Class III, Max1 = 60 kg, e1 = 10 g Max2 = 300 kg e2 = 100 g Max3 = 500 kg e3 = 500 g Unloaded instrument,indicated value in weighing range (WR) 1 = WR1 0,000 kg Loading with tare load, internal value = 0,448 kg, rounded and indicated value = WR1 0,450 kg 1) After releasing tare-balancing, indicated net value = WR1 0,000 kg Net Loading with net load, internal value = 312,753 kg, rounded and indicated net value = WR3313,000 kgNet 1)2) with automatic change over to weighing range 2, the tare-weighing value shall be rounded to the actual e of weighing range 3, rounded tare-weighing value = WR3 0,000g 2)3) Total loading, internal value = 313,201 kg rounded and indicated (if possible) gross value = WR2 213,000 kg 1)2) Possible printouts acc. to 4.6.11: a) 313,000 kg B (or G)313,000 kg N 0,000 kg T 2)4) b) 313,000 kg 313,000 kg N 0,000 kg T 2)4) c) 313,000 kg N 0,000 kg T 2) d) 313,000 kg N 2) e) 313,000 kg 2) (Note: A 0.15 percent overregistration of the net weight is introduced if tare is permitted to round to zero) An example of the same instrument with a required minimum tare of 1 e (tare is prohibited from rounding to zero) Loading with tare load, internal value = 0,448 kg, rounded and indicated value = WR1 0,450 kg 1) After releasing tare-balancing, indicated net value = WR1 0,000 kg Net Loading with net load, internal value = 312,753 kg, rounded and indicated net value = WR3312,500 kgNet 1)2) with automatic change over to weighing range 2, the tare-weighing value shall be rounded to the actual e of weighing range 3, rounded tare-weighing value = WR3 0,500g 2)3) Total loading, internal value = 313,201 kg rounded and indicated (if possible) gross value = WR2 212,500 kg 1)2) Possible printouts acc. to 4.6.11:a)313,000 kg B (or G)312,500 kg N0,500kg T 2)4)b) 313,000 kg 312,500 kg N 0,500 kg T 2)4) c) 312,500 kg N 0,500 kg T 2) d) 312,500 kg N 2) e) 312,500 kg 2) (Note: A 0.08 percent underregistration of the net weight is introduced if a minimum of 1 e tare is required to comply with net weight laws and regulations) Footnote 3: There is an extra “in” (. . . which is in in operation.) | P- We think the example in 4.6.12.3 is sufficient. No, in any case the rounding shall be correct. For example: If e = 10 g, an internal value of 0,444 kg shall be rounded to 0,440 kg and internal value of 0,445 kg shall be rounded to 0,450 kg. (Your example is printed in correct format at the last page of this summary) Your proposal corresponds with ours in 4.6.12.3, but in our example the tare value is not close to zero. However, the way of rounding in your proposal shall be the same as in 4.6.12.3. The indicated tare value in WR1 = 0,450 kg shall be rounded correctly to WR3 = 0,500 kg – not to 0,000 kg. But if WR1 = 0,240 kg then the rounding shall be WR3 = 0,000 kg. Your example would not be correct in the sense of R76. The correct values would be: Tare load = 0,448 kg indication WR1 0,450kg Tare balancing indication WR1 0,000kg Net load = 312,753 kg, indication WR3 313,000kg Net Total load = 313,201kg Gross (if displayable) indication WR3 313,000kg Printouts: a) 313,000kg B (or G) N 313,000kg N 0,500kg T This seems to be a mathematical contradiction, but all three weighing values (tare, net and gross) are correct rounded and within their mpe. A way out may be to distinguish one value as calculated value, e.g. “C 312,500kg B” Conclusion: No change because the example in 4.6.2.3 is sufficient and further examples will enlarge the R76 too much. P+ |
| 4.6.12.6 | US | This example does not appear to correspond with the definition of the term T.5.3.2 Calculated weight value since the tare value in the example is not identified as a preset tare as a possible printout. | P+ See new example in 2CD |
| 4.7 | CA | Preset Tare Device Measurement Canada suggests that wording be included in this section (or as appropriate) to prevent the use of long term stored tares in situations where their use is likely to result in incorrect measurement. For example, storing tare values for road vehicles, which are subject to weight changes from fuel consumption, different loading, different trailer combinations, mud & snow, etc. is likely to cause large measurement errors. Suggested Wording: 4.x.x - Preset tare shall only be used in cases where the tare value is stable, well known and unlikely to change over time. In all other cases, the actual tare weight must be determined for each transaction. Rationale Measurement Canada has conducted a study of those vehicle scale installations where stored tares are commonly being used. The stored tare values were often very far from the actual measured tare weights of the vehicle. Measurement Canada is currently considering a legislative change to address this issue and feels that it is likely an issue elsewhere as well. The ‘Scope’ of the Recommendation does not seem to indicate that it is applicable to ‘type evaluation’ only. The recommendation will be used as the basis for all evaluations if adopted in Canada. | P- Although we know the problem we think that this cannot be tested on type evaluation; it is rather a matter of honesty (or legal control) of the user and/or other person responsible for a correct weighing. P- The scope of this recommendation does not make proposals concerning users surveillance. See Chapter 1. |
| 4.7.1 | FR | We suggest to make it coherent with revision of R 51. | P+ Has been made coherent to R51 (DR 2005) by the amendment “…, and the maximum preset tare value shall not be greater than Max1”. |
| 4.7.1 | US | U.S. manufacturers have recently reported that rounding of tare weighing results are inconsistently applied when a tare weighing result is stored into a temporary memory while the operator is working with more than one customer at a time (e.g. vehicle weighing applications). One issuing authority stated that the tare weighing result became a preset tare since it was stored into a temporary memory location while the vehicle was being loaded and therefore rounded to the nearest scale division. During the time the vehicle is loaded, the operator can weigh other vehicles (this is called a weigh-in/weigh-out procedure in the U.S.). A different issuing authority stated that the tare value should remain stored to the internal resolution of the instrument since tare was determined on a tare weighing device. The U.S. agrees that a predetermined tare introduced into an instrument by means such as keystroke, code, and stored in memory for multiple transactions should be considered a preset tare value and rounded to the value of the scale division. The U.S. also agrees that a net weight can be calculated as the difference the gross and tare weights taken to the internal resolution of the instrument. However, we do not understand the reason for the rounding of tare weighing results that are automatically entered into a temporary memory, along with other customer and transaction information, when no rounding is required for price computing instruments. The U.S. recommends that tare weighing values automatically entered into a memory device should not be considered as preset tare values and requests that the Secretariat and the members of TC9/SC1 consider renumbering and amending sub clause 4.14.4.3 as follows: 4.14.4.3 4.20 Multi-vendor or multiple customer operation *) An instrument may be designed to be used by more than one vendor or to serve more than one customer at the same time, provided that the connection between the transactions and the relevant vendor or customer is appropriately identified. An instrument may perform this additional function only if all transactions performed by the instrument or by connected peripherals are printed on a ticket or label intended for the customer. This function shall not lead to confusion about the results of weighing and transaction. (Note: Paraphrased from 4.14.4 Special applications of a price computing instrument) *) 4.7.1 does not apply to tare weighing values in these instruments. | P- We think that there is as a misunderstanding of the preset tare function; therefore we tried to improve the terminology in T.5.3.1, and we hope that this solves the problem. We cannot, however, agree with the suggested amendment of 4.14.4.3 / renumbering into 4.20, because for such a significant change the support of TC9/SC1 members is unlikely. |
| 4.7.2 | CN | Change“… automatically onhy if the preset…”in 4.7.2 into“… automatically only if the preset…”. | P+ |
| 4.7.2 | ZA | Last Par:…may operate automatically only onhy if the preset tare value…. | P+ |
| 4.7.2 | UK | “onhy” should be “only”. | P+ |
| 4.7.2 | US | Typographical error in the last sentence. Preset tare devices may operate automatically onhy only if the . . . | P+ |
| 4.7.2 | CA | Spelling mistake “Preset devices may operate automatically onhy only if the preset tare…” | P+ |
| 4.7.3 | UK | Refers to the 4th paragraph of Section 4.6.5 not being required (for preset tare), but it should probably refer to the 2nd paragraph of Section 4.6.5, as the Notes should not be classed as paragraphs. | P+ see changes in 2CD |
| 4.10 | CN | In 4.10, we suggest adding the word “equivalent to gross zero” at the back of “…a negative net value”, i.e. “- from a greater to a smaller weighing range, when there is no load on the load receptor, and the indication is zero or at a negative net value equivalent to gross zero; the tare operation shall be cancelled, and zero shall be set to ± 0.25 e1, both automatically.” “- only from a greater to the smallest weighing range when there is no load on the load receptor, and the indication is zero or at a negative net value equivalent to gross zero; the tare operation shall be cancelled and zero shall be set to ± 0.25 e1, both automatically.” | P+ P+ |
| 4.11.5 | JP | As for the last sentence “… but only if the correct zero position has been automatically checked before”, what does this mean? Example of this case is requested to give. | The meaning is: If the instrument (or the display) is in the switched-off mode and is then loaded, the instrument will normally switch-on automatically and display the actual weight value. This cannot be permitted unless the operator or user of the instrument has checked the instrument for the correct zero position before loading. An alternative to zero checking by the operator is that the instrument itself checks itself automatically for the correct zero position in the switch-off mode and prevents the display of weight values if the zero position was not correct. |
| 4.11.5 | UK | Delete the sentence “In any mode and at any time it shall be possible to switch back to the normal weighing mode” as this is a basic requirement of the instrument. | P- We think this hint is important. |
| 4.13.4 | CA | Preset Tare Device The last sentence of this section should be amended as follows: Where a preset tare is associated with a price look up (PLU) the preset tare value may must be cancelled at the same time as the PLU is cancelled. Rationale If a PLU is cancelled and there was a tare associated with that PLU, then you would expect the tare to be cancelled as well. Failure to do so will ensure that the wrong tare value remains in the machine, leading to potential incorrect measurement on subsequent transactions. Agree with Secretariat’s response, however, if the preset tare is acssociated with the PLU, what rationale could there possibly be for keeping it after the PLU is cancelled? | P+ We agree but suggest, however, to use “should” instead of “must”, because there were no serious problems in the past, and “must” would probably not be supported by a majority of TC9/SC1 members. Exceptions should be possible, because one can think of special cases where the PLU is changed while the preset tare value remains. |
| 4.13.6 | AU | 4.13.6 Visibility Change to the following: All primary indications shall be indicated clearly and simultaneously to both the vendor and the customer. Refer also to 4, 4.2.1, 4.2.2, 4.6.5 and 4.4.4 which stress the need to avoid ambiguity in regard to the indications provided. On digital devices that display primary indications, the numerical figures displayed to the customer shall be at least 9.5 mm high, except in the case of any displayed tare or preset tare value which shall be at least 6.5 mm high. The reduced size allowance for the tare or preset tare value is intended to provide additional possibilities for distinguishing the tare or preset tare value from the net weight, to reduce ambiguity. The size of designations such as TARE, NET and the size of any zero indication device shall be at least 4 mm high. These sizes are intended as an absolute minimum – manufacturers are strongly advised to design to achieve larger sizes than this. On an instrument to be used with weights it shall be possible to distinguish the value of the weights. Explanation: Currently there is no specification regarding size of such designations so that manufacturers do not have a clear understanding of what is acceptable, and approval authorities do not have objective criteria for rejecting instruments where the size of such designations has been reduced to be ‘too small’. The readability of displays is suffering as a result and this needs to be remedied. In regard to the size of the indication of a tare value, whilst we would prefer that all indications be at least 9.5 mm, we feel that the need to ensure a clear distinction from the net weight value is of greater importance, and hence have proposed allowing a small size for the tare value. Note that this problem has arisen as instruments are increasingly incorporating preset tare devices, so that the value must be displayed in accordance with 4.13.4. | P+ But slightly changed because all applicable requirements of chapter 4 have to be observed (see also 4.13.6 ZA). P- 4.2.1 is also valid for this chapter: The primary indications shall be of a size, shape and clarity for reading to be easy, see our response to 4.2.1. |
| 4.13.6 | ZA | In terms of this clause an instrument with only one set of primary indications will be acceptable if both the vendor and the customer are on the same side of the instrument, our interpretation. According to the secretariat’s response on previous WD “Two sets are necessary” We propose the following to clear any confusion: All primary indications shall be indicated clearly to both the vendor and the customer, and in instances where the vendor and customer are on opposite sides of the instrument when the measurement takes place, primary indications on both sides shall be provided. | You are right, sorry for confusion. P+ We suggest, however, a shortened text with references to 4.13.1 and 4.14.1, see 2CD. |
| 4.13.11 | AU | 4.13.11 Self-service instrument Add the following: A self-service instrument (when in normal use) shall have a limited set of functionality and shall be designed such that it can be used easily by an unskilled operator. Potential for incorrect operation resulting in incorrect weighing shall be minimised. Explanation: It seems self-evident, however it is becoming apparent that sometimes the obvious does need to be spelt out. | P- No need to spell out what is really obvious and what the market regulates by itself. |
| 4.14.4.1 | AU | 4.14.4.1 Non-weighed articles Add an acceptable solution: Acceptable solution: A number of articles shown on the weight display may be distinguished from a weight by including a designation such as PC or PCS (indicating piece or pieces). The price for one article on the unit price display may include a designation such as EA (indicating each). Explanation: The use of a number in the weight display (for an unweighed item) has been seen. This is not entirely unambiguous and could possibly be taken for a weight. Manufacturers should be guided to options such as PC which cannot be confused with a weight value. | P+ We suggest, however, to restrict the acceptable solution to the number of articles only (see new acceptable solution in 4.14.4.1, 2CD). |
| 4.14.4.3 | US | U.S. manufacturers have recently reported that rounding of tare weighing results are inconsistently applied when a tare weighing result is stored into a temporary memory while the operator is working with more than one customer at a time (e.g. vehicle weighing applications). One issuing authority stated that the tare weighing result became a preset tare since it was stored into a temporary memory location while the vehicle was being loaded and therefore rounded to the nearest scale division. During the time the vehicle is loaded, the operator can weigh other vehicles (this is called a weigh-in/weigh-out procedure in the U.S.). A different issuing authority stated that the tare value should remain stored to the internal resolution of the instrument since tare was determined on a tare weighing device. The U.S. agrees that a predetermined tare introduced into an instrument by means such as keystroke, code, and stored in memory for multiple transactions should be considered a preset tare value and rounded to the value of the scale division. The U.S. also agrees that a net weight can be calculated as the difference the gross and tare weights taken to the internal resolution of the instrument. However, we do not understand the reason for the rounding of tare weighing results that are automatically entered into a temporary memory, along with other customer and transaction information, when no rounding is required for price computing instruments. The U.S. recommends that tare weighing values automatically entered into a memory device should not be considered as preset tare values and requests that the Secretariat and the members of TC9/SC1 consider renumbering and amending sub clause 4.14.4.3 as follows: 4.14.4.3 4.20 Multi-vendor or multiple customer operation *) An instrument may be designed to be used by more than one vendor or to serve more than one customer at the same time, provided that the connection between the transactions and the relevant vendor or customer is appropriately identified. An instrument may perform this additional function only if all transactions performed by the instrument or by connected peripherals are printed on a ticket or label intended for the customer. This function shall not lead to confusion about the results of weighing and transaction. (Note: Paraphrased from 4.14.4 Special applications of a price computing instrument) *) 4.7.1 does not apply to tare weighing values in these instruments. | P- See our response to 4.7.1 US P+ Instead of paraphrasing 4.14.4 we think, however, a simple reference to 4.14.4 could be sufficient, see 2CD. |
| 4.16 | CN | For this fixed values of weight in 4.16 mentioned –“Printing of labels with fixed values of weight, unit price and price to pay is allowed provided that the weighing mode is made inoperative.” When printing it is should expressed that the weight value is theoretical weight or the permissible deviation of weight. | P+ “Obviously” has been added in 2CD |
| 4.18.1 | CA | See also 3.9.1.1 Instruments liable to be tilted. The first sentence reads in part: “The instrument shall have appropriate means to indicate that ther limiting value has been reached or exceeded.” As in 3.9.1.1, the words reached or should be struck out and only exceeded left in. | P+, see also 3.9.1.1 |
| 4.18.1 | NL | The last part of the paragraph about the tests for the tilt sensor should be moved to clause 8 Last sentence: “The OIML Test Report ....” should be moved to Annex A. | P- See our response under “General” |
| 4.18.1 | PL | paragraph 7 line 2: one of double ’the’ should be deleted | P+ |
| 4.18.1 | UK | seventh paragraph: “...the the...”. The *footnote relating to “fish auction halls” is too specific. In addition, “in-door” should be the single word “indoor”, as “in-door” suggests inside that it is inside a door. | P+ P+ “halls with uneven floors” P+ |
| 4.18.2 | UK | This should be re-written as at present it implies that other mobile instruments need not have a levelling device. These instruments shall have a levelling device which can be operated easily without tools. | P+ We propose, however, a slightly changed wording, see 2CD |
| 4.19 | AT | Portable instruments | P+ |
| 4.19 | CA | The second paragraph specifies that the applicant must provide a description of the appropriate mounting surface; however, the appropriateness of such a mounting surface and other metrologically significant installation-dependent parameters (e.g. length and angle of approaches) should be left to the discretion of the national authority. This should be specified in the Recommendation. Canada agrees with Secretariat’s response. The two Notes in the 1 CD are very ‘wordy’. Would it simply be better to state “the load must be entirely supported by the load receiving element(s)”? Spelling 4.19 “Portable Weighing instuments instruments for weighing road vehicles” | P+ Will be considered, although wording has already been changed in the 1CD P- We understand your point of view, but in order to avoid possible misunderstandings the wording should be kept. P+ |
| 4.19 | CZ | This section does not clearly state the purpose of these instruments. Axle/Wheel wheighers are originaly designed for traffic control of road vehicles to determine not only the total load but also the axle loads. If only instrument using groups of associated axle or wheel load weighers that indicates only total load of a vehicle is allowed (and we understand the paragraphs of this section this way since it is not clear that also indication of axle loads is legaly allowed) such instrument is to be used for trade purposes to determine mass of a vehicle only similary as, let us say, common weighbridges. This probabaly should not be the purpose of this section. On the other hand we understand that to set up simple rules for this kind of instruments to be legaly controled is nearly impossible. From the reasons we think that treating of these instruments should depend on national legislations and consequently we propose to withdraw these instruments from the scope of this recommendation and only mention in this section that the use and the requirements are covered by national legislations. | P- There was a clear vote to integrate these instruments into R76. By both Terminology T.1.2.12 and Chapter 4.19 we think it is sufficiently clear that R76 covers only weighbridges and groups of associated non-automatic axle (or wheel) load weighers that determine simultaneously the mass of a road vehicle with all axles (or wheels) being simultaneously supported by appropriate parts of a load receptor. |
| 4.19 | NL | 1st paragraph about application (for approval or for verification?) and OIML certificate: This is not a requirement for the instrument (see title of clause 4). So this should be moved to chapter 8. | P+ see 2 CD P- see our response in “General” |
| 4.19 | UK | “instuments” should be “instruments”. | P+ |
| 5 | CA | 5 – Technical Requirements for Electronic Instruments Add the word “Technical” to the title to remain consistent with other sections. Canada agrees with Secretariat’s response. | P+ |
| 5.3.3 | NL | “shall be subjected to span stability test” (type test and/or verification?). This first sentence should be moved to clause 8. | Unambiguously for type approval testing only, not mentioned under 8.3.3. P- Please refer to our response under “General” NL |
| 5.4 | NL | These tests (type tests or verification ?) should be moved to clause 8. | P- See 5.3.3 NL |
| 5.4.3 | SI | In Table 10 some tests from B.3 are missing: surge, immunity to radiated EMF and conducted disturbances (instead of electromagnetic susceptibility), transients. | P+ |
| 5.5 | JP | Should the title “Additional requirements for software-controlled electronic deices” be “Additional requirements for software-controlled electronic devices for legal metrology”. | P- It is sufficiently clear that R76 applies to instruments under legal control only |
| 5.5.1 / 5.5.2 | CA | Devices with embedded software/Personal Computers … with loadable software The definitions for embedded software and loadable software are somewhat confusing. Clearly, a ROM is embedded, however is an EEPROM embedded or loadable? What about Flash RAM? Even newer PC’s use Flash RAM for their BIOS settings. The phrase ‘cannot be modified or uploaded via any interface’ also causes concern as all types of PROM can be changed, although some more easily than others – most require external programmers. Flash RAM however can be changed within the instrument/PC. Unfortunately, as the distinction between an instrument and a PC continues to become less clear, it will be more difficult to define each in legislation. Perhaps this is best left to each approval laboratory to define? To clarify then, all software that can be ‘secured’ is considered ‘embedded’ even if it can be changed after the securing provision has been by-passed? If this is true, then why not simply state in 5.5.1 that the software is ‘secured’ rather than ‘embedded’? The current wording seems to indicate that there are only hard wired instruments and PC’s, when in fact many newer devices are programmable (with sealing &/or audit trails), but would not be considered PC’s. Perhaps it is the different securing philosophy that is causing issue. Canada does not require that all parameters be protected against change, provided that the change is identificable as with an audit trail. The wording in the CD seems to require these devices to be treated as PC’s. | We don’t think so. “Embedded software” clearly means the software itself which can, of course, be stored in any memory component (ROM, RAM, Flash etc.). We do not quite understand the concern, because for the embedded software the consequence simply is that the hardware component (PROM etc.) in which it is stored must be appropriately secured so that after securing the software can no longer be modified or uploaded. Otherwise the software will no longer be “embedded” and para 5.5.2 would apply. P- We agree that the distinction between a “built-for-purpose” instrument and a PC continues to become less clear. We do not agree, however, that the definition should be left to the approval laboratory. We think with the addition to the Note under 5.5.2 it is now sufficiently clear that if the software is not “embedded” in the sense of 5.5.1 the instrument is considered as a “PC” so that 5.5.2 applies. It should be sufficiently clear from the heading of 5.5.2 that this para not only covers “PCs” but also “instruments with PC components, and other instruments, devices, modules, and elements with programmable or loadable legally relevant software”. “Embedded” (in 5.5.1) cannot be exchanged by “secured” because the software of “5.5.2 devices” must, of course, be adequately secured, too! |
| 5.5.1 | CECIP | Please add a third acceptable solution: displaying the software identification at start-up of the instrument. | P- The two acceptable solutions provided seem to be sufficient. In addition, displaying the SW identification only at start-up can normally not be considered sufficient, if there is no possibility to display the software identification also in the normal operation mode. |
| 5.5.1 | CN | Change“…by the instrument amd listed in the OIML ...”in 5.5.1 into“…by the instrument and listed in the OIML ...”. | P+ |
| 5.5.1 | JP | Should the title “Devices with embedded software” be “Devices with embedded software for legal metrology”? When software data related with legal functions are stored in ROM with hardware protection means, could the contents of the documentation required to submit be reduced from those listed? | P- R76 concerns instruments under legal control only No, because the list has already been reduced to the minimum. |
| 5.5.1 | NL | The necessary tests and the documentation to be submitted (for type tests or verification?) are no requirements for the device but a matter of tests/evaluation. So this information should be moved to clause 8. | P- Please refer to our response under “General” and 5.3.3 |
| 5.5.1 | ZA | Last sentence: …by the instrument and amd listed in the OIML Certificate. | P+ |
| 5.5.1 | UK | (line beginning “The software identification...”): “amd” should be “and”. Although this section refers only to “embedded software” which “...cannot be modified or uploaded...”, it then demands software descriptions, identifications (continuously or on demand), and securing measures. Is this necessary? | P+ P- No software description is required, but only limited descriptions of the legally relevant functions, software identification and securing measures are required. We think this is the absolute minimum that must be required for software-controlled electronic instruments and devices. |
| 5.5.2 | JP | According to the description in this section, even the instruments combined with measuring modules and general purpose PC’s could be understood to be legal measuring instruments, when they meet the requirements. With this expansion of the scope, is it possible to control those instruments legally? | We think there is a misunderstanding: R76 does not define whether an instrument is under legal control or not. This is defined by national laws or regulations. But if national regulations (based on R76) define that a NAWI is under legal control for a special application, then para 5.5.2 would apply for a “PC” or similar software-controlled device. |
| 5.5.2.1 Table 11 | CECIP | Category 3: Power supply as for category 2, other parts as for category 4 | P+ See changes in table 11 |
| 5.5.2.1 Table11 | NL | The necessary tests and the documentation to be submitted (for type tests or verification?) are no requirements for the device but a matter of tests/evaluation. So this information should be moved to clause 8. The word “pattern” should be replaced by “type” (refer to VIML). | P- Please refer to our response under “General” and 5.3.3 P- No, it is really meant “pattern” for testing and not “type” |
| 5.5.2.2 | JP | In item d, total 11 special software documents are required in addition to the documentation outlined in 8.2.1.2. The contents of those 11 documentations should be specified in more details. There may be some documentation that could not be submitted, as they are corporate secret. | P- The list is already rather long. It will not be possible to provide a complete list of all necessary documentation because it depends on the design of the special instrument. In most cases descriptions are required rather than secret documents like listings of the software code etc.. In addition, 8.2.1.2 requires that all documents shall be kept confidential by the approving authority without distributing it to others, except to the extent agreed with the manufacturer. |
| 5.5.3.2 | CECIP | The checksum or other signature of the legally relevant data defined in the note may not be used for the single data sets but for the complete data storage memory. Please delete the word “set” and modify the note to “…of the data stored”. Or alternatively modify to “… of the data set stored or of the complete storage memory.” | P+ See change in 5.5.3.2 |
| 5.5.3.6 | UK | What purpose is served by requiring the device that is capable of displaying or printing the data be subject to legal control. Surely it is better that the data can be displayed or printed on any suitable device, provided that the display or print command does not also cause the data to be deleted from the storage device. Delete this section. | P- If stored weighing results (legally relevant data) shall be verified with a display or printer, this device must be subject to legal control! It is the same principle as if weighing results of a weighing instrument are displayed or printed, it is evident that the instrument shall be subject to legal control. Otherwise the whole procedure is useless. |
| 6 | ZA | Our request to change the heading and first paragraph of chapter 6 which was made in our comment on the working draft was not accepted as the secretariat is not aware of misunderstandings in the past but this is what we are bringing to your attention. From the further reasons given we must deduct from your explanation that self- and semi-self indicating instruments are not included in clause 6 and therefore an instrument with a lever system, single load cell and digital indication may have knife edges, bearings and friction plates of any hardness as there is no requirements for these. This cannot be the case because clause 4.2.5 refers to chapter 6 for compliance of poises and there are many other instances where chapter 6 applies to self- and semi-self indicating instruments. This anomaly should then be resolved. | P- For self- and semi-self indicating instruments chapter 4 applies and instruments undergo type-approval testing. Therefore for knife edges, bearings and friction plates of such instruments the requirements of chapter 6 need not be applied Testing proves the suitability. |
| 6.1 | NL | The last sentence is no requirement but a test (type test or verification ?). So this should be moved to clause 8. | P- Please refer to our response under “General” NL |
| 6.3 | CH | Some remaining constructional design requirements should be removed (like for example in 6.3). | P- 6.3 contains requirements for certain simple instruments for submitting direct to initial verification (without type approval), and this provisions shall not be removed. |
| 7 | ZA | 3.9.3 refers to “between the lower and upper limit of the nominal voltage marked on the instrument (U)…..” In clause 7 there is not a requirement for the nominal voltage to be marked. Does this mean you don’t have to do the test if it is not marked? Suggestion: a) ….and upper limit of the nominal voltage marked on of the instrument (U)…. or b) In clause 7 make it compulsory if applicable for the nominal voltage (U) to be marked on the instrument. | P+ See also 3.9.3 ZA |
| 7.1.1 | CA | Additional marking is necessary: nmax is required to be marked for indicators, and power supply voltage is required to be marked for indicators and complete instruments. Canada defines Max in terms of mass (kg) while nmax is in terms of divisions (d). For instruments/modules, especially indicators, tested separately, nmax should be required information. | P- No such markings have ever been required as metrological markings both for complete instruments and indicators. P- See 7.1.5.3 |
| 7.1.2 | US | H): Typographical error in the first bullet: - maximum safe load in the form of Lim = . . . (if the manufacturer bhas provided for a maximum safe load of more than Max + T) | P+ |
| 7.1.4 | AU | 7.1.4 Presentation of descriptive markings (i) We do not agree that only showing the Class, Max, Min, e values on a manual command is sufficient. These paramaters can influence the acceptability of an instrument for various uses (according to national regulations) and hence we believe they should be visible to the user and the customer. (ii) We are concerned that if display of various items (markings and calibration codes) is provided on manual command, this is not sufficiently well defined. We have seen a trend for access to information (e.g. calibration code - but it is a trend which could extend to marking information) to be only accessible through a complex system of menus - the term 'manual command' could encompass this. This is unacceptable as neither an unskilled operator, nor a weights and measures official can be expected to negotiate their way through such complicated and differing arrangements. We would suggest that for information which may be acceptable to be accessed by a manual command (e.g. items other than in marking categories A & B, and also including a calibration code), the wording should be "accessed by a single simple manual command". | P+ This is almost in line with the new wording proposed in the 2CD: "The markings: Max, Min, e =, and d (if d unequal e) shall be shown at least in one place and permanently either on the display or near to the display in a clearly visible position..." Only the class will fall under the other provision, see our response to (ii). P+ Has been (almost) considered by the new wording: "All additional information as mentioned in B and G may be shown alternatively on a plate or simultaneously displayed by a software solution either permanently or accessed by a simple manual command." P- Concerning "a single simple manual command": This could lead to practical problems if you think e.g. of a menue function button that will open an "info menu" offering the additional information in question, which should be acceptable. |
| 7.1.4 | JP | As the means of sealing or security, lead is recommended. However, as lead is a hazardous substance, another way to express the recommendation should be considered, for example “…cap of suitable metal inserted …” in the second sentence of item c in 7.1.4. | P+ See 2 CD |
| 7.1.4 | UK | This now appears to allow as an alternative Max, Min, e (and d if appropriate) to be displayed on demand, and not permanently shown anywhere. Is this intended? In addition, can the word “near”, in relation to “near or on the display” be clarified? c): The first sentence of this section requires the use of at least one rivet made of red copper. Why? Plates which cannot be removed without being destroyed have been acceptable in the past, why should a requirement be introduced now that says such plates shall be riveted, screwed or otherwise secured using non-removable control marks. Even where plates have been riveted in the past there has been no requirement to use at least one red copper rivet. This appears to be overkill. Delete the words “with one of the rivets being of red copper or material having qualities recognised as similar” and replace with “or by other means such as being made of destructible material” | No. See new wording in 2CD which hopefully clarifies your point. Our intention was to say: Max, Min, e= and d= (if d(e) have to be shown at least in one position and permanently either on the display or near to the display in a clearly visible position. Automatically scrolling (but not on manual command) is considered as “permanently”. P- 1. This is an acceptable solution rather than a requirement. 2. The wording of this “old acceptable solution” has already been improved by introducing “non-removable control marks” instead of “red copper”. |
| 7.1.4 | US | The U.S. recommends that; - the presentation of 7.1.3 Additional markings, - the identification mark on separate but associated units, and - the presentation of special limits should be listed as acceptable software solutions to physical markings. The U.S. is unaware of any technical justification why the markings in 7.1.2. E (identification mark if separate but associated units), F (approval mark), and H (special limits) may not be displayed in lieu of physical markings. Except for the additional markings in 7.1.3, customers in direct sale transactions have no need to view the markings in E, F, and H. Additionally, the U.S. believes that the should not be required to be “simultaneously displayed.” The language should permit the presentations of markings to be scrolled or menu driven. As an alternative all applicable markings in B, E, F, G, and H may be simultaneously displayed by a software solution either permanently or on manual command. In this case the markings are considered as device-specific parameters (see T.2.8.4, 4.1.2.4 and 5.5). Additional markings in 7.1.3 may be simultaneously displayed by a software solution as an alternative to physical markings. An alternate proposal would be to recognize that the markings in E, F, H, and also D (serial number) could be provided by a software solution in the event that the markings are not located in a clearly visible location (e.g. under an easily removable load receptor, bottom of the instrument, etc.). Peripheral devices should not be held to the same marking visibility and grouping requirements in 7.1.4 as weighing instruments, main devices, modules, and other legally relevant devices. The examples of peripheral devices in T.2.3.5 are not manufactured with the intent that they be used exclusively or predominantly as a part of a measuring instrument or module. Their markings are not always located together or as clearly visible as they are for weighing instruments. For example, video displays, PC computers, and printers may have the name of the manufacturer on the front of the device with the type designation, serial number, and other information including (i.e., CE, FCC, UL, etc.) located on the back or bottom of the device. The U.S. recommends that the language in 7.1.4 be amended as follows to clarify that peripheral devices do not have the same visibility and grouping marking requirements as weighing instruments as follows: The descriptive markings shall be indelible and of a size, shape and clarity allowing easy reading. Except for the markings required for peripheral devices in 7.1.5.4, the descriptive markings They shall be grouped in one or two clearly visible places either on a plate or sticker fixed permanently to the instrument, or on a non-removable part of the instrument itself. In case of a plate or sticker which is not destroyed when removed, a means of securing shall be provided, e.g. a control mark that can be applied. | P- We do not think that we get worldwide acceptance of your proposal. However, we hope that the new wording (see 7.1.4 UK above) is an acceptable compromise for you. P- We think it is useful to have the information as far as possible grouped together. We do not see any practical problem to realise that. |
| 7.1.5.1 | UK | The word “viz” is slightly obscure. It would be better to replace it by some other word or phrase. | P+ |
| 7.2 | JP | As the means of sealing or security, lead is recommended. However, as lead is a hazardous substance, another way to express the recommendation should be considered, for example “…cap of suitable metal inserted …” in the second sentence of item c in 7.1.4. | P+ See 2 CD |
| 7.2 | SI | (acceptable solution) The sentence “These marks should be adequately durable for the intended use of the instrument, e.g. by means of a suitable protection” shall be deleted. This is not the attribute of the instrument and the producer of the instrument doesn’t have an influence to the quality of the marks. | P- R76 contains requirements to be observed not only by manufacturers but also by authorities like type-approval bodies and verification authorities. |
| 8.1 | CA | The third paragraph is misleading and wording should be changed as follows: “However, instruments according to 6.4 through 6.9 of this Recommendation shall not be subject to type approval according to this Recommendation, national legislation may provide for type and initial verification for particular instrument applications.” Canada worries that the first half of this paragraph may be interpreted to mean that all devices identified in sections 6.4-6.9 are automatically exempt from type approval. The second half which allows for national deviations is fine by us. | P- As there was no such request from any other TC9/SC1 member nor a vote to change this para, it seems very unlikely that the proposed change will find the support of a majority of TC9/SC1 members. |
| 8.1 | JP | The term, “in-service”, is used here and other places. This term was already proposed to be “service” by some countries in previous case. This word is to be revised to “service”. | P+ |
| 8.2.1 | UK | “moduls” should be “modules”. | P+ |
| 8.2.1.2 | AU | 8.2.1.2 Descriptive documents Comment: We do not believe that an OIML document should specifically refer to CE marking as this is specific to a limited range of member nations – the term conformity marks would seem sufficient – perhaps conformity and/or approval marks would be better. In addition whilst we recognise the need for confidentiality, the final statement goes way too far (it effectively suggests that the general characteristics at item 2 must be kept confidential for example). The approving authority may legally (perhaps not in Europe, but certainly elsewhere) need to be in a position to write and publish descriptions of the instrument operation (including perhaps extracts of the material provided) so that verification/certification personnel can identify the approved instrument (and detect unapproved equipment). In addition in a situation in which non-conformity has been detected an authority may need to present any or all of the material to legal authorities to initiate prosecution. Perhaps the sentence needs to say …shall normally be kept confidential by the approving authority, except to the extent that elements of the material may need to be included in approval documentation (with the agreement of the manufacturer/supplier). In addition material may need to be disclosed in the case of legal action (for example should non-conformity found). Also the list of descriptive documents does not actually say what is to be supplied, perhaps the following needs to be included prior to the list. Descriptive documents which detail and explain aspects of the instrument operation and design, including how various elements of this Recommendation are met shall be provided in addition to operational and technical manuals for the instrument. In the case of a family of instruments the descriptive documents shall address differences and similarities between members of the family. The following is a list of various aspects which should be addressed. | P+ See 2 CD where we have changed the wording under No 8.2.1.2 / point 4.2 into: “…securing elements; descriptive markings, identification, confornity and / or approval marks (7.1, 7.2)” P+ We propose, however, just to add to the last sentence of para 8.2.1.2 the following: “…, except to the extent agreed with the manufacturer.” P- The chapter contains already a very long list for necessary documents. |
| 8.2.1.2 | RO | We think that the documents needed to check the conformity of the instrument (where there are doubts or as provided at point 8.3 of the recommendation) should be made available in the annex to the OIML certificate. We understand the need of maximum confidentiality pending the placing on the market of a new type of instrument, but afterwards it is necessary to provide protection against non-compliant or fake instruments. We suggest to re-phrase as follows: “All documents of the weighing instrument (which are not published by the producer) shall be kept confidential by the approving authority, with the following exception: the documents needed to check conformity of the instrument to the approved type shall be published in the annex of the OIML certificate”. | P- See response to comment AU above |
| 8.2.1.2 | SI | Item 10: It is necessary to include the reference to 5.5.2.1 (Table 11). | P+ |
| 8.2.1.2 | UK | Number 10: It would be better to replace “acc.” by the full word. Also Section D.2.2 and D.3. | P+ |
| 8.2.2 | AU | 8.2.2 Type evaluation Replace the first two paragraphs with: The submitted documents shall be examined to verify that the descriptions indicate compliance with the requirements of this Recommendation. In addition checks shall be performed to establish that the instrument(s) supplied comply with the requirements of this Recommendation, and also in accordance with the documents submitted. The checklist in R76-2 shall be used to record the results of these checks. Reactions to significant faults need not be triggered. Explanation: The term ‘spot checks’ permits and indeed encourages a less than thorough pattern evaluation. The situation is developing where on the basis of a single evaluation of an instrument (just one in a family) an unlimited number of instruments may be distributed worldwide. If that single evaluation is not even to be carried out in a thorough manner then quite simply the majority of the requirements in this document may as well simply be eliminated. It is more important that the instrument does actually comply with the Recommendation rather than just the documents (note that it is stated elsewhere that the documentation may be in a draft form), we believe this clause should reflect that We have encountered a number of cases where authorities have apparently granted approval on the basis of documentation, where the actual instrument did not comply. | P+ We cancelled the word “spot” in 2 CD to follow the sense of your proposal. |
| 8.3 | AU | Comment: Matters of initial and subsequent verification are highly dependent on local legal systems, rules and procedures. In Australia certification (similar to initial verification) may be performed by other than ‘authorised national bodies’ or ‘the manufacturer’ – personnel of licensed companies can carry out this function. Such matters should be left to national regulations. Hence we believe the first paragraph should be changed to Initial verification may be performed by authorised personnel according to national regulations. It should also be recognised that in many cases the manufacture may not exist in the particular jurisdiction in which the instrument is to be installed. Much of the material in this section appears to be an attempt to write European procedures into the OIML documents – we do not support this. | P+ |
| 8.3 | CA | Suggest remove everything after authorized national bodies. This would include the manufacturer if they are an authorized national body. Editorial Change, last paragraph – “…provided that the instrument meets all of the requirements within this zone.” If a manufacturer’s “quality system for production is acknowledged for this task”, would that not make them ‘authorized”? Canada agrees with the Secretariat’s response. | P+ see 2 CD P+ We propose the following change: “…instrument meets the respective national or regional requirements.” |
| 8.3 | JP | When national regulations are identical, the products approved at initial verification should be mutually accepted. | P- We agree, but this proposal would go beyond the scope of R76, and there was no such vote. Of course, any country is free to accept your proposal under its national regulations. |
| 8.3 | NL | 5th par: Add after “The second stage shall be carried out at the place of use of the instrument.”: “If the instrument has a gravity compensation device all tests can be carried out in the first stage and in the second stage the gravity compensation can be set.” | P- Obviously, this is already possible now. It only requires a clear description in the respective Certificate. |
| 8.3 | PL | the last paragraph: We propose to complete the sentence: ’Instead of a place of use or gravity value at the place a gravity zone....’ | P+ We suggest, however, a slightly modified wording, see 2 CD. |
| 8.3 | ZA | Initial verification by accredited facilities such as verification laboratories is permitted in certain member countries of the OIML and it is proposed that the first paragraph be amended as follows; Initial verification may be performed by authorised national bodies, by the manufacturer itself, or by accredited organisations authorised or registered by authorised national bodies provided that- according to national rules – their management /quality systems for production or verification are acknowledged for this task. | P+ See response to AU and 2 CD |
| 8.3 | UK | The first sentence has no place in this Recommendation. Who carries out initial verification is a matter for national legislation not this Recommendation. Authorised national bodies and manufacturers with “acknowledged” (inappropriate term?) quality systems are merely examples of who may be permitted under national legislation, there are many other possibilities including for example independent third parties who may be economic operators. Delete this sentence. The first sentence of second paragraph of this section is also too restrictive. Who establishes conformity to the approved type and what criteria are used to evaluate their suitability to make such a judgement is a matter for national legislation. As worded the sentence restricts initial verification to bodies/ people that are the responsibility of another authorised body. Delete the words “under the responsibility of an authorised body” from the first sentence of the second paragraph. If the above is not accepted: In the first paragraph, replace “acknowledged” with “accredited [or certified]”. Also, after the word “production” add the words “and verification”. Also in the last paragraph, replace “data” with “parameters”. | P+ See response to AU and 2 CD P+ |
| 8.3.1 | AU | Comment: We do not understand the meaning of ‘The body in charge of the tests …’ perhaps this again reflects some procedures in the European system with which we are not familiar. | P+ See 2 CD |
| 8.3.2 | UK | As there is a comma at the end of the first bullet point, there should be commas at the ends of the second and third bullet points too. | P+ |
| 8.3.3 | CA | We do not understand the limitation to test loads ≥ 100 mg, and why the discrimination test is not to be performed on instruments with digital indications. The fourth paragraph mentions that the error limits shall be the maximum permissible errors (mpe) upon initial verification; however, Canada suggests that the error limit for the repeatability test should be the absolute value of this mpe. Canada agrees with the Secratariat’s response. | The test loads must be chosen acc. to A.4.4.1, but it is not important and not useful to include Min if Min < 100 mg. The discrimination test is not necessary, because the discrimination threshold at verification should be the same as at type approval. P+ because of 3.6.1 |
| 8.3.3 | FR | Third hyphen , read : “3.6.1 : repeatability (refer to A.4.10, 2nd paragraph). In initial verification of a mobile instrument, a test of tilting should be included | P+ Change to “(A.4.10, 3rd paragraph)” P+ See NL below |
| 8.3.3 | NL | Add one more item: - 4.19 tilt in case of mobile instruments. OIML certificates do not include descriptions of tests. So delete this sentence. | P+ See change in 8.3.3. P- In special cases the OIML certificates may and even should include descriptions of tests, e.g. the tilting test for mobile instruments. |
| 8.3.3 | UK | In the sentence “The authority responsible...” delete the word “from”. | P+ |
| 8.3.4 | CA | Stamping Suggest that the title of this section should be something other than ‘stamping’ to better reflect the intent of the section. ‘Initial Verification marking and securing’ would be acceptable. Would this entire section not be better situated under section 7.2 Verification Marks?? Canada agrees with the Secratariat’s response. | P+ However , we suggest the shorter title “Marking and securing” P- We suggest, however, to add a reference to 7.2 under 8.3.4, see 2CD. |
| 8.4.1 | ZA | Subsequent verification by accredited facilities such as verification laboratories is permitted in certain member countries of the OIML and it is proposed that clause 8.4.1 the first paragraph be amended as follows; Subsequent verification may be performed by authorised national bodies, or by accredited organisations authorised or registered by authorised national bodies provided that- according to national rules – their management /quality systems for production or verification are acknowledged for this task. | P+ However, we suggest in analogy to 8.3 a shorter and more general wording, see 2CD. |
| A | UK | Heading: Has “NONAUTOMATIC”, but elsewhere this is hyphenated as “non-automatic”. It should be consistent (preferably hyphenated). | P+ It is changed to “non-automatic” |
| A.4.1.1 | SI | We suggest slight modification of the last sentence in order to clarify its meaning: “For instruments of class (I) all necessary corrections of the mass of the test load with respect to influence factors shall be applied.” | P- See A.4.1.1 UK |
| A.4.1.1 | UK | A clarification of the sentence beginning with “For instruments of Class I...” is required. | P+ We add at the end of the sentence: “i.e. influence of air buoyancy” |
| A.4.2.1.1 A.4.2.1.2 A.4.2.1.3 | CN | Can the test for initial negative zero-setting range be omitted? Is it suitable that the test for negative zero-setting range should first load more than 5 percent of Max on receptor and initial zero-setting then test negative zero-setting range in A.4.2.1.2 and A.4.2.1.3. | P- We cannot agree with this proposal because we have made the experience that there can be a problem with the software which could only be detected by testing. |
| A.4.3 | CN | For A.4.3, whether or not add a kind of method to determine the zero point: c) For the weighing instruments with d(0.2e can directly numerate the error of zero. | P- This is not necessary here, because A.4.3 deals with the setting to zero before loading, not with the error of zero. |
| A.4.4.1 | CA | In brackets (Min only if Min ≥1mg) should read (Min only if e ≥ 1mg)? Canada agrees with the Secratariat’s response. | P+ However it must be changed to “Min only if Min ( 100 mg“ because of 8.3.3. Has been changed also in A.4.6.1 |
| A.4.6.1 | ZA | In case of 8.3 and 8.4 the practical test may be alternatively replaced by other appropriate procedures, e.g. by numerical or graphical considerations. Simulation of a tare-balancing operation by displacement (shifting) of the error limits (mpe) to any points of the error curve (= curve of weighing test results). Checking if the error curve and hysteresis are inside the mpe at every point. 3rd Par: The additional sentence adds to clarify the first sentence creates further confusion as it is not clear what the simulation actually entails Is the simulation a numerical or graphical consideration? We are not convinced that this is practical in the field and can guarantee accuracy of weighing. | P- The simulation may save time, an additional complete weighing test with tare may not be necessary. Numerical or graphical considerations are possible and are state of the art already. In case of doubt the full test can and shall be carried out. |
| A.4.6.1 | UK | The sentence “The steps shall include values close to Min (Min only if Min ≥ 1 mg)...” is ambiguous, as it could easily be read to mean that if Min ≥ 1 mg then onlyvalues close to Min are taken! For additive tare, is the term “3/3” the most suitable? Also, in the new text, the wording could be improved. | P+ See our response to A.4.4.1 CA. P- We are convinced that two tests are necessary, one at 1/3 and one at 3/3. |
| A.4.7 | CA | Eccentricity tests The statement “It is sufficient to apply the load only to the eccentric segment, not to the centre of the load receptor” assumes that the load is always a ‘known load’. If the load is an ‘unknown load’, it should be applied to the centre of the platform to establish a refence weight. Failing to do so may result in uncertainty in establishing the applicable limit of error. This is only applicable in the few instances where the load is very near the limit of error change value. Why does this statement need to be in the Recommendation? We suggest it be removed. | P- The test load shall be a known load, because “… the indications for different positions of a load shall meet the maximum permissible errors” (3.6.2). This means you have to determine the deviation against the nominal weight value on all positions and not just the deviations between center and corners. |
| A.4.7 | SI | It is necessary to replace the wording “bad results” in the last sentence of the fourth paragraph to something metrologicaly more appropriate, e.g. “However in case of exceeding mpe the test with zero error prior to each loading is necessary.” | P+ We changed it as proposed |
| A.4.7 | UK | The new sentence starting “However in case of bad results...” could be improved along the lines of “However, should a failure occur …”. | P+ See A.4.7 SI |
| A.4.7.4 | CA | Instrument used for weighing rolling loads (3.6.2.4) We must restate our concern with this section. Although the test may be valid, it appears incomplete. Individual ‘sections’ of the scale are never tested using this method. It would only work on a scale of 2-3 sections, for scales of 4 or more sections, the test pattern is incomplete. | P+: Yes we agree, because these instruments are similar to multirange instruments. See new wording in 2CD. |
| A.4.7.4 | FR | 1st line : read : “A rolling load shall be …”. | P+ |
| A.4.7.5 | CA | Eccentricity for mobile instrument This section contains a sentence which states “If operating conditions are such that no eccentricity can occur, eccentricity tests need not be performed.” Canada feels that this is a general statement regarding eccentricity and it should be applicable to all devices, not simply portable devices. This statement should be moved to section A.4.7. | P+ See 2CD, A.4.7 |
| A.4.8.2 | FR | For the same reason than those concerning 3.8, the first sentence shall be changed into : “This test applies only to type examination and to instruments with d ( 5 mg”. | P- Please refer to 3.8 FR |
| A.4.8.2 | NL | The word “replaced “ in the sentence: “One of the additional weights shall be replaced and a load …” could be misread as “exchanged”. Furthermore, it is not clear that when this additional weight is placed back the indication should increase again. We propose to change the wording to: “One of the additional weights shall be placed back (Indication increases one scale interval) and a load …” | P+ See new text in 2CD. P- The indication may, but need not, increase by one scale interval. |
| A.4.10 | UK | In the new last sentence, beginning “3 weighings on...”, change “3” to “Three”, and then to be consistent change “6” to “six”. | P+ |
| A.4.12 | CECIP | It is not acceptable that a detailed description of the stable equilibrium or any other function design is required. There may exist an unlimited number of different technical solutions from different manufacturers that the notified bodies will not be able to estimate all these solutions. There will be not guarantee that different notified bodies will come to the same objective decision. Any examination of a NAWI has to take the instrument as a “black box” and shall compare the reaction and accuracy of the instrument against the requirements. The technical way that was chosen by a manufacturer to meet a requirement is nonessential. There is also no sense to define a “worst case”. A.4.12 has only to define a test method, nothing else. Please delete the first part of A.4.12 completely. | P+: See 2CD, slightly modified. The functionality of the instrument must be sufficiently described anyway. |
| A.4.12 | SI | Clause 4.4.2 defines that in case of printing and/or data storage we compare the printed or stored weighing values with the final weighing value, but A.4.12 requires the comparison between printed or stored weighing values and the indicated value over a period of 5 seconds following print-out. We propose to rearrange sentences “In the case of printing or data storage, read the indicated value over a period of 5 seconds following print-out. Stable equilibrium is considered to be achieved when no more than two adjacent values are indicated, one of which being the printed value.” in a view of provisions of clause 4.4.2 to “In the case of printing or data storage, read the final weigh value. Stable equilibrium is considered to be achieved when the printed or stored weighing values do not deviate more than 1 e from the final weighing value (i.e. two adjacent values are allowed).” | P- It is not always possible to get the final weighing value during normal operation. |
| A.4.12 | UK | (line beginning “- Definition of...”): “...the the...”. | P+ |
| A.4.13 | UK | The sentence “There could be more or less...” should be re-worded to improve understanding. | P+ See 2CD |
| A.5.1.3 | UK | Health and safety considerations need to be taken into account when tilt testing high-sided or similar vehicles. Also, in the new paragraph, the text should read “… the load receptor comes into contact with the …” | P-: R76 considers only metrological requirements P+ |
| A.5.3.1 | CN | Whether or not change “-at a temperature of 5 °C, if the specified low temperature is ( 0 °C, …”in A.5.3.1 into “-at a temperature of 5 °C, if the specified low temperature is < 0 °C, …”, because 0° and 5° is quite near, it is not necessary. | P- The test at 5°C is necessary to check for condensation problems when coming from low temperature. This is important even if the low temperature is 0°C. |
| A.5.3.2 | NL | Temperature effect on no-load: This comment is based on combined Temperature effect on no-load indication and Static temperatures test. During this test zero-track and automatic zero-setting has to be switched off. However the purpose of this test is to determine the behaviour of the instrument during a change of temperature. Sometimes the instrument is being set to zero during the measurement. For example with a zero-drift of more than 9e, the instrument needs to be set to zero in order to perform the static temperature test at Max. In this case the complete Temperature effect on no-load indication has to be repeated. Since the error at zero-load is determined before and after the static temperature measurements it is not important if these zero-load values differ significantly (due to zero-setting) because the temperature is constant during this test. During the change of temperature the zero-setting and zero-tracking needs to be disabled. We would like to have the possibility to set the instrument to zero provided that the errors on no-load are determined before and after the setting to zero and the temperature is constant. Proposed addition: Zero setting allowed when before and after the zero-setting operation the error of the zero indication is determined and the temperature during this period has not changed. | P- This has never been a serious problem in the past and thus does not need to be regulated. The notified body will always be able to decide how to determine the zero error, with or without zero-setting. |
| A.5.4 | NL | Consider to perform this test only with 10e. | P- In contrast to the EMC tests the mpe apply here. |
| A.5.4 | UK | The sentence “The test consists...” should be re-worded to say that the appropriate variation should be applied. It currently implies that the EUT should be subjected to all of these variations. | P+ See 2CD |
| A.5.4.2 | NL | This is not in line with D11 for battery operated instruments For the normal (disposable) batteries, we do not see any justification for requiring 1,2 U for battery supply. | See also 3.9.3 NL P+ We agree, and have changed the wording so that this requirement is valid only for instruments that are powered with rechargeable batteries or batteries for road vehicles. In that case 1,2 U is easily possible just after charging. In addition “Battery operated instruments“ has been exchanged by “Battery powered instruments” in the entire document. |
| B | CA | Canada questions why this Annex is included in the recommendation. The preliminary note at the beginning of the Annex states that the latest draft revision of D11 has been taken into consideration in development of the Annex. This is cause for concern. If D-11 changes, R76 will no longer be in harmony with D-11. We suggest that R76 merely reference D-11 rather than incorporate it. | P- See our response under “General” CA. (D11 requires “implementation” into the specific Recommendation!) |
| B | CN | EMC (Electromagnetic Compatibility) tests consist of EMS (Electromagnetic Susceptibility) and EMI (Electromagnetic Interference). Tests for disturbances not only EMS (according with IEC 61000), but also EMI (according with CISPR 22 ). We suggest that the tests for disturbances of “Instrument for direct sales to the public” and “Hanging (Crane) Instrument “ should be added in EMI (according with CISPR 22). The International Standard is as follows: CISPR 22:1997 < Information technology equipment - Radio disturbance characteristics - Limits and methods of measurement > | P- EMC is used in accordance with OIML D11 (2004) and International (IEC / ISO) Standards which must be used as reference documents here. Other (regional) standards cannot be used in this context. We assume that by “EMI” you understand “electromagnetic emission”. This is, however, not under the scope of OIML (neither D11 nor R76). It might be regulated under national or regional legislation. P- We can also not agree for “direct sales scales” and “hanging crane” scales, because except for some tests (eccentricity and tilting) R76 requires these tests for all NAWIS. |
| B | JP | The numbering of referred standards should be reviewed and those should be referred in the text correctly. | P+ Both the Bibliography and all references in the document have been completely revised and updated. |
| B | UK | The very first note: The text “… conducted on the basis of most recent versions valid at the time of testing” could pose problems in actually determining if the most recent version was actually used. The sentiment is understood, but a practical solution is required. | P+: See also under “Contents” UK |
| B.1 | CA | This entire Annex should be limited to modules of any capacity, except weighing modules, and complete instruments having a Maximum capacity no greater than 1 000 kg. It is very unlikely that these tests could be performed on large instruments. If the Recommendation is to be applied consistently, how do you reconcile the fact that larger instruments may not be subjected to the tests as outlined in this Annex? We are thinking of large weighing element specifically – perhaps they are already exempt from these test? | Each test is difficult for instruments with a large Max. In nearly all cases it would be possible to perform the tests on an indicator with a small weighbridge. For high capacity weighing instruments one can also make use of the modular concept. These instruments are not exempt from these tests. If it is not possible to perform the tests on large instruments then modules of the instrument shall be tested separately. |
| B.1 | UK | Says “EUT” stands for “electronic instruments under test”, which doesn’t fit the acronym. However T7 on Page 21, and Section 3.10.4, say the usual “equipment under test”. | P+ We have deleted “(EUT)” in the heading of B.1, because “EUT” will now only be used for “equipment under test” (see also “Contents” UK) |
| B.3 | SI | Add “as it is defined in B.3.7.” at the end of the last paragraph. | P+ See 2CD |
| B.3.1 | UK | Now says that the “test generator shall be verified...”. Could “verified” be replaced by some other word, to avoid confusion? Similarly Section B.3.2, B.3.3 and B.3.4. | P+ See 2CD Changes also in B.3.2, B.3.3 and B.3.4 |
| B.3.3 | CECIP | Industrial customers force the manufacturers of instruments to meet not only the product related standards like R76 but also generic IEC standards for compliance with electromagnetic effects. Therefore OIML should avoid starting to copy any EMC standard existing in the world. That is not possible and useful. As Dutch NMI pointed out a surge testing is not essential for normal industrial applications. In extreme cases the general clause 4.1.1.1 or 5.1.1 comes into force. Please delete B.3.3 completely. | P- There was a vote to upgrade EMC testing acc. to the development of International Standards. In addition “Surge” is already included in other OIML-Recommendations concerning weighing instruments (e.g. R51); therefore it cannot be deleted here. |
| B.3.3 | JP | Maximum allowable variations: “The difference … shall not exceed e…” is too severe. “…the instrument shall detect and react to a significant fault” is short on specifics and different interpretation might be developed. The description for this point should be made clearer and more concrete. In addition, the definition for indoor and outdoor use should be also specified. | P- There is no reason to deviate from T.5.5.6, especially because the test is conducted with a small test load. Please refer also to the examples of non-significant faults given in T.5.5.6. In addition we think that with the introductory note left at the beginning of B.3.3, indoor and outdoor situations are now defined sufficiently clear. |
| B.3.4 | CN | In B.3.4, we propose to change “The time interval between successive discharges shall be at least 10 seconds.” into “The time interval between successive discharges shall be at least 1 second.”. | P- We should keep in line with D11 (2004) which requires at least 10 seconds. |
| B.3.5 | CECIP | It is not acceptable to separate different test severity levels for instruments used in household or in industrial applications. As long as there is no descriptive marking on an instrument informing the user about the allowed location all instruments have to be examined with the same test severity level. That means all instruments have to be examined with either with 10 V/m or all with 3 V/m. | P+ See 2CD |
| B.3.5 | UK | For retail applications it now appears to be saying that 3 V/m on one face only is acceptable. Is this correct? | This is, indeed, not correct. B.3.5 has therefore been corrected and is now in line with D11 (12.1) |
| B.3.6 | CECIP | It is not acceptable to separate different test severity levels for instruments used in household or in industrial applications. As long as there is no descriptive marking on an instrument informing the user about the allowed location all instruments have to be examined with the same test severity level. That means all instruments have to be examined with either with 10 V/m or all with 3 V/m. For information: In the European Standard EN61326 the test level was reduced from 10V to 3V with the following comment: The test level for the conducted RF test is lower than the level for the radiated RF test because the conducted RF test simulates the resonance condition at each frequency and is thus a more severe test. | P+ According to 61000-4-6 conducted and radiated field strength shall be equal. Following change in B.3.5 we have chosen 10 V/m. Remark: EN61326 is only a (regional) product family standard for electrical measuring, control and laboratory instruments. |
| B.3.6 | UK | “acccording” should be “according”. | P+ |
| B.3.7 | NL | Replace ISO 7637-2.3 by ISO 7637-2 (2004). Delete test pulse 1 and reconsider the application of test pulse 2b (see the final version 2004 of D11). | P+ See 2CD P+ See 2CD |
| B.3.7 | SI | For better clarity replace the heading “Special EMC requirements for road vehicles” with “Special EMC requirements for NAWI powered from vehicle battery”. | P+ See 2CD |
| B.4 | JP | For the test sequence, it is specified as “Automatic zero-tracking shall be made inoperative and automatic built-in span adjustment device shall be made operative”. As the same can be said to semi-automatic built-in span adjustment device, “Semi-automatic built-in span adjustment device shall be made operative before testing” should be added. | P- No change. This is a contradiction to A.4.1.8. The behaviour of the user (in this case releasing an adjustment) shall not be part of a test. It is not guaranteed that this device will ever be used so the worst case needs to be covered. An automatic built-in span adjustment device works fully automatically during normal use. |
| C.1 | UK | Line 3.9.3: “suply” should be “supply”. Line 4.7: Uses “preset-tare”, but elsewhere it is “preset tare” without the hyphen. Line 4.9: Has “removeable”, but this spelling is rare; “removable” is much more common and is used in the actual Section 4.9. | P+ P+ P+ |
| C.1 | US | The U.S. recommends that the list of applicable requirements include 4.5.1. Maximum effect (of range of the initial zero setting device). | P- No change, 4.5.1 is included in “4.5” |
| C.1.1 | CA | May conflict with section D.1.2.2 Indicator Accuracy Class must be same as weighing element. However, D.1.2.2 allows exception to this. Is this a problem? Suggest that wording be changed to state: “The indicator shall be of the same accuracy class or a higher accuracy class, than the weighing instrument it is intended to be used with.” | P+ See 2CD |
| C.1.1 | CN | We suggest changing “The indicator shall have the same accuracy class as the weighing instrument it is intended to be used with.” in C.1.1 into “The indicator shall have the same accuracy class or higher accuracy class as the weighing instrument it is intended to be used with.”(see also E.3.1). | P+ see C.1.1 CA |
| C.1.2 | CA | Conflicts with section D.1.2.3 Both sections require the same or higher number of verification scale intervals. This effectively means they must both have the same number of verification scale intervals. Suggest that the section D.1.2.3 need not be included. | P- No change. These chapters concerns several modules, e.g. an indicator with 6000 e can be used in an instrument with only 3000 e. |
| C.1.3 | CA | Conflicts with section D.1.2.4 Both sections require the same or larger temperature range than the other. This effectively means that the temperature range must be identical. One should supercede the other. | P- No change. These chapters concerns several modules, e.g. an indicator of class III with -10°C to 40°C can be used in an instrument with 0°C to 30°C. |
| C.1.6 | CN | For the problem of range of load cell impedance (C.1.6), whether the input impedance can be expressed as the maximum current of excitation. | P- No change, “impedance” is correct. |
| C.1.7 | SI | Start the first sentence with “Only indicators...” to stress that only 6-wire technology is allowed. | P+ See 2CD |
| C.2.2 | CN | For C.2.2: Table 12 B.3.1 Short time power reduction 1 high*) min Whether high*) is low*)？ It is not right that the impedance of the load cell only refer to the input impedance of it. For the great impedance, the interior impedance of signal source in indicator, that is, the output impedance of the load cell has the larger influence to indicator. The input impedance of one load cell is equal to the output impedance of it, whose expression is not mathematical. To limit the (equivalent) minimum input impedance is more suitable than to do the (equivalent) maximum input impedance. Here, is that ‘equivalent’ indicates impedance from the angel of indicator fit? | We understand your point of view, but in practice the tests B.3.1 to B.3.5 are applied with the same test equipment and the same load cell. For simplifying the procedure the test B.3.1 shall be done with the same test set up. |
| C.2.2 | JP | In Table 12, B.3.3 should be changed to B.3.4 and B.3.4 to B.3.5. Aren’t the tests in B.3.3, B.3.6 and B.3.7 required? | P+ See 2CD Yes |
| C.2.2 | SI | Revise Table 12 with EMC tests as defined in B.3 (surge, immunity to radiated EMF and conducted disturbances, transients). | P+ See C.2.2 JP |
| C.2.7 | UK | Sentence beginning “The manufacturer has to fix...: “than” should be “then”. Also remove “to” from “...assigned to”. | P+ |
| C.3 | JP | In C.3, item 4.12.3 listed does not exist. Item 4.14.10 may be 4.13.10. Please confirm. | P+ |
| C.3.1.1 | CN | On P111, in C.3.1.1 (e), Pi × 1e / 5K shall be revised into Pi × e / 5K; | P+ |
| C.3.1.1 | SI | An explanation is missing whether the procedure described from (a) to (f) refers to indicators with verification scale interval more, equal or less than 1 µV. It is also not clear if this procedure is the acceptable solution or the prescribed testing procedure. | P+ See 2CD |
| C.3.3.1 | SI | The first sentence is not complete and not understandable. | P+ |
| C.3.3.1 | UK | The wording of the first sentence needs to be changed. | P+ |
| C.3.3.1 | US | Typographical error: C.3.3.1 Scope Indicators intended for connection of strain gauge load cells employ the 4- or the 6-wire principle of the load cell connection. When 4-wire technology is used, lengthening cable the load cell cable or using a separate load cell junction box with an extra cable is not allowed at all. . . | P+ |
| C.3.3.2 | UK | Change the word “charge” to “value”. | P+ |
| C.3.3.2 | DK | Please add the following very important test: The sense circuit shall also be tested under conditions where the connections of the sense wires are interrupted one after the other. The difference in indication with sense and no-sense shall not exceed mpe or the exceeding shall be so large that it is evident for the user that something is wrong (more than 10% wrong indication). | P- There was no TC9/SC1 vote for introducing such a new test |
| C.3.3.2.1 | UK | “...in parallel to...” should be “...in parallel with...”. | P+ |
| C.3.3.2.2 | UK | Replace the word “neglectable” with “negligible”. | P+ |
| C.3.3.2.3 | UK | Replace “readjusted” with “re-calibrated” and replace “placed” with wither “adjusted” or “set”. | P- The indicator shall be adjusted, not calibrated P+ See 2CD |
| C.4.1 | FR | First sentence of the second paragraph, read : “… under “Identification of the certified type module”:”. | P+ |
| C.4.1 | UK | “...intervals,minimum...” should have a space after the comma. | P+ |
| D.2.2 | ZA | Last Par: We suggest to change the sentence from “An any case the indicating device should allow…” to “The indicating device shall allow indication to a higher resolution to determine the error, e.g. in a special service mode. | P+ Has been changed to: “The indicating device should preferably allow indication to a higher resolution…”, see 2CD. ” Shall” cannot be required, see D.2.2 UK. |
| D.2.2 | UK | Sentence beginning “An any case”: Is it really intended that a high resolution mode be compulsory? This sentence needs to be re-worded. | P+ See 2CD |
| D.4.2 | UK | Line numbered 2: “Weighing instrument” should be “module”, as in C.4.2. Line 7: “srew” should be “screw”. | P+ P+ |
| E | FR | In the list of necessary data to check the compatibility, add the humidity evaluation symbol (NH, SH, CH or no symbol) of the load cells. | P+ Has been added |
| E.1 | CA | There is a question mark “?” after the phrase “Connecting system, 6 wire system?” Does this belong here? Canada agrees with the Secratariat’s response. | P+ |
| E.1 | UK | Add “The” to the sentence “Following metrological...” to make “The following metrological...” Line starting “DL...”): “himself” should be “itself”. Replace “… and additional fix mounted constructions on the load receptor.” with “… and any additional construction mounted on the load receptor.”. | P+ P+ P+ |
| E.1 | US | Insert the word “the” at the beginning of the sentence as follows: The Ffollowing metrological and technical data of the weighing instrument are necessary for the check of compatibility: | P+ |
| E.2.1 | CN | The format of table 13 in E.2.1can be amended as follows: Accuracy Reference WI OIML R76 LC A A*), B B*), C C, D OIML R60 Changed: Accuracy Reference WI OIML R76 LC A A*), B B*), C C, D OIML R60 The format of table 14 in E.3.1 is amended the same as table 13. | P+ P+ |
| E.2.3.X | CA | Since OIMLR60 approved load cells can be approved with either steady-state humidity, cyclic humidity or no humidity testing, a requirement is needed to specify that only load cells having the “SH” classification can be used as a module for a non-automatic weighing instrument. This is briefly mentioned in the note to Table 7. Load cells with the “CH” classification should not be deemed as meeting the damp heat requirements of section B.2.2. <<Pending comments>> | P+ for NH load cells / P- for CH load cells We think, because of the note to Table 7, there is a misunderstanding that CH load cells are worse compared to SH tested load cells which is not the case! Normally CH tested load cells meet the requirements for SH tested load cells, because the CH test cycles include e.g. a kind of “durability test” which is more than just a simple “humidity test”. In order to avoid misunderstandings we therefore suggest to change the 1CD as follows: 1. Delete in Table 7 the second part of note (**): “... but not for CH tested load cells (pLC=1)” because Table 7 just presents examples and the factor pLC=0,7 applies to SH tested load cells only. 2. Add in 3.10.2.4 under “- Load cells ...” that only SH or CH tested load cells are allowed but no NH marked load cells. 3. Add a similar sentence under E.2 (new F.2): SH and CH allowed under the modular approach, but no NH load cells |
| E.2.5 | UK | Should be “...must equal or exceed...” (or could be “...must be equal to or exceed...”). | P+ |
| E.2.6 | CN | On the issue of relation between load cell (nLC) and weighing instrument (n) (E.2.6) Temperature ranges and the evaluation of stability against humidity and creep of load cell(s) (LC) must meet the requirements for the weighing instrument (WI). For each load cell the maximum number of load cell intervals nLC shall not be less than the number of verification scale intervals n of the instrument. Since OIML R60 International Recommendation on Load Cell points out: The classification of load cells into specific accuracy classes is provided to facilitate their application to various mass measuring systems. In the application of this Recommendation, it should be recognized that the effective performance of a particular load cell may be improved by compensation within the measuring system with which it is applied. Therefore, it is not the intent of this Recommendation to require that a load cell be of the same accuracy class as the measuring system in which it may be used. Nor does it require that a measuring instrument, giving indications of mass, use a load cell which has been separately approved. | Within the compatibility test it is stated, that any module must have the same or a better accuracy than the weighing instrument. A correction of a load cell (e.g. linearity) using an appropriate indicator is not the intention of the modular approach. |
| E.2.7 | CN | On Page 121, in E.2.7, the sentence of “electrical data with regard to the weighing instrument ” shall be the contents of E.3.5 | P+ |
| E.3 | US | There is no compatibility check for IZSR between separately tested indicators and weighing instruments (combination of load cell and connecting elements). The IZSR of the separable indicator may be configured so that it is larger than the IZSR of the weighing instrument. | P- See A.4.2.1.1 |
| E.3.1 | CN | On expressing method of accuracy class of indicator (E.3.1) In order to be different from accuracy class of weighing instrument, whether accuracy class of indicator can be expressed by Class 1, Class 2, Class 3 and Class 4 separately. | We understand your point of view but there is no vote for changing the symbols. Otherwise both weighing instruments and indicators are considered in the same OIML Recommendation. Furthermore the metrological characteristics are identical in all accuracy classes with the exception of pi for indicators. Therefore we think this does not lead to a confusion. |
| E.3.5 | JP | For items “(mV) Measuring range minimum voltage” and “(mV) Measuring range maximum voltage”, the symbols may be necessary. In addition, the definition for C in the equation of Δu is short. | P+ See 2CD P+ See list of abbreviations (T.9) |
| E.5 | CECIP | For a digital load cell the maximum cable length is not of interest. See the definition of a digital load cell in T.2.2. | P+ See 2CD |
| E.6.1 / E.6.2 | CA | Temperature range is shown as -10°C ( 40°C. This is clearly not desirable and must be corrected. Clarification: Typographical error. Symbol is incorrect ( (divided by) should be – (dash) Also found in section E.6.2 | P+ Has been changed to “–10 °C to +40 °C” in order to avoid further problems with ASCII symbols |
| E.6.1 | CN | On Page 126 and Page 128, -10°÷40° shall be revised into -10°~+40° | P+ See CA |
| E.6.1 | UK | Is it worth having an example which shows that the modules are not compatible? | We think that the existing two examples are sufficient. |
| F | CA | Canada would like to see ‘Software’ (this entire Annex) removed from R76 and kept as a second document (D-SW?) which would be referenced. In this way, the relevant tests would not need to be duplicated by each Recommendation. This would make ongoing maintenance and updating of all documents much easier and would eliminate the concern of developing conflicting requirements in the future. | P- With the reasoning given in our response to CA comment under “General” |
| F.2.2.3 | UK | “spotcheck” should be “spot check” (two words). | P+ |
| F.2.4 | UK | “spotcheck” should be “spot check” (two words). | P+ |
| Bibliography | NL | /5/ The new OIML D 11 and ISO 7637-2 have been published in the meantime. | P+ The respective requirements, testing procedures, references, and the Bibliography of the 2CD have been completely revised and updated according to the new OIML D11 document. |