1

DEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR BIOLOGICS EVALUATION AND RESEARCH

BLOOD PRODUCTS ADVISORY COMMITTEE

This transcript has not been edited or corrected, but appears as received from the commercial transcribing service:  Accordingly the Food and Drug Administration makes no representation as to its accuracy.

Friday, July 23, 2004

8:00 a.m.

Gaithersburg Holiday Inn

2 Montgomery Village Avenue

Gaithersburg, Maryland 20877

2

PARTICIPANTS

Kenrad E. Nelson, M.D., Chair

Linda A. Smallwood, Ph.D., Executive Secretary

Pearline K. Muckelvene, Scientific Advisors

& Consultants Staff

MEMBERS:

James R. Allen, M.D., M.P.H.

Kenneth Davis, Jr., M.D.

Donna M. DiMichele, M.D.

Samuel H. Doppelt, M.D.

Jonathan C. Goldsmith, M.D.

Harvey G. Klein, M.D.

Suman Laal, Ph.D.

Katherine E. Knowles,

Acting Consumer Representative

D. Michael Strong,

Non-Voting Industry Representative

TEMPORARY VOTING MEMBERS:

Liana Harvath, Ph.D.

F. Blaine Hollinger, M.D.

Katharine E. Knowles

Matthew J. Kuehnert, M.D.

Susan F. Leitman, M.D.

Keith C. Quirolo, M.D.

George B. Schreiber, Sc.D.

Donna S. Whittaker, Ph.D.

3

C O N T E N T S

PAGE

Update on West Nile Virus, Hira Nakhasi, Ph.D.             6

IV. Hepatitis B Virus Nucleic Acid Testing (NAT)

for Donors of Whole Blood:

A. Introduction and Background,

Gerardo Kaplan, Ph.D., Laboratory

of Hepatitis and Related Emerging

Viruses, DETTD, OBRR, FDA                    28

B. Serological Course of Hepatitis B,

F. Blaine Hollinger, M.D.,

Baylor College of Medicine                   32

C. Preclinical and Clinical Data for

HBV MP NAT, Steven Herman, Ph.D.,

Roche Molecular Systems                      51

Allan Frank M.D., M.S.,

Roche Molecular Systems                      65

Open Public Hearing:

Michael Busch, Blood Centers

of the Pacific                               103

William Andrew Heaton, Chiron                  121

Sherrol McDonough, Gen-Probe                   129

Richard Smith, NGI                             136

Harvey Alter, AABB                             144

IV. Hepatitis B Virus Nucleic Acid Testing (NAT)

for Donors of Whole Blood:

E. Committee Discussion and

Recommendations                             154

V. Current Trends in Plasma Product Manufacturing

A. Introduction and Background,

Mark Weinstein, Ph.D., Associate

Deputy Director, OBRR, FDA              223

B. Presentation, Jan M. Bult, CEO,

Plasma Protein Therapeutics

Association                             225

Open Public Hearing:

Patrick Schmidt, CEO, FFF Enterprises          258

4

1                      P R O C E E D I N G S

2             DR. SMALLWOOD:  May I ask all advisory

3   committee members to, please, take your seats?

4   Welcome to the second day of the Blood Products

5   Advisory Committee meeting.  Yesterday I read the

6   conflict of interest statement that applies to this

7   meeting, however, we have a new process now and we

8   will read a conflict of interest statement for each

9   day.

10             So, if you will indulge me, I will read

11   that at this point.  This brief announcement is in

12   addition to the conflict of interest statement read

13   at the beginning of the meeting yesterday, and is

14   part of the public record for the Blood Products

15   Advisory Committee meeting on July 23, 2004.  This

16   announcement addresses conflicts of interest for

17   topic V.

18             Drs. Liana Harvath, Blaine Hollinger,

19   Matthew Kuehnert, Susan Leitman, Keith Quirolo,

20   George Schreiber, Donna Whittaker and Ms. Katherine

21   Knowles have been appointed as temporary voting

22   members for this meeting

5

1   Dr. Michael Strong is participating in this meeting

2   as the non-voting industry representative, acting

3   on behalf of regulated industry.  The Food and Drug

4   Administration has prepared general matters waivers

5   for the special government employees participating

6   in this meeting who required a waiver under Title

7   XVIII, United States Code 208.

8             In addition, there are regulated industry

9   and other outside organization speakers making

10   presentations.  These speakers have financial

11   interests associated with their employers and with

12   other regulated firms.  They were not screened for

13   these conflicts of interest.  I would just like to

14   remind everyone participating to, please, make

15   known, if you have not already done so, any

16   affiliation you may have and your status with that

17   affiliation prior to speaking.

18             Our committee chairman, Dr. Kenrad Nelson

19   has joined us this morning, and we also have Dr.

20   Blaine Hollinger who will also be part of the

21   committee this morning.

22             I just wanted to announce to those who

6

1   were not here yesterday that the next date, which

2   is tentative however pretty much firm, for the next

3   Blood Products Advisory Committee meeting will be

4   October 21st and 22nd, 2004.

5             At this time I will turn over the

6   proceedings of the meeting to the chairman, Dr.

7   Kenrad Nelson.

8                    Update on West Nile Virus

9             DR. NELSON:  Thank you, Dr. Smallwood.  I

10   will try to keep awake after the 24-hour airplane

11   ride.  I came in last night but I feel really

12   pretty good and I am very interested in the topic

13   today so I think that will help.

14             The first topic is an update on West Nile

15   virus by Hira Nakhasi.

16             DR. NAKHASI:  Good morning.  I just want

17   to give you an update, as Dr. Kenrad Nelson

18   mentioned, on the West Nile epidemic and donor

19   testing which is happening now, in 2004.  First I

20   will try to wrap up last year's things and then

21   come up to 2004.

22             Next slide, please.  The topics which I

7

1   will update you on are, as I said, last year's

2   epidemiology and the investigational West Nile

3   testing outcome of that, and some of the

4   transfusion-transmitted cases, and then the trigger

5   for the ID-NAT testing.  Then I will update you on

6   the West Nile donor and product management

7   recommendations with the recent revelations we have

8   got.  Then I will update you on the 2004 epidemic

9   and investigational West Nile testing, and also our

10   efforts in-house on the panel development and other

11   scientific issues--you know, the variation among

12   the strains of viruses infectivity of these

13   studies.

14             Next slide, please.  If you summarize in

15   one slide the last year's epidemic, it really

16   basically sums up that we had approximately 1000

17   [sic] cases or, to be precise, 9862 cases, human

18   cases, and 264 deaths.  And, the proportion of the

19   West Nile meningitis/encephalitis was 29 percent,

20   whereas, the fever was 69 percent in the human

21   cases.

22             Forty-six states, including Washington,

8

1   D.C., were endemic, and donor testing started, as

2   all of you know, in July of 2003, using two

3   investigational NAT testing.  In some cases, a

4   small proportion started in the middle of June.

5   Despite this testing, I think these two

6   investigational NAT testing--these are minipool and

7   the two tests were the Gen-Probe test and the Roche

8   test, and Roche tested, as you know, in pools of 6

9   and the Gen-Probe test involves a pool of 16.

10             Despite testing, there were some

11   transfusion-transmitted cases and CDC had

12   investigated a total of 23 cases.  They were

13   confirmed by NAT and IgM reactivity and also by

14   follow-up of both the donor and the recipient.  Out

15   of the 23, 6 were confirmed cases.  Only 4/6, you

16   may recall, had very low viremia, around 0.1

17   pfu/ml.  Eleven cases did not confirm; 3 were

18   inconclusive because of the follow-up situation;

19   and 3 were open investigations.

20             Next slide, please.  As I said, since it

21   started on July 1 of last year, screening using

22   minipool NAT and IND, all geographic regions of the

9

1   U.S. were screening at that time.  With that, what

2   happened 1000 units of West Nile infected blood

3   donors were interdicted after screening

4   approximately 8 million donations.  So, I think it

5   was a very, very vast improvement over the year

6   before when there was no testing.  The last

7   positive donation was reported in the middle of

8   December in 2003.

9             Despite this testing, as you see, the

10   majority of cases were interdicted, more than 75

11   percent, but there was a small percentage which

12   went through because, as you know, this was done in

13   minipool NAT.

14             Next slide, please.  This slide is Mike

15   Busch's slide where he showed why we were missing

16   some of these cases, and we knew that minipool NAT

17   sensitivity was such.  The areas, you know, where

18   the wrap-up takes place when--you know, he calls it

19   stage I, II, III, IV and V, and in stage I and II

20   they are ID-NAT positive but minipool NAT negative,

21   IgM negative.  So, it could be plus/minus.  So,

22   during that stage they become IgM positive but they

10

1   become minipool negative and they are still ID-NAT

2   positive.  So, this region and this region were the

3   ones where they went through.  But, you know, these

4   were IgM negative and these were IgM positive so

5   the question is what is the infection of these

6   types of samples.

7             Next slide, please.  So there was a

8   potential for transmission of West Nile through

9   minipool NAT negative blood of low viremia in some

10   patients.  Therefore, what happened at that time is

11   that limited prospective ID-NAT testing started in

12   high incidence areas.  If you remember last year,

13   Colorado, Kansas and certain other areas, and

14   Nebraska were hot spots and ID-NAT was triggered at

15   that time, and the trigger was based on if the

16   preceding the rate of 1/200 minipool NAT positive

17   rate of 1/250, then they would start testing with

18   ID-NAT testing.  Also, what happened at that time

19   is that there was voluntary withdrawal of the

20   frozen transfusables in the high incidence areas

21   before the ID-NAT was initiated by some blood

22   establishments.

11

1             Next slide, please.  There was also

2   another initiative started at that time.  The

3   initiative was to go back to do the retrospective

4   study on the minipool NAT negative samples and test

5   them by ID-NAT to find out how many we missed.  It

6   would also let us know what was the low level of

7   viremic high incidence samples in high incidence

8   areas where minipool NAT did not pick them up.

9             The other purpose of the study was also to

10   identify samples which are like minipool NAT low

11   titer, minipool NAT negative but ID-NAT positive

12   for infectivity studies.  I told you that we do not

13   know whether those samples are still infectious at

14   low levels, and what is the level of infectivity.

15   So, these samples would be tested in various animal

16   models including non-human primates.  Also, the

17   purpose of these samples is to really find out the

18   relative clinical sensitivity of various West Nile

19   investigational testing.  I will report in a minute

20   what is happening with the infectivity state.

21             Next slide, please.  Based on the

22   observation that we had minipool testing and we

12

1   missed some of the samples because the viremia was

2   low, and also in the ID-NAT testing in the high

3   incidence areas--based on those studies and based

4   on the logistics issues, the question was what

5   should be the trigger for ID-NAT, and also logistic

6   issues such the availability of adequate resources,

7   recruitment, reagents and trained technologists.

8             So, the discussion about the trigger for

9   ID-NAT was held in collaboration with the AABB task

10   force.  By the way, we are very indebted to the

11   AABB task force for the biweekly meetings almost

12   throughout the year, and weekly meetings with the

13   task force during the epidemic to update us and

14   jointly discuss the strategies for how to go

15   forward with the testing performance, as well as

16   the epidemic.

17             So, based on that discussion, which was

18   held in February, the recommendations were the

19   following for the ID-NAT trigger:  It was discussed

20   that we should monitor reactive rates by zones

21   daily, enrolled 7 days when the epidemic was

22   starting, which was usually, you know, around the

13

1   beginning of July and early June even and this year

2   even May some cases were found.  The trigger was

3   that if you have 2-4 cases in any geographic

4   area--that is the blood collection, and the

5   frequency of 1/1000.  This was based on the fact

6   that every 1/4 would be missed by minipool NAT and

7   require ID-NAT.  This was the study done by ARC and

8   BSL and they found out that that would be the

9   trigger.  And, you go back to minipool NAT only

10   when you see ID-NAT reactivity and you don't find

11   zero cases in a consecutive 3-4 day period or the

12   rate is less than 1/1000.  So, that was the trigger

13   because, you know, we wanted to be prepared this

14   year because last year it was on an ad hoc basis to

15   start ID-NAT testing in those hot areas.  So, we

16   wanted to be prepared this year if these areas

17   become hot so that we get the logistics present

18   there so we can start without interruption of the

19   ID-NAT testing.

20             Next slide, please.  Now we come to 2004,

21   where are we now?  As of July 20, which is a couple

22   of days back--as you see, every week the numbers

14

1   keep changing.  Last week there were 108.  This

2   week it is 182 human cases out of which there were

3   4 deaths.  There are 2 from Arizona, 1 from Texas

4   and 1 from Iowa.  Out of total infections, 74

5   percent of cases are neuroinvasive West Nile

6   illness and 26 percent cases are West Nile fever.

7   At the moment there are 35 states endemic for West

8   Nile.  This slide has been kindly provided by Jen

9   Brown, from CDC, and other slides which I will

10   mention later.

11             The total number of presumptive West Nile

12   viremic donors reported to the CDC ArboNet--that is

13   why I highlighted this, is 23.  There are more

14   cases than that but, as you know, there is a delay

15   in reporting to the ArboNet from the health

16   departments.  So, using minipool NAT as well as

17   ID-NAT in select areas, starting on May 4.  Out of

18   these 23 presumptive West Nile viremic donors, 21

19   are from Arizona.  The majority are from the

20   Maricopa county near Phoenix, in Arizona; 1 from

21   New Mexico and 1 from Iowa.  But this is the tip of

22   the iceberg.

15

1             Next slide, please.  This slide, again, is

2   provided by Jen.  You can see the distribution of

3   the West Nile, both the animal, avian and mosquito

4   infection, which is in this color, and the blue

5   color shows you the human cases.  You can see it is

6   very high in Arizona and California.  I am telling

7   Mike Strong that it is creeping up in Washington

8   soon.  So, he has been telling me we don't see

9   anything and I said, well, wait and watch!  As you

10   remember, in 1999, how this started and how it is

11   spreading and, you know, it just keeps on going.  I

12   hope it will end up in the ocean sometime.

13             Next slide, please.  This is just to give

14   you how early the human cases can be detected.  As

15   you see from the slide, the earliest one was in

16   April.  So, you know, there is an expansion of this

17   epidemic, it looks like.  We were told in the

18   textbooks it is mostly in August and September or

19   late July but you can see it as early as April now,

20   and last year we saw it as late as December, in the

21   middle of December.  So, you know, it is almost a

22   year-round activity now.

16

1             Next slide, please.  Thanks to all the

2   blood establishments and testing establishments, I

3   got these data from several folks and I will

4   acknowledge them as I speak.  The total number,

5   according to my calculations but this may not be

6   right, is 61 presumptive viremic donors reported,

7   starting in May, 2004.  As I said, some of them are

8   reported to ArboNet and some of them are not.  So,

9   it is not in addition to that; it is inclusive of

10   the ArboNet reports.  ARC has told me--Sue Stramer

11   gave the data from June 16 to July 20, 7 hard

12   cases.  Again, this is also in the Arizona area.

13   But she says no region has their ID-NAT trigger.

14             Mike Strong gave me this data from Roche.

15   There are 2 positive confirmed by ID-NAT--around

16   300,000 donations screened.

17             BSL, Sally Caglioti and Mike Busch told me

18   that there are 23 confirmed, out of which 16 came

19   from minipool NAT and 7 came from ID-NAT, confirmed

20   positives.  There are 14 pending and he was saying

21   that some of them are ID-NAT and would have been

22   missed by minipool NAT.  Also, some of them are low

17

1   viremic and also there are some which are IgM

2   positive.  The denominator is around 400,000.

3             Gen-Probe, Leanne Kiviharju, gave the

4   data.  These are non-ARC data but I am not sure--I

5   sent an email to Leanne--whether this is also

6   non-BSL but I am not sure; maybe we can find out

7   from here, but 21 confirmed positive and 7 are

8   pending.  I am glad that you guys sent me several

9   slides.  I was basically trying to summarize what

10   the presumptive donors are and, you know, I really

11   appreciate your sending extra slides.

12             The Department of Defense, Ron Hagey sent

13   me the data which has 8 confirmed out of 62,774

14   since January of 2004.

15             So, you know, this is the majority of the

16   screening going on at this time and there may be a

17   few cases which have not been reported yet, but

18   this is where we stand as of today.

19             Next slide, please.  I just wanted to sort

20   of briefly remind you that FDA is still continuing

21   to work closely with the test kit manufacturers and

22   we would like to facilitate implementation of these

18

1   tests and expediter test licensure.  I just want to

2   remind you that we issued two guidances in October,

3   2002 and May, 2003.  There are 3 INDs for West Nile

4   minipool-NAT.  One is from Roche, one from

5   Gen-Probe and one from ARC.  This is public

6   information.  FDA is continuing to work with the

7   AABB task force.  I think that has been a

8   wonderful, wonderful collaboration with the AABB

9   task force and the people on the task force are

10   really helpful in doing this project together, and

11   with the CDC, NIH help, and to monitor the epidemic

12   and monitor the testing.

13             Next slide, please.  Both ARC and BSL did

14   a study, which is unpublished observation.  We had

15   a small discussion at the task force on what they

16   found out in some of the viremic donors when they

17   followed up.  They wanted to find out what is the

18   rate of the disappearance of RNA when they convert

19   IgM and IgG.  As you remember, in the last years

20   before the testing started the literature was that

21   it can go as long as 28 days of viremia.  But from

22   their studies, and I don't want to go into detail

19

1   here because these are unpublished and, you know, I

2   don't want to divulge information--the gist of that

3   was that what they found out in both cases is that

4   the viremia may last up to 49 days in one case and

5   39 days in the ARC study, and in the BSL 49 days,

6   and West Nile RNA may go coexist with IgM.

7   Therefore, this sort of started us thinking.  In

8   the guidance document we put 28-day donor deferral

9   and so we may have to rethink the deferral for

10   that.

11             Next slide, please.  We have not discussed

12   it with the AABB task force but we will be

13   discussing with the task force that, you know, the

14   integration of West Nile testing information.  We

15   are thinking about maybe 56-day deferral for West

16   Nile diagnosis of symptoms, including headache and

17   fever, or 14 days after symptom resolution if it is

18   more than 56 days.  Potential reinstatement of

19   donor deferral for West Nile symptoms only

20   following 30 days without symptoms, and negative by

21   West Nile IgM or ID-NAT.  Again, this is current

22   thinking.  We have nothing in the works yet but we

20

1   have internal discussions, and we will discuss it

2   at our next regular AABB task force before we come

3   up with a recommendation.  Dr. Alan Williams is

4   spearheading this initiative.

5             Next slide, please.  With regard to our

6   activities in-house, as I mentioned last year also,

7   we are still working on the panel development.  The

8   purpose is to monitor sensitivity of assays to

9   detect viral nucleic acid antibodies, and also

10   trying to isolate and characterize West Nile

11   strains from human samples during 2003 and 2002

12   epidemics. The purpose of this study, which is done

13   by Dr. Maria Rios in our group--and all these

14   studies actually really are done by Dr. Maria Rios'

15   group--is the genetic variation of viral strains;

16   detection by currently available West Nile assays.

17   The purpose is to really see if there is any

18   genetic variation and also infectivity studies

19   using animal models.  Currently, the samples have

20   been identified which could be used for infectivity

21   studies.  However, there are logistic issues about

22   the animals, baboons, which are being worked out

21

1   with the Southeast Medical Center.  I guess the

2   task force is working on that.  Hopefully, we will

3   get some information by fall and we will be set to

4   do those studies.

5             Next slide, please.  Briefly, they have

6   two isolates, NY99 in 2002, which have been

7   characterized by genetic sequencing which I can

8   show you in a minute.  The viral infectivity is

9   determined by in vitro studies using cell lines and

10   primary human blood cell cultures.  Final panel

11   specifications are being established through the

12   collaborative studies, and the range of

13   concentration ranges between 1000-5 copies/ml.

14             Next slide, please.  Just a piece of

15   information here that Maria was kind enough to

16   provide to me.  You know, she did the comparison of

17   the human 2002 strain and the NY99 flamingo isolate

18   and then passed through the Vero cells.  She found

19   there were 20 nucleotide mutations and one

20   insertion.  The mutations are distributed all

21   across the region which result in 5 amino acid

22   substitutions.  She is characterizing more isolates

22

1   and she already has 6 from 2002, 11 from 2003 and 6

2   from 2004.  So, the purpose is to really compare

3   and to see what the differences are and how those

4   differences impact on our tests.

5             Next slide please.  The outcome of the

6   panel testing--six laboratories participated in

7   that.  She tells me there were no false-positive

8   results reported.  More variability in detection

9   was found towards the lower end of the viral

10   concentration, i.e., 80 percent of the time

11   detected 100 copies/ml member but all laboratories

12   detected 100 percent of the time the panel members

13   of 500-1000 copies/ml.  Further testing is going to

14   define the consensus copy number.

15             Next slide, please.  This is the important

16   slide.  I would like to thank all the people who

17   really helped to make this talk possible.  Jennifer

18   Brown, whom I have always been bugging to provide

19   the slides.  Thank you, Jennifer.  Dr. Sue Stramer,

20   Dr. Mike Busch and Sally Caglioti, Dr. Mike Strong,

21   Leanne Kiviharju, Roland, Maria and all these

22   people--whoever I send an email they are kind

23

1   enough to respond quickly.  Also my colleagues at

2   the FDA, Maria Rios, Alan Williams, Dr. Epstein,

3   Martin Ruta, Indira Hewlett--always helping in this

4   whole project and, last but not the least, the AABB

5   task force.  I am really, really grateful to them

6   for providing all the information and helpful

7   discussion.  Thank you very much.

8             DR. NELSON:  Thank you.  Any questions or

9   comments?  Yes?

10             DR. GOLDSMITH:  Do you have additional

11   data on the level of viremia in these samples that

12   you have been studying?  What is the maximum level

13   of viremia?

14             DR. NAKHASI:  Which samples are you

15   talking about?

16             DR. GOLDSMITH:  The ones that you

17   recovered from the viremic donors.

18             DR. NAKHASI:  From the viremic donors, I

19   don't know.  Maria, do you know what the levels

20   are?

21             DR. RIOS:  Between 10

5 and 106 is the high

22   level of viremia that we have found.  Are you

24

1   asking for the range of viremia or the high level

2   of viremia?

3             DR. GOLDSMITH:  I was just curious about

4   the high but it is fine to give the range.

5             DR. RIOS:  It varies.  It varies.  The

6   assays, in general, that use lower volumes do not

7   detect them.  Assays that have higher volume and

8   high throughput detect, but do not give accurate

9   quantitation, to 10

6 copies/ml.

10             DR. NELSON:  One of your slides had 23

11   positives with 16 by minipool and 7 by ID.  Were

12   those 7 not detectable by minipool or was it just

13   that ID screening was triggered and they weren't

14   tested by minipool?

15             DR. NAKHASI:  I think they came for the

16   ID-NAT testing.  Is that true?  Yes.  You know, in

17   BSL they had already started ID-NAT testing in

18   Maricopa County.  The trigger had started earlier.

19             DR. NELSON:  So, they were negative by

20   minipool?

21             DR. STRONG:  No, the trigger was activated

22   and they started doing ID screening so they haven't

25

1   gone back yet, I think, to see if those would have

2   been picked up by minipool.

3             DR. BUSCH:  Actually, 7/12 that were

4   picked up in the region that had been converted to

5   ID-NAT, 7 of them had been fully worked up and 5 of

6   those 7 are negative at 1:16 dilutions so they

7   would have been missed by minipool.  Of those 5, 1

8   of them is antibody negative and 4 have IgM and

9   IgG.

10             DR. NELSON:  Has anybody looked at the

11   characteristics of the donors that have low levels?

12   Are there host factors that might influence whether

13   somebody has high level or low level?  I know one

14   feature may be antibody but in those that are

15   antibody negative, I wonder if there are any donor

16   characteristics that influence the level of

17   viremia.

18             DR. BUSCH:  Sue has looked at that I think

19   more formally and there wasn't any correlation.

20   These are representative donors of the donor pool

21   in terms of the viremics, non-viremics and low

22   viremics.  I think it is just by chance.  This

26

1   phase of early viremia is completely asymptomatic.

2             DR. RIOS:  It may have some inherited

3   characteristics that limit the viral replication.

4   The reason why we think that is because we have

5   performed some in vitro studies with human primary

6   macrophages and there is a great variability not

7   only in the day of the viral peak, but some

8   individuals can have a very steady and low titer

9   that doesn't progress to peak.  So, that indicates

10   that some inheritance variability may interfere

11   with replication.

12             DR. NELSON:  That is interesting.  Other

13   questions?

14             DR. LAAL:  Unless I misunderstood, I

15   noticed that in 2003 we had a majority of your

16   isolates from people who had fever, and about

17   one-quarter were from neuroinvasive cases.  In 2004

18   it is reversed.

19             DR. NAKHASI:  Yes, that is an important

20   point.  I discussed it with the CDC folks and they

21   said, you know, don't pay attention to that because

22   the fever cases were--you know, this year they are

27

1   paying more attention so some of the fever cases

2   were not real fever cases.  You are right, you saw

3   the switch.

4             DR. LAAL:  But then in the isolates that

5   you are picking up now for the genetic studies, are

6   you carefully making sure that you look at both

7   types?

8             DR. NAKHASI:  Maybe Maria can say; I don't

9   know.

10             DR. RIOS:  The isolates that have been

11   studied so far don't come from patients.  Actually,

12   that is the effort we are going to move towards

13   now.  They are identified through the blood

14   screening.  So, in order to evaluate if there is

15   any isolate that may not be picked up by the blood

16   screening we need to acquire samples from cases

17   that are non-blood donors to investigate this

18   possibility.

19             DR. NELSON:  Yes, Mike?

20             DR. STRONG:  Just a quick comment on the

21   donors.  In the studies that were done last year,

22   many of the donors that were interviewed, in fact,

28

1   were symptomatic either shortly before or shortly

2   after their donations but the screening questions

3   just didn't pick them up.

4         IV. Hepatitis B Virus Nucleic Acid Testing (NAT)

5                    for Donors of Whole Blood

6             DR. NELSON:  Thanks.  The next topic is

7   hepatitis B virus nucleic acid testing for donors

8   of whole blood.  Dr. Gerardo Kaplan will introduce

9   this and give us background.

10                  A. Introduction and Background

11             DR. KAPLAN:  Good morning.

12             [Slide]

13             I am Gerardo Kaplan, Chief of the Lab of

14   Hepatitis and Related Virus Emerging Agents.  I am

15   with the Office of Blood, and I will introduce for

16   you the hepatitis B virus n nucleic acid testing

17   for donors of whole blood.

18             [Slide]

19             The general agenda for this meeting is

20   that after the introduction and background, Dr.

21   Blaine Hollinger will give us an update on the

22   serology of hepatitis G.  This will be followed by

29

1   two presentations from the Roche Molecular Systems

2   and their preclinical and clinical data in support

3   of their application.  Finally, I will come back to

4   give you the FDA perspective on hepatitis B MP-NAT

5   and present the questions for the committee.  I

6   understand that there will be a break and then a

7   public hearing.

8             [Slide]

9             So, the issue is that the FDA is seeking

10   the opinion of the committee on the performance of

11   the Roche COBAS AmpliScreen HBV test in minipools

12   of 24 samples to screen blood for transfusion by

13   nucleic acid testing, and its proposed intended use

14   as an alternative to hepatitis B surface antigen

15   testing in conjunction with testing for antibodies

16   to hepatitis B core antigen.

17             [Slide]

18             In support of their claims, Roche

19   Molecular Systems performed a clinical trial.  The

20   study objectives of the clinical trial were to

21   determine whether the COBAS AmpliScreen test, which

22   is a minipool of 24 samples of plasma from

30

1   volunteer blood donors, can detect hepatitis by DNA

2   in hepatitis B surface antigen-anti-core negative

3   window period cases.  This is the primary objective

4   of the study.  In hepatitis B surface

5   antigen-positive donors, who are acutely infected

6   or chronic carries, and in persons previously

7   exposed to hepatitis B as the secondary objective.

8             [Slide]

9             During the clinical trial, Roche Molecular

10   Systems identified two window period cases in about

11   600,000 volunteer whole because donations screened

12   by hepatitis B NAT using their minipools of 24

13   samples.

14             RMS, Roche Molecular Systems, claims that

15   the use of the COBAS AmpliScreen HBV test in

16   conjunction with the anti-core test would reduced

17   the residual risk of transfusion-transmitted

18   hepatitis B, and also that this test, the COBAS

19   AmpliScreen, could be used as an alternative to the

20   commonly used hepatitis B surface antigen donor

21   screening test.  I would like to point out that

22   blood in the U.S. is currently being tested by

31

1   surface antigen and core antibodies.

2             [Slide]

3             So, we would like to request comment of

4   the committee on three questions.  Basically, the

5   firs would be do the sensitivity and specificity of

6   the Roche COBAS AmpliScreen hepatitis B test in

7   minipools of 24 samples support licensing of the

8   assay as a donor screen?

9             [Slide]

10             If so, assuming continued use of screening

11   tests for anti-hepatitis B-core, do the data

12   support use of the Roche COBAS AmpliScreen

13   hepatitis B test in minipools of 24 samples to

14   screen blood for transfusion as an equivalent

15   alternative to the surface test?  If not, which

16   studies would be required to validated such a new

17   test?

18             The third question is do the data support

19   use of the Roche COBAS AmpliScreen hepatitis B test

20   on minipools of 24 samples to screen blood for

21   transfusion as an added test in conjunction with

22   licensed donor screening tests for hepatitis B

32

1   surface and anti-core?

2             Now I would like to introduce Dr. Blaine

3   Hollinger.  He will give us an update on hepatitis

4   B serology.

5               B. Serological Course of Hepatitis B

6             DR. HOLLINGER:  Thank you, Gerardo.  It is

7   always a pleasure to come to these meetings and I

8   always tend to learn more than I think anyone here

9   because there is always so much information.

10             [Slide]

11             The talk today is about serology, and

12   serology can be confusing.  It is confusing to our

13   GI residents and fellows in trying to determine

14   what happens after hepatitis B infection.  The

15   clinical and serologic changes that occur following

16   infection represent a complex interaction between

17   the host, the virus and specific antigens.  If you

18   have an understanding of these items, I think it

19   becomes very easy to understand what happens.

20             There are certain changes that occur after

21   infection with all viruses that seem to be very

22   similar.  Viruses are either enveloped or

33

1   non-enveloped.  Hepatitis B virus happens outcome

2   have a lipoprotein envelope that contains hepatitis

3   B surface antigen as the envelope protein.  There

4   are also some other particles in this serum which

5   are seen in excess of the infectious virion, and

6   these are the small particles here, and these

7   tubular forms which actually come off, in many

8   cases, from the hepatitis B virion.

9             The inside, or the nucleocapsid of the

10   virion is comprised of the hepatitis B core

11   antigen, and this encloses a relaxed, circular,

12   partially double-stranded DNA molecule.  Another

13   antigen which is seen also is the hepatitis B

14   antigen, but this is not part of the virion in any

15   way.  It is a secretory protein, about 16-17 kd in

16   size, that is secreted during active infection, and

17   is representative usually of high concentrations of

18   virus in the bloodstream.  We don't really know

19   what is the reason for this particular antigen.

20             [Slide]

21             This is the HBe genome.  It has four open

22   reading frames.  One of them is responsible for the

34

1   hepatitis B-surface antigen and is comprised of a

2   small, a medium and a large protein.  A second one

3   is the core gene which expresses the core-antigen,

4   the hepatitis B core-antigen and also the e-antigen

5   as well.  There is a DNA polymerase that is

6   important for the replication of this virus, and

7   there is an x-gene here which is very important for

8   initiation and probably maintenance of the

9   infection.  It is probably also weakly oncogenic as

10   well.  Now, once one understands these things, you

11   can then start to see what happens after an

12   infection.

13             Next slide, please.  This slide shows what

14   happens in the typical course--now, I am going to

15   talk about typical things here.  Certainly all of

16   us know about the exceptions and we have all seen

17   them at one time or another, but I think it is

18   important to deal right now with the typical

19   changes that can occur.  Again what happens, you

20   have a hepatocyte that is infected.  There is a

21   very short eclipse phase in which we would not see

22   virus in the hepatocyte.  Then viruses start to be

35

1   produced and are transported out of the cell.

2   Then, one sees initially HBV DNA in the

3   bloodstream.  It precedes the expression of

4   hepatitis B-surface antigen which then starts to

5   occur here very early in the course of the disease

6   and peaks usually during the acute phase of a

7   clinical illness.

8             What happens then in this non-cytolytic

9   infection is that there are usually some holes that

10   are punched in the plasma membrane of a hepatocyte.

11   ALT, which is in the cytosol, aminotransferase

12   enzyme which is in the cytosol of hepatocytes, are

13   released into the bloodstream.  So, the ALT is

14   elevated in these patients and then follows the

15   clinical symptoms of anorexia, fatigue, jaundice,

16   etc.

17             Usually about the time the ALT begins to

18   appear, HBeAg and IgM anti-HBc are found.  IgM anti

19   HBc appears in high titer or high concentration in

20   individuals as a response to the synthesis of

21   nucleocapsids and, therefore, represents active

22   viral replication.  The IgM rises during the early

36

1   courses of infection and then starts to disappear.

2   Usually by 4-6 months or so, or 6-9 months, the IgM

3   anti-HBc disappears, although it may be present for

4   up to 2 years in individuals, if you use very

5   sophisticated research tools.

6             Usually the HBV DNA disappears in the

7   acute infection at the time that a hepatitis

8   B-surface antigen disappears.  So, you have a

9   pattern here.  HBV DNA first; HB surface antigen,

10   and the HBsAg disappears and usually about that

11   time HBV DNA is gone, or at least is

12   non-detectable.

13             Again, another pattern appears.  When

14   HBsAg disappears it is replaced by its antibody,

15   anti-HBs.  Now, we know that probably anti-HBs is

16   produced very early in the course of the infection,

17   even probably back in this area here, which results

18   in serum sickness, vasculitis or other things, but

19   you don't detect it because there is so much

20   hepatitis B surface antigen available so it is

21   non-detectable.  But eventually the surface antigen

22   goes; anti-HBs then is detected.  When HBeAg

37

1   disappears in the active infection, then anti-HBe

2   replaces it.  Anti-HBe lasts in most patients for a

3   long period of time but in about a third of the

4   patients disappears after about 6 months of

5   follow-up.

6             Of course, the IgM anti-HBc--there is a

7   switch-over between the IgM anti-HBc and IgG

8   anti-HBc.  What is important to recognize is that

9   the assay, the total anti-HBc assay, detects not

10   only IgM but IgG antibodies.  It is not an IgG

11   test; it detects both IgG and IgM.  So, whenever

12   IgM is present you ought to see the total anti-HBc

13   test positive.

14             Next slide, please.  This is an older

15   slide I have but I think it helps us a little bit

16   understand the relationship of all these assays

17   that we are looking at.  The EIA assay and those

18   that are being developed can detect down to about a

19   tenth of a nanogram of HBsAg.  HBsAg circulates up

20   to 200 mcg or more per ml of blood.  These tests

21   really have an upper level of sensitivity of about

22   1 mcg/ml.  So, usually they are at their upper

38

1   limits of detection for most infections.

2             The second test that came along them were

3   the hybridization assays, the liquid phase and the

4   filter hybridization assays.  These could detect

5   down to a tenth to 1 pg of genomic equivalence in

6   the blood.  At least, at 24 hours if you have the

7   autoradiographs stay for maybe up to 5 days, you

8   can detect anywhere from 2- to 10-fold more virus

9   in the blood.  So, they are detecting about this

10   level here, somewhere between here and here, and 1

11   pg is equivalent to about 300,000 copies/ml.  You

12   will see 280,000, 330,000 but let's just talk about

13   300,000 copies/ml.  So, 1 pg is here, 300,000,

14   30,000, 3,000 and 300 down at the femptogram

15   levels.

16             The PCR assays, which are somewhere

17   between 700-7,000 times more sensitive than

18   hybridization assays, can detect down to 50 copies

19   or less per ml.  So, they are down in this range,

20   here.  As I said, this is 300,000, 3,000 and 300

21   copies/ml.  This would be 30 copies/ml.  These are

22   usually with nested PCRs or other assays, nucleic

39

1   acid--the tests that are available today.

2             So, you see the differences between these

3   assays.  One might ask why is it then that the EIA

4   tests are detecting HBsAg very closely to the HBV

5   DNA.  Part of the reason for that is that there are

6   anywhere from 1,000 to 10,000 times more--at least

7   10,000 times more hepatitis B surface antigen

8   particles per infectious virion, and there may be

9   even more naked HBV particles.  So, you have a lot

10   more HBsAg produced initially than you do

11   infectious virions.  That is the reason I think

12   that there is not a great deal of difference, at

13   least initially, in the detection of these

14   particles.

15             Next slide, please.  So, during the course

16   of infection HBV DNA is detected from 2-5 weeks

17   after infection and up to 40 days before the HBsAg

18   is detected.  However, there is a wide coefficient

19   here of variation but the mean of that is around

20   6-15 days, and there are some exceptions to this.

21   It rises slowly during the course of infection at a

22   relatively low level, say 10

2 to 104 copies/ml

40

1   during the seronegative period, that is, before the

2   HBsAg appears.

3             Next slide, please.  HBsAg, on the other

4   hand, appears 1-3 weeks before the ALT becomes

5   abnormal or 3-5 weeks before the onset of symptoms

6   or jaundice.  It reaches a peak during the acute

7   stage of the disease and then it declines to

8   undetectable levels within 4-6 months.

9             Next slide.  IgM anti-HBc is indicative,

10   as I said, of ongoing viral replication and appears

11   at the onset of ALT abnormality at high

12   concentration.  It is present but undetectable in

13   some chronic infections and may reappear,

14   therefore, during reactivation of HBV during

15   chronic infection.

16             Now, one of the questions that often

17   arises is when you have acute hepatitis how do you

18   determine that this is not a reactivation of

19   chronic disease compared to an acute disease?  That

20   has always been an issue.  For all practical

21   purposes, the IgM anti-HBc test can be utilized to

22   determine that.  It is at high concentrations.

41

1   Whereas, patients who have reactivation of their

2   hepatitis B virus infection from chronic disease

3   often have a relatively low concentration of IgM

4   anti-HBc in their test.  We use a sample to cut-off

5   ratio and it is usually somewhere between 1-3 in

6   those cases, whereas in acute disease it is at its

7   maximum.

8             In addition, the other way, if you have

9   the means to do this, is that the IgM has a 19S

10   sedimentation coefficient during acute infection.

11   It has a 7S sedimentation coefficient during

12   chronic infection.  That is another way that one

13   could use to determine acute from chronic.

14             Next slide, please.  The anti-HBs is a

15   neutralizing antibody occurring during recovery and

16   after infection, and in most cases indicates

17   protection against reinfection.  However, it may

18   become undetectable in up to 20 percent of patients

19   after several years.

20             Next slide.  This slide sort of shows what

21   happens in a patient who gets acute disease and

22   gets over it.  The anti-core is a very strong

42

1   antibody, high concentration of antibody compared

2   to anti-HBs.  Over time, as I said, in about 20

3   percent of patients this antibody will disappear.

4   Then you are left with a person who has only an

5   anti-HBc response.

6             Next slide.  One of the ways that one can

7   tell that that is present is you can do an HBsAg

8   vaccine challenge.  You can give them a vaccine.

9   About 2-4 weeks later you have them come back in

10   for a quantitative anti-HBs test and if the

11   response is greater than 10 milli-international

12   units/ml, that is indicative that this person

13   probably had a prior infection and was immune.  It

14   is unusual for a susceptible person or a person

15   with a low-level infection to respond to the

16   hepatitis B surface antigen by 2-4 weeks.  Probably

17   less than 3 percent in 2 weeks will respond to the

18   vaccine at this particular level.

19             Next slide.  I have talked a little bit

20   about chronic infection.  The difference is that

21   arbitrarily, say, if the HBsAg is still positive

22   after 6 months or the person has ALT abnormalities

43

1   for a period of time, then they probably have

2   chronic infection.  In addition, the HBsAg is at

3   very high concentrations through that period of

4   time.  It doesn't start to decline.  So, if you

5   could dilute these samples as a patient is being

6   followed along, you will find that the HBsAg

7   concentration is declining.  Just remember that the

8   upper level of their reference is about 1 mcg/ml so

9   you do have to do some dilutions in order to see

10   this happening.  The IgM does disappear but remains

11   usually positive at very low levels.  With research

12   tools you can detect it in most patients.  It just

13   isn't positive in the way the test is done, which

14   is to dilute the sample about 1:1000 or 1:2000.

15   That is why you don't see it in these individuals.

16   That is how the test is configured.  Part of that

17   is because of prozone effects and other things

18   which could make a positive test negative if you

19   did not dilute.

20             E antigen is usually present initially,

21   but over time about 5-10 percent of patients will

22   seroconvert from HBeAg positivity to anti-HBe

44

1   positive.  During that period of time they usually

2   show reactivation of their disease.  Often it is

3   just an elevation of the liver enzymes but they may

4   become jaundiced and they may develop symptoms that

5   look just like acute disease.  Until we understood

6   this, this was one of the reasons we used to think

7   that 10 percent of patients who had acute disease

8   would become chronically infected.  Now that we

9   know that many of these patients were chronically

10   infected to start with, we now feel that probably

11   less than 3 percent of the immunocompetent patients

12   who develop acute disease will actually become

13   chronically infected.  Of course, the anti-core and

14   HBsAg remain positive in most patients.

15             Next slide, please.  Let me just go

16   through, finally, some of the findings that we just

17   talked about.  This is HBVd and a surface antigen,

18   anti-HBc and anti-HBs.  You can see in the

19   pre-seroconversion window period--we have a couple

20   of windows here because we used to talk about a

21   window period between the end of the HBsAg phase

22   and the development of anti-HBs in which the IgM

45

1   anti-core is present during acute disease.  But

2   here I am talking about the pre-seroconversion of

3   antibodies or the seronegative, if you will, window

4   period.

5             Then, in the early acute infection both

6   the HBVd and a and HBsAg are positive.  The

7   anti-core and the anti-HBs are negative.  The

8   presence of HBsAg is indicative of a hepatitis B

9   infection, either acute or chronic.  Just the HBsAg

10   alone--you can't tell.  But the presence of all

11   three of these together in the absence of anti-HBs

12   indicates an HBV infection, either acute or

13   chronic.  Eventually you will lose these markers

14   and you are left with anti-HBc and anti-HBs, which

15   means a previous infection with immunity.

16             Next slide.  The other findings are in

17   individuals who have a solitary anti-HBc test

18   present.  They have no anti-HBs and no HBsAg.  Some

19   of them may contain HBV DNA and these represent

20   low-level carriers which do not have enough surface

21   antigen present in the blood to be detectable by

22   the current assays.

46

1             The other thing is that if this is

2   negative, this could also indicate an early

3   convalescent period in which this anti-HBc will be

4   IgM.  It could represent an HBV infection in the

5   remote past, as we talked about, in which the

6   anti-HBs has disappeared and you could do a vaccine

7   challenge, or it could be in many cases a

8   false-positive reaction.  Most of these reactions

9   show anti-core at a very low level, again, a sample

10   to cut-off ratio somewhere between 1-3.  The

11   anti-HBe is usually absent in these individuals as

12   well, and they don't respond to the vaccine

13   challenge.

14             Then you have a group that have no

15   markers.  This usually excludes an HBV infection in

16   this case.  Then you have individuals with anti-HBs

17   only.  This represents a vaccine type response.  It

18   is really unusual to see something else happening,

19   or you shouldn't see anything other than anti-HBs

20   in a patient who gets the vaccine because it only

21   contains HBsAg.

22             Next slide.  This is my final slide.  It

47

1   looks at discordant or unusual hepatitis B

2   serologic profiles requiring further evaluation.  I

3   have probably seen all of these at some time or

4   another in my career.  For example, I have a

5   patient right now who has HBsAg positive only, who

6   is anti-core negative, with 900 million copies or

7   IU of virus per ml of blood.  He has been like that

8   most of his life.  It happens in some young

9   children who get hepatitis B and become chronically

10   infected.  It is seen in other adults, particularly

11   in Asians.  But it is a really unusual phenomenon.

12   So, that is a possibility but really rare.

13             Finding of all these three markers in one

14   patient does occur in about 5-10 percent of

15   patients.  It is mostly in drug users or people who

16   have been exposed to different genotypes of the

17   virus.  It could be found in people who have been

18   vaccinated and have made antibodies to a different

19   genotype, and it usually is in patients who have

20   more serious liver disease.

21             Anti-HBs or anti-HBc positive only--we

22   have talked about those phenomena.  The are not so

48

1   uncommon.  HBsAg negative individuals who are HBeAg

2   positive--you shouldn't really see that but I have

3   seen at least two cases that have been HBeAg

4   positive, HBV DNA negative but do not have the

5   presence of surface antigen present.

6             You could have patients who are positive

7   for both e antigen and anti-e at the same time.

8   That has to do with the nuances of the test and

9   probably nothing else.  And, you can find patients

10   who are anti-core negative--I mean, you should not

11   find a patient who is anti-HBc negative but IgM

12   anti-HBc positive.

13             Well, I hope now you are a little bit more

14   familiar with the serology of this disease.

15   Antigens are followed by antibodies.  HBV DNA

16   precedes the development of antigens very early in

17   the course and usually remains throughout the

18   disease entity itself.  Thank you very much.

19             DR. NELSON:  Thanks, Blaine.  Are there

20   any questions?  Harvey?

21             DR. KLEIN:  Blaine, there have been a

22   number of reports of people who have low levels of

49

1   HBV DNA and have anti-HBs.  Do you know if any of

2   those are infectious?  Have you ever seen one?

3             DR. HOLLINGER:  Well, it is not whether we

4   have ever seen one.  The issue I guess is that we

5   have this question all the time about occult

6   infection in HBV DNA even though there are no

7   markers, usually in people who are very sick and

8   would not be donors, for example, liver cancer

9   patients and immunocompromised individuals.  So,

10   you wouldn't find these as donors anyway.  It is my

11   belief that most of these issues of finding nucleic

12   acid in the presence of antibodies probably means

13   it is neutralized and it is not infectious.  You

14   see that with RNA found very late in the course of

15   the disease.  In the presence of antibody it

16   doesn't appear to be infectious biologically,

17   clinically or in trials.

18             So, I think that finding anti-HBs at a

19   reasonable concentration and HBV DNA would probably

20   not indicate somebody is infectious.  But you just

21   have to do these studies, and no one has really,

22   say, taken a chimpanzee and looked at these.  I

50

1   think Prince did so a few years ago and the data

2   was that they were not infectious.

3             DR. NELSON:  You didn't mention the

4   genetic variations in the surface antigen.  Could

5   those lead to a false-negative surface antigen test

6   in people who were actually infectious?

7             DR. HOLLINGER:  Yes, Kenrad, it is a good

8   question.  There are about six or seven genotypes

9   of this disease, from a to h, but because they all

10   have a group a antigen present, they all seem to be

11   detectable by the usual tests which are available

12   today.  Now, could it happen?  Sure, I think it is

13   possible but I think it would be unusual.

14             DR. NELSON:  One other comment too, we

15   have studied hepatitis B in injection drug users

16   and it is extremely common for them to have

17   hepatitis B core antibody only, without surface, at

18   very high levels.  They are actually truly infected

19   and it is up to 20 percent of injection drug users.

20   It is possible that co-infection with hepatitis C

21   may play a part in that because they are all

22   infected with hepatitis C as well.  It is a very

51

1   complex interaction.  But seeing hepatitis B core

2   antibody only in somebody who actually is a drug

3   user and comparing drug users with others without a

4   history, it is like 20 percent of the drug users,

5   maybe 1 or 2 percent in gay men who have had

6   hepatitis B infection.

7             DR. HOLLINGER:  Yes, you see it not only

8   in drug users but you see it in volunteer donors

9   who are found to be anti-HCV positive and then you

10   go do some other tests and they are found to be

11   anti-HBc positive only.  They don't respond well to

12   the vaccine.  And, we know that HCV interferes with

13   the replication of hepatitis B virus so that is a

14   possibility.

15             DR. NELSON:  Other comments?

16             [No response]

17             Thanks.  Dr. Herman, from Roche?

18         C. Preclinical and Clinical Data for HBV MP NAT

19             DR. HERMAN:  Thank you very much.  It is a

20   pleasure to be here to describe to you the results

21   of our non-clinical and clinical performance

22   studies on our COBAS AmpliScreen HBV test.

52

1             [Slide]

2             I am going to begin the talk with a

3   description of our non-clinical performance

4   studies, and Dr. Frank will conclude with a

5   description of our clinical study results.

6             [Slide]

7             Here is an outline of the talk.  I am

8   going to give an overview of the COBAS AmpliScreen

9   system and then I am going to describe the

10   different portions of the non-clinical performance

11   studies and give you their results.

12             [Slide]

13             The COBAS AmpliScreen system is designed

14   for screening of minipool samples or individual

15   donor samples.  The samples are put onto the

16   Hamilton microlab pipetting robot which prepares

17   the pools.  Then the pooled or individual samples

18   are processed to extract nucleic acid with the

19   generic sample processing method.  There are two

20   variations of the method, the multiprep method for

21   the pooled samples and the standard method for

22   individual donor samples, and I will describe them

53

1   in the next slide.

2             Aliquots of the extracted material can be

3   processed in parallel with our three different

4   COBAS AmpliScreen tests, the test for HBV and the

5   licensed tests for HCV and HIV.  The analysis is

6   done on the COBAS Amplicor analyzer which automates

7   the nucleic acid amplification and detection

8   processes.  Then, the software for the analyzer and

9   the software for the pipetting machine interact

10   with the data output management system.

11             [Slide]

12             This slide describes the two sample

13   preparation methods, the multiprep method which is

14   used for the minipools, and the standard method

15   that is used for individual donor samples.  For the

16   minipool method, 1 ml of the pool is centrifuged at

17   23,500 g to enrich for the viral targets, HCV, HIV

18   and HBV.  After the centrifugation, 900 mcL of the

19   supernatant is discarded, leaving behind the pellet

20   and 100 mcL of the supernatant.  The nucleic acids

21   are extracted by adding a k-atropic lysis reagent

22   which contains the internal control.  After a brief

54

1   incubation, isopropanol is added to precipitate the

2   nucleic acids.  The nucleic acids are collected by

3   centrifugation.  The pellet is washed with ethanol

4   and resuspended in a specimen diluent, and then an

5   aliquot of the specimen diluent can be used in each

6   of the AmpliScreen tests.

7             The standard specimen processing method

8   for individual samples is almost identical.  All

9   the green steps are identical and the differences

10   are highlighted in red.  So, the starting specimen

11   volume is 200 mcL instead of 1 ml, and there is no

12   concentration step.  The k-atropic lysis reagent

13   with the internal control is added directly to the

14   200 mcL specimen and then the remaining steps are

15   the same.

16             [Slide]

17             Now I am going to describe the analytical

18   sensitivity studies that were performed.  These are

19   the results using the multiprep procedure.  We used

20   the World Health Organization HBV DNA international

21   standard and we prepared dilutions of it between

22   100 and 3 IU/ml and analyzed 120 replicates of each

55

1   dilution, at Roche, using 2 different kit lots.

2   This shows the results.  So, we had 100 percent hit

3   rate at 100, 30 and 10 IU/ml and a 95.8 percent hit

4   rate at 5 IU/ml.  The predicted 95 percent limit of

5   detection, using the PROBIT statistical method, is

6   4.4 IU/ml.

7             [Slide]

8             This slide shows the results using the

9   standard specimen processing procedure.  In this

10   study we analyzed dilutions of the World Health

11   Organization standard from 300 to 10 IU/ml.  There

12   was 100 percent hit rate on 100 replicates at 300

13   and 100 IU/ml, and 99 percent on 30 IU/ml, and 97

14   percent at 20, and 95.8 percent at 15 IU/ml.  The

15   predicted 95 percent limit of detection, using the

16   PROBIT method, is 16 IU/ml for the standard

17   specimen processing method.

18             [Slide]

19             This slide describes the results on a

20   panel prepared by CBER.  The original panel

21   contained 3 members at 100, 10 and 0 copies/ml, and

22   at Roche we prepared 4 additional dilutions from

56

1   the 100 copy/ml member at 50, 25 5 and 2.5

2   copies/ml.

3             This slide shows the results using the

4   multiprep specimen processing method and on 6

5   replicates, 100 percent hit rate at 100 copies and

6   at 10 copies/ml, then 67 percent and 42 percent

7   hits at 5 and 2.5 copies/ml.

8             Using the standard specimen preparation

9   method, 100 percent hits at 100 copies/ml with 6

10   replicates and 92 percent and 75 percent at 50 and

11   25 copies/ml.

12             [Slide]

13             This slide shows the results of our

14   analysis of the performance of the test on HBV

15   genotypes A through H.  These samples were obtained

16   from commercial sources, and we analyzed up to 25

17   individual isolates for each genotype.  The samples

18   were quantified using the COBAS Amplicor HBV

19   monitor test or, in the case of genotypes G and H,

20   monitor tests that we have in development.  Then

21   the samples were diluted to approximately 2-3 times

22   the limit of detection for the multiprep method and

57

1   the standard prep method, and then they were

2   analyzed using the COBAS AmpliScreen test.  So,

3   with the multiprep method and the standard method

4   all the isolates yielded positive results.  So, the

5   test detected HBV genotypes A through H.

6             [Slide]

7             This slide summarizes the results on 40

8   commercially obtained seroconversion panels using

9   the multiprep method and the standard prep method.

10   The slide shows the number of days that DNA was

11   detected prior to detection of surface antigen

12   using a licensed surface antigen test, the Ortho

13   surface antigen test system 3.  The orange bars

14   show the results with the multiprep test, and for

15   that analysis the samples were diluted 24-fold to

16   simulate minipool testing.  The blue bars show the

17   results with the standard prep on undiluted

18   samples.  The panels are sorted by the number of

19   days prior to antigen that DNA is detected with the

20   multiprep method.  So, using the multiprep method,

21   in 38/40 panels HBV DNA was detected prior to

22   surface antigen, and in 2/40 panels HBV DNA was

58

1   detected in the same bleed as surface antigen.

2             Using the standard prep method on neat

3   samples in 39/40 panels HBV DNA was detected prior

4   to surface antigen, and in 1 panel HBV DNA was

5   detected on the same day as surface antigen.  I am

6   going to show you the results on one of the panels,

7   this panel, 39.

8             Before that, with the multiprep method, on

9   the 38 panels on which DNA was detected prior to

10   surface antigen, DNA was detected an average of 17

11   days earlier, and with the standard method DNA was

12   detected an average of 22 days earlier.

13             [Slide]

14             Here are the results on panel 39.  This

15   panel had 14 bleeds.  This column shows the results

16   with the Ortho Surface Antigen Test System 3, and

17   you can see that this turned positive on bleed 12

18   at day 143.  This column shows the results with the

19   minipool test, and you can see that it was positive

20   on the first and third bleed and then positive

21   again at the 11th bleed on day 113.

22             On the previous slide, on this panel it

59

1   was described as having DNA detected prior to

2   antigen 30 days in advance this one bleed.  Even

3   though DNA was detected intermittently earlier than

4   that, we made a conservative interpretation and

5   only counted the number of bleeds that DNA was

6   continuously detected prior to surface antigen.

7             In this column, the results with the

8   standard preparation method on undiluted samples is

9   shown.  There were 11 bleeds prior to detection by

10   surface antigen and the HBV DNA was positive in 7

11   of them but did not yield a positive result on

12   bleed 11.  So, in the previous graph the standard

13   prep method was described for this panel as being

14   detected on the same day as surface antigen.

15             [Slide]

16             This just illustrates that even though DNA

17   was detected intermittently earlier, on multiprep

18   method we only counted 1 bleed, which was 30 days

19   in this case.

20             [Slide]

21             This slide summarizes the results on the

22   40 seroconversion panels.  Using the multiprep

60

1   method on samples diluted 24-fold to stimulate

2   pooling, DNA was detected prior to surface antigen

3   in 38/40 panels and on the same bleed in 2/40

4   panels an average of 17 days prior to surface

5   antigen.  Using the standard procedure on undiluted

6   samples, DNA was detected prior to antigen in 39/40

7   panels an average of 22 days earlier, and on the

8   same bleed on the 40th panel.

9             [Slide]

10             This slide summarizes the results of the

11   analytical specificity testing and the testing of

12   potentially interfering substances.  For the

13   analytical specificity we looked at 38

14   microorganisms and cell lines, and all these

15   samples yielded negative results with the

16   AmpliScreen and valid internal control results.

17   There was no cross-reactivity observed.

18             For the potentially interfering

19   substances, we looked first at 95 clinical

20   specimens with different diseases in patients

21   infected with these viruses or with autoimmune

22   disease or with yeast infections, and multiple

61

1   specimens of each.  These specimens were tested

2   both with the addition of HBV target and without.

3   So, the testing with the addition of HBV target was

4   looking for interference.  No interference was

5   observed.  Testing without HBV target was looking

6   for cross-reactivity and no cross-reactivity was

7   observed.

8             Then we looked at these various

9   potentially interfering substances.  Again, they

10   were tested with and without the addition of HBV

11   target.  On the tests with HBV target there was no

12   inhibition and on the tests without HBV target

13   there was no cross-reactivity.

14             [Slide]

15             This slide and the next look at our

16   reproducibility study.  This slide shows the

17   results of the multiprep procedure.  The study was

18   conducted at 3 clinical sites using 3 lots of

19   reagent.  The study was conducted over 5 days with

20   2 operators at each site.  Each day the operators

21   analyzed the 6-member panel, blinded 6-member panel

22   consisting of 2 negative samples and 4 samples at

62

1   low and moderate target concentrations.

2   Essentially the same results were obtained with the

3   3 kit lots and at the 3 sites.  So, there was 1

4   false-positive result with lot 3 and that happened

5   at site 3.  At the low copy samples the percent

6   hits were nearly identical among the 3 lots and

7   among the 3 sites.

8             [Slide]

9             This slide shows the reproducibility study

10   with the standard procedure and the same outcome.

11   Essentially the same results between the 3 lots and

12   between the 3 lights on the negative and the

13   positive samples.

14             [Slide]

15             Finally, this slide describes the

16   performance of the AmpliScreen test alone or in

17   combination with the licensed anti-core test on 918

18   antigen-positive clinical specimens.  I am showing

19   the results with the multiprep test on samples

20   diluted 1:14 to simulate pool testing and with the

21   standard method on undiluted samples.  So, looking

22   at the results of the AmpliScreen test alone with

63

1   the minipool method, 871 of the 918 samples were

2   DNA positive for sensitivity of 94.9 percent.

3   Using the standard method on undiluted samples, 898

4   samples were positive for sensitivity of 97.8

5   percent.  These 20 samples that were not detected

6   by the standard test were among the 47 that were

7   not detected by the minipool test.  All the NAT

8   negative samples were sent out for anti-core

9   testing and all of them were positive by the

10   anti-core test.  So, using the multiprep method, 1

11   sample had insufficient volume for the testing, a

12   sample that was positive with the standard sample

13   prep and negative with multiprep so we couldn't get

14   a result on that so it was excluded from the

15   sensitivity analysis.  So, for the 917 samples with

16   NAT positive results and anti-core results on the

17   NAT negative samples the sensitivity was 100

18   percent for NAT and core combined relative to

19   antigen.  On the 1900 samples with evaluable

20   results using the standard method, again, the

21   sensitivity was 100 percent.

22             [Slide]

64

1             To summarize the non-clinical performance

2   studies, the analytical sensitivity was determined

3   on the World Health Organization standard.  With

4   the multiprep method the limit of detection at 95

5   percent probability is 5 IU/ml and with the

6   standard method the limit of detection at 95

7   percent probability is 15 IU/ml.  The performance

8   on the CBER panel with the multiprep method, 100

9   percent hits at 10 copies/ml and with the standard

10   method 92 percent hits at 50 copies/ml.  Genotypes

11   A through H were detected by the test.

12             [Slide]

13             On the 40 seroconversion panels using the

14   multiprep method on samples diluted 24-fold, HBV

15   DNA was detected prior to antigen in 38/40 panels

16   an average of 17 days or earlier, and with the

17   standard method on undiluted samples DNA was

18   detected in 39/40 panels prior to antigen an

19   average of 22 days earlier, and there were no

20   panels in which antigen was detected prior to DNA.

21   There was no cross-reactivity or interference

22   observed in the analytical specificity and

65

1   potential interfering substances studies.  And, the

2   sensitivity on antigen-positive specimens, in

3   combination with the anti-core assay, was 100

4   percent.

5             [Slide]

6             I would like to acknowledge all the people

7   who did this work, Yuanfeng Yang and his group at

8   Roche Molecular Systems did an excellent job

9   developing the assay and performing these studies.

10   Larry Pietrelli coordinated the clinical studies

11   and, in this part of the talk, he coordinated the

12   reproducibility studies.  They were conducted at

13   the Community Blood Center of Greater Kansas City,

14   Royal Blood Centers, Minnesota and the Gulf Coast

15   Regional Blood Center.

16             Now I think I will turn the microphone

17   over to Dr. Frank who will describe the results of

18   the clinical performance study.

19             DR. FRANK:  Thank you, Dr. Herman.

20             [Slide]

21             It is my pleasure to be here today to

22   describe to you a prospective study to evaluate the

66

1   screening of plasma pools from volunteer blood

2   donations for the presence of HBV DNA.

3             [Slide]

4             A clinical trial summary will focus today

5   on two areas of primary focus of improved safety of

6   blood donations by HBV minipool-NAT.  The key focus

7   in that particular area will be looking at the

8   potential window period or pre-seroconversion areas

9   of HBV.  A secondary focus will be replacement of

10   surface antigen by HBV minipool-NAT.  To focus on

11   that, we will look primarily at the areas of

12   samples that were surface antigen positive that

13   were picked up by minipool-NAT and surface antigen

14   positive picked up by anti-core.

15             This clinical study was initiated in

16   August of 2002 at 5 U.S. sites, and was completed

17   in April of 2003.  A total of 704,902 specimens

18   were tested with HBV minipool-NAT, of which 581,790

19   were included in the sensitivity and specificity

20   analysis.  The differential between the two numbers

21   occurred because specimens were excluded if no

22   electronic results for surface antigen and/or

67

1   anti-core were available from the site.

2             [Slide]

3             If you look at this chart, which is a

4   schematic of the results listing the analytes

5   anti-core, surface antigen and DNA--if you can't

6   see in the back of the room, the red means they

7   were HBV negative for that particular analyte; the

8   green means they were positive for that particular

9   analyte.  The totals are on the far right-hand

10   column, the numbers that had that criteria.  As you

11   can see, the majority of the donors were, as

12   expected, negative for all three.

13             [Slide]

14             We were primarily interested in looking at

15   this group down here, and those were enrolled in

16   our follow-up study, the protocol for which I will

17   be describing briefly.

18             [Slide]

19             Of these 4 eligible for the follow-up

20   study, the 2 areas we wanted to look at are

21   primarily the last row, which is the potential

22   window cases, that is, those cases which were HBV

68

1   minipool DNA positive but negative for core and

2   surface antigen.  This represents the type of

3   pattern one would expect in pre-seroconversion

4   window cases, and then the potential surface

5   antigen false-positive results, that is, surface

6   antigen positive but core negative and HBV minipool

7   negative as well.

8             [Slide]

9             I mentioned the follow-up protocol that

10   was utilized and this is the follow-up protocol for

11   those columns that we followed up, the test index

12   donation was tested by alternate NAT, provided by

13   the National Genetics Institute.  If alternative

14   NAT was positive, the result was quantitated and

15   then the subject was offered enrollment in a

16   6-month follow-up study for weekly draws times 4

17   then monthly draws to complete 6 months of

18   follow-up.  The analytes tested during follow-up

19   are as listed on your screen.

20             [Slide]

21             The primary objective was to improve the

22   safety by use of HBV minipool-NAT for detection of

69

1   window period cases.

2             [Slide]

3             As stated previously there were 23 donors

4   that were HBV DNA positive, surface antigen

5   negative and core negative.  Of these 23, 14 were

6   enrolled in the follow-up study.  It should be

7   noted that 9 donors declined follow-up and for the

8   purposes of our calculations for sensitivity and

9   specificity these were presumed to be false

10   positive and were negative on additional index

11   testing as well.

12             Of the 14 enrolled donors, 2 were

13   confirmed window period cases, the details of which

14   I will be showing next, and 12 were shown to be

15   valse positive on HBV due to persistently negative

16   anti-core, persistently negative surface antigen

17   and persistently negative HBV DNA, and negative by

18   alternate NAT index specimens.

19             [Slide]

20             This window period case is a 26 year-old

21   male repeat blood donor with no known risk factors.

22   His HBV DNA was positive for a quantitative level

70

1   of 2000 on index donation.  On day 17 post index

2   donation the surface antigen was positive, and on

3   day 48 anti-core was positive.

4             [Slide]

5             The next window case is a subject who is a

6   49 year-old female repeat blood donor, healthcare

7   worker, with a history of HBV vaccination.  This is

8   an interesting case because she had a negative

9   anti-surface antibody result 2 months prior to her

10   index donation when she donated blood previously.

11   What is interesting is between that time, where

12   this result 2 months prior was negative, and here

13   index donation was 2,340--now, this normally

14   wouldn't be done during her blood screening but she

15   was HBV DNA positive and she quantitated out at 200

16   copies/ml, which would be an infectious dose.  So,

17   this is a window period case of an individual who,

18   on follow-up, never did convert for surface antigen

19   and did, on day 22, convert for anti-core.  What

20   this likely represents is that somewhere between

21   the 2 weeks prior to her index donation when this

22   was negative and the result of her index donation,

71

1   she became infected with HBV and had a nice immune

2   system response to her previous HBV vaccination.

3             [Slide]

4             There are additional window cases I would

5   like to show to you today, and these were detected

6   by sites continuing to use HBV NAT following

7   conclusion of our clinical study.  The dates for

8   those 3 sites were April, 2003 to the present, and

9   there are 3 additional window period cases that

10   were detected in the one million donations

11   screened.

12             [Slide]

13             The first case is a 27 year-old male

14   repeat blood donor who reported high risk male

15   sexual partners.  He had HBV DNA positive on

16   donation with a quantitative level of 61,000.  On

17   day 7 his surface antigen turned positive and on

18   day 36 his core was repeat reactive.

19             [Slide]

20             This case is a window period case of a 29

21   year-old male repeat donor who has no appreciable

22   risk factors for hepatitis B, other than a possible

72

1   acupuncture treatment for 8 weeks prior to his

2   donations, weekly treatments.  His HBV index was

3   positive at a quantitative level of 2,300.  His

4   surface antigen turned positive 7 days later and

5   his core turned positive 28 days later.

6             [Slide]

7             This last case is still on follow-up but

8   this is a 50 year-old male repeat blood donor,

9   positive on index donation with 37,000 copies/ml.

10   On follow-up he remains positive on HBV DNA and the

11   traditional blood banking serologies are still

12   pending.

13             [Slide]

14             In summary, I thought it was interesting

15   to look at the 5 cases that HBV minipool-NAT picked

16   up.  If you look at the left-hand column, and as I

17   have noted in the footnotes of your slides, all of

18   these are repeat blood donors who could normally be

19   expected to be donating blood regularly to the

20   blood pool, and all donated blood units that would

21   have been missed by the current blood banking

22   algorithms.  All of them had infectious dose

73

1   concentrations on their index donations.  In fact,

2   even the lowest dose at 200 copies/ml, if the unit

3   were divided appropriately, would be sufficient to

4   infect everyone in this room.

5             [Slide]

6             Well, what is the yield for HBV

7   minipool-NAT?  If you look at the clinical study

8   yield alone we had 2 cases that I showed you out of

9   approximately 0.7 million donations for a yield of

10   1/350,000.  If we look at the continuing site data,

11   which I also showed you for the 3 cases in

12   approximately 1 million donations, the yield is

13   similar, 1/330,000.

14             Well, what does this mean in perspective

15   of what we are already doing in blood banking to

16   protect the nation's safety of the blood supply?

17   The yield for HBV minipool-NAT is approximately

18   equal to the yield for hepatitis C minipool-NAT and

19   the yield for HBV minipool-NAT is far greater than

20   the yield for HIV minipool-NAT.

21             [Slide]

22             If we look at this data that was presented

74

1   by Mike Busch in Paris of this year, this

2   emphasizes what I have just stated, that for HBV

3   the yield is approximately 1/340,000, based upon

4   his presentation.  For hepatitis C, 1/230,000, but

5   for HIV, 1/3.1 million.  Although these aren't

6   quite comparable, West Nile Virus was included for

7   because there aren't a plethora of antibody and

8   antigen tests available yet implemented for West

9   Nile Virus but that yield is 1/5,000.

10             [Slide]

11             This brings us to a sensitivity and

12   specificity calculation.  You can see how we

13   assigned HBV status, on the screen here, for

14   positives, negatives, the HBV status unknown for

15   those that were surface antigen negative but core

16   positive.

17             [Slide]

18             So, again, in summary, here are the totals

19   and the results of the initial testing.  These are

20   the potential surface antigen false-positive

21   results.  These are the potential window cases.

22   This last row I have gone into detail on, the row

75

1   above, the 4 and the 3 will be going into the

2   secondary objective.

3             [Slide]

4             The final status determination is here, on

5   the fir right-hand column.  The ones that could not

6   be determined are yellow, for those of you in the

7   back of the room.  The green is HBV positive.  The

8   red is HBV negative.  The window cases that were

9   proven in the clinical study are shown here as 2.

10   The remaining 21, which will be accounted for in

11   sensitivity and specificity calculations, were

12   counted as HBV positive but final status HBV

13   disease negative.

14             [Slide]

15             To that point, our specificity

16   calculations, if you look at the total number of

17   subjects for which we have data--the difference

18   between the numerator and the denominator is the 21

19   cases that I just spoke about a second ago to yield

20   a specificity calculation of 99.9964 percent, with

21   the ranges noted on your screen.  For sensitivity

22   we looked at the positive AmpliScreen minipool HBV

76

1   results in specimens with positive HBV status and

2   divided by the total number of specimens with

3   positive HBV status.  Again, these calculations are

4   for HBV minipool-NAT alone and the sensitivity here

5   is 89 out of 105 or 84.7 percent, with a range of

6   74.4 to 91 percent.

7             [Slide]

8             So, the primary objective conclusion is

9   that COBAS AmpliScreen HBV test has identified

10   individuals in  pre-seroconversion window period

11   that would otherwise have been undetected by the

12   blood banking system.  The COBAS AmpliScreen HBV

13   test is suitable for blood screening with minipool

14   strategies presented today, and these data indicate

15   that HBV minipool-NAT will increase the blood

16   safety.

17             [Slide]

18             Let's turn now to the secondary objective,

19   and that is consideration for the HBV minipool-NAT

20   and anti-core for replacing surface antigen and

21   anti-core.

22             [Slide]

77

1             The index donations of interest for this

2   discussion then would be those that were anti-core

3   negative/ surface antigen positive and we will

4   discuss those first.  Those donors fitting this

5   category were 7 out of the total number for which

6   we have data, or 0.0012 percent.  Those that had

7   core-negative/surface antigen positive HBV DNA

8   negative were 3.  Of those 3, 1/3 enrolled and

9   converted to anti-core.  The other 2 remaining

10   donors declined the follow-up study.  These were

11   determined to be HBV final status positives.  For

12   DNA negatives, that is DNA negative, core negative,

13   surface antigen positive donors, 2 out of these 4

14   enrolled in follow-up.  Two of the remaining

15   declined to enroll in follow-up but we do have

16   additional data that has been available and I will

17   make available and show to you today.  All 4 of

18   these were assessed to be HBV negative.  Let's look

19   at some of those details.

20             [Slide]

21             This first subject was HBV DNA negative,

22   surface antigen positive, core non-reactive.  On

78

1   follow-up this subject remained core non-reactive

2   throughout the follow-up study.  On surface antigen

3   the subject has non-reactivity on follow-up, and

4   HBV DNA remained negative, as did the other

5   follow-up indicators.

6             [Slide]

7             This would indicate a case of an initial

8   surface antigen test which did not repeat on

9   follow-up, and all other negative serology results

10   indicate this patient had an initial surface

11   antigen positive that is consistent with

12   contamination or carryover.

13             [Slide]

14             The next case is a little bit more

15   interesting.  This case is surface antigen positive

16   again on index donation, HBV negative, core

17   non-reactive.  It should be noted that the surface

18   antigen was positive and repeat reactive throughout

19   follow-up but that the surface antigen was negative

20   on neutralization for all of these testing time

21   points.  The donor also states a history of HBV

22   vaccination.  Of note, the core never did react

79

1   throughout the course of follow-up and HBV DNA was

2   negative throughout the course of follow-up.  The

3   anti-surface antibody was positive and that we

4   attribute to his vaccination status, and the repeat

5   reactive, again, was negative on neutralization

6   throughout.

7             In summary for this case, the anti-surface

8   antibody is explained by the donor's vaccination

9   history.  And, the persistent surface antigen EIA

10   repeat reactivity, which is negative on

11   neutralization, combined with all the other

12   negative results such as anti-core, is consistent

13   with a false-positive surface antigen result due to

14   cross-reactivity.

15             [Slide]

16             We have additional data for the 2 donors

17   who declined the follow-up protocol.  Of the

18   remaining 2 donors, 1 donor retested 1 month

19   following his index donation and was negative for

20   surface antigen, anti-core and anti-surface

21   antibody.  The other donor was retested a year and

22   a half post index and was negative for surface

80

1   antigen and for core antibody.

2             [Slide]

3             This slide may look somewhat familiar.

4   This is a look at the total HBV infected, HBV

5   surface antigen positive specimens.  In our study

6   there were 103 specimens fitting that

7   categorization.  If we look at sensitivity from

8   that perspective alone, HBV detected 87 of the 103

9   specimens and was negative on 16, for a sensitivity

10   as a stand-alone of 84.5 percent in this clinical

11   study.  However, the proposal of HBV minipool-NAT

12   plus anti-core antibody yielded positive in all 103

13   of 103 HBV positive cases and missed none of the

14   cases, for a sensitivity of 100 percent.  This

15   parallels the non-clinical data shown to you

16   earlier today by Dr. Herman.

17             [Slide]

18             What are the secondary objective

19   conclusions?  All 105 HBV positive donors were

20   identified by anti-core and HBV minipool-NAT

21   combined.  These 105 include 2 window case donors

22   that would have been missed by anti-core and

81

1   surface antigen combined, and 103 surface antigen

2   positive cases that were identified by the HBV

3   minipool-NAT and anti-core as an alternative.

4   These data then suggest that HBV minipool-NAT

5   combined with anti-core provide an alternative to

6   combined surface antigen and anti-core screening.

7             [Slide]

8             In summary conclusion today, I offer that

9   the COBAS AmpliScreen HBV test improves the safety

10   of blood transfusion.  I offer that COBAS

11   AmpliScreen HBV test combined with anti-core may be

12   considered as replacement for surface antigen

13   combined with anti-core.  This is supported by the

14   clinical trial data, and also supported by the

15   non-clinical trial data, specifically

16   seroconversion panel data shown here today and the

17   other HBV positive clinical specimens shown here

18   today by Dr. Herman.

19             [Slide]

20             With that, I would like to acknowledge the

21   blood centers that participated in this large

22   study, as well as the Roche Molecular personnel.

82

1   Thank you.

2             DR. NELSON:  Thank you very much.

3   Questions or comments?  Yes. Harvey?

4             DR. KLEIN:  On the two window period cases

5   that you detected with HBV alone, what was the

6   antigen assay that was used for those?

7             DR. FRANK:  In the two in the clinical

8   study, the first window period case was Ortho Test

9   System 2, and all remaining four window cases that

10   were presented were Abbott Auzyme.

11             DR. KLEIN:  And, do you by any change have

12   any data with head-to-head comparisons with PRISM?

13             DR. FRANK:  I don't believe we ran this

14   test against any test that was not approved by the

15   FDA.

16             DR. KLEIN:  Even in Europe?

17             DR. HERMAN:  We analyzed the

18   seroconversion panels against the PRISM and the

19   Ortho and I can show that data if you would like to

20   see it.  That was the European version of the PRISM

21   test that has the Paul Ehrlich Institute approval.

22   Do you want to see that data?

83

1             DR. NELSON:  Yes, I think so.

2             [Slide]

3             DR. HERMAN:  These are the same 40

4   seroconversion panels that I showed you previously

5   and the same NAT results but compared to the Paul

6   Ehrlich Institute licensed Abbott PRISM test, and I

7   guess this is the lot number.  Again, these are

8   sorted by the results with the multiprep test on

9   samples that were diluted 1:24-fold.  Those are the

10   orange bars.  The blue bars represent the results

11   with the standard test.

12             [Slide]

13             With the multiprep method DNA was detected

14   an average of 14 days prior to antigen, and with

15   the standard method DNA was detected 19 days prior

16   to antigen.

17             [Slide]

18             This slide has the summary.  Compared to

19   the European PRISM test, with the multiprep method

20   DNA was detected prior to antigen in 34 of the 40

21   panels, and with the Ortho DNA was detected prior

22   to antigen in 38 of the 40 panels.  On the other 6

84

1   panels DNA and antigen were detected in the same

2   bleed.  Using the standard method compared to the

3   PRISM test, DNA was detected prior to antigen in 38

4   of the 40 panels, and on the same bleed in the

5   other 2 panels.

6             DR. NELSON:  Thank you.

7             DR. HOLLINGER:  I just have a question.

8   We talk about the mean and the median between the

9   appearance of one antigen for nucleic acid before

10   the other one.  But these are not tests that are

11   done every day so the panels do not include--and to

12   say 14 days or 19 days is a little misleading I

13   think in some cases because theoretically if you

14   draw blood today and you drew one 200 days from

15   now, and that is all you had, you would say, well,

16   it would detect things 200 days before one was

17   positive and one was negative.  So, you know,

18   unless samples are collected every single day for a

19   period of time, this is somewhat misleading.  So,

20   give me some feeling about the days between each

21   one of these sample collections.

22             DR. HERMAN:  It varies between the panels.

85

1   So, the only data that I have available on the

2   slides here is the one panel I showed you where the

3   interval between bleeds I think was as few as a few

4   days and as large as 30 or 40 days.

5             DR. HOLLINGER:  It is not your fault in

6   doing this; it is just that, unfortunately, that is

7   how panels are set up.  But I think all of us have

8   to understand that it may be much less than that.

9   It won't be more than that but it certainly could

10   me much less than that.

11             DR. HERMAN:  That is absolutely correct,

12   and that is a limitation of these studies so one

13   shouldn't interpret this to conclude that if

14   testing was done every day these would be the

15   numbers.  This is an approximation, subject to the

16   limitations of the interval between the bleed dates

17   and the different panels.

18             DR. NELSON:  Actually, often

19   epidemiologists would assume a median--a mid-point

20   seroconversion between the two days.  So, if you

21   had a sample that was taken on day zero and day 14,

22   you would assume that the one that the one that was

86

1   positive on day 14 seroconverted on day 7.  That

2   would be one way, a different analysis of the data

3   than probably what you have done.

4             DR. EPSTEIN:  Could I get you to clarify

5   whether the apparent window period samples

6   themselves were tested with HBsAg assays with known

7   sensitivities of 0.1 ng/ml?

8             DR. HERMAN:  Let me defer to my

9   colleagues.  The window period cases from the

10   clinical study?

11             DR. EPSTEIN:  Correct, and you say you had

12   five of them.  Right?  Two from the trials that

13   analyzed sensitivity and specificity and the

14   additional three--you have five available samples

15   that could have been retested with the most

16   sensitive available HBsAg assays.

17             DR. STRONG:  I have to declare my conflict

18   here because I am one of the clinical test sites,

19   but we have been able to test four of the five.

20   The most recent one hasn't been tested, but the

21   four have been tested on the Ortho system 3, the

22   new licensed test.  It picked up one of the four

87

1   but missed three.

2             DR. KUEHNERT:  Could you clarify what

3   those are?  You said five.

4             DR. STRONG:  The five window cases that

5   are described here.  The first four were tested and

6   it picked up the one that you would expect it to

7   pick up, which is the high copy number sample.  The

8   fifth sample, which hasn't been tested, also has a

9   relatively high copy number and you might expect it

10   to be picked up as well, with 37,000 copies.

11             DR. NELSON:  Are there blood donor

12   policies or recommendations for the interval

13   between receiving hepatitis B vaccine and donation?

14   Because that is a source of a false-positive

15   infection for the surface antigen.  Are there any

16   policies on that?

17             DR. EPSTEIN:  FDA has no policy for donor

18   deferral after HBV vaccination.  I am not certain

19   whether AABB has a voluntary policy, but I think

20   for non-live vaccines they would also not recommend

21   deferring.  Someone else would need to clarify

22   that.

88

1             DR. KLEINMAN:  I think that is right, Jay,

2   except that there have been some instances where

3   people have had false-positive surface antigen so

4   it is sort of unspoken that you might want to defer

5   people for a week or so, but there is no absolute

6   AABB requirement to do that.

7             DR. LEITMAN:  Can I ask a different

8   question.  I am not sure I am following what is

9   going on in the 16 donors who were HBsAg positive,

10   anti-core positive but multipool-NAT, COBAS

11   negative.  Were those individuals individual donor

12   NAT-negative?  In other words, are they infectious

13   or are they not infectious?  If they are not

14   infectious, do they just happen to have a lot of

15   HBsAg particles circulating but not infectious

16   virion?

17             DR. FRANK:  I have to apologize, I can

18   only hear parts of the question because of where I

19   am standing.

20             DR. LEITMAN:  I don't think I can do that

21   again.  The 16 donors who were HBsAg positive,

22   anti-core positive but NAT-multipool negative, were

89

1   they retested in individual donor NAT to see if

2   they have low level viremia or DNA nucleic acid, or

3   are they, in fact, true HBsAg positive but not

4   infectious because they don't have intact virion

5   but just have particles circulating?

6             [Slide]

7             DR. FRANK:  You are talking about these

8   16?

9             DR. LEITMAN:  Those 16, yes.

10             DR. FRANK:  The final status here is that

11   they were HBV positive.  The question you are

12   posing is are these 16 patients infectious.  They

13   have surface antigen positivity and core

14   positivity.  I have to defer to my colleague as to

15   whether or not we tried to quantitate with NGI the

16   HBV minipool NATs.

17             DR. PIETRELLI:  I am Larry Pietrelli, from

18   Roche Molecular Systems.  Yes, on those 16 donors a

19   sample was supposed to be tested by alternate NAT

20   and also by individual donation.  One donor was not

21   tested by either assay.  Of the 15 remaining, 10

22   were positive by ID-NAT, of which 7 sere also

90

1   positive by alternate NAT.  Two were negative.  One

2   was not done.  Of the 5 that were negative by

3   ID-NAT, 2 were positive by alternate NAT; 2 were

4   negative; and 1 was not tested.  The viral

5   concentration was low, as expected.

6             DR. LEITMAN:  Thank you.

7             DR. KUEHNERT:  I have a follow-up question

8   about those 16.  They were core antibody positive.

9   Is there any way to quantitate that level of core

10   antibody positivity?  Were any of these at an

11   equivocal level?  I am just sort of getting at, you

12   know, whether you had greater numbers whether some

13   of these might have been close to a threshold for

14   missing them by core antibody positivity.

15             DR. FRANK:  I don't have that data.  I

16   don't know whether my colleagues have access to

17   that data today.

18             DR. PIETRELLI:  No, we did not collect

19   that information but we could go back to the sites

20   to see what the S to CO was.

21             DR. HOLLINGER:  I have a technical

22   question.  Is there are a reason you used 24,000

91

1   times g for pelleting?  I know others have used

2   40,000 times g for pelleting for an hour at the

3   enhancement stage, so to speak.  Also, why did you

4   use 24 samples in your minipool or multipool?

5             DR. HERMAN:  I can answer the first

6   question about the centrifugation.  It is a

7   limitation of the instrument that is widely

8   available, which is a table-top centrifuge.  It is

9   not quite an ultracentrifuge so it is more

10   convenient and easier to use.  That was a choice of

11   what equipment would be the most suitable for the

12   multiprep procedure for all three virus targets.

13             DR. HOLLINGER:  And the other question?

14             DR. HERMAN:  The choice of pool size, I

15   will defer to Mike.

16             DR. STRONG:  The pool size is one that was

17   selected for the HIV, HCV trials which were

18   concluded a couple of years ago and were licensed.

19   So, it was just to continue with the logistical way

20   in which laboratories operate to be consistent with

21   the other testing, and it is what the software does

22   for us.

92

1             DR. NELSON:  It is interesting that there

2   are many places, in the Orient particularly, where

3   core antibody testing is not done because of very

4   high rates of positivity.  So, that would cast some

5   limitation.  Even though the NAT assay would be

6   useful under those circumstances, it wouldn't

7   perhaps be as useful as continuing using the core

8   antibodies as is done in the U.S. and Europe.  You

9   had a question, Donna?

10             DR. DIMICHELE:  Thank you.  I was just

11   wondering if anybody can estimate, based on these

12   data or based on your data, what the residual

13   risk--in using your assay, if your assay was to be

14   implemented with core antibody testing, what do you

15   estimate the residual risk of transmission of HBV

16   would be through transfusion?

17             DR. FRANK:  That is a good question, and I

18   can say that we have shown the yield today but in

19   terms of discussion of residual risk, Dr. Kleinman,

20   would you like to comment?

21             DR. KLEINMAN:  Yes, Steve Kleinman.  I am

22   a consultant to Roche Molecular Systems on this

93

1   project.  The residual risk estimate is dependent

2   upon what you think the risk is now prior to

3   initiating HBV NAT.  We have estimated that risk

4   through the REDS study on two occasions, one that

5   was published in 1996 and then we kind of confirmed

6   the same numbers about six years later as about

7   1/60,000.  The risk estimate is based on the

8   incidence window period model and the data for that

9   model with HBV are not as strong as they are for

10   HIV and HCV.  So, other people have estimated the

11   residual risk now prior to NAT as being lower than

12   1/60,000, being about 1/200,000.

13             At any rate, if you accept the 1/60,000

14   risk estimate, and then you take a look--the

15   question is how much does this new assay decrease

16   the window period.  If you draw on data that was

17   published by Robin Biswas and colleagues on their

18   panel testing of decrease of the window period, you

19   get about a 10-day decrease of the window period in

20   that study, and in this study it looks like the

21   window period decreases by 17 to 20 days based on

22   the panel.  But, as Dr. Hollinger mentioned, that

94

1   is a maximum case because you don't have closely

2   spaced samples.

3             At any rate, if you take a 10-day

4   lessening of the window period and you apply it to

5   what we think the current window period is, about

6   59 days, and you work this through with the window

7   period and the incidence of HBV, you would probably

8   get that the residual risk is still 1/200,000 to

9   300,000 even with this test being implemented.

10   Actually, it may even be 1/100,000.  So, you are

11   removing some of the risk but because you are only

12   cutting the window period back 10 days out of a

13   potential 59 days, the likelihood is that most of

14   the risk still remains even with minipool-NAT and

15   that in order to decrease that risk further you

16   have to go to either smaller minipools or

17   individual donation NAT, which would buy you more

18   in terms of lowering the window period.  So, you

19   are making an incremental step in decreasing the

20   risk; you are not decreasing the majority of what

21   we think is the residual risk.

22             DR. NELSON:  These data are heavily

95

1   dependent on what the window period is, which may

2   vary in different populations.

3             DR. STRONG:  Yes, they are dependent on

4   the incidence and what you think the current window

5   period is, and that is something that there hasn't

6   really been good data on.  We have estimated the

7   current window period as 59 days based on some

8   older work from the TTVS study and then, through

9   REDS, we made a revised estimate of the window

10   period through mathematical modeling to be about 45

11   days now, the infectious window period.  So, these

12   are estimates.  We don't really know what the

13   current window period is so it is a little bit

14   difficult.  We don't know exactly how much we

15   shortened it so that is why I think you have to say

16   that we can't make the kinds of precise estimates

17   that we have made about HCV and HIV.

18             DR. DIMICHELE:  Can I just ask you a

19   follow-up question then?  What is the

20   estimate--let's say given that residual risk, what

21   is the estimate of the impact of vaccination,

22   widespread vaccinations now going on in the younger

96

1   population with respect to when they become blood

2   donors?  Does anybody know how the risk of

3   transmission will be decreased just by vaccination,

4   widespread vaccination alone?

5             DR. KLEINMAN:  I don't know the answer to

6   that.  I will answer kind of a related question,

7   which is what is the risk to recipients?  As more

8   recipients get vaccinated, then they should be

9   immune to a challenge, you would think

10   theoretically even if they were exposed.  However,

11   most of our recipients are in the older age group,

12   60 and above, and some have compromised immune

13   systems.  So, I don't think that vaccination is

14   going to protect much of the recipient population.

15   Your question is different, whether it makes donors

16   less likely to acquire new HBV infection and I

17   don't know that anybody has worked that through

18   mathematically.

19             DR. ALTER:  Harvey Alter, NIH.  The

20   mathematical models have been extremely useful but

21   there has always been a little discomfort with HBV

22   in that you just don't see the cases as frequently

97

1   as you would have predicted from the model.

2             It seems to me now that if you make the

3   assumption that the real risk would be coming from

4   HBV DNA positive donors, you could actually

5   determine what the actual risk is now by taking the

6   best or maybe all the different HBV DNA assays and

7   testing these samples as single donations.  If you

8   are picking up 1/350,000 now in a minipool format

9   you will pick up even more in an individual donor

10   format.  That theoretically would be the risk.  If

11   we can exclude those donors, that risk would be

12   removed.  It seems too simplistic to be true but I

13   would like to hear a counter argument.

14             DR. BUSCH:  Mike Busch.  I wanted to

15   comment on the vaccine issue.  In this study, I

16   think of the yield cases one of the yields that is

17   in the original clinical trial, clearly by history

18   vaccinated person, had an anamnestic anti-surface

19   response, low viremia in the setting of the index

20   sample.  I think one of the additional yield cases

21   also had anti-surface on index.  It is interesting

22   that the data from Taiwan is showing that when they

98

1   have applied NAT to donor screening--this is a

2   massively vaccinated population--they are seeing as

3   these young people, you know, enter adolescence and

4   in their 20s, they are seeing a fairly high rate of

5   the same kind of phenomenon.  These are essentially

6   vaccine breakthrough infections.  So the vaccine,

7   rather than being sterilizing--still people are

8   getting exposed and infected and, in a sense,

9   boosted from a wild exposure.  In their setting the

10   vast majority of their NAT yield are people who

11   have been vaccinated and have anti-surface and

12   never develop surface antigen, never develop

13   chronic infection.  So, the critical question of

14   whether these could transmit I think is still on

15   the table.

16             DR. STRONG:  Another point to follow-up on

17   what you were commenting on, on the value of

18   anti-core, I think this clinical trial also

19   demonstrated, as has been published, a small yield

20   of DNA positives from the anti-core positive group,

21   roughly in the same area of 5.5 percent or

22   thereabouts.  Also, I think it demonstrated a

99

1   difference in specificity with the anti-core assays

2   between the two licensed tests.  So, there is

3   clearly a large number of false-positive anti-core

4   donors that we have excluded on the basis of

5   false-positive reaction.  I think those of us who

6   have been involved in this trial, we have been

7   anxious to try to push the sensitivity of the assay

8   to allow us to reenter a lot of those donors

9   because it is a very large number.

10             DR. HOLLINGER:  I just want to follow-up

11   briefly on what Mike said.  I mean, this shouldn't

12   really surprise us very much, if you look at most

13   of the vaccine trials and follow patients for ten

14   years or so you will see anti-core appearing in

15   these patients.  I just mentioned that anti-core is

16   a marker of viral replication.  So, somewhere along

17   the line these patients will probably have DNA in

18   the bloodstream.  The issue is whether it is

19   infectious or not and that is, of course, I think a

20   very critical issue with it.

21             Then I just wanted to ask a question

22   either of Harvey or Steve about how many cases are

100

1   recorded now of transfusion-transmitted hepatitis

2   B?  I don't see very many.  I don't want to get

3   back to this argument we had about 20 or 30 years

4   ago when we said there was no hepatitis C or

5   non-A/non-B hepatitis when really there was a lot.

6             DR. KLEINMAN:  No, I think that is exactly

7   right, Blaine.  There are very few post-transfusion

8   hepatitis B cases reported.  As you know, it is a

9   question about whether those people become

10   symptomatic.  So, there may be asymptomatic cases.

11   Some may go on to be chronic carriers.  That is

12   certainly possible.  Secondarily, even if a case

13   occurs, is it recognized by the clinician?

14             It is certainly a correct statement that

15   we don't see the number of HBV clinical cases

16   reported as post-transfusion cases as the

17   mathematical models would predict.  So, then you

18   have two potential explanations.  One, the

19   mathematical models are wrong and these cases don't

20   occur or, two, the cases occur and they are not

21   recognized and reported.  I don't think we have a

22   way to sort out which one of those hypotheses is

101

1   correct.  Harvey's point is that we would have to

2   take a different approach to estimate risk, that

3   is, we would have to use the most sensitive DNA

4   assay available, hoping we could pick up one

5   copy/ml or even less.  By that, we could basically

6   make the equivalence that if we have a DNA-emic

7   specimen that equates to an infectious specimen,

8   and we could do that kind of study that was done 15

9   years ago or I guess 17 or 18 years ago for HIV,

10   but that would be an extremely laborious and

11   expensive study to do.

12             DR. NELSON:  I think we will take a break

13   at the moment and then come back for the open

14   public hearing.  Maybe 20 minutes.

15             [Brief recess]

16             DR. NELSON:  The meeting is now in session

17   again and we will start with an open public hearing

18   on the topic presented of HBV DNA testing.  Before

19   we start I will read this statement.

20             Both the FDA and the public believe in a

21   transparent process for information gathering and

22   decision making.  To ensure such transparency at

102

1   the open public hearing session of the advisory

2   committee meeting, FDA believes that it is

3   important to understand the context of an

4   individual's presentation.

5             For this reason, FDA encourages you, the

6   open public hearing speaker, at the beginning of

7   your written or oral statement to advise the

8   committee of any financial relationship that you

9   may have with the sponsor, its product and, if

10   known, its direct competitors.  For example, this

11   financial information may include the sponsor's

12   payment of your travel, lodging or other expenses

13   in connection with your attendance at the meeting.

14   Likewise, the FDA encourages you at the beginning

15   of your statement to advise the committee if you do

16   not have any such financial relationships.  If you

17   choose not to address this issue of financial

18   relationships at the beginning of your statement,

19   it will not preclude you from speaking.

20             With that, Mike Busch?

21                       Open Public Hearing

22             DR. BUSCH:  Thank you.

103

1             I wanted to just briefly touch on three

2   items.  One is just one slide sort of summarizing

3   the different context of some of the issues that

4   were raised in the questions as to the value of HBV

5   minipool-NAT both relative to the other NAT systems

6   and ID-NAT.  Then in most of my comments I want to

7   focus on the issue of can we drop either anti-core

8   surface antigen and, if not or if the data doesn't

9   support it at this point, what further studies

10   would be useful to address that.  I have Roche's

11   computer here--

12             DR. NELSON:  That sounds like a conflict

13   of interest!

14             [Slide]

15             DR. BUSCH:  The committee does have this.

16   This is a published analysis of cost-effectiveness

17   and I am not going to talk at all about sort of

18   cost-effectiveness but I just wanted to show one

19   slide that sort of emphasizes the relevant clinical

20   impact of these three viruses.  We tend to talk

21   numbers as if each of these viruses is equivalent

22   and the NAT yield is equivalent, but what this is

104

1   showing is the number of quality life years lost

2   per transmission.

3             What you can see is that for HIV it is

4   about 7 quality life years to the average 60 years

5   old transfusion recipient.  Transmission of HIV

6   would cost that person 7 quality life years; HCV,

7   0.6 and HBV, only 1.6.  This is because most

8   recipients are older people who naturally resolve

9   the acute infection of HBV, and even if chronic it

10   rarely evolves to significant morbidity/mortality

11   over the persistent life span of the average

12   recipient.  So, you can see that essentially HBV is

13   one-quarter as clinically important as HCV and 44

14   times less important than HIV.

15             The other point I wanted to make has to do

16   with this issue of going from current serologic

17   testing to minipool versus serologic testing to

18   ID-NAT, and how much of the yield that would be

19   accomplished by going to ID-NAT is detected by

20   minipool-NAT.  For HIV we had very sensitive

21   antibody tests in place so NAT by minipool closes

22   the window by about 5 days and ID closes it by

105

1   another 5.  So, we picked up about 50 percent of

2   what could be picked up but the risk is very low.

3             With HCV, because of the long plateau

4   phase, 95 percent of the window period is detected

5   by minipool-NAT.  With HBV, because of the slow

6   ramp-up, at least the model estimates and some of

7   the empiric data like in the Biswas paper would

8   suggest that minipool is only picking up about 23

9   percent or a quarter of what could be detected by

10   ID.

11             Next slide.  The issue of anti core and

12   could we get rid of anti core--there are three

13   studies that have been either fully published or in

14   abstract form, actually the Roche data, that have

15   looked at high numbers of donations that were

16   anti-core only reactive--so surface antigen

17   negative, anti-core positive--and subjected then

18   HBV NAT.  Each of these studies, as you can see,

19   has given almost identical rates both on a per

20   anti-core and, if you extrapolate, on a

21   per-transfused unit basis of detecting DNA in

22   donations that were surface antigen negative.  So,

106

1   about a quarter to a half percent in all these

2   studies.

3             What is interesting is that in each of

4   these studies there was quantitation done in one

5   context or another, and all of these viremic

6   anti-core reactive units, surface antigen negative

7   have very low viral loads.  In Roche's data, for

8   example, 92 percent based on previously presented

9   work of their viremic anti-core only were negative

10   by minipool but only detected by ID-NAT.  So, these

11   are very low viral loads that really would require

12   ID-NAT to detect.

13             [Slide]

14             This is just a summary and I am not going

15   to walk through this but with respect to low

16   viremic anti-core reactive donations or anti-core

17   only's there are a number of studies that have

18   documented transmissions, rare but there are

19   transmissions from these low viremic units.  Again,

20   I don't have time to go into these but we also know

21   that in liver transplantation from anti-core

22   reactive, surface antigen negative donors very

107

1   frequently transmits the virus.  So, these people

2   who have low viremia do harbor infectious virus

3   and, obviously, the anti-surface status is a

4   factor.  But I personally don't think there is any

5   prospect for dropping anti core.

6             [Slide]

7             Next slide.  This was alluded to earlier.

8   There was one study where purposeful transfusion of

9   these anti-core reactive, DNA positive, surface

10   antigen negative units into chimps was done.  This

11   was by Fred Prince and was published.  These did

12   not transmit.  Three patients who had remote

13   hepatitis B, persistent anti-core, negative

14   antigen, low-level DNA.  However, the caveat is

15   that all three of these donors had anti-surface and

16   the volume inoculated into chimps was quite small,

17   1 ml.  So, I think larger studies with larger

18   volumes may be indicated to potentially explore the

19   infectivity of these units.

20             [Slide]

21             Next slide.  In terms of dropping antigen,

22   which I think is the more likely and I believe

108

1   optimistically we will get there, I think clearly

2   for these viruses like HIV, HCV, HBV where there is

3   not only the acute phase but there are chronic

4   carriers and some of these chronic carriers have

5   low-level infectious viremia the optimal paradigm

6   really is the combination of a front-end sensitive

7   direct assay and a serologic test for antibody.

8   For HBV we have both antigen and NAT which could be

9   viewed as redundant, analogous to P24 antigen in

10   HIV NAT.  Clearly, on the front end we know, and

11   saw from Roche and other studies, that HBV NAT is

12   clearly more sensitive than antigen during the

13   acute phase.

14             Next slide.  But the problem is, as Blaine

15   summarized, HBV is really an unusual virus in that

16   it produces in chronic infections amazingly high

17   levels of circulating antigen in the absence of

18   infectious particles.  Therefore, this excess

19   antigen can occasionally be detected in the absence

20   of detectable DNA by minipool or even ID-NAT.  I

21   want to present a recent study on that.  Therefore,

22   to really be sure that we are not taking a step

109

1   back and dropping antigen when we have NAT, we need

2   to do large studies and I think Roche's data is

3   very impressive but whether it is enough is for the

4   committee to decide.  But these studies need to

5   look at these issues in a different context,

6   obviously, the front-end window period anti-core

7   negative as well as chronic carriers.  But also I

8   think they need to reflect populations with

9   different endemnicity and different routes of

10   transmission; populations that have been vaccinated

11   because we are, as I mentioned earlier, seeing

12   vaccine breakthrough infections, and how would they

13   play out in terms of antigen versus DNA; and also

14   genotypes and mutants.

15             Next slide.  Just to contrast, one of the

16   big differences in acute phase versus chronic is

17   the relationship between the DNA load and the

18   antigen reactivity.  This is one of a number of

19   studies.  This is a Japanese study.  During this

20   acute phase these were all anti-core negative

21   front-end infections.  They were actually all

22   detected, 181 units detected as negative by the

110

1   Japanese particle agglutination antigen test.

2   Actually, 105 of these were reactive by the PRISM

3   assay.  So, only a subset, about 40 percent of

4   these remained negative.  But what you are seeing

5   here is a very nice linear regression relationship

6   between the reactivity in the antigen test and DNA

7   load.  That is typical of the front-end acute

8   viremic phase.

9             Next slide.  I want to summarize now a

10   paper that is in press in collaboration with Mary

11   Kuhns and Steve Kleinman from the REDS group, where

12   we took 200 antigen positive, anti-core positive

13   donations in REDS.  We first subjected them to

14   quantitative HBV DNA using the Amplicor Roche

15   assay.  This assay has a sensitivity of 400 copies,

16   and 64 percent of the antigenic units had DNA above

17   400 copies.  Those samples that were negative, the

18   72, were taken through a more sensitive assay that

19   had a sensitivity of about 65 copies, which is

20   probably comparable to minipool type NAT, and 12

21   remained negative even at that level.  Those 12

22   samples were taken to a very high sensitivity,

111

1   essentially 1 copy/ml assay, high volume, and still

2   there were 6 that were negative.  So, we ended up

3   with still 3 percent negative even after taking it

4   through sequential increasing sensitivity assays.

5   So, these are the chronic carriers in whom DNA is

6   undetectable with at least one time point and one

7   high sensitivity assay.

8             Next slide.  This is showing the

9   relationship between antigen reactivity and viral

10   load.  Unlike that nice linear relationship, here

11   there it is essentially a scatter plot.  These are

12   the 3 that were negative by all NAT.  These are the

13   ones that were detected only by the high

14   sensitivity.  These are the Amplicor negative but

15   single input, not as sensitive, positive.  And,

16   these are the quantifiable.  You can see that

17   relationship is not observed in chronic carriers.

18             Next slide.  That is people who have

19   anti-core.  I think Roche did a nice job of sort of

20   saying that if we keep anti-core the issue is not

21   that problem with anti-core positives, antigen

22   positives having absence of DNA.  It is really are

112

1   there donations that are antigen positive that lack

2   anti-core that are truly infectious and not

3   detectable by NAT?

4             In the REDS program and in other studies

5   varying 1-5 percent of antigen positive donations

6   lack anti-core.  Now, when these are worked up most

7   of these are window phase seroconverters so they

8   have high-level DNA very consistent with acute

9   infection.  But in a proportion of DNA negative,

10   and we saw those cases in Roche, the issue is

11   studying enough of those cases to sort out are they

12   true HBV infections, that either there is a

13   mutation not detected by NAT or individuals who

14   have failed to form anti-core, chronic carriers who

15   have low-level DNA and have failed to form

16   anti-core.  Do they lack contamination, which I

17   think most of them are?  Are they persistent,

18   non-specific antigen reactivity, and Roche showed

19   us an example of that?  Or, are they possibly

20   recent vaccinees?

21             Next slide.  This is three studies, two

22   published papers and some unpublished data from

113

1   REDS that have looked for these people who are

2   antigen positive, anti-core negative and either

3   lack or have very low-level DNA.  I am not going to

4   go through these in detail but in Japan one study

5   identified three such donors who had no history of

6   vaccination and who remained low-level PCR-reactive

7   with only a very high sensitivity enhanced assay;

8   remained anti-core non-reactive, so never

9   converted; and, again, remained antigen positive

10   and very low-level PCR-reactivity for over a year.

11             The French group, Couroce's group reported

12   two donors from endemic countries as well, with 0.1

13   percent of their antigen-positive donors who had

14   this pattern of persistent antigen without

15   anti-core and very low-level HBV DNA, and no

16   history of vaccination or immunosuppression and no

17   evidence of mutations to explain the failure of

18   this person to form anti-core.

19             In REDS--this is a cross-sectional study

20   so we weren't able to do follow-up, but we had 20

21   antigen-positive donors who lacked anti-core.

22   Sixteen of these were positive for DNA and were

114

1   probably typical acute phase, but 4 were negative.

2   Three of these were probably contaminations.  They

3   had no other serologic markers, probably represent

4   surface antigen carryover.  But one had anti-e and

5   could represent an atypical carrier.

6             So, just a caution that there are some

7   studies, particularly from endemic settings where

8   there may be people who have antigen in the absence

9   of anti-core and low-level or absent DNA, and I

10   think we just need to study these further.

11             Next slide.  Just in conclusion, I do

12   think that we would see a small incremental yield

13   by minipool-NAT but the clinical impact of that I

14   think needs to be considered as well as the yield,

15   and we also need to view the context of what would

16   we get were we to get all the way to ID-NAT.

17             I am optimistic that we will eventually be

18   able, with ID-NAT or very small pools or high

19   sensitivity HBV DNA, to get rid of antigen,

20   however, I think this is going to require further

21   studies.  Particularly, we are capturing units

22   daily that could be worked up to accrue large

115

1   numbers with high volume of the plasma components

2   from these various sort of atypical patterns.

3   Particularly relevant is that I do think the

4   combination of NAT anti-core is the ultimate goal.

5   Particularly relevant are these samples that are

6   antigen positive, anti-core negative, and really

7   both studying these samples and particularly

8   enrolling and following the donors, as Roche did,

9   to determine whether these are false positives or

10   contamination.  Thank you.

11             DR. NELSON:  Thank you.  Comments?  Yes,

12   Harvey?

13             DR. KLEIN:  Mike, we heard from Dr. Alter

14   earlier that, with the problems in trying to

15   calculate what the residual risk is, it might be

16   reasonable as a start to do a study simply by

17   taking 400,000, 600,000 specimens and using a

18   single unit detection system along with the

19   serology.  I would just like to know what your

20   opinion on doing that kind of a study might be.

21             DR. BUSCH:  Yes, I think that is a good

22   idea.  I mean, we have the Biswas paper included

116

1   comparing assays on window period panels neat

2   versus minipool.  That is kind of these model

3   estimates that minipool picks up only a quarter of

4   what might be detected neat.  But, yes, I think

5   such a large-scale study either with neat--you

6   know, one of the issues with Roche's system is

7   their multiprep high extraction volume assay--we

8   saw the data that was comparing the multiprep pool

9   of 24 versus the standard prep neat, but had they

10   applied their high sensitivity multiprep neat,

11   which is what was done in the Biswas paper, they

12   would have seen a greater incremental closure than

13   was documented by ID-NAT.  So, doing a head-to-head

14   comparison study of a very small pool with proposed

15   minipool, I expect, would double or greater than

16   double the yield.

17             DR. KLEIN:  That might give us at least an

18   idea of what the risk is now.  It seems as if no

19   one is very confident about that.

20             DR. NELSON:  One other issue too I

21   think--I know you didn't talk about the

22   cost-effectiveness in any detail but your initial

117

1   slide assumes that it is a 60 year-old person that

2   is transfused but there are kids with thalassemia

3   and pediatric, depending on the life expectancy,

4   etc. of the patient, I think those are different

5   issues.  Yes?

6             DR. KUEHNERT:  On that slide you were

7   reviewing surface antigen positive, core antibody

8   negative donors in the last bullet of the REDS.

9   What is the denominator of those tested?

10             DR. BUSCH:  Unfortunately, I don't have

11   that number.

12             DR. KUEHNERT:  The other question I had is

13   you had 20 antigen positive and 16 DNA positive by

14   the standard PCR assay.  So, the PRISM assay was a

15   false negative.

16             DR. BUSCH:  These were people who were

17   antigen positive, anti-core negative and 16/20 had

18   ample levels of DNA and were probably acute phase.

19   They were picked up but at that point it was

20   Auzyme's.  So, they were picked up by  Auzyme and

21   those 16 would have been sort of, if you will,

22   run-of-the-mill acute hepatitis B.

118

1             DR. KUEHNERT:  So, 16/20 were DNA positive

2   originally.

3             DR. BUSCH:  Yes.  They were all antigen

4   positive and 16 were DNA positive.

5             DR. KUEHNERT:  And then 4 that were

6   negative.  Okay, thanks.

7             DR. ALLEN:  Mike, thanks for that very

8   nice overview.  Obviously, the committee and the

9   FDA, when they are considering how to place these

10   things, is looking at just one test at a time.

11   What is the implication of adding this into the

12   whole system that is already there in a blood bank

13   laboratory with or without discontinuing the

14   surface antigen test?

15             DR. BUSCH:  Well, it is certainly

16   possible.  Certainly, at sites that are doing Roche

17   today the system is optimal and they have actually

18   already done the extraction and they are just

19   taking a residual aliquot of the resuspended pellet

20   lysate and directing it to an automated system, you

21   know, to get the results.  So, it is a moderate

22   workload impact.  The other platform, the Gen-Probe

119

1   Chiron platform, you know, that test is I believe

2   in clinical to try that version assay.  Especially

3   if you are deriving it off the current pool, I mean

4   that sort of becomes the challenge.  You know, to

5   do it on individual donation samples while we are

6   still doing the licensed pooled assay for HIV/HCV,

7   that becomes, you know, a substantially greater

8   task.  To go to smaller pools--I mean, the licensed

9   assays for HIV/HCV are approved for up to a

10   particular pool size so it is possible to go to

11   smaller pool sizes.  You are still retaining the

12   licensed status of those other assays, which I

13   personally think is the more appropriate way to go

14   to get to pools of 4 or 6.

15             DR. ALLEN:  And in terms of integrating a

16   test into the laboratory software system for--

17             DR. BUSCH:  I think that is not a huge

18   problem.

19             DR. ALLEN:  And do you foresee that

20   availability of this test will enable algorithms

21   that will enable entry of donors that are currently

22   deferred who clearly are not infected in any way

120

1   with hepatitis but have some false-positive tests?

2             DR. BUSCH:  Right, there has been a lot of

3   discussion about anti-core reentry and, certainly,

4   the current expectation is that that would require

5   that the donors be negative on anti-core on a

6   follow-up period so they can't be the real

7   HBV-exposed population or the old false positives.

8   They not only need to be negative on current test

9   of record or improved anti-cores, but they also

10   need to be DNA negative with a single input assay,

11   potentially with a single input assay run on

12   multiple reps in order to make sure everybody is

13   confident that these people do not have occult

14   viremia.

15             DR. ALLEN:  Thank you.

16             DR. NELSON:  Thanks, Mike.  Donna?

17             DR. DIMICHELE:  I just want to actually

18   comment on your comment, and that was sort of the

19   basis of my previous comment.  You know, in the

20   pediatric populations that are getting chronically

21   transfused, I just want to say that all of them

22   have been vaccinated ever since this vaccine has

121

1   been licensed, and we pay a lot of attention to

2   documenting primary serological response so that,

3   you know, their risk is primarily of getting

4   reinfected bit certainly not chronically infected

5   unless they happen to be immunosuppresed because

6   they have already been HIV infected or HCV

7   co-infected.  But, for the most part, you know, in

8   the chronically transfused sickle cell thalassemic

9   and hemophilia populations, they have been

10   vaccinated and monitored for a very long time.

11             DR. NELSON:  Thank you.  Dr. Andrew

12   Heaton, from Chiron?

13             DR. HEATON:  Thank you for the opportunity

14   to present early Ultrio or Triplex test results as

15   part of our worldwide introduction.

16             As you may be aware, Chiron Gen-Probe

17   received CE approval in January of this year.  We

18   are beginning to see significant implementation on

19   a worldwide basis.  We anticipate, by the end of

20   the year, the availability to those countries

21   covered by that application to be able to perform

22   individual donor testing.  We also have reduced

122

1   pool size software under review at the moment in

2   order to allow our customers a wide range of

3   testing options in terms of sensitivity.

4             Next slide, please.  As I mentioned, CE

5   approval of our Ultrio test occurred in January,

6   and it allows the availability of HBV NAT testing

7   to European countries.  As I will show you in a

8   minute, the high prevalence countries are those who

9   have adopted this first.  There has been

10   significant early adoption and I will show you some

11   of the early yield cases that we have experienced.

12             Next slide, please.  As you might expect,

13   the countries which have adopted the Ultrio test,

14   dots marked in green, represent the southern

15   European countries who have a higher prevalence of

16   HBV.  The second group of countries where we have

17   conducted or are conducting in-country evaluations

18   include the more eastern countries in Europe who

19   are joining the European Union and, therefore, will

20   become the subject to blood safety directive; and

21   then a number of very high prevalence countries in

22   the East.  So, there is significant worldwide

123

1   introduction of HBV NAT testing and it is clearly

2   related to the prevalence in the countries

3   involved.

4             Next slide, please.  We have had 2 yield

5   cases.  The first yield case occurred in Spain, in

6   the Madrid Community Blood Center, where the sample

7   was tested in pools of 8 and they got their

8   positive after 11,000 units had been tested.  They

9   have now tested 13,000 so, if you are running the

10   clock, about 1/13,000 is the incidence rate.  The

11   donor was hepatitis B core antibody positive and

12   the HBV quantitation, using an in-house

13   quantitative method, was 88 IU/ml and 650 copies/ml

14   when the test was performed by NGI.

15             In Italy, individual donor testing

16   identified 1 positive yield case in 9,732.  The

17   donor was also hepatitis B core antibody positive.

18   In-house testing revealed 11 IU and it was tested

19   out at 215 copies by another reference method.

20             Next slide, please.  In terms of yield

21   then, or predictive yield, really the world falls

22   into three major categories.  The first of those is

124

1   the low incidence countries which use hepatitis B

2   core antibody testing, where the predicted yield

3   may be expected to be very low.  The second is

4   really the rest of the world excluding the United

5   States and France, which are the only two countries

6   that now do core antibody testing as routine, where

7   you would expect the yield to be associated with

8   the incidence of hepatitis B core antibody

9   positivity and the yield, which is about 0.5

10   percent DNA positive in that population or around

11   1/40,000.  Then, in the very high prevalence

12   countries in the East we have somewhere around 20

13   and 70 percent hepatitis B core antibody positive

14   where you would expect a much higher yield.

15             So, if you average out the first two yield

16   studies, we are running about 1/10,000, which is

17   slightly higher than we would have predicted.  It

18   is interesting in that respect that Canada, which

19   has just started HB core antibody testing on a

20   North American type donor population, noted 1.5

21   percent hepatitis B core antibody positive in a

22   donor population that would be expected to be very

125

1   similar to the U.S.

2             Next slide, please.  The current questions

3   then as we go through the implementation process,

4   there is a rapidly growing tissue bank sector in

5   the United States which is looking for nucleic acid

6   testing to assure the quality and safety of their

7   components.  An issue that badly needs attention

8   there is a definition of sample standards because

9   sometimes what the tissue banks call a blood sample

10   is really neither blood nor a sample, and it

11   certainly isn't subject to consistent standards.

12             In another area, the blood banks have

13   asked us to look carefully at hepatitis B core

14   antibody false-positive reentry and correct

15   classification of hepatitis B core antibody and

16   they would like to see HBV NAT testing used to

17   assist in preventing the false classification of

18   donors who are not truly positive for the hepatitis

19   B core antibody.

20             There is certainly a need for infectivity

21   studies of NAT negative, HBsAg positive donations.

22   We have received requests from a couple of centers

126

1   to work with them on reviewing these issues,

2   particularly those centers in the East.

3             Finally, we have made dHBV available under

4   IND and, indeed, 4-5 blood centers have completed

5   our clinical trial and are now using dHBV on a

6   routine basis.

7             The worldwide implementation key issues

8   then are higher than expected yield that will

9   nearly always be core antibody positive.  We now

10   need to look at the rate of DNA positivity in a

11   B-core positive population.  In the U.S. that is

12   only about 0.5 percent but in Taiwan it is reported

13   to be as high as 10 percent.

14             We need to look at the implications of

15   vaccination.  We have already seen one donor who

16   silently seroconverted from anti-HBs to anti-HBc

17   without having a period of HBsAg positivity.  So,

18   the implications of bringing this test to make it

19   available on a highly vaccinated population will

20   clearly be significant.

21             Lastly, in the East we are facing the

22   issue of wide-scale vaccination in a highly endemic

127

1   population, a population which is now loaded then

2   with vaccine-escape mutants.  In Taiwan they report

3   about 1/1,000 vaccinees seroconverts to HB-core

4   antibody per year.  So, there is a significant

5   ongoing breakthrough in a community situation.

6             Finally, as we complete our clinical trial

7   we will be looking to carefully evaluate the

8   relationship between test pool size and the desired

9   sensitivity.  Thank you.

10             DR. NELSON:  Thank you, Dr. Heaton.

11   Questions?  Comments?  Yes, Susan?

12             DR. LEITMAN:  I have a question on slide

13   5.  We have seen that you can't measure residual

14   risk because it is so low that you don't see it

15   clinically and there is not a large enough

16   prospective study to detect that.  So, it is all

17   based on mathematical models.  We heard of

18   1/60,000, by Dr. Kleinman on behalf of REDS.  What

19   is published I think on behalf of American Red

20   Cross is 1/100,000 or 1/120,000, by Dr. Dodd.  And,

21   you have 1/300,000.

22             DR. HEATON:  Yes, the 1/300,000 was based

128

1   on the result reports the HBV DNA testing in the

2   U.S.  The bulk of numbers are really based on the

3   level of hepatitis B core antibody and the

4   frequency with which hepatitis B core antibody

5   only's are HBV DNA positive.  The point I was

6   making here is that we are experiencing a much

7   higher rate in our highly endemic European

8   countries in any of these models that are

9   predictive.

10             DR. SCHREIBER:  On that same slide those

11   numbers that you showed for the endemic countries,

12   I think those are prevalence rates with the new

13   cases incidence rates and not residual risks.  The

14   residual risks would actually be at least a factor

15   of 10 times lower.  So, you are talking about

16   something on the order of 1/110,000 and then with

17   the adjustments that were made in the other you are

18   roughly talking about multiplying those by another

19   2.5.  So, you are talking about 1/300,000 or so for

20   Spain, if my rough math is right, so that Spain did

21   not come out a hell of a lot different, I don't

22   believe, than the United States.

129

1             DR. HEATON:  I accept that, and in

2   discussions with experts in Taiwan, one of the

3   points they made there is that so many of their

4   recipients have, in fact, been previously infected

5   with HBV that it is extremely difficult to model

6   out the risk because you have such a high basic

7   resistance level.

8             DR. NELSON:  Next is Dr. Sherrol

9   McDonough, from Gen-Probe.

10             DR. MCDONOUGH:  Good morning.  I will

11   continue the discussion of detection of hepatitis B

12   DNA using NAT, and I will be talking more about the

13   Procleix assay.

14             [Slide]

15             As you can see, it was developed for core

16   detection of HIV-I, HCV and hepatitis B in pools of

17   up to 16 down to individual donor testing.  In

18   addition, there are 3 discriminatory assays to

19   allow Ultrio reactors to be discriminated.

20             This assay is based on the same

21   technology, same equipment as assay procedures that

22   are used in the licensed Procleix HIV/HCV assay.

130

1   There are a large number of common reagents so we

2   made as few changes as we could.  And, the assay

3   reagents and the formulations that are used in the

4   licensed semi-automated Procleix format are the

5   same as those on the fully automated Procleix

6   Tigris system.

7             Next slide.  So, I am talking about the

8   assay type shown on the left, the multiplex Ultrio,

9   as well as discriminatory HBV, HIV and HCV, and I

10   am talking about the semi-automated Procleix system

11   as the licensed system and the fully automated

12   Procleix Tigris system.  The two rows at the top

13   are what I will be talking about today, focusing on

14   HBV detection.

15             Next slide.  To mention some of the

16   features of the Procleix Tigris system, after setup

17   this system is designed to process up to 500

18   samples in a 10-hour period or 1,000 samples in a

19   14-hour period.  For the multiplex Ultrio assay

20   that represents 1,500 results in 10 hours.  That is

21   500 for each of the analytes, or in 10 hours that

22   would be 3,000 results.  Then, of course, the

131

1   specimens are tested in pools of up to 16.  That

2   would be 24,000 results in 10 hours or 48,000 in 14

3   hours.

4             Next slide, please.  The validation that

5   is ongoing for this assay includes the following:

6   the clinical evaluation is taking place at 5

7   clinical sites, with 3 reagent lots with both the

8   Ultrio and all 3 discriminatory assays and on the 2

9   different instrument platforms.  We have included

10   specificity studies in pools of both 16 as well as

11   IDT, and several different populations are being

12   studied for clinical sensitivity.  Of course, there

13   is a reproducibility study.  The items that are

14   highlighted here, specificity, seroconversion

15   panels or reproducibility are what I will describe

16   today.

17             Next slide, please.  In-house analytical

18   testing includes the following.  You can see

19   multiple operators, reagent lots and instrument

20   platforms.  They include analytical sensitivity,

21   genetic variants, the donor, donation factors,

22   interfering substances, etc., specimen collection

132

1   handling and the validation of use of the assay on

2   cadaveric specimens.  Today I will only be

3   describing the analytical sensitivity of the assay.

4             Next slide, please.  Analytical

5   sensitivity is demonstrated in this slide.  This

6   was performed by taking known concentrations of HBV

7   and preparing panels at various international unit

8   per/ml concentrations.  Each group of bars

9   represents a different concentration.  The first is

10   45 IU/ml and then we are going down 3-fold, 15 IU,

11   5 IU, etc.  The different bars represent either the

12   Ultrio assay on the semi-automated system or the

13   Procleix Tigris system, as well as the

14   discriminatory HBV assay on both semi and fully

15   automated systems.

16             Before you spend too much time scrunching

17   your eyes, I will just say there is no statistical

18   difference between either assay or either

19   instrument platform for any of the copy levels.

20             Next slide, please.  Of course, another

21   way of looking at clinical sensitivity is to look

22   at seroconversion panels.  We have had a lot of

133

1   discussion about seroconversion panels already

2   today.  This is a set of panels that we have

3   analyzed.  What I am showing here are the days

4   earlier that you can detect hepatitis B DNA

5   compared to surface antigen.  The surface antigen

6   test was the PRISM using the European cut-off.  The

7   results are shown for the Ultrio multiplex assay at

8   dilutions of 1:16 as well as neat, and on both

9   instrument platforms.  The final set are

10   discriminatory HBV results on both platforms.

11             So, if you focus on the highlighted

12   numbers at the bottom, the median days earlier that

13   we can detect HBV at 1:16 dilution is 6 or 7 days;

14   neat, 17 days on either instrument platform; and

15   the discriminatory assay shows 15 or 14 days

16   earlier.

17             Next slide.  I will focus on the word

18   preliminary.  This is not final but this is a

19   snapshot in time of where we are with specificity

20   testing.  The first row shows the results with the

21   Ultrio multiplex assay on the semi-automated system

22   testing pools.  Specificity is 99.5 percent and

134

1   that overlaps with the specificity observed with

2   the licensed HIV/HCV assay.

3             The next slide shows the results from the

4   Ultrio multiplex test again on the semi-automated

5   system, now testing IDT, and we see the same

6   specificity of 99.5 percent.  Finally, the results

7   with the Ultrio multiplex assay on the Tigris

8   system, testing individual donor samples, is 99.8

9   percent.

10             Let me address the hatchmarks on the first

11   two rows.  The first hatchmark reminds me to tell

12   you that there was a potential HBV yield case in

13   that population.  The second row with the 2

14   hatchmarks--there were another 2 potential HBV

15   yield samples from that population.  All 3 of these

16   samples were Ultrio reactive, discriminatory HBV

17   reactive, as well as alternative NAT reactive from

18   a different specimen.

19             Next slide, please.  These are the

20   preliminary discriminatory HBV specificity results.

21   Testing IDT, the specificity is 99.9 percent.

22   This, again, includes one of the yield cases

135

1   identified in the previous slide.  We are calling

2   that potentially yield case number 3.

3             Next slide.  The reproducibility study has

4   been done on both instrument platforms.  It

5   includes Ultrio, discriminatory assays, 3 clinical

6   sites, 3 reagent lots, 2 operators per site, and

7   both negative and low panel positive samples.

8             Next slide.  The constituents of each

9   panel member are shown at the left.  We have both

10   HIV, HCV and HBV positive panel members, as well as

11   negatives and co-infected samples.  The agreement

12   in this study was 100 percent for every panel

13   member, and the total variability is shown on the

14   right column at 15.2 percent or lower.

15             Next slide, please.  In conclusion, the

16   Ultrio and discriminatory HBV assays show similar

17   performance on the semi-automated Procleix and

18   fully automated Procleix Tigris systems.  We

19   believe that the automated system will allow IDT

20   testing to be performed much more easily in the

21   laboratory.

22             HBV was detected in seroconversion panels

136

1   on an average 14-17 days and 6-7 days earlier than

2   PRISM HBsAg in neat and 1:16 diluted samples

3   respectively.  There are 3 potential HBV yield

4   samples to date that were identified in the

5   specificity study.  Also, we have an additional 6

6   potential yield cases from our high risk study.

7   Again, each of these samples was Ultrio reactive,

8   discriminatory HBV reactive and alt-NAT reactive on

9   a different sample.  Thank you.

10             DR. NELSON:  Thank you, Dr. McDonough.

11   Questions or comments?

12             [No response]

13             Thank you.  Next is Dr. Richard Smith,

14   from National Genetics Institute.

15             DR. SMITH:  Good morning.  Richard Smith,

16   from National Genetics Institute.  I believe that

17   NGI has a financial relationship with probably

18   every company in this room so I will just do a

19   blanket statement.

20             [Slide]

21             This morning I will tell you about HBV NAT

22   IND for whole blood donations, just started at NGI

137

1   for the American Red Cross, in June of this year.

2   In this study, only those donations that have been

3   found to be repeat reactive for the anti-HBV core

4   antigen antibodies are tested for HBV DNA.

5             Next slide.  There are 3 reasons listed in

6   our IND for performing this testing.  First, the

7   results are being used as a donor counseling tool

8   for anti-core repeat reactive donors.

9             Next.  Second, we plan to analyze the

10   results of this study to help design another IND

11   for establishing a donor reentry algorithm and for

12   potentially false-positive anti-core reactive

13   donors.  The new study would also involve a second

14   licensed anti-core assay, hopefully, with improved

15   sensitivity.

16             Next.  Finally, to limit the viral load in

17   plasma pools destined for further manufacture.

18   Currently, anti-core repeat reactive donations are

19   included in these pools and we would expect some

20   number of these to be positive for HBV DNA.

21             Next slide.  The NAT assay used for this

22   study is the same assay being used for the IND

138

1   established to screen source plasma donors for HBV

2   DNA, and is performed on the same platform as the

3   FDA-approved HCV and HIV PCR assays that have been

4   used for source plasma donor screening for the past

5   several years.  DNA extraction for HBV is performed

6   with an initial sample volume of 0.1-2 ml,

7   depending on the desired level of sensitivity and

8   the volume of the sample that is available.

9             For HBV, we use 4 distinct HBV specific

10   primer pairs in separate amplification reactions,

11   and we perform duplicate reactions when increased

12   sensitivity is required for a total of up to 8

13   separate amplification reactions per assay.  If

14   just one of those reactions is positive the assay

15   is considered positive.  Each reaction also

16   includes an internal control specific primer so

17   that the success of each reaction can be

18   independently determined.  Detection of specific

19   amplification products are performed by southern

20   blot analytical on every reaction.

21             Next slide.  This slide shows what the raw

22   data of one of NGI's HBV NAT assays looks like.

139

1   The underlying image is a scan of an HBV specific

2   southern blot and this membrane represents analysis

3   of just 1 of the 4 HBV specific primer pairs used

4   for each HBV assay.

5             After scanning, the same membranes are

6   rehybridized with an internal control specific

7   probe and bands appear in all positions,

8   representing successfully extracted and amplified

9   material.  A reaction such as the one depicted

10   here, in position 14, would be considered invalid

11   and negative for the target but invalid so it would

12   have to be repeated.  Membranes also include

13   diluted amplified material, shown here in the last

14   position, to ensure adequate southern blot

15   detection sensitivity.

16             Next slide, please.  The testing and

17   algorithm being used for this particular study

18   first involves combining up to 16 individual

19   samples into one pool, and then performing an assay

20   with a 95 percent detection cut-off of 0.9 IU/ml,

21   which translates to a 85 percent detection cut-off

22   of 14.4 IU/ml on an individual sample basis.

140

1             Next slide, please.  When a pool of 16 is

2   found to be positive for HBV each individual sample

3   is tested with an assay with a 95 percent detection

4   cut-off of 1.6 IU/ml or

5   if the sample volume is limited we can perform an

6   assay with only 100 mcL but involving twice the

7   number of amplification reactions and, therefore,

8   producing the 95 percent detection cut-off of 9

9   IU/ml.  With an anti-core reaction rate of

10   approximately 0.6 percent, we are expecting to test

11   approximately 42,000 donations per year for the Red

12   Cross for HBV DNA.  The next slide shows the

13   results we have obtained since starting this

14   project, just in June of this year.

15             Next slide, please.  To date, we have only

16   reported results for 1,255 anti-core reactive

17   donations; 900 have been found to be not

18   implicated, which means that the pool they were in

19   was negative; 302 were tested individually and

20   found to be negative for HBV DNA, 50 of them found

21   positive for HBV DNA and 3 had insufficient volume

22   for testing.

141

1             Of the HBV DNA positive donations, we have

2   HBV surface antigen data for 37.  Data for the

3   other 13 is forthcoming but was unavailable to me

4   at the time I was preparing this presentation.

5   And, 22/37 are antigen positive and neutralizing,

6   and 15 are antigen negative.  So, so far, about 40

7   percent of the HBV DNA positive donations are

8   antigen negative, or about 1.6 of all the anti-core

9   reactive donations screened.

10             Next slide, please.  We have gathered some

11   additional data on the DNA positive donations.  The

12   average antigen reactive HBV DNA positive donation

13   has an antigen neutralization value of 98.8

14   percent, with a range of 89.3 to 100 percent.  The

15   average viral load in these samples is 70 million

16   international units/ml, with some having very low

17   viral loads, in the range of 10

2 and some as high

18   as 10                                       8.

19             Next slide.  In contrast, the average

20   viral load in the antigen-negative DNA positive

21   donations is just 70 IU/ml, one million times lower

22   than the antigen positives, with a range of below

142

1   the sensitivity of the quantitative assay up to

2   just 150 IU/ml.

3             Next slide, please.  Prior to initiating

4   the current study we performed a pilot study on

5   anti-core reactive donations.  This study was

6   blinded, with no knowledge of whether the donors

7   tested anti-core reactive on a previous donation.

8   So, the population most likely included a mixture

9   of donors that would be ineligible for inclusion in

10   a donor reentry study.

11             In this study only antigen-negative

12   donations were tested for HBV DNA.  Of the 3,000

13   anti-core reaction, surface antigen negative

14   donations, 19 were found positive for HBV DNA using

15   an individual PCR test, with a 95 percent detection

16   cut-off of 9 IU/ml.  Viral loads for the 19

17   positive donations were determined and the range

18   observed for these was also from below cut-off up

19   to 150 IU/ml, which corresponds to about 500

20   copies/ml.  The "greater than" sign on the bottom

21   actually should be a "less than" sign.  Those are

22   all less than the cut-off of the quantitative

143

1   assay.  The next slide compares these data with

2   findings from the REDS study, published in 2003 by

3   Steve Kleinman and others in The Journal of

4   Transfusion.

5             Next slide, please.  I was going to thank

6   Sue Stramer for this slide but maybe I should thank

7   Mike Busch, or maybe Mike should thank Sue--I am

8   not sure.  Data from the REDS study, conducted over

9   a much longer period, also shows a very low rate of

10   the HBV DNA positivity in anti-core reactive

11   donations.  DNA positive donations from the REDS

12   study also had very low viral loads.  These are the

13   same results that you have seen presented in

14   multiple talks earlier today.

15             As part of the study currently being

16   carried out, we will enroll antigen negative, HBV

17   DNA positive donors for follow-up to determine the

18   predictive value of the test.  We will also follow

19   2 time anti-core reactive donors that test HBV DNA

20   negative upon initial screening.  That is it, just

21   a short but sweet presentation.

22             DR. NELSON:  Thank you.  Comments?

144

1             DR. SCHREIBER:  One question, one of your

2   objectives seemed to indicate that you are

3   interested in trying to reduce the viral load in

4   plasma pool.  I think you need to be clear that you

5   are talking about recovered plasma since I believe

6   that all source plasma collected is already

7   subjected to HBV testing, NAT testing, in the

8   United States, probably worldwide.

9             DR. SMITH:  Yes, you are right.  That is

10   absolutely correct.  Really this was to help the

11   plasma fractionaters justify the use of recovered

12   plasma that wasn't being HBV NAT tested.  Likewise,

13   the source plasma units that they use are not being

14   anti-core tested.  So.

15             DR. NELSON:  Thank you.  Harvey Alter?

16             DR. ALTER:  Thank you.  Could I have my

17   first slide, please--oh, I don't have any slides!

18             [Laughter]

19             DR. NELSON:  How can you talk without

20   slides?

21             DR. ALTER:  I am substituting for Steve

22   Kleinman because he has a relationship with Roche

145

1   and I realized today that I am on the scientific

2   advisory board for Roche, that I have to admit to

3   although, after reading this statement I suspect

4   that will be very short-lived!

5             [Laughter]

6             Lastly, I have this horrible bronchitis

7   and I may not make it through this so I am going to

8   give my quick overview and then read the statement.

9   Also, I am representing both the AABB and the

10   American Blood Centers.

11             The three issues that we saw were, one,

12   the company provides sufficient evidence for

13   licensure.  Secondly, if the test is licensed,

14   should it be mandated or optional?  Thirdly, can

15   the test replace surface antigen in its current

16   minipool format?  So, in case I break down or you

17   fall asleep, our position is yes, no and no.

18             I am going to skip the first paragraph,

19   which tells you how great AABB is, and then go on.

20   HBV remains the most common clinically important

21   viral infection recognized from transfusion since

22   the control of HIV and HCV infections through

146

1   improved donor selection and serologic and NAT

2   screening.  The data presented by Roche Molecular

3   Systems from its IND study of HBV NAT in minipools

4   of 24 samples are an important contribution to the

5   ongoing improvement of donor testing.

6             AABB sees three issues of primary

7   importance to the blood community to be addressed

8   by BPAC and the FDA.  First, is the Roche HBV assay

9   approvable as a donor screening test?  Second, if

10   approvable, shall its implementation be required in

11   blood collection facilities?  A third question is

12   whether a claim for HBV NAT in minipools to replace

13   HBsAg testing should be granted.

14             Regarding the first question, the data

15   that were available for review by the AABB's

16   Transfusion-Transmitted Diseases Committee indicate

17   that the Roche minipool HBV NAT assay appears to

18   perform adequately in terms of analytical

19   sensitivity and specificity, and generates

20   incremental yield of NAT positive specimens over

21   current serologic tests.  This suggests that the

22   assay may be approvable in the currently proposed

147

1   minipool NAT context, but its efficacy should be

2   greater if it were applied to individual donations

3   or significantly smaller minipools.

4             The second question is more difficult to

5   answer.  The minipool-based assay under

6   consideration appears to yield between

7   1/250,000-300,000 positive donations that are

8   negative on currently licensed tests for HBsAg and

9   anti-HBc.  This rate is similar to the yield rate

10   for HCV minipool NAT, and substantially higher than

11   that for HIV NAT.  It is comparable to or slightly

12   higher than predicted by the Biswas et al. study in

13   a comparative study of NAT and serologic assays.

14             As suggested from the data on the

15   evolution of markers of HBV infection, these

16   donations tend to contain low copy numbers of HBV

17   genome, and incomplete data suggests that some

18   HBsAg assays, either available or under development

19   for evaluation by FDA, maybe able to interdict some

20   of these yield donations.  These include HBsAg

21   tests from Abbott, Ortho Clinical Diagnostics and

22   Genetic Systems.  It is critical for the accurate

148

1   analysis of the true impact of HBV minipool NAT

2   that samples from these current yield cases, and

3   those identified in the future by HBV NAT assays,

4   be tested not only by the currently licensed

5   serologic tests, but also by the developmental

6   tests that are likely to be licensed in the future.

7   Studies of the infectivity of yield cases are also

8   desirable, and particularly or units that

9   concurrent HBV DNA and anti-HBs in the absence of

10   detectable HBsAg and anti-HBc, as seen on two of

11   the yield cases in the Roche trial.

12             Thus, despite measurable yield,

13   introduction of HBV minipool NAT will offer only a

14   minuscule increment in transfusion safety compared

15   to currently required tests for HBsAg and anti-HB

16   core.  The result of this low incremental yield,

17   coupled with low rates of chronic infection and

18   clinical disease after HBV transmission, renders

19   the marginal cost-effectiveness of HBV NAT in

20   minipools very poor.  There is a reference to

21   Jackson's paper in Transfusion.  This

22   cost-effectiveness will decline further into the

149

1   future as a larger and larger proportion of the

2   population has vaccine-induced immunity to HBV

3   infection.

4             Regarding the third question, current data

5   are not robust enough to support elimination of

6   either serologic marker.  It is possible that HBV

7   NAT will eventually allow discontinuation of HBsAg

8   screening, but this will require a larger data set

9   including parallel testing by HBV DNA, likely on

10   individual donations rather than minipools, anti-HB

11   core, and HBsAg, using maximally sensitive antigen

12   assays.

13             In summary, minipool HBV NAT is an

14   expensive new screening assay that offers little

15   clinical benefit and that will not be offset by

16   discontinuation of any current testing.  More

17   sensitive HBsAg tests are available now and more

18   will become available in the foreseeable future.

19   More specific anti-HB core tests will also become

20   available.  Based on these considerations, AABB

21   does not support--and that would be American Blood

22   Centers as well--a requirement for the use of NAT

150

1   in minipools for blood donor screening at this

2   time.  Rather, if HBV DNA NAT is licensed, its use

3   should be optional.  The requirement for HBV NAT

4   testing should be reconsidered when technology

5   allows for individual unit testing.  Thank you.

6             DR. NELSON:  Questions for Harvey?  Yes?

7             DR. STRONG:  I am curious about your

8   clinical opinion since there are yield cases that

9   will be missed if we don't do that testing, and you

10   are recommending that we don't have to do NAT

11   testing.  Do you see that as a significant clinical

12   risk?

13             DR. ALTER:  Well, we really struggled with

14   this.  You have so many variables that you have to

15   measure.  First of all, we think the test is

16   licensable.  Secondly, if licensable, it will be

17   easy for places that have Roche testing in place

18   but it is going to be a major, major implementation

19   for people who are using other systems.  Is the

20   yield sufficient to warrant that, knowing that

21   within a reasonable short time there will be

22   alternate NAT assay and there will be improved

151

1   surface antigen assays?  I guess an interim measure

2   would be that the FDA could set a time interval in

3   which to implement this test that was sufficiently

4   long that non-Roche establishments could possibly

5   catch up.

6             But when you get to the issue of if you

7   pick up one or two cases a year, does that make it

8   mandatory to do this test, that is a very difficult

9   thing.  I think this decision is similar to that

10   for third generation anti-HCV assays versus second

11   generation assays where the yield was considered to

12   be sufficiently small that blood banks were allowed

13   to use either second or third generation testing,

14   even though it was known that third generation

15   testing was slightly more sensitive.  It is a

16   complex issue.  Would I want to go on a stand?  I

17   don't know.

18             DR. STRONG:  Just a follow-up, it sounds

19   like part of that is logistical, which I certainly

20   understand as a blood tester.  But I am still

21   interested in whether or not you think there is a

22   significant clinical risk for those that are

152

1   missed.

2             DR. ALTER:  There is some clinical risk.

3   You know, if this was HIV I would be talking

4   differently but HBV generally is a recoverable

5   disease in 95 percent of the people and the other 5

6   percent has this long duration of usually

7   non-progressive disease.  So, in the current

8   setting it is a very, very small margin of

9   significant clinical disease.  Nonetheless, we

10   don't want to transmit disease so that is a tough

11   issue.  You know, my feeling is we need to get to

12   individual testing as soon as possible and then we

13   should we should be doing HBV DNA testing on NAT

14   when the yield will be higher.

15             DR. EPSTEIN:  Harvey, AABB has suggested

16   that it feels, based on its own review of data made

17   available, approval would be reasonable but the

18   test should remain optional.  What would be the

19   practical consequence if FDA acted to approve such

20   a test and was silent on a recommendation?  Would

21   it not be a chaotic situation in the marketplace

22   where some entities might implement it?

153

1             DR. ALTER:  Well, yes.  I mean, FDA has

2   always tried to license two tests at once sort of

3   to get around this issue.  It will create some

4   chaos if it is licensed and it is mandated

5   immediately.  We feel that Roche has proven its

6   case that it is a good assay and it has yields.  I

7   can't say it shouldn't be licensed.  They got there

8   first and they shouldn't be punished for that.  But

9   for those people who are not using the Roche assay

10   it will create chaos if it is mandated immediately.

11             Then the other issue is the PRISM issue

12   and when will that be licensed, and will that be

13   implemented, and will that decrease the yield even

14   further and make the need for HBV DNA less in a

15   minipool format?  So, there are a lot of things

16   riding here, and when do you go from minipool to

17   individual testing or to mini-minipool?  So, there

18   are really three things going on simultaneously and

19   I am glad I am not you.

20             [Laughter]

21             But I think one possible scenario would be

22   to license, to leave it optional for the time

154

1   being, to re-address this at some time interval.

2   The second option is to license it and make it

3   mandatory but give sufficient lag time that it will

4   be reasonable that the alternate assay will be

5   available also.

6             DR. NELSON:  Thank you.  the open public

7   hearing is closed, I am told.  Could you restate

8   the questions for the committee?

9             DR. HOLLINGER:  Just a question, is there

10   not a statement from Red Cross?

11             DR. NELSON:  Not on the list here.

12             DR. PAGE:  Red Cross has chosen not to

13   make a statement on this issue.

14             DR. HOLLINGER:  Do you have a reason why

15   not?

16             [Laughter]

17             Committee Discussion and Recommendations

18             DR. KAPLAN:  I will try to wrap up this

19   session and then given the FDA perspective and then

20   go through the questions.  What I will try to do is

21   to just put together some of the elements that Dr.

22   Blaine Hollinger gave us in his great talk, with

155

1   some of the points from the clinical trial that Dr.

2   Herman and Dr. Frank, from Roche, explained to us,

3   and also bring up some of the points that Dr. Busch

4   and Dr. Alter brought to this equation.  Basically,

5   most of the things I am going to say you have

6   already heard so I will try to wrap it up.

7             [Slide]

8             The Roche clinical trial for HBV NAT, 059,

9   was done in about 600,000 volunteer whole blood

10   donations.  The index donations were tested for

11   surface anti-core with their minipool NAT test.

12   These are the concordant results.  Basically, all

13   three markers are positive.  However, as you might

14   expect, there are many discordant results here

15   which are important for analysis of this trial and

16   licensing.  Basically, you can have the possibility

17   that the surface antigen is positive and the other

18   two markers are negative.  So, what this trial did

19   is they performed alternative NAT, quantitative NAT

20   and then they followed up the donors.

21             Another possible discordant result is that

22   the surface is positive and DNA is positive.  What

156

1   they did is an alternative NAT and quantitative NAT

2   and then they followed up the donors.

3             The third possibility is that the surface

4   and anti-core are positive, and what they did was

5   an alternative NAT and quantitative NAT.

6             Another possible discordant is that only

7   the core is positive and all the others are

8   negative.  So, what they did is the alternative

9   NAT, individual donation NAT, quantitative NAT

10   anti-surface, IgM anti-core and then they followed

11   up.

12             The other possibility is that the core was

13   positive and DNA was positive and then they

14   performed alternative NAT, quantitative NAT and

15   they followed up.

16             The last discordant possibility is that

17   the DNA was positive.  So, what they did was

18   alternative NAT, quantitative NAT and then they

19   followed up the donors.

20             [Slide]

21             So, I will give you an idea of the

22   numbers.  About 600,000.  As expected, most of them

157

1   were negative for the 3 markers.  What they found

2   in trial 059 is that 84 donations were positive for

3   the 3 markers.  They found 4 positive for the

4   surface; 3 positive for surface and DNA; 16

5   positive for surface and anti-core; 2,988 positive

6   for core; 1 for core and anti-core DNA; and 23 for

7   DNA.  I will focus o the last one.  This is pretty

8   important because it gives you the yield of the

9   whole trial that was discussed several times here.

10             From these 23, what they found is that 21

11   samples were false-positive samples and only 2 were

12   window period samples.  So, the whole yield would

13   be 2 in about 500,000 to 600,000 donations.  That

14   will imply about 3-4 per million.  If you have 15

15   million donations a year you have around 60 units

16   that can be interdicted with this test.

17             Several considerations were raised here

18   regarding which of these donations could progress

19   to severe clinical outcome after transfusion, and

20   that is a consideration that we want the committee

21   to address.

22             The other group that I would like to point

158

1   out to you is related to the safety of the

2   implementation of the company claim regarding

3   alternative use to surface.  I would mainly like

4   you to focus on these 4 cases that were only

5   surface positive.  Upon follow-up, 2 out of the 4,

6   as the company mentioned, were negative in

7   follow-up in the trial.  Basically, 1 was a

8   false-positive surface antigen and it was

9   vaccinated.

10             Two were not followed up in the initial

11   clinical trial, however, the company just showed

12   you some of the data and basically both cases did

13   not seroconvert.  So, for these 2 cases that were

14   not followed up, the alternative and ID-NAT or the

15   index samples were negative and after follow-up,

16   after the trial was finished, they did not

17   seroconvert.  So, I think we have a good take here

18   that these 4 were false positive for infection.

19   From this, 3 that had both markers, only 1 was

20   followed up and was positive, and the others were

21   not followed up.

22             So, regarding the sensitivity in general

159

1   of this test, there were 107 donations that were

2   positive for surface; 100 donations were surface

3   anti-core positive; and 7 of them were anti-core

4   negative.  As a note, Roche, indeed, accounted for

5   the 4 surface only donations in support of the

6   replacement claim.  However, these numbers are

7   small and several points were raised by Dr. Busch

8   and Dr. Alter, and we are asking the committee to

9   consider whether with this small number of

10   discordants this data is enough to support the

11   Roche claim to use it as an alternative of the

12   surface.

13             [Slide]

14             I would like to talk a little about the

15   surface and anti-core positives, and in this case

16   they found 16.  This pertains to the sensitivity of

17   the assay.  So, when they did ID-NAT they found 10

18   positives and 5 negatives using their own ID test.

19   Using the alternative NAT, a test by NGI that was

20   described by Dr. Smith, none were positive and 4

21   were negative.  These are not exactly the same.

22   There is no exact agreement here.

160

1             When they quantitated the 16, 3 of them

2   were 100 copies/ml; 1 was 200 copies/ml; 1 was 700

3   copies/ml; 1 was 1,200 copies/ml; 1 was 2,600

4   copies/ml; and 1 was 5,900 copies/ml.  The rest

5   were below the detection level.  The sensitivity of

6   this then is that 12/16 were detected by ID-NAT and

7   alternative NAT; 3/16 were negative by NAT and

8   alternative NAT and 1 was not tested.  A note which

9   I would like to make is that although 16 donations

10   were minipool NAT negative, all were anti-core

11   reaction.  This indicates a sensitivity issue but

12   not a safety issue.

13             [Slide]

14             So, I would like to present the questions

15   to the committee.  The first question is do the

16   sensitivity and specificity of the Roche COBAS

17   AmpliScreen hepatitis B test in minipools of 24

18   samples support licensing of the assay as a donor

19   screen?

20             DR. NELSON:  Thank you, Dr. Kaplan.

21   Comments?

22             DR. KAPLAN:  I will go through the three

161

1   questions and then come back to the first one.

2             [Slide]

3             Second question, if so, (2(a), assuming

4   continued use of screening tests for anti-core, do

5   the data support the use of the Roche COBAS

6   AmpliScreen hepatitis B test in minipools of 24

7   samples to screen blood for transfusion as an

8   equivalent alternative to the surface antigen test?

9             2(b), if the data do not support use of

10   the Roche COBAS AmpliScreen hepatitis B testing

11   minipools of 24 samples as an equivalent

12   alternative to HBsAg to screen blood for

13   transfusion, what additional data would be required

14   to validation such use?

15             [Slide]

16             The third question, do the data support

17   the use of the Roche COBAS AmpliScreen HBV test on

18   minipools of 24 samples to screen blood for

19   transfusion as an added test in conjunction with

20   licensed donor screening tests for HBsAg and

21   anti-core?

22             And now we will go to the first one.

162

1             DR. NELSON:  Comments?  Questions?

2   Anybody?  Yes, Susan?

3             DR. LEITMAN:  I have a question for the

4   FDA.  Is there precedent in the past for this

5   committee recommending licensure of a screening

6   assay without mandating its use?  Dr. Alter

7   mentioned third generation HCV EIA.  Is that the

8   only example?

9             DR. EPSTEIN:  Of course, we did recommend

10   HCV EIA screening; it is just that we were neutral

11   whether the EIA-3 should replace EIA-2.  We let

12   them coexist on the market.  CMV EIA is approved as

13   a screen and its use is voluntary.  When we first

14   reviewed the Abbott P24 test we ended up approving

15   it as a diagnostic but not a donor screen even

16   though studies had been done on its use for donor

17   screening.  Of course, we had a different set of

18   results.  So, I think that really the CMV is the

19   best example.

20             DR. ALLEN:  Let me just follow-up on that.

21   Again, as I understand the question, the committee

22   is being asked for a recommendation only on

163

1   supporting licensing and not on its use?

2             DR. EPSTEIN:  That is correct.  In other

3   words, question one is should it be approved with a

4   claim as a donor screening test.  Questions two and

5   three have to do with what recommendations for use

6   FDA might make or not make.  So, question two is

7   could you replace HBsAg testing, treating them as

8   equivalent alternatives?  That is the EIA-3 versus

9   EIA-2 model for HCV.  Then question three is if we

10   do not regard it as an equivalent alternative,

11   should we consider recommending it as an additive

12   test?  That is where we ended up in 1996 with P24

13   where we recommended it as a screen that was

14   additive.  Of course, there was no issue of

15   replacing a test at that point.

16             DR. KLEIN:  Well, not being a lawyer, I

17   don't really care much about precedent.  It seems

18   to me what we saw this morning, in my opinion, is

19   that it is a sensitive assay that has been

20   developed, that has been applied in the screening

21   format that picks up additional cases of hepatitis

22   B.  I think I find no reason why it shouldn't be a

164

1   licensed screening assay.

2             DR. GOLDSMITH:  Are we just considering

3   this for screening blood donations?  In other

4   words, donors of source plasma would not be

5   subjected to this test?

6             DR. KAPLAN:  Yes, at this point we are

7   only considering the blood donations.

8             DR. EPSTEIN:  Let me just clarify that

9   source plasma donations are not screened for anti

10   core, and we have seen data that DNA NAT will miss

11   anti-core positives, some of which would have DNA

12   on an ID-NAT test or alternative test.  So, no one

13   contemplates, and the company has not requested a

14   label to screen source plasma.  So, this is blood

15   for transfusion.

16             DR. NELSON:  I agree with Harvey.  I can't

17   see any reason that it shouldn't be licensed, and

18   this may not be as severe clinical outcome as HIV

19   but it is associated with chronic disease--cancer,

20   cirrhosis, etc. in some people and I think we

21   should try to prevent it.

22             There are some sub- populations in the

165

1   U.S. that have much higher rates of both prevalence

2   and incidence from HBV.  Although there were

3   600,000 donors screened, there was no attempt, I

4   don't think, to find the highest risk donors to

5   screen, or anything like that.  So, my guess is

6   that applied to the entire donor system the yield

7   will be somewhat irregular; maybe not entirely

8   predictable by the data that are available.  I

9   mean, in some populations may be more useful than

10   others in terms of picking up higher numbers of NAT

11   positive surface antigen negative donors.  Anybody

12   feel that this shouldn't be licensed?

13             DR. HOLLINGER:  I guess the question is

14   really, at least in my mind, not that the assay is

15   going to pick up some samples I think prior to HBs

16   antigen positivity.  There is a lot of information

17   we don't have to make that decision, as I look at

18   it, anyway.  One, we don't know what the carrier

19   state potential is in the people who receive this

20   kind of blood.  We don't know what the disease

21   potential is in patients receiving this type of

22   blood.  We don't know the outcomes--I think is a

166

1   whole problem here.

2             We are trying to make a decision about

3   doing a licensing assay on something we have very

4   little, if any data on in the seronegative window

5   periods regarding a disease outcome.  Even if you

6   forget the disease outcome and look at it

7   straightforward with what we know in ordinary

8   circumstances, the fulminant hepatitis and you will

9   see numbers like three-tenths percent or

10   four-tenths percent of people will get the disease,

11   I think that is very high and I think in reality it

12   is much less than that.  So, if you get one person

13   out of 200,000 donors who gets one of these units

14   of blood and four-tenths percent of 1/200

15   individuals eventually who will get this will

16   develop fulminant hepatitis, that is looking at

17   something--if my evaluation is simple, it would be

18   like one out of 40 million units of blood that

19   might result in fulminant hepatitis.  I may be

20   erroneous in that assumption, but the risk is

21   really very small.

22             On top of that, even people who get

167

1   chronic disease, which is probably less than 3

2   percent--now, some of these are immunosuppressed

3   patients and recipients so we can't probably use

4   that number appropriately but, certainly, half the

5   recipients will die over a certain period of time,

6   many of them around 60 years of age or so, getting

7   blood.  If they do get this disease, chronic liver

8   disease, to progress to cirrhosis takes decades, on

9   an average about 10 years a stage and if there are

10   4 stages, about 40--30 to 50 years probably before

11   you reach cirrhosis.  Even people who get

12   cirrhosis, 80 percent are still alive and doing

13   well 10 years after the diagnosis of cirrhosis is

14   made, and probably longer than that but that is

15   where the data leads us.

16             So, even with hepatitis B, those numbers

17   are probably very similar.  So, I think this is a

18   little different than the HIV situation that we are

19   dealing with here.  So, while the assay certainly

20   is a good assay and I have no problem with that,

21   the issue is whether it should be a donor screening

22   assay at this juncture.  I mean, I could see that I

168

1   would approve the assay because it can be useful to

2   any blood banking situation.  The issue is whether

3   we have enough information to make the

4   recommendations as a donor screening tool at this

5   stage.

6             DR. NELSON:  You have certainly raised

7   some of the unknowns about natural history.  The

8   other side of the coin is that there are, if you

9   will, worse case scenarios for hepatitis B.  It is

10   not an innocuous virus.  It is certainly more

11   significant than HAV and others that can be

12   transmitted by transfusion.  Chronicity varies,

13   certainly by immunosuppression, age and infection.

14   If these are kids that are transfused you are going

15   to see some liver cancer and some cirrhosis.  So,

16   given this, I just can't see that this is totally

17   innocuous.  It is unknown how severe the outcome

18   is, but it is potentially significant.

19             DR. KLEIN:  you know, I agree with you,

20   Blaine, entirely but I guess the question in my

21   mind is that if it is not licensed, then if I have

22   patients in my cancer hospital and they are

169

1   primarily bone marrow transplant patients or

2   patients who are immunosuppressed, then I can't use

3   this as a screen unless it is licensed unless I get

4   an IND.  If it is licensed I have the choice of

5   using it in my high risk population as others do.

6             So, I guess what I am saying is I would

7   like to see it licensed although not necessarily

8   mandated, and I think there is a difference there.

9             DR. HOLLINGER:  I was under the impression

10   that makes it very difficult.  I could agree with

11   that but I thought Jay was saying this is a real

12   difficulty to license something without mandating

13   it.

14             DR. KLEIN:  Well, I guess that would be

15   his problem.

16             [Laughter]

17             DR. EPSTEIN:  We can license the test and

18   be silent whether its use is recommended.  But I

19   think that that will create a chaotic situation.

20   Just as in the other direction, if we license it

21   and we recommend use within some period, even if it

22   is a fairly long period, since it would then become

170

1   available very quickly to users who have the right

2   platform, it would also be a chaotic situation.

3             In other words, you know, I think we have

4   disruption to our system whichever direction we go,

5   assuming that it is a licensed product, whether we

6   are silent on recommended use or whether we

7   recommend the use, either way.

8             DR. STRONG:  This is a bit of deja vu for

9   me.  We had this very debate on our medical staff

10   because we are one of the clinical trial sites.

11   When the trial came to an end, do we continue or

12   not continue since we have the platform in-house?

13   So, we spent half a day in discussions, many of the

14   same points being made.  Now, in Seattle, of

15   course, we do have a patient population with a lot

16   of bone marrow transplants, a lot of cancer

17   patients, a large clinical children's hospital,

18   etc.  Our medical staff, which is made up of 8

19   hematologists and a couple of lab medicine people

20   were unanimous to continue to test.  In fact, the

21   three new cases are all our yield cases.  So, they

22   feel justified in having continued testing.

171

1             DR. ALLEN:  If, as Dr. Epstein says, there

2   is no recommendation to mandate its use if it is

3   licensed, I understand the forces of the

4   marketplace and it is likely to drive it very

5   quickly and fairly uniformly.  There is the cost

6   factor.  There are rising healthcare costs and the

7   pressure to try to maintain those as well as

8   possible, not add any unnecessary increase.  There

9   are certainly liability concerns and

10   considerations.  I think the marketplace will speak

11   to this and that is the way that our system works.

12             I think the question before us is very

13   simple, do the sensitivity and specificity in the

14   minipools, as the data presented, support licensing

15   the assay as a donor screen?  I think Dr. Klein's

16   first comment spoke very clearly to that.

17             DR. STRONG:  I think there is a mixed

18   feeling in the marketplace, much as has been voiced

19   here.  Of the five clinical sites that were

20   involved, two have stopped testing and three have

21   continued.  So, it is a debatable issue.

22             DR. NELSON:  Are we ready to vote on

172

1   question one?  Linda?

2             DR. SMALLWOOD:  Question one, do the

3   sensitivity and specificity of the Roche COBAS

4   AmpliScreen HBV test in minipools of 24 samples

5   support licensing of the assay as a donor screen?

6   Dr. Allen?

7             DR. ALLEN:  Yes.

8             DR. SMALLWOOD:  Dr. Davis?

9             DR. DAVIS:  Yes.

10             DR. SMALLWOOD:  Dr. DiMichele?

11             DR. DIMICHELE:  Yes.

12             DR. SMALLWOOD:  Dr. Doppelt?

13             DR. DOPPELT:  Yes.

14             DR. SMALLWOOD:  Dr. Goldsmith?

15             DR. GOLDSMITH:  Yes.

16             DR. SMALLWOOD:  Dr. Klein?

17             DR. KLEIN:  Yes.

18             DR. SMALLWOOD:  Dr. Laal?

19             DR. LAAL:  Yes.

20             DR. SMALLWOOD:  Dr. Harvath?

21             DR. HARVATH:  Yes.

22             DR. SMALLWOOD:  Dr. Hollinger?

173

1             DR. HOLLINGER:  No.

2             DR. SMALLWOOD:  Dr. Kuehnert?

3             DR. KUEHNERT:  Yes.

4             DR. SMALLWOOD:  Dr. Leitman?

5             DR. LEITMAN:  Yes.

6             DR. SMALLWOOD:  Dr. Quirolo?

7             DR. QUIROLO:  Yes.

8             DR. SMALLWOOD:  Dr. Schreiber?

9             DR. SCHREIBER:  Yes.

10             DR. SMALLWOOD:  Dr. Whittaker?

11             DR. WHITTAKER:  Yes.

12             DR. SMALLWOOD:  Ms. Knowles?

13             MS. KNOWLES:  Yes.

14             DR. SMALLWOOD:  Dr. Nelson?

15             DR. NELSON:  Yes.

16             DR. SMALLWOOD:  Dr. Strong, your opinion?

17             DR. STRONG:  Yes.

18             DR. SMALLWOOD:  The resulting votes are 15

19   yes votes and 1 no.  The non-voting industry

20   representative agrees with the yes vote.

21             DR. NELSON:  Can we move on to the second

22   question?

174

1             DR. KAPLAN:  Question 2(a), assuming

2   continued use of the screening tests for anti-core,

3   do the data support use of the Roche COBAS

4   AmpliScreen HBV test in minipools of 24 samples to

5   screen blood for transfusion as an equivalent

6   alternative to the HBsAg test?

7             DR. NELSON:  Comments?  Questions?

8             DR. KLEIN:  Since I want to get to lunch,

9   I guess I will start off--

10             [Laughter]

11             I would address this in a similar way to

12   what Blaine said earlier.  I don't think that these

13   are equivalent.  I think we have seen data to

14   suggest that there may be more sensitivity in acute

15   cases and we don't know too much about chronic

16   cases, and we have a very small amount of data at

17   this point in time.  So, I certainly don't think

18   that I would accept that as an equivalent

19   alternative.

20             DR. ALLEN:  I concur with that assessment

21   and I think certainly we have had speakers this

22   morning during our discussion periods who have

175

1   spoken to the need for additional data, the type of

2   data that could be collected.  I think this is very

3   exciting and promising and look forward to the

4   collection of additional data so that we can make

5   further decisions down the road.

6             DR. LEITMAN:  I just want to say that it

7   is a relatively simple answer.  We saw data

8   presented to tell us that it is no an equivalent

9   alternative, not yet.  We saw the data that says it

10   is not.

11             DR. HOLLINGER:  I think the test is

12   probably a better test in general, I should say it

13   is a more sensitive test for detection at the

14   seronegative window period time.  But, again, I

15   agree.  I think we need more data.  There are some

16   benefits to redundancy.  If you look at some of the

17   reproducibility assays you will see some false

18   negatives by log; you will see some false negatives

19   by site, and so on.  So, there has always been

20   benefit to redundancy with it is the HBs antigen

21   test or the anti-HBc test, and so on.  Until we

22   have enough information with testing down the line

176

1   in several assays and a variety of other things, I

2   think it is too early to make that determination at

3   this time.

4             DR. KLEIN:  We also have several

5   generations experience with HBsAg testing and I

6   would be very reluctant to step away from that

7   right now on the basis of a few hundred or even a

8   few thousand assays that we have seen here.

9             DR. NELSON:  One question for the blood

10   bankers in the audience and the panel is when a

11   surface antigen screening test is positive is a

12   neutralization assay performed routinely?  Because

13   I saw in some of the data that were presented that

14   they talked about surface antigen screening and

15   then they went back and said it was neutralization

16   positive or negative.  I thought that maybe that

17   was part of the routine.  It is, I guess.

18             DR. ALLEN:  But am I correct that that is

19   not to exclude a unit of blood, obviously; it is

20   only to give information back to the donor?

21             DR. KLEIN:  Well, the unit of blood is

22   going to be excluded anyway.

177

1             DR. NELSON:  Harvey?

2             DR. ALTER:  Just to reiterate what Harvey

3   Klein said, you know, this test is the apple pie of

4   blood banking.  It was the first test we introduced

5   for specific hepatitis screening.  It has really

6   served us incredibly well for 35 years, and one

7   would have to have overwhelming evidence to drop it

8   at this point.  I think certainly the evidence is

9   not overwhelming.  And, I just want to remind you

10   that we are still doing syphilis testing.

11             [Laughter]

12             DR. ALLEN:  We still have syphilis too!

13             DR. KLEINMAN:  I just wanted to comment

14   that surface antigen tests--obviously, there are

15   multiple licensed ones on the market and we have

16   seen that one recently introduced test has had

17   tremendous specificity problems.  My remarks really

18   aren't predicated on whether the test should be

19   dropped but in those cases having an HBV NAT will

20   help resolve some of the interpretation of a newly

21   introduced hepatitis surface antigen test that is

22   not performing in the way we expected it to.  So,

178

1   there are different caveats.

2             DR. SCHREIBER:  We also know from the

3   Biswas paper that the improved antigen tests are

4   going to have incremental yield so that should

5   decrease the incremental yields of the NAT test

6   introduction.  So, I think that is not a

7   replacement.

8             DR. KLEINMAN:  I do have a sort of a

9   conceptual question to the FDA though.  I mean, we

10   know that there is a test out there in the pipeline

11   that is supposed to be more sensitive, but it seems

12   kind of odd to compare a test that is pending

13   licensure to one that isn't licensed yet.  I mean,

14   it seems like a kind of odd bar to set that we

15   think that newer tests for another analyte are

16   going to be better, therefore, we should use that

17   information in deciding about whether a current

18   test should or should not be used or should replace

19   one.  I mean, I can argue it both ways.  If we know

20   that something better is coming along, we can't

21   disregard that data and, yet, it is not an

22   FDA-licensed test that we are comparing to.  So, it

179

1   is a sort of dichotomous way to think about things

2   and I just wonder how FDA thinks about this because

3   they know more about whether the tests coming

4   through the pipeline, in fact, are more sensitive

5   and, in fact, are going to be licensed.

6             DR. NELSON:  As I recall the Biswas paper,

7   certainly the individual NAT tests were still more

8   sensitive than any of the pipeline tests for

9   surface antigen, and it wasn't that the surface

10   antigen tests in the pipeline were more sensitive

11   than the NAT so these are two different assays.

12   Here is Robin.  He can tell us about it.

13             DR. BISWAS:  All I wanted to say is that

14   you have to deal with really the here and now.  You

15   have a test which we are asking you questions

16   about, how we should manage it, whether we should

17   license it or not and whether it should replace the

18   HBsAg.  Keep in mind that technology improves and

19   changes.  For example, the study which has been

20   referred to several times today, remember that that

21   was published a year ago, over a year ago, and it

22   was actually done two or three years before that.

180

1   So, these tests do change.  In a nutshell, we

2   compared the then investigational tests to the at

3   that time currently licensed tests.  We compared

4   pooled NAT as it was then to the currently licensed

5   tests at that time.  And, we compared single unit

6   NAT as it was at that time to the currently

7   licensed tests at that time.  I should also say

8   that some of the investigational NATs a year ago,

9   two years ago, three years ago have, in fact been

10   licensed.

11             DR. NELSON:  In the interest of lunch, are

12   we ready to vote on this question?

13             DR. HOLLINGER:  I am just a little

14   confused about this question.  I mean, the question

15   is asking should this assay replace HBsAg.  Is that

16   correct?  If that is what it is then, of course, I

17   would say no.

18             DR. NELSON:  That is not the way, it says

19   is it an equivalent alternative, and in some ways

20   it is better and in others we don't have enough

21   data.

22             DR. HOLLINGER:  If the question is asking

181

1   should it replace the HBs antigen obviously, for

2   me, I would say no.  But that is not sort of the

3   way it is worded.

4             DR. NELSON:  It is worded is it an

5   equivalent alternative.

6             DR. KAPLAN:  Yes, because the company

7   claims it could be used instead of,

8   basically--instead of the surface, you could use

9   the surface or you could use the NAT.  That is the

10   claim.

11             DR. HOLLINGER:  So, basically it says

12   should it replace the HBs--I like the word

13   "replace" myself.

14             DR. NELSON:  Jay?

15             DR. EPSTEIN:  First of all, we are asking

16   the committee to comment on the strength of the

17   science and leave the policy decision to the FDA.

18   That said, the implication of this question is the

19   science base for considering replacement of HBsAg

20   by minipool NAT.  So, I think you have the question

21   right, it is just to clarify that we are asking how

22   strong is the science.

182

1             DR. BUSCH:  It is in the context of

2   anti-core screening.  So, as Roche presented, the

3   only concern is if anti-core failed to detect

4   something.

5             DR. NELSON:  Right, that is true.  If

6   there were no core testing it would be a different

7   scenario, right, which, in much of the world, is

8   the case.

9             DR. BISWAS:  I just wanted to say for the

10   record that I think I said HBV NAT investigational

11   when I meant HBsAg investigational.

12             DR. NELSON:  Okay.  Linda, we are ready.

13             DR. SMALLWOOD:  Question number 2(a),

14   assuming continued use of screening tests for

15   anti-HBc, do the data support use of the Roche

16   COBAS AmpliScreen HBV test in minipools of 24

17   samples to screen blood for transfusion as an

18   equivalent alternative to the HBsAg test?  Dr.

19   Allen?    DR. ALLEN:  Not at the present time, no.

20             DR. SMALLWOOD:  Dr. Davis?

21             DR. DAVIS:  No.

22             DR. SMALLWOOD:  Dr. DiMichele?

183

1             DR. DIMICHELE:  No.

2             DR. SMALLWOOD:  Dr. Doppelt?

3             DR. DOPPELT:  No.

4             DR. SMALLWOOD:  Dr. Goldsmith?

5             DR. GOLDSMITH:  No.

6             DR. SMALLWOOD:  Dr. Klein?

7             DR. KLEIN:  No.

8             DR. SMALLWOOD:  Dr. Laal?

9             DR. LAAL:  No.

10             DR. SMALLWOOD:  Dr. Harvath?

11             DR. HARVATH:  No.

12             DR. SMALLWOOD:  Dr. Hollinger?

13             DR. HOLLINGER:  Gosh, I feel more

14   comfortable with this one now; I am not left out

15   hanging.  No.

16             DR. SMALLWOOD:  Dr. Kuehnert?

17             DR. KUEHNERT:  No.

18             DR. SMALLWOOD:  Dr. Leitman?

19             DR. LEITMAN:  No.

20             DR. SMALLWOOD:  Dr. Quirolo?

21             DR. QUIROLO:  No.

22             DR. SMALLWOOD:  Dr. Schreiber?

184

1             DR. SCHREIBER:  No.

2             DR. SMALLWOOD:  Dr. Whittaker?

3             DR. WHITTAKER:  No.

4             DR. SMALLWOOD:  Ms. Knowles?

5             MS. KNOWLES:  No.

6             DR. SMALLWOOD:  Dr. Nelson?

7             DR. NELSON:  I guess you can't vote maybe

8   so I will vote no.

9             DR. SMALLWOOD:  Dr. Strong, your opinion?

10             DR. STRONG:  Despite my conflict, I guess

11   I have to vote no.

12             DR. SMALLWOOD:  The results of voting for

13   question number 2(a), unanimous no.  The non-voting

14   industry representative agrees with the no vote.

15             DR. NELSON:  Do you want to read the next

16   question?

17             DR. KAPLAN:  Yes, 2(b), if the data do not

18   support use of the Roche COBAS AmpliScreen HBV

19   testing minipools of 24 samples as an equivalent

20   alternative to the surface to screen blood for

21   transfusion, what additional data would be required

22   to validate such use?

185

1             DR. NELSON:  This one isn't yes or no.

2             DR. ALLEN:  In the interest of time, I

3   would suggest that I think there have already been

4   a lot of suggestions made earlier during the

5   discussion.  In particular, I think the

6   presentations by Dr. Alter and Dr. Busch addressed

7   this.  Unless the FDA wants additional, more

8   specific discussion, perhaps we can move on.

9             DR. NELSON:  I guess just more data.

10   Since blood donors are routinely screened for

11   surface antigen, you know, larger numbers could be

12   screened for DNA.  That would seem to be one way to

13   approach it.  Harvey?

14             DR. KLEIN:  I would just like to see a

15   larger field experience, such as we would get if

16   this were a licensed assay and some people, such as

17   in Seattle, who are using it.

18             DR. KLEINMAN:  Yes, I wanted to pose one

19   question though because Mike, in his slides,

20   suggested that studies should not only be done in

21   blood donor populations but should be done in

22   subpopulations who acquire infection through

186

1   different means, and done in vaccinated recipients,

2   and I am not sure that I agree with that.  I don't

3   know whether I agree with that or not at this point

4   because we are really looking at the assay in the

5   donor context, and if you find that there are

6   discrepancies in other contexts, if those people

7   don't become donors I don't know whether it is

8   important or not.  Clearly, you need to give some

9   direction to the manufacturers, or FDA eventually

10   needs to give some direction to the manufacturers

11   about whether studies in donors are sufficient or

12   whether they have to go and do their studies in

13   other population groups.  So, I think a little bit

14   of discussion on that might be helpful.

15             DR. NELSON:  Except that the donor

16   population is somewhat heterogeneous.  It is made

17   up of many subpopulations and I would see some

18   justification--when we looked at hepatitis

19   B-infected IV drug users we found a different

20   serologic pattern than was common among other

21   subpopulations.  Whether or not that would

22   translate to different HBsAg dynamics I don't know

187

1   because we were looking at core.  But there are

2   population variations and, you know, most donors,

3   thankfully, aren't drug users but some are.

4   Because of that, I think understanding this better

5   in subpopulations--I can see some justification for

6   that.

7             DR. KLEINMAN:  Kenrad, I don't doubt that

8   the dynamics might be different and the findings

9   might be different.  I am really posing the

10   question whether if you find different results in a

11   population of IV drug users, if they contradict the

12   results you get in screening of millions of blood

13   donors, would that be enough to persuade you that

14   you still don't want to drop the test if the

15   dynamics are different.  Because if it is not

16   enough to persuade you, then why require it?  I

17   mean, the studies are interesting from a scientific

18   point of view and really need to be done, but they

19   wouldn't be a requirement for data to drop the

20   test.  So, that is the question I am posing.

21   Scientifically, it should be studied but whether it

22   needs to be studied for getting a claim is really

188

1   the issue I would like to hear discussed.

2             DR. NELSON:  Well, I would agree when you

3   talk about millions but there are small, hopefully

4   small subpopulations in a donor population that

5   comes from a different universe.

6             DR. BUSCH:  The reason I kind of argued

7   that that data are needed is I do think that the

8   donor pool here does have a sampling of these

9   atypical cases in endemic infections.  For example,

10   the vaccine breakthrough infection that Roche

11   detected are seen routinely in Taiwan now.  By

12   participating in studies there we would better

13   understand what is going on, get volumes from these

14   people and be able to look at infectivity.  There

15   are two cases in France of surface antigen positive

16   donors persistent who lacked anti-core were both

17   people from Africa who were probably infected based

18   on the work perinatally.  So, it may be they never

19   formed anti-surface because they were essentially

20   colonized in utero to the antibody.

21             So, I agree that the data that is relevant

22   for the U.S. decisions is large-scale donor data

189

1   but I think the numbers of these unusual cases that

2   trickle up in our setting are small and we can

3   better study them and understand them by these

4   collaborative studies.

5             DR. ALTER:  I just want to go on record

6   that, as much as I love this test, the concept here

7   is correct.  It is just premature.  I think

8   ultimately HBV DNA will replace HBsAg testing when

9   we go to smaller pools or individual testing and

10   when the data are larger, when we have a greater

11   database.  But I think the database will support

12   going from this test to the DNA.

13             DR. GOLDSMITH:  Just to add, I think that

14   this test will apply to blood donors and non-blood

15   donors, disease state subjects, that we might get

16   additional clinical information that might add to

17   the validity of the test.  So, it would be useful

18   for the manufacturer to apply the test outside the

19   primary donor group.

20             DR. DIMICHELE:  That is exactly what I was

21   going to say so I would like to second that.

22             DR. STRONG:  Just one comment, I mean, we

190

1   have continued to screen, as have two other

2   centers, so we are collecting more data.  I think

3   these kind of cases that are surface antigen

4   positive only are going to be very hard to find.  I

5   mean, they just don't happen, at least with the

6   numbers that we have which are approaching a couple

7   of million now.  They are either false positives

8   when they are surface antigen alone or they are

9   positive along with anti-core.  So, those surface

10   antigen only are going to be hard to come by.

11             DR. BUSCH:  But they could be found

12   routinely from the Red Cross Blood Systems.  What

13   we are talking about is taking surface antigen

14   positives that are anti-core negative in programs

15   that aren't doing NAT studies, and those units are

16   being found all the time and if they are just

17   worked up in some sort of national collaborative

18   study by the different NAT assays, in follow-up I

19   think we could generate a larger data set fairly

20   quickly.

21             DR. STRONG:  I agree we can generate a

22   larger data set.  I am just saying with our limited

191

1   experience we have thus far, all of those have

2   turned out to be either false positives or core

3   positives.

4             DR. KAPLAN:  So, would you say Phase IV in

5   a high risk population to get the correct numbers?

6             DR. BUSCH:  Yes, I mean, we routinely

7   screen for antigen and anti-core, and the units

8   that are antigen reactive and anti-core negative

9   are the problematic cases.  I don't know how many,

10   but we pick up a moderate rate of those and if we

11   were to simply put into place a protocol that

12   captured those units, and made them available to

13   Roche and any other company and enrolled those

14   donors, in a blood donor context we could capture

15   these cases.

16             DR. KAPLAN:  Yes, but it would require a

17   large number of these cases so you are talking

18   about a prolonged Phase IV, or what is your idea

19   about the number of total cases?

20             DR. KLEINMAN:  Well, we can address that

21   because when we looked at data from 1996, I think

22   it was, in REDS we documented surface antigen

192

1   positive in the absence of anti-core in about

2   1/100,000 tested donors, and then we did some DNA

3   studies, which have remained unpublished.  If you

4   look here, in the Roche study, it was 7/600,000, so

5   about the same.  So, I think this finding with

6   stable HBsAg assays occur about 1/100,000 donors.

7   Now, the Ortho version 3 test that has some

8   specificity problems, this would not be a good test

9   by which to find such donors unless they fix those

10   problems because we know that that is the exact

11   profile of false-positive surface antigens so we

12   would be overwhelmed by samples that don't actually

13   fit the bill.  But I think with stable assays,

14   either with the Abbott Auzyme or presumably once

15   Ortho's assay gets fixed, we are looking at a rate

16   of 1/100,000.

17             So, if you had a program to say we are

18   going to try to do this in the large blood

19   collection systems across the U.S., so let's say

20   American Red Cross, BSL and a few others, so now

21   you have 8 million units or so a year, 9 million

22   units, 1/100,000 so what does that come out to be?

193

1   Ninety.  So, 80 or 90.  We might be able to accrue

2   80 or 90 units a year.  Even if the estimates are

3   half of that we might be able to accrue 40 units a

4   year.  If we could get those donors in for

5   follow-up and capture those plasmas and capture the

6   samples, test by various surface antigen assays,

7   various DNA assays, we might be able to accrue a

8   larger sample set, not in a short amount of time

9   but not in too long a period of time.

10             DR. NELSON:  This was an essay question,

11   which I think we have already answered.  So, let's

12   move on to the third one.

13             DR. KAPLAN:  The third question is do the

14   data support the use of the Roche COBAS AmpliScreen

15   HBV test on minipools of 24 samples to screen blood

16   for transfusion as an added test in conjunction

17   with the licensed donor screening tests for HBsAg

18   and anti-hepatitis B core?

19             DR. NELSON:  Yes?

20             DR. SCHREIBER:  Based on all the

21   discussions, can we add the word "voluntary" after

22   "added?"  Added voluntary tests?

194

1             DR. KUEHNERT:  I thought the spirit of

2   this question was a required test.  That is what I

3   thought was being asked.  Which is it?

4             DR. EPSTEIN:  We are not asking you to

5   comment on the policy question, which is should we

6   recommend additive testing.  We are asking you to

7   comment on the strength of scientific data.  In

8   other words, are these valid results?  What is

9   their magnitude?  Are there caveats?

10             DR. KUEHNERT:  I am sorry, how is question

11   one and three different?

12             DR. EPSTEIN:  Question one is whether the

13   sensitivity, specificity, reproducibility or

14   comparable to other tests for hepatitis B virus

15   that we have approved as screens.  So, does it meet

16   the standard of screening tests?  Question two was

17   if it does, is it sufficiently better as an advance

18   that it should replace HBsAg?  We heard your

19   answer.  Question three is really is there an

20   additive yield that we feel is meaningful?  You

21   don't directly have to answer the policy question

22   of should we go on to recommend it.

195

1             DR. NELSON:  Yes, this doesn't include a

2   "win."  It just includes do the data suggest that

3   added to the surface antigen and core antibody,

4   would it be an improvement or would it be

5   scientifically valid.

6             DR. EPSTEIN:  I mean, committee members

7   can express caveats if they wish.

8             DR. ALTER:  But, Jay, if there wasn't an

9   additive yield we wouldn't have voted yes to the

10   licensure.  So, the two are really--

11             DR. EPSTEIN:  I actually disagree with

12   that, Harvey.  I mean, it could have had no

13   additive yield and still be an adequate screen.

14             DR. ALTER:  So, we can't ask the question

15   we really want to ask.

16             [Laughter]

17             DR. EPSTEIN:  Well, I think committee

18   members can comment.  I mean, they are here to

19   advise the FDA and they can certainly express their

20   thinking about what we ought to do, but the

21   scientific advisory committees are intended to be

22   advisory on the science and we are trying to ask

196

1   you the science question.

2             DR. NELSON:  Well, it seems to me that the

3   data are pretty clear on this, that adding the NAT

4   test to the surface antigen and core true positives

5   were detected.

6             DR. KLEIN:  Well, I would like to add a

7   caveat because I certainly agree more with what I

8   heard Blaine say at the very beginning, and that is

9   that while there is no question in my mind that you

10   picked up additional donors who might transmit, we

11   really don't have any evidence that the public

12   health impact is going to be significant, for all

13   of the reasons that were stated earlier.  So, I

14   think we are too early in the game to say there is

15   going to be a public health impact of requiring

16   such a test, and I think I would leave it at that,

17   not recommending to the FDA whether they should

18   require it or not but saying that, in my mind, I

19   don't think the public health benefit would push me

20   to require it.

21             DR. NELSON:  I am not convinced that the

22   public health benefit or lack of benefit has been

197

1   adequately studied.  I think that what has been

2   adequately studied--not adequately but what we have

3   some data on is that NAT testing will pick up some

4   infectious units in the context of surface antigen

5   and core-antibody, all the testing that is being

6   done.  Coupled with what we know about hepatitis B,

7   I think it is likely that there is some public

8   health benefit.  The magnitude of it is unclear I

9   think.

10             DR. KLEIN:  Just looking at the wording

11   here, it says did the data support use?  I mean,

12   everyone can express their opinion on whether they

13   think there would be additional yield with this

14   test but it asks specifically about the data

15   presented.  But it is very difficult to answer this

16   question because there is what is on the paper and

17   what people think it means if there is a yes vote.

18   So, that is what I think is very difficult for me

19   in voting on this.

20             DR. ALLEN:  I think Dr. Epstein is very

21   clear that we are basing our decision as a

22   scientific decision or recommendation on the data

198

1   presented, and not making a policy decision for the

2   FDA or even really a policy recommendation.  I am

3   comfortable with that as long as that is the

4   interpretation.  I have no doubt that this will

5   prevent some hepatitis B infections because it does

6   move into the window period; that there will be a

7   relatively small number of infections eliminated.

8   The clinical impact of that I think is a little

9   more difficult.  Certainly, in children, in people

10   who are immunocompromised I think the clinical

11   impact for those individual people could be

12   significant although the numbers are going to be

13   relatively small and probably decreasing over time

14   as our donor population more and more becomes those

15   who have been immunized earlier.

16             I am a little conflicted in terms of

17   exactly what the cost-benefit value is going to be,

18   but I don't think we are being asked to make that.

19   Very clearly, if you want to close the window

20   period as much as possible, given this test, you

21   would do individual testing rather than minipools

22   of 24.  And, I think these are some of the

199

1   questions that need to be investigated by blood

2   collection centers, by public health officials and

3   others as we move forward with a licensed test in

4   place and see how it best fits into the total

5   testing scheme.  And, I would certainly encourage,

6   as additional data become available, to reopen the

7   question of whether we then need three tests or can

8   move to a combination of two tests based on hard

9   data in actual situations and careful evaluation

10   over a period of time.

11             DR. NELSON:  Steve?

12             DR. KLEINMAN:  Judging by the comments of

13   the panel, I feel like the wording of the question

14   is not being interpreted the same by everybody on

15   the panel.  By that, I mean the specific words

16   support the use.  I think you could interpret the

17   words "support the use" as do the data indicate

18   that there is increased yield from this test.  If

19   you interpret it that way you might answer the

20   question one way.  The other way you could answer

21   the question of "support the use" is not only do

22   the data show there is increased yield but do I

200

1   personally think the test should be done if I were

2   making the decision.  That means not only do they

3   support increased yield but I would then take into

4   account all the other factors and vote according to

5   that.

6             I don't think everybody on the panel is

7   even trying to answer the same question right now

8   so I think you need to know whether you are asking

9   the question do the data support that an increased

10   yield would be achieved by implementing this test

11   in conjunction with surface antigen and anti-core,

12   or is the question do the data support the fact

13   that we recommend that this test be used?  They are

14   different questions and I don't think they are

15   being interpreted the same by everybody on the

16   panel.

17             DR. NELSON:  There was no data presented

18   on the natural history of infection.  There were a

19   few opinions but there was no data.  We have to

20   deal with the data that was presented so I think

21   the first of your questions is the way we should

22   vote.  That is my own opinion.  I am not sure that

201

1   all the potential hepatitis B infections are

2   benign.  You know, there is no data on that and

3   certainly we have people with a lot of clinical

4   experience but we need better data on this I think

5   if we are going to interpret it more broadly.  Yes?

6             DR. HEATON:  There is one key issue that I

7   would like to add to the discussion, putting on my

8   advisory committee for blood safety and

9   availability hat.  I made the point that there is

10   an enormous reimbursement implication of this

11   recommendation.  If it is recommended that this

12   test does improve public safety and the BPAC would

13   support that, then it becomes a much better case

14   for the blood centers to get reimbursement.  In the

15   absence of such a recommendation, for the average

16   blood center which might wish to implement this

17   test the hospitals will view this as entirely

18   discretionary and may, in fact, push back on

19   covering the cost.

20             From one of the manufacturer's

21   perspective, I can tell you that in pretty well

22   every case nucleic acid testing detects this virus

202

1   before the protein is present.  So, in terms of

2   safety, whilst you may argue about the level of

3   improved safety, on the front end of the case you

4   get good detection.  So, there is a reimbursement

5   issue here as well as a medical issue.

6             DR. NELSON:  Harvey?

7             DR. ALTER:  Yes, I think Steve framed the

8   issues very well.  If you break it down into two

9   separate questions.  For the first question, do the

10   data support that the test does something positive,

11   you can't vote no on that if you voted yes on

12   number one.  Number one is sort of a minimalistic

13   thing saying it can be licensed even if it is

14   equal, but everybody feels it is better so you have

15   to vote yes on number three.

16             But I think the way it is worded now, if

17   people outside this room are going to see that

18   question and then see the vote, it is going to give

19   the other, the part two impression, that is, that

20   the committee is recommending that the test be

21   used.  So, I think either the wording needs to be

22   more specific or question three should be dropped.

203

1   You got the sense of the group and you don't want

2   our opinion on policy.

3             DR. HOLLINGER:  This is just sort of an

4   aside, but if you want public safety in this

5   country you vaccinate everybody.  The other two

6   diseases we talked about have no vaccines

7   available; we have an excellent vaccine for

8   hepatitis B.  We are spending a lot of money for

9   donor testing and screening on this issue which

10   could be spent for vaccine policies of the

11   population in this country.  Then there is no risk

12   if you immunize everyone.

13             DR. NELSON:  Providing everybody responds

14   and gets the vaccine, and all those kind of things.

15   Jay?

16             DR. SIEGEL:  I think, given the difficulty

17   of a consensus interpretation of the question, it

18   might be better not to have votes but to hear the

19   committee members' individual discussions of how

20   they see the issue of additive testing framed.

21   Clearly, we are saying something about additive

22   testing and people can just comment on what they

204

1   think about a scenario in which minipool HBV NAT

2   would be an additive test to HBsAg, and just avoid

3   up and down votes.

4             DR. LEITMAN:  Everybody said the data

5   suggests that at the front-end of an infection in

6   the seronegative window period this clearly adds

7   something, adds something small, and I heard

8   somewhere between one-fourth to one-sixth of all

9   cases undetected by current methods would be

10   detected by this assay.  So, it is not the

11   majority; it is the minority.  And, we heard

12   suggestions and data that one would head towards

13   the majority with individual donor NAT and that

14   that might be substantially equivalent to and

15   better than in all cases HBsAg but we are not there

16   yet.

17             So, Jay phrased it as is there an additive

18   yield?  Yes, there is clearly additive yield but I

19   have seen nothing to suggest that that additive

20   yield translates to something real in terms of

21   better outcomes in the patients that we transfuse.

22   I have no idea what the risk of those very low

205

1   viremia window period donations is in the absence

2   of that.  Given all the substantial changes and

3   upheavals required to implement a new test, I would

4   wait until there is individual donor NAT data or

5   something that moves me further to do that.

6             I was also very moved by Dr. Hollinger's

7   comments.  My oldest child is 16 and all three of

8   them have been vaccinated.  So, if you get

9   pediatric care, you are vaccinated.  I don't know

10   that there is a public health recommendation to

11   vaccinate the American adult population though.

12             DR. NELSON:  Well, there is for pregnant

13   women, screening and vaccination, and for the

14   military and for healthcare workers.  But this

15   country is very diverse.  We get people from all

16   over the world who, you know, sleep in subways and

17   what-have-you.  Some of them become blood donors,

18   unfortunately.  I am all in favor of vaccination

19   but I am not sure--it is going to take a while to

20   solve this problem by vaccination, and it is

21   primarily viewed I think as pediatric, pregnant

22   women and certain subpopulations with higher

206

1   exposure.  So, you know, I think it is a great

2   vaccine and it is very important but we are not

3   quite there yet.

4             DR. HOLLINGER:  The other thing that I

5   thought was interesting in the discussion today is

6   the fact that the assays, even very sensitive

7   assays, pick up something differently.  I think

8   there was some testing where one alternative test

9   picked up things that another test did not pick up.

10   That, of course, is an issue that has me somewhat

11   concerned in looking at false-negative tests, not

12   so much false-positive but false-negative assays.

13   That is what we are really worried about, not the

14   false positives.  A lot of false positives also

15   were in there--the number is small considering the

16   number of people that are tested so, I guess, in

17   that way it is not a real big issue.  But I think

18   that is another interesting part of what we have

19   been discussing.

20             DR. NELSON:  I guess we are not going to

21   vote on this.  Yes?

22             DR. DIMICHELE:  I would just like to make

207

1   one comment, and that is that one of the statements

2   that was made earlier on the licensing question was

3   that the licensing would give people and blood

4   banks specifically options with respect to testing

5   in high risk populations.  I think we can't forget

6   those high risk populations.  In my opinion, the

7   data does support its use in picking up some

8   additional potentials for infections and I think we

9   have to recognize that we have both congenital and

10   acquired immunosuppressed populations.  We have the

11   cancer populations and even in pediatrics,

12   unfortunately, there is a population that isn't

13   covered by vaccine and that is our neonatal

14   population.  Much like we do tests, neonatal blood

15   for CMV and making sure neonatal blood is CMV

16   negative because of the risk in that particular

17   population, given the risk of hepatitis B in the

18   neonatal population, for instance, I could see this

19   as being a very important screening tool for

20   neonatal blood.  So, I think it is very important

21   to consider it as an additional tool because, if

22   not widespread use, certainly for selective

208

1   populations I think it would be extremely

2   important.

3             DR. ALLEN:  I would certainly hope, based

4   on what has been said today, that there will be

5   blood centers, once the test becomes licensed, that

6   will implement appropriate laboratory and

7   epidemiological studies that will enhance our

8   knowledge in a variety of ways.  I would hope that

9   other blood centers that are collecting blood from

10   populations that tend to have a higher incidence

11   and prevalence of hepatitis B infection might

12   choose to implement this as a routine.  And, I

13   think we certainly need to encourage the situation

14   where independent blood centers in areas of the

15   country where there is very little hepatitis B

16   incident infection, that they not feel compelled to

17   implement the test unless local data suggest that

18   it would be helpful.

19             DR. STRONG:  That is an interesting

20   comment because of the three test sites that have

21   continued to test, one of them is Minnesota which

22   probably has the lowest B activity in the country.

209

1   They have decided to test.  But I think it is

2   affected by, one, ease of implementation because we

3   already have it; two, the patient population we are

4   serving, and in our situation, in Seattle, it is

5   the cancer patients, the bone marrow transplant

6   patients, a large hemophilia population, and our

7   medical staff didn't have trouble with that

8   particular decision but I am sure it would be

9   complicated if we had to implement a whole new test

10   system.

11             DR. GOLDSMITH:  I guess we could also then

12   encourage the use of smaller minipools as they

13   proceed with their research, maybe with 12 or 6,

14   and there would still be some benefit of numbers,

15   or single donors to see if there is an additional

16   benefit that would actually improve the public

17   health.

18             DR. HOLLINGER:  And at the same time, in

19   terms of looking at these things in follow-up, it

20   would be important, as I think Jim mentioned, to

21   get a lot of other information, epidemiological

22   information.  We saw they picked up two people that

210

1   were within that window period.  I would like to

2   know more about those two people, their background,

3   their epidemiology, something of that nature, what

4   happened to them clinically, and so on.  We need

5   that kind of information to know if our

6   questionnaires are working and other things too.

7             DR. NELSON:  Yes, the sentinel county

8   study that the CAC has done has shown in the last

9   few years that the heterosexual transmission of

10   hepatitis B is one of the major routes.  The

11   incidence has actually declined overall in these

12   sentinel counties but heterosexual transmission--of

13   the incident cases that occur are probably the most

14   important in the general population.  These are

15   potential donors, some of them.  Let's break until

16   about 1:45.

17             [Whereupon, at 1:00 p.m., the proceedings

18   were recessed for lunch, to resume at 1:45 p.m.]

211

1             A F T E R N O O N  P R O C E E D I N G S

2             DR. NELSON:  In discussions with Dr.

3   Epstein after the conclusion, he would like to

4   either have a vote or a further statement--and I

5   hope we can make it brief--on the previous

6   question, question three, on hepatitis B NAT

7   testing.  Jay?

8             DR. EPSTEIN:  Yes, I appreciate that there

9   was difficulty with getting a common understanding

10   of the question.  This is question three to the

11   committee on HBV minipool NAT.  What I would like

12   to do is try to clarify what exactly is FDA asking

13   and let you determine, Kenrad, with you feel the

14   committee can vote on the question with that

15   understanding.  If not, it might be helpful to the

16   FDA if we could have individual committee members

17   simply state their views.

18             So, on the question of what do we mean,

19   first of all, I need to explain that we are asking

20   the committee for scientific advice.  We understand

21   that in any ultimate decision there are factors

22   such as feasibility of implementation, the

212

1   logistics, time scales and cost-effectiveness of

2   intervention.  But we are not asking this committee

3   those questions.  There will be a process through

4   the DHHS, which has an advisory committee on blood

5   safety and availability, to ask more global

6   questions about public health significance and

7   cost-relatedness of any FDA decision.

8             So, it is with the understanding that the

9   question is narrowed to a scientific point of view,

10   putting aside issues of cost-effectiveness, and the

11   implication of that is that the committee members

12   should not feel that if they vote that there is

13   significant benefit that that will in any way

14   obligate the FDA to make a policy decision in favor

15   of additive testing.

16             Then, what we are asking is whether

17   additive testing with HBV minipool NAT as described

18   would provide a benefit, a benefit to blood safety,

19   a benefit to patient health.  We understand that

20   there may be different points of view on the degree

21   of such benefit, but it is hard to avoid the

22   question because, after all, the committee has

213

1   basically has advised that the data support

2   approval as a screen; that the data do not appear

3   to be sufficient at this point in time for use of

4   the minipool test to replace HBsAg.  But, if FDA

5   were to approve it, what exactly do we say about

6   it?  And, the question will hang in the air whether

7   it ought to be recommended as an additive test.

8             So, the question won't go away.  FDA will

9   ultimately have to make a policy stand.  So, what

10   we are then saying is, well, we would like to be

11   advised by this committee whether, on scientific

12   grounds, we are or are not looking at health

13   benefit.  I think part of what has concerned the

14   committee is that there are really two dimensions.

15   One is an additional pickup rate of positive units.

16   The other is the significance of transfusing any

17   such units in terms of medical outcome.  You know,

18   we assume that you will be considering both factors

19   if you decide to vote on the question.

20             Again, we are asking you in this question

21   whether additive testing is a benefit, and we would

22   like you to answer that question with the

214

1   understanding that that is putting aside the

2   questions of economics and logistics.

3             DR. SMALLWOOD:  The proceedings of an

4   advisory committee require that after we have

5   closed discussion on an issue that we do not reopen

6   it because we have not made it public that we are

7   doing something differently.  We had previously

8   stated that we would not vote on question number

9   three and to open it up again, after individuals

10   have left, even the principals that are involved

11   here, would not be fair.  So, we would be going

12   against the advisory committee procedures.

13             DR. NELSON:  Principles don't have a vote.

14   There may be one committee member that has gone.  I

15   don't see why we couldn't--

16             DR. SMALLWOOD:  Excuse me, it is the

17   discussion, Dr. Nelson--

18             DR. NELSON:  It is not a discussion now.

19   We are only going to vote.  The discussion is over.

20             DR. SMALLWOOD:  Excuse me, this is an open

21   public meeting.  Under the rules for an open public

22   meeting we have to announce what we are going to do

215

1   in advance.  We have closed out the discussion.  We

2   closed out the discussion on this issue before we

3   went to lunch, and I am just quoting what the rules

4   are.

5             DR. EPSTEIN:  Is it permissible to have

6   individual expressions of opinion?

7             DR. SMALLWOOD:  Dr. Freas?

8             DR. NELSON:  Is the Supreme Court in

9   session?  Do we need to see what Justice Thomas

10   thinks about it?

11             DR. FREAS:  I really believe that it is

12   okay.  This meeting was announced in the Federal

13   Register.  There is nothing other than a statement

14   of clarification and I believe that if we went

15   around the table quickly and the committee members

16   made their comments, that would be permissible.  I

17   would prefer no votes at this time.

18             DR. SMALLWOOD:  To be clear, there will be

19   no votes but there will be committee comments

20   coming from individual members.

21             DR. NELSON:  It will take longer but do

22   you want to announce the names so they can make

216

1   comments then?  Dr. Doppelt?

2             DR. DOPPELT:  I think that there is some

3   benefit clearly that you can pick up some cases in

4   the window period.  It isn't a lot of cases that

5   you pick up though you do pick up some.  So, you

6   know, I think most people would probably say if it

7   didn't cost anything, what the heck, just do it.

8   But it does.  So, not commenting on the issues of

9   cost benefit, but just from the simple fact of what

10   does the data show, yes, you pick up cases and

11   there is benefit.

12             DR. GOLDSMITH:  I think the information we

13   saw does support the idea that there were a couple

14   of cases found but the testing algorithm that was

15   used hasn't been studied long enough to see if this

16   is the most powerful way to find additional cases;

17   that the 24 samples used for screening is probably

18   not the best way to use this new technology.

19   Therefore, if I had to vote, I would probably say

20   no at this time about this question, that the

21   sponsor should do additional work to find out the

22   real utility of this test, and also do additional

217

1   clinical work to find out what the outcome is of

2   potential transmissions.

3             DR. KLEIN:  For the record, I was not

4   considering either logistics or cost in my comments

5   this morning.  I think the data that was presented

6   clearly demonstrate that this assay will detect

7   some otherwise undetected donors infected with

8   hepatitis B virus.  From what we know about

9   hepatitis B virus, that would very likely infect

10   recipients.  The amount of public health benefit I

11   believe, based on those two statements, is finite.

12   It is real but it is very small.

13             DR. HARVATH:  I concur with the comments

14   that have just been made.  It is what I would have

15   said.

16             DR. KUEHNERT:  I think there is, you know,

17   the data presented and it is all we have to go on.

18   It seems to indicate that you pick up a small

19   number of additional cases.  I think you also have

20   to consider whether in implementing the test you

21   clarify or further muddy the picture as far as, you

22   know, possible false positives, and how those then

218

1   get reconciled.  So, to me, that all needs to be

2   added into the decision processes.  You might be

3   able to detect some additional cases but also what

4   is the full picture as far as the impact on false

5   positives?  You have to resolve those false

6   positives as well.  So, this question, to me, I

7   sort of have to consider those things as well.  So,

8   to me, I am not sure that it is a benefit when you

9   add all that in.

10             DR. NELSON:  I thought that the results

11   were fairly impressive, that it could actually pick

12   up additional cases and narrow the window period

13   particularly in the incident cases, and also pick

14   up some that were core positive who would have been

15   deferred with current screening.  So, when I look

16   at the international situation I see this as even a

17   more important application where core antibody

18   testing is not--I realize that is not part of the

19   U.S. algorithm, but I think that preventing

20   hepatitis B, which may be benign and self-limited

21   in many adults who have a normal immune system,

22   nonetheless, this is an oncogenic virus and it is

219

1   not all normal adults that are transfused, but

2   sometimes immunocompromised children.  So, I would

3   support it and I think it should be implemented

4   when it can be.

5             DR. ALLEN:  I would agree with the prior

6   comments.  It seems to me that the evidence is

7   fairly clear that this test, even in a minipool of

8   24, does narrow the window during the

9   seronconversion of early infection state and,

10   therefore, would result in a very slightly safer

11   blood supply, looking at the nation's blood supply.

12   I think the impact for individual people who might

13   otherwise receive those units is variable.  It

14   could be significant in some people.

15             In terms of the implementation, I think

16   there are many issues that need to be examined,

17   including cost benefit, the impact on the ability

18   of blood bank laboratories to implement the test

19   easily and successfully.  I think we need quite a

20   bit of additional research, as was discussed,

21   including examining the relationship between the

22   NAT test and the surface antigen test and, once we

220

1   have additional data, what the optimal combination

2   should be for routine screening in blood centers.

3   I think we need to examine the issue of whether

4   smaller minipools or even individual donor testing

5   provides any additional benefit versus the cost.

6             DR. DIMICHELE:  Given the data that was

7   presented earlier today and the discussion that

8   followed, I believe that there is strong enough

9   evidence for the FDA to give strong consideration

10   to adding this test to the complement of tests

11   already used for hepatitis B detection.  I believe

12   that the preliminary data mandates further

13   research, the nature of which has already been

14   stated.

15             DR. DAVIS:  I think the test will pick up

16   cases that would not have been discovered

17   otherwise.  I don't know that those cases are

18   enough to justify adding the test.  I would have

19   voted no.

20             DR. HOLLINGER:  I agree with much of what

21   has been said here.  One thing that has not been

22   added, I personally think the clinical impact will

221

1   be insignificant overall.  We have not talked even

2   about treatment, and there are some good nucleotide

3   and nucleoside analogs which are available which

4   work very well in acute viral hepatitis B in

5   stemming fulminant hepatitis, and so on.  So, even

6   if someone were to get clinically sick, the

7   probabilities are that they would have good

8   response to therapy and that would make this even

9   less significant.

10             DR. QUIROLO:  Well, I am a pediatrician

11   and all the people that I transfuse are children.

12   So, all those people would have hepatitis B for

13   decades if they were to contract it from a

14   transfusion.  Not only that, I take care of

15   minority children and immigrants and most of those

16   people are not immunized, the families aren't, so

17   the child would transmit that to his parents and

18   his siblings.

19             So, I think that the test is worth

20   considering.  I also don't see it as a stand-alone

21   test.  It says in conjunction with those tests.  I

22   don't think you can use it without using those two

222

1   tests; it is not a stand-alone test the way it is.

2             DR. WHITTAKER:  I think the data support

3   the use of this test, however, I think additional

4   testing is needed in smaller pools before we add it

5   as a mandated test for blood donor testing.

6             MS. KNOWLES:  I think also the data

7   support using this test as an added test.  I think

8   this is a part of the evolving armamentarium of

9   screening tools that we really need to keep pushing

10   on forward.

11             DR. STRONG:  If I were allowed to

12   comment--

13             [Laughter]

14             --we made the decision locally, as I have

15   already stated, with our medical staff that we will

16   continue to use it.  I think it has provided us

17   some additional benefits.  One, for example, is

18   that with the anti-core positive donors, and we

19   have a large number of false positives, but

20   remember that we are allowed to have two positive

21   anti-core results and we have had some pickups with

22   the first time anti-core that would be DNA

223

1   positives.  So, that would obviously be a

2   circumstance in which we would want to defer that

3   donor on a first time anti-core result.  So, it

4   adds that additional benefit.

5             It clearly has helped us pick up some

6   window cases, I think to some degree to our

7   surprise.  Our yield is more like 1/150,000, but

8   because we have had the three cases recently that

9   number will be diluted with time.  But the staff

10   certainly feels that it has had an added benefit.

11   I think there still needs to be data collected in

12   terms of the hepatitis B surface antigen issue and

13   core, clearly, will need to be continued.

14             DR. NELSON:  Now maybe we can move on.

15   Jay, is that satisfactory?

16             DR. EPSTEIN:  That is fine.

17             DR. NELSON:  The next item is current

18   trends in plasma product manufacturing,

19   introduction and background with Dr. Weinstein.

20        V. Current Trends in Plasma Product Manufacturing

21                  A. Introduction and Background

22             DR. WEINSTEIN:  Thank you, Kenrad.  I am

224

1   going to curtail my talk severely and just bring

2   out the important points.  I know that people have

3   to go at three o'clock and we do want to hear the

4   industry presentation as well as further comments

5   from the audience here.  So, let me just comment

6   here that the plasma fractionation industry has

7   undergone a number of notable changes, significant

8   changes within the last year.

9             [Slide]

10             We have heard about closures of plasma

11   collection facilities.  We have heard about layoffs

12   of personnel.  Some of the major plasma

13   fractionaters have reduced their plasma collection

14   significantly by a couple of million liters.  We

15   have also heard about the consolidations and

16   restructuring of the industry.  Alpha Therapeutics

17   has been absorbed by Grifols and Baxter.  Beyer has

18   put its fractionation plant up for sale.  Aventis

19   has been purchased by CLS to form ZLB-Behring, and

20   Octapharma is the new player on the U.S. market.

21             The question is do we have to fear chronic

22   shortages because of these changes?  What is

225

1   industry doing?  What is the state of the industry?

2   I called Jan Bult, representing the Plasma Protein

3   Therapeutics Association, to give us an overview of

4   where the industry is.

5                           Presentation

6             MR. BULT:  Thank you, Dr. Weinstein, for

7   this introduction.  I am going to, hopefully, tell

8   you in my presentation that we don't see a

9   near-term threat for supply, just to immediately

10   respond to one of your opening comments, and in

11   this presentation I would like to explain why.

12             [Slide]

13             What I am going to cover today is changes

14   since '98, the last time when we were facing

15   serious shortages in the United States; a little

16   bit about our data collection and how we

17   fractionate and therapies; what has happened

18   recently; what is going on with supply; tell you a

19   little bit about economics of plasma fractionation;

20   and then conclude with our statements.

21             [Slide]

22             Well, I think the most important thing

226

1   that has changed since 1998 is that we have

2   immediately implemented monitoring systems to

3   assess changes in supply dynamics.  That

4   information is available on a monthly basis.  It is

5   available on a website.  We do believe that

6   industry as a whole is much better positioned to

7   meet consumer demand.  You have seen from the list

8   from Dr. Weinstein that we have new companies

9   entering the market.  In addition to Octapharma

10   that you mentioned, also Grifols or Provitas has

11   entered the U.S. market with their immune

12   globulins.  So, we have companies now with broader

13   product portfolios.

14             But it must also be said that a lot of

15   things have changed, especially when it comes to

16   economics.  In order to guarantee long-term

17   viability of industry, we have responsibility and a

18   need to address the current industry economics.

19             [Slide]

20             Coming back to April, '98 when we were

21   dealing with shortage of immune globulins, which

22   was also the subject for a congressional hearing

227

1   and a lot of media attention, we have made a

2   commitment, and the commitment was that we will

3   start a data collection effort, originally on a

4   quarterly basis so that all parties will understand

5   the dynamics of supply.

6             What has happened in the meanwhile is that

7   we quickly moved to monthly reporting.  That

8   information is available since six and a half

9   years.  It is provided on an Association website

10   and is publicly available.  But also the situation

11   that we have right now is that we use a traffic

12   light system, green, yellow and red.  Green means

13   there is sufficient supply of inventory.  Yellow

14   means let's be careful, and red means we have a

15   problem.  We haven't had a red situation.  We had

16   two months of a yellow indication and what we have

17   done, we immediately informed the patients about

18   the situation so we have been very transparent and

19   open about the developments.

20             [Slide]

21             A couple of things that need to be said

22   are that we can, as an association, talk about

228

1   publicly available information.  Some of that has

2   already been shown by Dr. Weinstein.  But we need

3   to be careful because this is a very concentrated

4   industry and, because of that, we cannot sit

5   together with individual companies and agree on

6   what is going to happen with production trends.  We

7   have to respect anti-trust laws.

8             Also, we have to accept that certain

9   things are unpredictable.  We would like to know

10   everything in the world but that is simply not

11   possible so unpredictability is a factor that we

12   have to take into account.  We can only comment

13   about things that are within our control.  We are

14   only part of the entire chain.

15             But, as I said before, we do not see a

16   near-term threat to supply.  What we will see is

17   that individual companies will take actions to make

18   sure that the economic situation that we are facing

19   is going to be addressed, and I will explain to you

20   in the presentation later on why.

21             [Slide]

22             Don't be frightened by this slide, but I

229

1   have to do this.  I just want you to understand

2   what we can and what we cannot do.  The points that

3   are pointed out in red are things that are simply,

4   by law, illegal.  We cannot as an industry agree to

5   limit production or to reduce or increase

6   inventories.  We cannot coordinate output.  We

7   cannot allocate capacity.  We cannot determine

8   quotas.  We canon talk about continuation or

9   discontinuation of products, and we cannot limit

10   supply of particular products.

11             This is important because you can have

12   those conversations with individual companies but

13   once you are together you can simply not talk about

14   that.  And, if you have a lawyer in the room, they

15   will immediately stop that discussion.

16             [Slide]

17             Now, having said that, we want to be

18   cooperative.  We want to share as much information

19   as we can.  Therefore, I am going to tell you what

20   we believe the situation is.  Let's focus on some

21   of the products that are made from plasma.  As you

22   can see, I focus here only the major products,

230

1   alpha-1, Factor VIII, immune globulins, albumin.

2   As you can see, these are proteins that are

3   available in plasma in different amounts.  What you

4   see here are the quantities estimated industry

5   aggregate, but that may differ from company to

6   company and differ in the technology that is being

7   used.

8             [Slide]

9             In the fractionation process a combination

10   of temperature, pH, time and alcohol is going to

11   separate the different fractions that are used to

12   manufacture the different therapies.  I am not

13   going to go in detail here.

14             [Slide]

15             What you have seen, and Dr. Weinstein

16   already pointed that out, is that in the last

17   couple of years we have seen consolidations.  We

18   have seen companies making a decision to divest

19   their interests in the plasma industry.  We have

20   seen closures of plasma collection centers.  We

21   have also seen some closures of fractionation

22   plants.  As a result of that, there is a reduction

231

1   of the plasma that is used for manufacture, and

2   also we have seen massive layoffs, people that are

3   losing their jobs.  Of course, everybody would like

4   to avoid that but this is the reality of today.

5             [Slide]

6             The good news is that we have seen new

7   companies entering the market.  We have seen new

8   product approvals.  We have seen upgrades and

9   build-outs in the facilities.  We have seen better

10   technology.  We have seen technology nowadays that

11   allows us to obtain products with a higher yield

12   which means that you need less plasma, and we see

13   utilization of both source and recovered plasma as

14   starting material.

15             [Slide]

16             Going back to 1998, the problem that we

17   were facing is that the supply did not meet demand.

18   if you look at this graph, which basically points

19   out the true distribution of immune globulins in

20   the United States since January, 1998, the bottom

21   line is that on an annual basis, in 1998, we were

22   supplying about 14,000 kg into the U.S. marketplace

232

1   and today it is between 26,000 and 27,000.  I think

2   it is a very impressive improvement and it is a

3   clear demonstration that the commitment that the

4   industry made to act upon the shortage was turned

5   into practice.

6             [Slide]

7             What we need to understand is the

8   economics associated with the manufacture of these

9   therapies.  The collection of plasma, which is used

10   as a starting material, is not a static process.

11   It is a very dynamic process and it is dependent on

12   the type of products which you need what the driver

13   is for the total volume of collection plasma.  As

14   you can see on the left side of this slide, in the

15   past the quantity of albumin was the one that

16   really determined how much plasma was needed.  What

17   we see today is that it is the volume of immune

18   globulins that determines basically how much plasma

19   needs to be collected.

20             [Slide]

21             So, what I am going to do now, I am going

22   to try to explain to you with a model how this

233

1   really works.  In this model, I am going to explain

2   to you five possible therapies, immune globulins,

3   albumin, Factor VIII, alpha-1 and other proteins.

4             [Slide]

5             Here is how it basically works.  If you

6   look at this slide, you can see that a company or a

7   country or global there is a certain capacity that

8   you can use to manufacture your products.  If you

9   look at the bottom column, it shows you in this

10   particular case the sales of immune globulins.  As

11   I mentioned before, the driver for collection is

12   plasma, so you need more plasma to manufacture

13   immune globulins but you are also manufacturing

14   albumin.  What you can see in the yellow column is

15   the volume of albumin that is being sold in this

16   particular model.  Then, companies sell Factor

17   VIII, other proteins and alpha-1.  On the left side

18   you can see there is a revenue per liter if you

19   calculate what the revenue for the industry is.  As

20   you can see in the grey line, there is a cost to

21   manufacture and basically what you can see here is

22   that in order to recover the cost of manufacture,

234

1   you need to sell multiple products.  It is

2   impossible to make up your cost with the sale of

3   one product.  This may differ from company to

4   company based on their technology and so on.  But

5   the reality is that some companies do not

6   manufacture Factor VIII.  You can imagine that the

7   cost of manufacture then needs to be recovered by

8   less products.

9             [Slide]

10             The other thing that we need to understand

11   is that if we manufacture immune globulins, here,

12   as a result of the fractionation process,

13   automatically you also end up with more albumin,

14   more albumin than can be sold on the marketplace.

15   What that means is that this volume goes in

16   inventories.  There is no need for it; there is no

17   demand for it but there is a cost to manufacture.

18             So, the simple question that one could ask

19   is, well, why not resolve any discussion about the

20   potential shortage in the future just by making

21   more?

22             [Slide]

235

1             Well, this is what is going to happen

2   then.  What you can do is you can make more, but

3   you can see this is what is being sold.  So, now

4   you are going to manufacture more products which

5   are not needed on the market.  They end up in

6   inventories.  You also end up with more albumin and

7   you are losing money.  And, this industry is not in

8   a position to continue to lose money.

9             [Slide]

10             Look at what happened if we go back to the

11   year '98-99, and you see what the revenue per liter

12   was for the entire industry.  You see a significant

13   reduction of about 50 percent in income of the

14   industry as a result of multiple factors.  That is

15   the reason why industry needs to make adjustments

16   and ensure that it is healthy or becomes healthy

17   again.

18             [Slide]

19             One of the reasons is that this industry

20   is totally different than the pharmaceutical

21   industry.  In many cases we are compared to the

22   pharmaceutical industry but if you just look at the

236

1   cost of manufacturing, two-thirds of the production

2   costs are the result of the cost of source

3   material, the testing and manufacturing compared to

4   about 20 percent for the pharmaceutical industry.

5             As a result of that, if you take into

6   account total cost, I have seen other slides that

7   basically said that the profit margin of this

8   sector is somewhere between 5-10 percent if

9   everything works well.  But if you go to the

10   pharmaceutical industry, it is about 20 percent.

11   It is this profit margin we need in order to make

12   further investments in researcher and development

13   and that margin is simply not there.  So, we need

14   to be very, very cost conscious for the future.

15             [Slide]

16             But there are implications of differences.

17   One is that what people tend to forget is that we

18   are not making products for diabetes patients or

19   high blood pressure.  We are talking about a

20   relatively small number of patients that are using

21   these therapies which means that the cost needs to

22   be shared by a limited number, which is totally

237

1   different than what we are seeing in the

2   pharmaceutical industry.  Nevertheless, when we

3   talk about public policy, public health policy, we

4   are many times caught in measures for

5   pharmaceuticals because we are considered to be the

6   same.

7             We are providing biologicals.  The

8   technology that is being used is very specific,

9   very complicated.  We are certainly no generic

10   manufacturers that has a different product life

11   cycle development because those therapies are used

12   for decades, in contrast to the pharmaceutical

13   industry where 10, 15, 20 years maximum is a life

14   cycle.  And, we are working in a regulatory

15   environment which is the most stringent within this

16   sector.  We are not complaining about it.  What we

17   are saying is that this is the case and should also

18   be recognized when it comes to compensation.  The

19   working capital, the up-front investment is

20   enormous.  It takes about 7-8 months to get your

21   product out of your manufacturing site after

22   collecting your plasma--enormous investments.

238

1             The clinical trial requirements are very

2   complicated because we have a small number of

3   patients.  It is not that easy to get large number

4   of trials, especially when you do it in a

5   randomized, placebo-controlled fashion.  It is

6   complicated.

7             Then, what we see is an enormous pressure

8   on reimbursement and, if this is going to continue

9   like we have seen in the last couple of years, this

10   may lead to further industry consolidation, further

11   company exits--we have already seen that

12   sufficiently.  Then, it will reduce the choice of

13   therapy, which is not a good thing.

14             [Slide]

15             If we look at the revenue factors, and I

16   am not going to talk about the details but just to

17   point out, because I know this is not on the

18   charter of the BPAC but I think it is important for

19   you to understand what is going on, what we see,

20   especially when we talk about the reimbursement

21   areas, is if you look at the lower expenditures you

22   might think about issues like prospective payment

239

1   and the sales price as a basis for reimbursement,

2   or competitive bidding.  While they may sound nice,

3   there are really important differences between the

4   products you cannot use--change immune globulins

5   from one patient to the other, what happens then

6   is, once this is done on the federal level, we see

7   also state average sales prices used as a basis for

8   reimbursement.  We see other issues like prior

9   authorization.  We see single source contracts, and

10   we see cluster pricing.

11             With these examples, what we see is that

12   the private payers are going to use that to develop

13   their own measures and talk about what is the

14   coverage, what is the reimbursement, and you have

15   all kinds of managed care plans.  I am not going to

16   go into further detail here because I believe this

17   is an issue that most likely needs to be addressed

18   with the committee on blood safety and

19   availability, but it is important to realize that

20   what we can do is look at another factor which is

21   to take away, in my view, wrong perceptions.

22             [Slide]

240

1             If we talk about reimbursement of

2   therapies, a lot of people think that reimbursement

3   is what the manufacturers get.  That is absolutely

4   not the case.  It is not the case.  The

5   manufacturer is the one who has to put in the

6   testing, the plasma collection and fractionation

7   and there is an enormous amount of added value in

8   this chain but if you go back to the reimbursement,

9   it goes to the providers, it goes to the

10   distribution chain and then, at the end, it is what

11   the manufacturers get.  If you make an analysis

12   about what the margins are, I think there is some

13   room for some further investigation.

14             [Slide]

15             Now, what can we do?  If you look at the

16   costs, that is an area where we can have a

17   meaningful discussion.  I think it is important to

18   look at a couple of areas, the regulation

19   requirements; we can look at team biologics; we can

20   look at clinical trials.  Let me give you one

21   example of a regulation requirement that, in our

22   view, needs some reconsideration.

241

1             About 20 years ago, 25 years ago, there

2   was a requirement that every donor needs to be

3   weighed before plasma collection to determine the

4   amount of volume that could be drawn from the

5   donor.  In that period, when AIDS developed, there

6   was also a significant weight loss seen as an

7   indicator to point out that you needed to do

8   further investigation.  But everybody knows that

9   with the introduction of testing and more knowledge

10   about HIV, this is a requirement that doesn't make

11   a lot of sense.

12             So, FDA changed that, however, it is still

13   in the inspection protocols when centers are being

14   inspected and what happens now is that if you have

15   a donor that has a deviation in weight of 10 lbs or

16   more, you need to do an enormous amount of

17   investigation.  Well, look at me.  I would like to

18   lose 10 lbs.  It doesn't mean I am not healthy; I

19   am maybe even healthier.  So, I think this is a

20   kind of requirement that we should consider, and I

21   think that it would make sense if we could come

22   together with a group of experts, collection

242

1   experts, industry experts, FDA experts to see

2   whether we can make an analysis and find these kind

3   of requirements that don't really add a lot of

4   safety.

5             If you look at team biologics, I think FDA

6   and team biologics do a tremendous job going in to

7   make sure that companies are in compliance, and I

8   don't have to repeat the past but I think that we

9   have all learned a lesson and I honestly can say

10   that where we are today with quality and safety is

11   enormous if you compare that to ten years ago.

12   Maybe it is time to reconsider what needs to be

13   done, whether there are some redundancies in

14   systems that can be removed as well.

15             Then, when we talk about clinical trial

16   requirements, the difficulty is that because we

17   have a small number of patients, to set up the real

18   big trials is very difficult.  Sometimes we see a

19   discrepancy between products that are already

20   approved in Europe and are not approved in the

21   States because the requirements are different, or

22   vice versa.  The way the information is submitted

243

1   is different.  There is no consistency in format.

2   I understand from a conversation earlier today that

3   this may be subject to see whether there is any

4   room for harmonization.  These are the kind of

5   things that we need to look at.

6             [Slide]

7             So, what we need to do is we need to

8   continue with having a clear monitoring system, as

9   I explained in the beginning, while we continue to

10   measure inventory and distribution.  We do have an

11   emergency plan for supply in place.  In 1998 we

12   made a commitment.  If you look at the data you

13   will see that every month there is an emergency

14   supply available.  In case patients need it

15   immediately, there is always a quantity made

16   available by the industry to make it available for

17   patients, but also healthcare professionals because

18   we are not dealing with patients alone.  Healthcare

19   professional organizations are very important.

20             Then what we do, we have regular

21   communications with all stake holders.  We have

22   four meetings per year, here in the United States,

244

1   where we meet with consumers; where we meet with

2   physicians; where we meet with regulators; where we

3   have open dialogue about all these developments and

4   in case we see a deviation that requires attention

5   we have to form to do so.

6             [Slide]

7             In conclusion, we do not see a near-term

8   threat to the availability of plasma therapies.  I

9   think that is a most important message.  We believe

10   that the current inventories and also the change

11   through technology that allows us to use higher

12   yields will most likely offset reductions in the

13   throughput that we have seen earlier from Dr.

14   Weinstein.

15             We will remain vigilant because it is our

16   task to make sure that this industry is healthy.

17   The health of the patients and the health of the

18   industry go together.  We can't do it without each

19   other, and we cannot afford to continue to

20   manufacture products and lose money.  That is not

21   in the interest of the industry but, most

22   importantly, not in the interest of the patients.

245

1   We will make sure, as an industry group, that we

2   make corrections.  If we do that, then we can

3   guarantee further access, choice and continue with

4   the profits that are necessary to continue to work

5   on innovation.  Thank you very much for your

6   attention.

7             DR. NELSON:  Thank you.  Any comments or

8   questions?

9             DR. DIMICHELE:  Thank you for that.  That

10   was very informative.  I wanted to ask a question

11   about the difference in the sort of profit margin

12   structure between biologics and pharmaceuticals.

13   My concern is that many of the plasma industry

14   companies or the plasma manufacturers are now part

15   of conglomerates that include pharmaceuticals.  To

16   what extent will the economics of the

17   pharmaceutical industry threaten the companies or

18   the parts of the companies that produce plasma

19   products?

20             MR. BULT:  It is very difficult to go into

21   company specific situations but, again, going back

22   to 1998 where we had a situation in the United

246

1   States where basically the market was supplied by

2   four companies and the American Red Cross, three of

3   those four companies no longer exist.  One is gone,

4   another is taken over, another is in the process of

5   divesting.  So, I believe we have gone through a

6   process of adjustments.

7             What I do see is that companies are very

8   aware of the economic situation and the corrections

9   that need to be made, and one of the things is that

10   as a result of the commitment that we made in '98

11   to produce more and more and more, which was good,

12   now we have a situation where on some occasion you

13   may even talk about over-supply, and if we have

14   over-supply that is not really helpful.  So, I can

15   see the point of some of the companies, in order to

16   make these economic adjustments, will tighten

17   supply to balance demand and supply.  I think that

18   is what is going to happen.

19             DR. ALLEN:  Looking at the series of

20   slides you showed on the cost to manufacturer and

21   then the revenue generated by the different

22   products, I was a little surprised to see the

247

1   immune globulins right at the bottom.  I would not

2   have predicted that.  In terms of the cost to the

3   manufacturers, is there a significant difference

4   between the intravenous immune globulin products

5   and the injectable immune globulins?

6             MR. BULT:  You should not draw any

7   conclusion about the order of the therapies.  The

8   only conclusion that you should draw is that if you

9   talk about fractionated plasma that is needed the

10   driver is immune globulin and, from that point of

11   view, the place in that graph is absolutely the

12   correct place.  The vertical access is absolutely

13   not an indicator of what the true revenue is

14   because that differs from company to company.  What

15   was important to me is to share with you the

16   concept and the model to help everybody understand

17   how it works and how complex it is.

18             DR. KLEIN:  I always worry a little bit

19   when I see fractionation plants closing down in

20   terms of response time.  I pretty much have a good

21   idea of how long it takes to put collection sites

22   and license them but if you could give us an idea,

248

1   first of all, of how long it takes to build and

2   qualify a fractionation plant, and where we stand

3   right now in terms of fractionation capacity

4   worldwide.

5             MR. BULT:  The numbers that we have always

6   used in the past are that in order to get a plant

7   in function after going from the drawing board is

8   about five years, building it up and going through

9   the validation processes, etc., etc.  When you talk

10   about collection, yes, of course it is true that

11   when you close your centers, and if a situation

12   occurs where you need to have more plasma, well, I

13   haven't spoken to the individual companies but I

14   would expect that they have not closed the most

15   efficient centers.  I would expect that they have

16   closed centers where there were some questions

17   about efficiency.  In case we need more plasma it

18   is easier to work with efficient centers to improve

19   that and increase it than start all over again.

20             The other thing is that if you look at the

21   technology that is being used, and I mentioned that

22   in the presentation, most of the companies that I

249

1   know use the technology that obtains higher yields,

2   which means that you use less plasma.  So, that is

3   another important contributing factor.

4             DR. KLEIN:  Do you have an idea about the

5   fractionation capacity worldwide in terms of

6   current needs and what there is in reserve?

7             MR. BULT:  I don't have those numbers in

8   my head but maybe there is someone in the audience

9   that could help me out here and give me a number,

10   if that is available, but I don't have it at hand

11   here.

12             DR. KLEIN:  I guess what I am asking is if

13   we needed a ton of anthrax immune globulin or a

14   whole host of other things that one could guess at,

15   would we be in a position to manufacture it?

16             MR. BULT:  Well, I think if we start

17   talking about the anthrax issue, I mean that opens

18   up a whole--

19             DR. KLEIN:  I was only using that as an

20   example.

21             MR. BULT:  Yes.  I do believe that if we

22   talk about existing products where you need to

250

1   expand capacity, that is easier than to develop new

2   products where you have to go through the whole

3   process.  So, if we talk about existing products,

4   that is easier to do.  But, again, I have no

5   indication that there is a near-term threat to

6   supply or availability.  I think that is the most

7   important message.

8             DR. KUEHNERT:  Looking at your slides on

9   supply trends and also distribution in the

10   handout--and maybe this is more of a question for

11   FDA also, but how are not only supply but demand

12   trends followed?  I mean, what system is there to

13   determine sort of what the supply is in relation to

14   demand in the United States on this products?

15             MR. BULT:  Let me make a couple of

16   comments about the slide that you have seen.  The

17   numbers that you have seen on the slide are what we

18   call a true supply into the market.  Let's take one

19   example if you go to the first dot on the slide,

20   which is January, '98, the way this was constructed

21   with the help of experts on economics, they have

22   taken the production in the month of January; they

251

1   looked at the inventory on February 1st, and

2   deducted the inventory of January 1st.  So, what

3   that means is that you have the true supply in the

4   market.

5             Now, what you have seen cannot be a true

6   amount because there are too many fluctuations,

7   especially when you look on a month to month basis.

8   What we do is we provide this information to FDA so

9   FDA has full insight on all the information company

10   by company and what you see here is aggregated

11   information.  But if you combine that information

12   we believe we can say, with a high degree of

13   confidence, that the annual gross of immune

14   globulins is somewhere between 4-6 percent per

15   year.

16             DR. HOLLINGER:   What is the expiration

17   time of IVIG, for example, how long?

18             MR. BULT:  Sorry, I couldn't hear the

19   question.

20             DR. HOLLINGER:  The expiration time for

21   IVIG?  How long?

22             MR. BULT:  I think I see a sign in the

252

1   back, three years.  That is the shelf life.

2             DR. HOLLINGER:  Three years?

3             MR. BULT:  Yes.

4             DR. HOLLINGER:  How about specialty IVIG

5   products, like HcIG, HbIG, new bioterrorism

6   products and so on that might be manufactured for

7   the military, and things of that nature?  And

8   albumin?  I didn't see the numbers for albumin but

9   I can tell you it seems like albumin is being used

10   a lot more than it has in the past, medically.

11             MR. BULT:  Yes.  Coming back to your

12   question, it is a while since I, myself, was

13   working in the industry but I recall from those

14   days that there was no difference between the

15   shelf-life of hyper-immunes and polyvalent immune

16   globulins.  What you do see is that it is dependent

17   on the regulations of the country where you

18   distribute.  Sometimes you can have an expiry data

19   that goes for months and you have 36 months.  There

20   are other countries where it is only twice a year.

21   So, you may have a difference of 6 months but that

22   is just a technical detail.  When it comes to

253

1   albumin, I think albumin goes up to 5 years, if i

2   am correct but I have seen different stability

3   numbers.

4             DR. STRONG:  I was interested in your

5   comment about harmonization.  Those of us who

6   provide plasma for recovered plasma, we are

7   noticing an increasing conflict between European

8   standards and American standards.  Even within

9   Europe there are country to country variations, for

10   example, a requirement for ALT testing in some and

11   not others.  What is your forecast for how

12   harmonization is going, and do you think that there

13   is a role here for us to do more?

14             MR. BULT:  You know, if you would have

15   asked medication that question two years ago I

16   would have said everybody wants to harmonize if you

17   do what I do.  But today I would say I see much

18   more willingness among regulators to really start

19   to realize what the differences are.  I already

20   mentioned before that in a conversation with Dr.

21   Weinstein we spoke about a potential workshop on

22   clinical trial requirements.  We have a dialogue

254

1   with regulators all over the world, from the United

2   States, from Europe, from Australia, to identify

3   those issues to see how we can help each other to

4   understand and work together to create better

5   harmonization.

6             A part of the problem that we have is that

7   this sector is not part of the ICH process, the

8   International Conference of Harmonization.  This is

9   true for the pharmaceutical sector.  Because of the

10   specific nature, we were excluded.  But I see far

11   more willingness of regulators to work with us and

12   to find ways to harmonize.

13             DR. WEINSTEIN:  I should mention that we

14   are having actually a workshop on that very topic

15   to study standards for recovered plasma that will

16   be held on August 31, September 1st at the NIH.  We

17   have invited speakers from Europe, from Australia,

18   from Canada to get international perspective on it.

19   We are looking for the scientific basis of these

20   standards.  We are looking for practicality and the

21   need for adopting certain standards.  So, we are

22   moving ahead on that.

255

1             DR. ALLEN:  Two questions, building on the

2   harmonization issue and also looking at what I

3   anticipate will increasingly become an enhanced

4   global demand for these kind of products, much of

5   it supplied by the plasma, I would assume,

6   collected in the United States and European

7   countries or certainly in North American and

8   Europe.  What is the anticipation in the industry

9   for the growth of global market sales, and are they

10   going to be able to meet that both in terms of

11   plasma collection as well as in terms of

12   manufacturing tank capacity?

13             MR. BULT:  If you look at the global

14   numbers, what I do know is that if you listen to

15   the consumer organizations you know that there is

16   an enormous amount of patients in the world who are

17   diagnosed.  For example, hemophilia patients, only

18   25 percent of the world population receive any form

19   of treatment.  If you look at immunodeficiency as

20   an example, here in the United States it is very

21   well developed but there are other parts of the

22   world where the awareness is not as high and the

256

1   diagnostic ability is not as high as here in the

2   States.  The same is true for alpha-1 for example.

3             It all depends on other factors such as

4   the economic situation in a country.  There are

5   certain countries that are crying for help but

6   there is no funding available to make that

7   available.  But if you look worldwide, as I

8   mentioned, for immune globulins worldwide there is

9   about a 4-6 percent increase for those countries

10   where immune globulins are being used.  For the

11   rest, it is dependent on individual companies and I

12   cannot talk about these issues with individual

13   companies because of the anti-trust laws.

14             DR. ALLEN:  The second question I have is

15   what percentage of the Factor VIII concentrate used

16   in the United States and Europe today is

17   recombinant?  You mentioned new technologies making

18   allowing more efficient extraction of fractionation

19   products from plasma.  To what extent do you

20   visualize that new technologies in other arenas,

21   particularly recombinant products perhaps, will

22   replace the need for plasma-derived products?

257

1             MR. BULT:  Well, I originally had a slide

2   for this presentation which identified which

3   therapies could be replaced with alternative

4   technology, like recombinants and currently it is

5   Factor VIII, Factor IX and Factor VII.  If I focus

6   on Factor VIII, I would say that in the United

7   States the last number I have seen is about 75

8   percent, 80 percent use for recombinant Factor

9   VIII.  When I go to Europe it is about 50 percent.

10   If I go to Japan it is about 50 percent.  If I go

11   to less developed countries there is a high use of

12   plasma-derived, mainly for cost reasons.

13             DR. DIMICHELE:  Just to say, yes, indeed,

14   most of the product that is used, or a large

15   proportion of it for hemophilia A and B is

16   currently recombinant.  But one of the things that

17   has become really apparent to those of us who treat

18   people with bleeding disorders is that there is a

19   vast array, you know, of other bleeding disorders,

20   most of which are more rare for which there is not

21   adequate therapy, and any therapy that is

22   potentially available, maybe with the exclusion of

258

1   Factor VII, has pretty much been plasma derived.

2   Also, within hemophilia A and B there are several

3   reasons to still need plasma-derived products with

4   respect to bypassing agents and certain

5   therapeutics.

6             So, one of the concerns that we have

7   certainly in the bleeding disorders field is that

8   the plasma industry continue to produce sufficient

9   supply and that we work with regulators and the

10   plasma industry to develop new technologies, new

11   treatments for the rare disorders.  And, you know,

12   with respect to modifications of clinical trials,

13   requirements, orphan drugs are the sort of things

14   that would sort of make these therapies accessible.

15   So, there is a lot more that we have to do.  The

16   whole problem of clotting disorders hasn't been

17   solved by recombinant technology and we have a lot

18   more work to do, and it still involves

19   plasma-derived therapies.

20                       Open Public Hearing

21             DR. NELSON:  In the open public hearing

22   Mr. Patrick Schmidt wanted to comment.

259

1             MR. SCHMIDT:  Good afternoon.  I will

2   speak very quickly because I want to make sure I

3   don't make me miss my plane.  Good afternoon,

4   distinguished ladies and gentlemen of BPAC.  My

5   name is Patrick M. Schmidt and I am the president

6   and CEO of FFF Enterprises.  We are the nation's

7   largest distributor of intravenous immune globulin

8   and human serum albumin.  Approximately 70-80

9   percent of the hospitals in the U.S. rely on the

10   distribution and management expertise of our

11   company as the primary source for these critical

12   care products.  Countless other physician offices,

13   home care companies and patients rely on us as

14   well.

15             From this unique perspective we see a new

16   marketplace reality unfolding.  The new marketplace

17   reality is characterized by rising prices,

18   declining inventories and a necessity to increase

19   vigilance on our industry's ongoing ability to meet

20   the needs of the patients we serve.

21             The plasma derivative market is very

22   complex.  It is characterized by steadily growing

260

1   demand for most products, met by the challenges of

2   a very difficult, highly regulated and

3   capital-intensive manufacturing environment.  We

4   encourage all participants in the marketplace to

5   recognize and understand this dynamic nature.  We

6   have a saying within FFF once you think you

7   understand the marketplace, you end up realizing

8   you don't because it changes so rapidly.

9             The economics of plasma fractionation is a

10   study of supply and demand.  An example, just in

11   the past five short years, the price of albumin has

12   declined from the mid-50s to the mid-teens.  During

13   the same period of time IVIG has vaccinated from

14   the low-30s to the upper-50's and back again.

15   While tremendous savings have been enjoyed by the

16   healthcare provider, those who understand the

17   financials associated with running a plasma

18   derivative business knew that in the long-term not

19   all manufacturers would survive this downward

20   spiral.

21             If you remember back to 1988, the two

22   largest group purchasing organizations in the

261

1   United States, Premier and Novation, each had

2   contracts with seven different suppliers of IVIG

3   and albumin.  They were Alpha, Baxter, Centeon,

4   ARC, Bayer, Immuno and Novartis.  Four of those

5   seven suppliers no longer exist.  They have been

6   replaced by new business combinations or new market

7   entrants like ZLB-Behring, Grifols or Octapharma,

8   as has been discussed before.  Sort through this

9   myriad of changes, mergers and acquisitions,

10   consolidations and plasma collection center

11   closures, and you are left with a new marketplace

12   reality--fewer suppliers.

13             The new look of today's marketplace will

14   feature three meta-suppliers that have global

15   reach, as was touched on earlier.  They have the

16   complete portfolio of products from the precious

17   alpha-1 therapies to the workhorse albumin

18   products.  Today's marketplace economics dictate

19   that the necessity that the manufacturers generate

20   revenue from multiple products per liter.  I

21   believe the days that manufacturers fractionate

22   plasma for the purpose of producing a single

262

1   product, like IVIG, are gone.  It just doesn't make

2   economic sense any longer.

3             Worldwide cuts are necessary from my

4   perspective and will result in a tightened supply

5   situation that will ultimately lead to shortages in

6   the IVIG marketplace and perhaps albumin market,

7   which is far less conceivable at this time.

8             This environment demands a supply solution

9   and distribution model that can manage rapidly

10   changing dynamics, a distribution model that

11   provides stability for both supply and pricing.  We

12   must provide consumers a responsible distribution

13   channel that delivers the most reliable IVIG and

14   albumin supply in the United States today

15   regardless of the manufacturing changes.

16             In anticipation of a tightened supply, we,

17   at FFF, have already place four lyophilized IVIG

18   products on supply guideline status.  That means we

19   have put filters in our internal ordering systems

20   that prevent customers from ordering quantities

21   that far exceed their previous utilization

22   patterns.

263

1             But please be aware--it is very important

2   to emphasize this--there is plenty of IVIG on the

3   marketplace today to meet the needs of the current

4   patient demand.  Please also be aware that this

5   situation is highly likely to change in the future

6   and a great deal of collaboration that currently

7   exists between the manufacturers and distributors

8   of these products must continue to avoid 1998-like

9   shortages.  Of course, the FDA is proactively

10   seeking to understand and mange this industry's

11   complex challenges and that is why we are here

12   today.

13             What role can the FDA assume to help

14   manage the transitioning marketplace and protect

15   the continuity of care?  We have three brief

16   suggestions we wish to make to help accomplish

17   that.  They are, first, prioritize or expedite any

18   manufacturing process change applications that

19   improve yields.  Second, engage in discussions with

20   other federal agencies, like CMS--and I realize

21   that may not be the purview of this committee, to

22   ensure that the appropriate reimbursement rates for

264

1   these critical products exist.  Often the greatest

2   barrier to access to products is economic.  Lastly,

3   consider implementing a pedigree requirement for

4   certain plasma products.  Six different IVIG

5   preparations exist on nationally published lists of

6   the 31 drugs most susceptible to counterfeiting.

7   IVIG products have been counterfeited, diverted and

8   subject to the voracious pricing policies of the

9   secondary grade markets.  Whatever you decide, I

10   encourage you to consider that the only ongoing,

11   viable solution must require the utmost in

12   distribution channel responsibility and

13   collaboration, managing the plasma product supply,

14   and necessitates the collective commitment of the

15   manufacturer, the distributor and the public

16   representatives, yourselves.

17             As our company has been involved in past

18   challenging marketplaces and markets surrounding

19   plasma derivatives, we remain ready and willing to

20   share data and perspective as the new market

21   reality crystallizes.  So, thank you very much for

22   the important work you do for our industry and for

265

1   the opportunity to address the committee today.

2   Thank you very much.

3             DR. NELSON:  Thank you.  Are there any

4   questions or comments?  Harvey?

5             DR. KLEIN:  I am not quite sure I

6   understood one of your statements.  If there is

7   plenty of IVIG why would you have filters in your

8   distribution system not to sell people as much as

9   they either want or need?

10             MR. SCHMIDT:  Well, because we are seeing

11   our ability to replenish at the same rate we were

12   purchasing before, and we are seeing declining

13   inventory so our ability to replenish is going

14   down.

15             DR. KLEIN:  So, I guess there isn't plenty

16   of IVIG.  Is that not correct?

17             MR. SCHMIDT:  At this time there is plenty

18   of IVIG.  I think the question that hasn't been

19   answered is what is the near-term, how long is the

20   near-term.  While we have plenty of IVIG, we are

21   beginning to proactively monitor the situation so

22   we don't get into an exacerbated situation of what

266

1   has happened in the past with decentralized, if you

2   will, management of these products.  We have 6,000

3   hospitals in the U.S. and if they believe there is

4   going to be a shortage, they train management to

5   increase their levels and possibly exacerbate

6   shortages.

7             DR. NELSON:  Jay?

8             DR. EPSTEIN:  Thank you, Patrick.

9             MR. SCHMIDT:  Thank you very much.

10             DR. EPSTEIN:  Could you just comment why

11   do you think that IVIG products have been more

12   susceptible to counterfeiting than some other

13   products?  What is the driver?

14             MR. SCHMIDT:  I believe the driver, and

15   the criteria for being included in the list, is the

16   fact that one or two of the drivers have already

17   been counterfeited.  There was an issue of gamunin

18   [?] that was counterfeited.  Then, IVIG at times

19   has been highly sought after and people will do

20   almost anything to get their hands on IVIG in a

21   real tight marketplace because it is a critical

22   care drug and people, unfortunately, die if they

267

1   don't have it.  We have experienced times in our

2   business where people call and order IVIG without

3   any consideration of cost.  I would also like to

4   say we have never taken advantage of that situation

5   but there people who have.

6             DR. GOLDSMITH:  I probably should make

7   this statement from there because I put my IDF hat

8   on for a minute.  I guess from the perspective of

9   the Immune Deficiency Foundation, I would like to

10   recognize the members of industry for being so

11   frank today, coming forward and speaking about what

12   they are doing, what their plans are.  I think I

13   heard some promises there that they would maintain

14   supply that was kind of laced into all the bad

15   news.

16             I want to also recognize the fine work

17   that has been done by the FDA in terms of bringing

18   three new IVIG products to the U.S. market in the

19   last year, which is really an outstanding

20   accomplishment and shows that regulatory wheels can

21   grind fast when they need to, to respond to this.

22             I guess I would like everybody to

268

1   recognize the fact that IVIG is a life-saving drug

2   for people with immune deficiency and it must be

3   available on the U.S. market.

4             DR. NELSON:  Thank you.  Dr. Carol Lee

5   Koski, University of Maryland?

6             DR. KOSKI:  Thank you.  I am actually a

7   neurologist that heads the Neuromuscular Division

8   at the University of Maryland.  I also am on the

9   medical advisory board for the Guillain-Barre

10   Foundation.  Although certainly immunodeficiency I

11   think is a significant indication for intravenous

12   immunoglobulin, I think it is also important to

13   remember that there are actually other neurological

14   patients, including those with Guillain-Barre

15   syndrome, chronic inflammatory demyelinating

16   polyneuritis and also myasthenia gravis where

17   immunoglobulin is actually used to prevent and

18   limit their paralysis.  It allows these patients to

19   not only survive but to actually go on being

20   productive members of society, able to support

21   their families and do their daily activities of

22   living.  So, don't forget the patients that may not

269

1   have an FDA indication for that particular disease.

2             DR. NELSON:  Harvey?

3             DR. KLEIN:  I just want to know are you

4   getting all of the IVIG that you need for your

5   patients at the University of Maryland?

6             DR. KOSKI:  Yes, we do.  But one of the

7   things that I will also tell you is that during the

8   shortage I did not have any problems in getting

9   intravenous immunoglobulin, but that was partially

10   because of the fact that our hospital pharmacy was

11   able to buy very large lots and was one of the

12   major suppliers.  But I very frequently had

13   patients that were referred to me from community

14   hospitals, that probably would have been treated in

15   that community hospital, because they simply didn't

16   have the intravenous immunoglobulin.

17             DR. GOLDSMITH:  May I make a quick

18   comment?  One of the things we need to watch--I

19   agree that there are certain neurological diseases

20   that are have clearly benefitted from IVIG and I

21   think the FDA, when it responded to the crisis, did

22   point out that there was a large amount of the IVIG

270

1   that was going for off-label use and it was up to

2   60 percent.  These are estimates made by large

3   medical centers, and part of the way we dealt with

4   the crisis was by interacting with practitioners at

5   big medical centers and indicating to them that

6   they need to form some kind of filter for how they

7   were using immune globulins.  We did not want to

8   curtail use for patients that had legitimate

9   diseases where there is a consensus that they

10   respond to IVIG.  But we identify, the community at

11   large identified a large number of patients or

12   patient groups that were being treated for causes

13   for which there was no evidence--so, something we

14   need to watch in the future in terms of

15   availability is off-label use, and make sure that

16   that off-label use is based on evidence that the

17   IVIG works for those patients.

18             DR. NELSON:  That is an important issue.

19   It is a therapy that could be, and I guess has

20   been, abused to some extent.  Yes, Dr. Koski?

21             DR. KOSKI:  Just to really reaffirm that

22   the particular indications that I was talking about

271

1   actually are really supported very heavily,

2   particularly Guillain-Barre syndrome, with

3   absolutely outstanding trials that comprise almost

4   600 patients, that were done in multicenters across

5   really Europe and North America.

6             I agree with you that there are perhaps

7   very, very rare disorders where some of the

8   randomized, controlled trials, such as stiff person

9   syndrome, comprise only about 15 patients.  But

10   some of those trials are really still extremely

11   good, with very high efficacy.

12             DR. NELSON:  But there have been some

13   proposed used for which the data were pretty shaky.

14   At one point they consumed a fair amount of the

15   supply, as I recall.  So, I think the point is well

16   taken that this is an important commodity,

17   particularly if it is in short supply its use

18   should be supported by good scientific data.

19             Well, thanks.  I guess that is the end of

20   the session.  I guess October 21 and 22 will be the

21   next meeting.  Thanks.

22             [Whereupon, at 3:10 p.m., the proceedings

272

1   were adjourned.]

2                              - - -