November 21, 1997

Contract # TW-0928

SPECIAL SPECIFICATION

SECTION 11640 -S

MISCELLANEOUS LABORATORY EQUIPMENT

Page

PART 1 - GENERAL

1.01	Summary	2

1.02	Submittals and Substitutions	2

1.03	Quality Assurance	3

1.04	Delivery, Storage, and Handling	3

1.05	Project Conditions	3

1.06	Warranty	3

PART 2 - PRODUCTS

2.01	Flammable Storage Cabinets	3

2.02	Tall Metal Storage Shelving 	4

2.03	Balance Tables	4

2.04	High Purity Water Units	5

2.05	Modular Gas Cabinets	5

PART 3 - EXECUTION

3.01	Installation	8

3.02	Adjusting	8

3.03	Cleaning and Protections	9

SPECIAL SPECIFICATION

SECTION 11640-S

MISCELLANEOUS LABORATORY EQUIPMENT

THIS SECTION USES THE TERM "ARCHITECT."  CHANGE THIS TERM AS NECESSARY TO MATCH THE ACTUAL TERM USED TO IDENTIFY DESIGN PROFESSIONAL AS DEFINED IN THE GENERAL AND SUPPLEMENTARY CONDITIONS.

PART 1 - GENERAL

1.01	SUMMARY

A.	Section Includes:

1.	Flammable Storage Cabinets

2.	Tall Metal Storage Shelving

3.	Balance Tables

4.	High Purity Water Units

5.	Modular Gas Cabinets

B.	Related Sections:

1.	Division 15:  Furnishing and installation of plumbing utilities and final

connections to service fittings.

2.	Division 15:  Furnishing and installation of exhaust ductwork and equipment, and final connection to exhaust devices.

3.	Division 16:  Furnishing and installation of electrical utilities and final connections to casework and accessory equipment.

1.02	SUBMITTALS and SUBSTITUTIONS:

A.	Product Data:  Submit manufacturer's data for each component and item of laboratory equipment specified. Include component dimensions, configurations, construction details, joint details, and attachments. Indicate location, size and service requirements for each utility connection.

B.	Manufacturer's operating and maintenance manuals.

C.	Submit the above in accordance with section 1300-S.

D.	Substitutions on any of the listed manufacturers in this specification shall be performed per the requirements of section 1600-S.

1.03	QUALITY ASSURANCE

A.	Manufacturer shall have at least 5 years experience in the fabrication of the specified equipment and shall have 10 installations of equal or larger size.

1.04	DELIVERY, STORAGE, AND HANDLING

A.	Delivery:  Schedule delivery of equipment so that spaces are sufficiently complete that material can be installed immediately following delivery.

B.	Handling:  Protect finished surfaces from soiling or damage during handling and installation. Keep covered with polyethylene film or other protective coating.

1.05	PROJECT CONDITIONS

A.	Do not deliver or install equipment until the following conditions have been met:

1.	Windows and doors are installed and the building is secure and weathertight.

2.	Ceiling, overhead ductwork, and lighting are installed.

3.	All painting is completed and floor tile is installed.

1.06	WARRANTY

A.	Provide a 1-year warranty against defects in materials and workmanship.

PART 2 -  PRODUCTS,    Substitutions: only under provisions of Section 1600-S.

2.01	FLAMMABLE STORAGE CABINETS

A.	Description:  Design and construct in accordance with OSHA Regulations and the requirements of NFPA 30, National Fire Protection Association, Flammable and Combustible Liquids Code. Cabinets shall limit the internal temperature at the center 1 inch from the top to not more than 325 degrees Fahrenheit when subjected to the fire test for storage cabinets as described in NFPA 30.

B.	Casing:  Bottom, top, back, door, and sides of cabinet shall be at least 18 gage sheet steel, double walled with 1-1/2 inch air space. Joints shall be welded airtight. Provide with adjustable zinc plated leveling legs.

C.	Door:  Provide with continuous piano hinge and a 3-point latch arrangement with door sill raised at least 2 inches above the bottom of the cabinet to retain spilled liquid within the cabinet. Door shall be self closing.

D.	Ventilation:  Cabinet shall not be ventilated unless required by local authorities. Vent openings shall be sealed with plugs, inside and outside, supplied by the manufacturer.

E.	Ventilation:  Provide a 2 inch threaded vent connection on each side of the cabinet for supply and exhaust ventilation air. Vents shall have flash arresters.

F.	Shelving:  Provide with three 18 inch deep, full width adjustable shelves.

G.	Grounding:  Cabinets shall be grounded.

H.	Identification:  All flammable storage cabinets shall be marked with conspicuous lettering:  FLAMMABLECKEEP FIRE AWAY.

I.	Finish:  Finish as specified for metal laboratory casework except interior shall be triple epoxy coated. Color selected by the Owner's Representative.

2.02	TALL METAL STORAGE SHELVING

A.	Manufacturers:  V-Grip Shelving/Storage Systems as manufactured by Equipto Shelving; Aurora, Illinois  60507. 765-7 Series, or approved equal.

B.	Size:  7 feet high with 5 adjustable shelves with top and bottom shelves fixed. Widths and depths as indicated on plans.

C.	Construction:  Shelving shall be a modular assembly of 13 gage steel uprights and 18 gage reinforced adjustable shelves painted in stock colors. Units shall have reinforced backs and ends.

2.03	BALANCE TABLES

A.	Manufacturer:  Waller Brothers Stone Company, McDermott, Ohio; VWR Scientific, PO Box 13007, Station K, Atlanta, Georgia  30324, (404) 262-3141; Fischer Scientific, 3315 Winton Road, PO Box 11666, Raleigh, North Carolina  27604; or approved equal.

B.	Size:  35 inches long by 30 inches high by 24 inches wide.

C.	Construction:  3 inch thick black epoxy impregnated sandstone slabs. Vertical supports shall be held in place by a metal pipe with end flange. Provide two 1/2 inch thick rubber isolators for placement between leg slabs and floor.

2.04	HIGH PURITY WATER UNITS

A.	Manufacturers:  Millipore Corporation, Bedford, Massachusetts or Barnstead/Thermolyne Corporation, Dubuque, Iowa.

B.	Type:  Wall mounted unit capable of producing 1.5 liters of Type I, reagent grade water meeting CAP Type I and ASTM Type I Specifications, 10 to 18 megohm-cm resistivity at 25 degrees Celsius.

C.	Construction:  Unit shall consist of a prefilter cartridge, an activated charcoal cartridge, two deionization cartridges, and a final filter. Provide with pump, inlet pressure regulator, resistivity sensor, and all valves, piping and controls for automatic operation.

D.	Electrical:  120 volt, 60 Hertz, single phase, 1.0 amp.

2.05	MODULAR GAS CABINETS

A.	Manufacturers:

1.	Spectra Gases, Inc.

277 Coit Street

Irvington, New Jersey 07111

(800) 932-0624

2.	Semi Gas Systems, Inc.

Division of Hercules Inc.

625 Wool Creek Drive

San Jose, California 95112

(800) 227-7468

3.	Scott Specialty Gases

Route 611

Plumsteadville, Pennsylvania 18949

(215) 766-0320

B.	Cabinet Construction:

1.	Each gas cabinet shall be constructed of a minimum of 12 gage steel. The cabinet shall be welded and coated with a corrosion resistant epoxy paint.

2.	Door:  Full length door with automatic closer and a flush mounted door handle with lock.

3.	Window:  Manufactured of either 1/2 inch wire reinforced safety glass or 0.25 inch Lexan. The window shall be hinged with automatic closer and key lock.

4.	Adjustable cylinder brackets and straps shall be provided for each cylinder the gas cabinet can hold.

5.	Shelves:  Vertically adjustable to accommodate different sizes of standard gas cylinders.

6.	Color selection by the Owner's Representative from one of the manufacturer's standard colors.

C.	Cabinet Accessories:

1.	Sprinkler Head:  UL approved, rated at 165 degrees Fahrenheit.

2.	The inlet air supply in the bottom of each door shall have a filter screen and a fusible link to close it off upon the entry of high temperature air.

3.	Each gas cabinet shall have a rubber mat on the floor of the cabinet.

4.	All wiring and tubing in the cabinets shall be standardized and color coated.

5.	Provide bulkhead connectors for all lines entering or exiting the cabinet that will enable a continuous tube run without breaks.

6.	Exhaust air shall enter at the bottom of the cabinet and sweep the complete interior of the cabinet to the smooth nonrestricting gas transition at the top of the cabinet. A 6 inch vent connection shall be provided at the top of each cabinet for connection to the building exhaust system.

D.	Piping Purge Assembly:

1.	All piping in the gas cabinet shall by Type 316L stainless steel electro-polished tubing with an internal surface finish better than 15 RA. All permanent connections are to be butt-welded. Removable components in the process line to have VCR connections. Nupro 1/4 turn diaphragm valves shall be used and shall have color-coded handles. Process line valves shall have red handles, purge line valves shall have green handles, and vent valves shall have white handles.

2.	A flow limiting orifice shall be permanently installed in the CGA connection of the pigtail. The pigtail shall be a deep purge configuration on the process gas panel. Design shall allow connection to cylinders ranging in height from 48 to 56 inches. Purge nitrogen shall be introduced into the process system at the CGA connection. The pigtail assembly shall be securely anchored to the panel and supported at the CGA connection by a roof-mounted chain. The CGA connection shall be protected when not in service with a stainless steel dust cap and a securing chain. A purge line check valve shall be provided to prevent decontamination. Provide 316L stainless steel electronic regulators for the process panel.

3.	Provide high pressure vent check valve to prevent back-contamination of the process gas-wetted surfaces and low pressure vent check valve to prevent backflow into the process lines. Provide vacuum venturi generator to generate 20 to 25 inches Hg vacuum at the CGA connection. Provide dual in-line venturi generator nitrogen check valves to prevent back-contamination into house nitrogen system.

4.	A 100 psig safety relief valve to protect against over-pressurization of process line. Excess flow device shall be provided. In-line high pressure filter shall be provided. Process gas check valve shall be provided. Provide pneumatically actuated automatic shutoff valve on process pigtail. All major components shall be removable from the mounting panel without detaching the panel from the cabinet.

5.	The process panel shall have a high pressure transducer and low pressure transmitter for closed loop control. All valve openings and closings shall be verifiable from independent pressure sensors. Provide pneumatic door interlocking system to prevent cabinet door from being opened unless the system has been safely purged and the process gas cylinder valve has been closed. Startup of the system shall be prevented unless the door is closed and latched. Critical alarms shall be hardwired in addition to going through the microprocessor.

6.	Provide simple autopurge operation with 3 push buttons which shall provide a simple way of purging the panel.

E.	Leak Detection System:

1.	Each room shown on the plans with gas cabinets shall be equipped with a leak detection monitor and alarm.

2.	Air sampling tubing shall connect the gas cabinets to a central monitoring station. An audible and visual alarm shall be transmitted from the central monitoring station back to the affected room.

3.	The central monitoring system shall have the capability to sense all gases contained within cylinders in the gas cabinets.

4.	Install central monitoring unit with back plane 6 inches from wall to allow for tubing and wiring connections.

5.	Unit shall be as manufactured by MDA Scientific, Inc. or approved equal.

F.	Testing:

1.	All piping and tubing systems shall be pressure tested and leak checked using a helium gas spectrometer leak detector with inboard ad valve set testing performed at 1x10-9 cc/second of helium limit, and outboard testing performed a 1x10-7 cc/second helium limit. The entire system shall be tested so that all features are properly operating.

PART 3 -  EXECUTION

3.01	INSTALLATION

A.	Miscellaneous laboratory equipment shall be installed in strict accordance with the manufacturer's recommendations, including recommended operating and servicing clearances.

B.	Mechanical hookups under Division 15.

C.	Electrical hookups under Division 16.

D.	Check project for conditions that affect work. Do not begin installation until unsatisfactory conditions are corrected.

3.02	ADJUSTING

A.	Repair or remove and replace defective work as directed by Owner's Representative upon completion of installation.

B.	Adjust moving or operating parts to function smoothly.

3.03	CLEANING AND PROTECTION

A.	Clean shop finished equipment, touch up as required, and remove and refinish damaged or soiled areas.

B.	Cover equipment for protection against soiling and deterioration during remainder of construction period.

C.	Equipment shall be protected before, during and after installation. Damaged materials due to improper protection shall be cause for rejection.

END OF SECTION

# Comments

BEFORE EDITING SPEC....

1. Run English or Metric macro:

Alt-F10 English or Alt-F10 Metric

2. Remove Spec Writer Notes:

Alt-F10 Delnotes

This comment will not print and may be deleted.